10884 | 1220648-2025-26925 | 21682705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21682705&pc=OZD | 02/27/2025 | 06/26/2025 |  | 03/24/2025 |  |  | 02/27/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608638 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Gastrointestinal Hemorrhage | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Additional information received from the clinical site that the patient had a bedside colonoscopy today to further address continual gastrointestinal (gi) bleeding. | The user facility reported reported 67 year old male patient on impella 5.5 support had a slight hematoma/swelling above the insertion site. four days later the physician performed a surgical wound exploration, washout at the right axillary insertion site with drain placement. the patient survived the procedure. |  |  |  |  | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26925 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26925. g4 pma/510(k)number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26925. h6 health effect - clinical code 4582 was reported incorrectly on manufacturer device report 1220648-2025-26925. new code was added health effect - clinical code. | B5 revised with additional information received from the clinical site h6 health effect - clinical code revised to include additional failure mode should the device or any new information be received, a supplemental mdr will be filed. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
10885 | 1220648-2025-26917 | 21683533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21683533&pc=PYX | 02/27/2025 | 03/24/2025 |  | 03/24/2025 |  |  | 02/27/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025590534 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella rp encountered low flows. the position was verified and was proper. the rp was replaced and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10886 | 1220648-2025-26924 | 21683878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21683878&pc=OZD | 02/27/2025 | 04/01/2025 |  | 03/24/2025 |  |  | 02/27/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544885 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 22 year old male patient had an attempted impella 5.5 implant and during the implant, there were several attempts to place impella 5.5 via graft with 23 axillary sheath. physician stated ¿artery way too small¿. the pump was explanted. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since it was unable to implant impella due to vascular anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26924 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26924.
10887 | 1220648-2025-26919 | 21684375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21684375&pc=OZD | 02/26/2025 | 04/07/2025 |  | 03/24/2025 |  |  | 02/26/2025 | 03/24/2025 |  |  | 07/31/2025 |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442678 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Additionally, the staff pushed the aic and bed over a bump. subsequently, the impella alarmed impella stopped. the perfusionist dropped the p level to p0 and then back up to p7 and the impella restarted. | The user facility reported a 55 year old male patient was on impella cp support and during support, the patient had hemolysis and renal failure. hemolysis was confirmed by 2 plasma free hemoglobin tests. patient on continuous veno-venous hemofiltration (cvvh) and not making any urine. patient was successfully weaned and explanted. |  |  |  |  | The pump was returned for investigation. no physical damage was noted on the returned product. small biomaterial was reported on the outflow cascade. the pump was run as intended in a bucket of water. as per the clinical report, hemolysis was noted from the start of the case and improved after the pump repositioning. the patient was on continuous renal replacement therapy (crrt) throughout the case run. additionally, a temporary pump stop occurred during automated impella controller (aic) when the button was pressed while it was bumped. the data log shows the impella position as unknown, along with alarms and suction during the early pump start when hemolysis was reported. also, later on 03/07, a pump stop was reported with alarm 115 and an impella plug connect alarm notification. the cause of the hemolysis issue was most likely improper positioning of the device, since the repositing the device resolved the issue. the cause of the renal failure issue was not determined due to insufficient clinical details. the cause of the temporary pump stop issue was user-related, as the pump connection accidentally came off during transport. instructions for use impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ hemolysis: section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b1 revised from the initial the initial manufacturer device report number 1220648-2025-26919 to include product problem that was inadvertently omitted. b5 revised from the initial the initial manufacturer device report number 1220648-2025-26919 to include additional information that was inadvertently omitted. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26919 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26919. h6 medical device problem code 1503 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26919. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26919. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10888 | 1220648-2025-26896 | 21673400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21673400&pc=OZD | 02/24/2025 | 04/04/2025 |  | 03/23/2025 |  |  | 02/24/2025 | 03/23/2025 |  |  |  |  | 03/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616042 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Disability | Required Intervention |  |  |  |  | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. intra-aortic balloon pump was initially placed for cardiogenic shock and hemodynamic support. the patient was subsequently transitioned to impell 5.5 and venoarterial extracorporeal membrane oxygenation (va ecmo) support. later in the day the patient was started on continuous renal replacement therapy. two days after start of support, the patient experienced an acute right middle cerebral artery perfusion zone infarct. neurology consulted and electroencephalogram performed showed no seizure activity. per reports, neurology stated there were signs of early forming stroke on the computed tomography of the head performed prior to va ecmo and impella placement. systemic anticoagulation remained on hold. following, an echocardiogram showed improved heart function with ejection fraction 30%-35%. approximately three days later the impella 5.5 device was explanted with a survived outcome. it was reported the patient's neurological status had improved. the patient withdrew to pain lower left extremity and follows commands with movement on right side. | [5 empty] | The investigation of stroke/cerebrovascular accident (cva) has been completed. the returned product was visually inspected and no abnormalities were found. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10889 | 1220648-2025-26897 | 21673403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21673403&pc=OZD | 02/24/2025 | 11/25/2025 |  | 03/23/2025 |  |  | 02/24/2025 | 03/23/2025 |  |  |  |  | 03/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support, and during an attempt to reposition, placement signal not reliable alarm triggered with no aorta placement signal or left ventricle waveform available on screen. the decision made to replace the impella cp device with a new one which was successfully completed. the patient was reported to be stable following the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. data log review confirmed the placement signal not reliable alarm. at the time of the alarm, snr drops to zero, contrast oscillates, lamp increases, gain and light both drop. none of the optical signal metrics return for the rest of the case. returned product failed the 5s check. when connected to the console, a controller error alarm triggered. there was a kink noted in the catheter just distal to the kink protector. however, the pump failed laser continuity with light leaking on the patient cable side inside the red pump plug. upon further inspection, it was found that the patient cable could rotate freely and slide back and forth within the kink protector. this aligns with the clinical narrative that the alarm triggered while the pump was being repositioned. when the kink protector was cut open and inspected, it was able to be removed from the patient cable easily. there was evidence of glue residue, however, when compared to a control pump without bond issues, it did not demonstrate the same tearing of the kink protector to demonstrate proper adhesion between the two parts. based on data logs and returned product, the root cause of the optical signal issue was determined to be a damaged fiber most likely due to insufficient adhesive in the manufacturing process. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d8/d9 device returned to manufacturer updated. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10890 | 1220648-2025-26898 | 21673408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21673408&pc=OZD | 02/24/2025 | 03/23/2025 |  | 03/23/2025 |  |  | 02/24/2025 | 03/23/2025 |  |  |  |  | 03/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611188 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and following the implant, after approximately eight hours of support. the patient began to experience hemolysis. as a result of the patient's extreme coagulopathy issue and profound hemolysis, the decision was made to explant the pump. hemostasis achieved via perclose x2 and 40 minutes of manual pressure over site. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10891 | 1220648-2025-26899 | 21673411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21673411&pc=OZD | 02/24/2025 | 04/29/2025 |  | 03/23/2025 |  |  | 02/24/2025 | 03/23/2025 |  |  |  |  | 03/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613692 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella 5.5 device to a patient with cardiomyopathy was unsuccessful. the device was unable to pass through vessel due to calcium, and the case was aborted. there was no report of adverse effects to the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of delivery issue is determined to be patient condition because the 5.5 was unable tot be delivered due to calcification in the vessel. d6a and d6b were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26899. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10892 | 1220648-2025-26900 | 21673413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21673413&pc=OZD | 02/24/2025 | 05/07/2025 |  | 03/23/2025 |  |  | 02/24/2025 | 03/23/2025 |  |  |  |  | 03/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596896 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention implanted with an impella cp device for mechanical circulatory support experienced a positioning issue that required explant with device replacement. the longer 14fr x25 cm sheath was placed due to calcified anatomy. the impella cp device was placed halfway in the body and could not pass the calcium with sheath and dilator. dilator removed, shockwave advanced passed impella sheath and treated vessel. the dilator and wire placed back into sheath and able to advance. impella was placed over the guidewire, crossed the valve, and the pump was started. however, the pump suboptimal position. the physician attempted to reposition, however, was unable to get out of papillary muscle. the physician, therefore, elected to remove and replace the impella cp device which was successfully completed. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the pump positioning was poor after delivery as it was within the papillary. the team attempted to reposition and was unable to get out of papillary muscle and it came across the valve so then explanted the pump. second pump placed with good position. the cause of the positioning issue was most likely use related as the pump was originally delivered too deep into the papillary but the replacement pump was able to achieve good position. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10893 | 1220648-2025-26889 | 21671517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21671517&pc=OZD | 02/25/2025 | 04/23/2025 |  | 03/22/2025 |  |  | 02/25/2025 | 03/22/2025 |  |  |  |  | 03/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593633 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when impella cp started in auto the aorta (ao) diastolic pressures were low. suction alarms were present and not clearing along with placement signal low. impella was repositioned. no changes in impella ao pressure which wasn¿t match ao pressures on patients monitor or cuff pressure. it was recommended to exchange the impella but the physician wanted to proceed with the percutaneous coronary intervention (pci) procedure. after the pci, impella was wean and explanted. there was no harm to the patient. | [5 empty] | The investigation of positioning issues has been completed. the failure mode (fm) was changed from positioning issue to optical signal issue because the pump was reported to be in good position and repositioning did not resolve the placement signal low alarms or the ao pressure not matching the patient's ao monitor. the physician was concerned it could be the sensor and because everything else appeared fine. product and data were not returned for review. the root cause of optical signal issue was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10894 | 1220648-2025-26890 | 21671524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21671524&pc=OZD | 02/25/2025 | 04/11/2025 |  | 03/22/2025 |  |  | 02/25/2025 | 03/22/2025 |  |  |  |  | 03/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1853697 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller had a controller error alarm. aic was restarted and rechecked, but the alarm occurred again. this occurred during training therefore no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the data analysis confirmed the reported failure mode. the console was tested and after the power cycling the console 10 times and running an optical test pump for 2 minutes each time, unable to reproduce the controller error. the supply voltage to the optical bench at j3 was 4.9v, the average voltage to the lamp was 4.8v, and the lamp resistance was 4.5o. no issues were found with the console hardware. root cause: the root cause for the controller error was not determined due to the inability to reproduce. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10895 | 1220648-2025-26891 | 21671543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21671543&pc=OZD | 02/25/2025 | 03/22/2025 |  | 03/22/2025 |  |  | 02/25/2025 | 03/22/2025 |  |  |  |  | 03/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609717 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, there was a pump stop. high motor current was seen prior to the pump stop. the pump was explanted and there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10896 | 1220648-2025-26892 | 21671573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21671573&pc=OZD | 11/29/2024 | 03/22/2025 |  | 03/22/2025 |  |  | 02/25/2025 | 03/22/2025 |  |  |  |  | 03/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025559513 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute bleeding in the right middle and lower lobes with an "unknown (undetermined)" relationship to the impella device used. heparin was held and a stat bronchoscopy was performed and large clot noted to be obstructing the endotracheal tube that was unable to be aspirated with the bronchoscope. subsequent bronchoscopy was performed and copious blood secretions were noted on the right. multiple lavages of cold saline and epi given which appeared to have slowed the bleeding. patient received multiple units of blood products and heparin and other anticoagulation was continued to be held. the patient recovered with no sequelae. additionally, loqi reported that the patient endured left cerebellar stroke with an "unknown (undetermined)" relationship to the impella device used. the patient recovered with no sequelae. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10897 | 1220648-2025-26893 | 21671606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21671606&pc=OZD | 02/25/2025 | 03/22/2025 |  | 03/22/2025 |  |  | 02/25/2025 | 03/22/2025 |  |  |  |  | 03/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600837 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported had a 78 year old patient with a cp pump that failed delivery to the left ventricle. the native anatomy precluded placement and so the team instead placed a intra-aortic balloon pump. the pump was explanted and the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
10898 | 1220648-2025-26894 | 21671632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21671632&pc=OZD | 02/24/2025 | 03/22/2025 |  | 03/22/2025 |  |  | 02/26/2025 | 03/22/2025 |  |  |  |  | 03/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025601311 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient developed hemolysis following impella 5.5 implant. support continued for two days until the pump was explanted in response to the condition. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10899 | 1220648-2025-26895 | 21671717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21671717&pc=OZD | 02/26/2025 | 04/11/2025 |  | 03/22/2025 |  |  | 02/26/2025 | 03/22/2025 |  |  |  |  | 03/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599049 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66 year old female on cp support ongoing for an acute myocardial infarction. the patient was being considered for coronary artery bypass graft (cabg) and valve disease versus percutaneous cardiac intervention. after 18 hours, the pump was explanted due to hemolysis concerns and escalation to the 5.5 pump. the patient survived the procedure. | [5 empty] | The investigation of hemolysis has been completed. the returned product was inspected and there were no physical abnormalities. there was no biomaterial observed around the impeller. the pump was tested in the lab for hemolysis and it passed with an mih of 7.32. the data logs show no motor current spikes and there were no positioning related alarms or continuous suction alarms. the root cause of the hemolysis issue was not determined as there were no abnormalities on the returned product, or in the data logs, and limited clinical information was provided g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26895. | The impella was received for evaluation, upon completion of the investigation, a supplemental mdr will be filed. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
10900 | 1220648-2025-26866 | 21661802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21661802&pc=OZD | 02/23/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 02/23/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577163 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after placement of an impella cp, a suction alarm occurred. troubleshooting was unsuccessful. the impella cp was exchanged and there was no harm to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10901 | 1220648-2025-26867 | 21661820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21661820&pc=OZD | 02/26/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 02/26/2025 | 03/21/2025 |  |  | 09/30/2025 |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445464 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced atrial fibrillation requiring cardioversion. the patient presented to the emergency department with st-segment elevation myocardial infarction and had two stents placed in the left anterior descending artery (lad). the patient went into ventricular tachycardia arrest, was shocked and returned to spontaneous circulation. an intra-aortic balloon pump (iabp) was placed, and the patient was sent to the unit. however, the patient became worse and was placed on an impella cp device. the iabp was removed. three days after start of the impella mcs, the patient went into atrial fibrillation with rapid ventricular response requiring cardioversion. the plan was noted to optimize the lad stents, decannulate ecmo and if all goes well explant the impella cp device which was completed two days later. the patient had improved ejection fraction of 55%. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10902 | 1220648-2025-26868 | 21662346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21662346&pc=OZD | 02/23/2025 | 04/23/2025 |  | 03/21/2025 |  |  | 02/23/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599947 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pain | Sepsis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after reaching for the bedside table during impella 5.5 support, the patient developed increasing pain and swelling at their access site. a lidocaine patch was applied and oxycontin was given to manage the pain. support continued uninterrupted with the implanted pump. roughly ten days into support, the patient developed an infection of unknown origin. antibiotics were given, but the patient became septic and began to decompensate. the decision was made to withdraw care and the patient passed away. thus, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of pain at insertion site and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10903 | 1220648-2025-26869 | 21663094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21663094&pc=OZD | 02/24/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 02/24/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025598695 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed via the 14fr sheath. the team had placed the pump for the support of a patient admitted in with acute myocardial infarction and cardiogenic shock. the team did not explant/peel away the 14fr sheath but rather left in the femoral. the leg was noted to be warm and perfused. additional details noted the patient died due to multi-organ failure. the leg was weak but perfused all the time. given the slight reduction in perfusion there is not enough information to exclude the impella device as an associated factor to the event outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10904 | 1220648-2025-26871 | 21663902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21663902&pc=OZD | 02/22/2025 | 04/23/2025 |  | 03/21/2025 |  |  | 02/22/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573828 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs) and the device became malpositioned requiring repositioning in which the sterile sleeve was damaged in the process. the patient would subsequently expire. the patient came as an outpatient for an endoscopy. during recovery the patient went into pulseless electrical activity arrest and received three rounds of cardiopulmonary resuscitation with return of spontaneous circulation. the patient was cannulated for extracorporeal membrane oxygenation (ecmo) and impella implanted for vent. approximately two days after start of the support an impella in aorta alarm triggered. echocardiogram confirmed, and the physician was able to reposition the pump back across the aortic valve into the left ventricle measuring 5.0 cm. in the process, the sterile sleeve was damaged (no action was taken regarding the damaged sleeve per the report). the patient continued on support, but prognosis was "poor." six days later, the patient had a flat electroencephalography and pupils were non-reactive. the patient was planned to undergo brain death testing. two days later, however, it was noted the patient quickly decompensated despite being maxed on pressers and impella mcs, coded and was unable to be revived after cardiopulmonary resuscitation. medical safety review of the event noted there is not enough information to exclude the impella as an associated factor to the event given the malposition of the device. the pump moved, was displaced from its intended position and cause is unknown. | [5 empty] | The investigation of positioning issues and product damage (sleeve) has been completed. the pump was not returned for investigation. as per the clinical information provided, the sterile sleeve was damaged while repositioning the pump. the cause of the product damage (sleeve) issue was not determined since product/images were not returned. as per the clinical information provided, the patient was rolled over which led to the pump getting pulled into the aorta. therefore, the root cause of the positioning issue was most likely patient movement. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26871 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 medical device problem code 4008 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26871. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10905 | 1220648-2025-26872 | 21663998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21663998&pc=OZD | 02/23/2025 | 05/07/2025 |  | 03/21/2025 |  |  | 02/23/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025578576 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient suffered from a small stroke during impella 5.5 support. it was noted that a clot had been observed prior to impella insertion, but was no longer visible during implant. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of stroke/cva has been completed. the pump was cut from the catheter when it was returned in the box for investigation and has no impact on the investigation. there were no abnormalities with the pump and the catheter upon investigation. borescope images were taken and found no biomaterial presence inside the cannula. stroke/cva: "strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. -timing of the stroke in relation to impella support (during or post-implant). -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed. -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. -response to any previous anticoagulation or antiplatelet therapies. -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿ d9 device returned to manufacture was erroneously selected no, however the device was returned at the time of initial manufacturer device report number 1220648-2025-26872. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26872 as it is unknown if the initial reporter also sent the report to the food and drug administration. h3 was erroneously entered on the initial manufacturer device report number 1220648-2025-26872. h6 code 4114 was selected entered on the initial manufacturer device report number 1220648-2025-26872. h10 narrative erroneously stated the device was not returned on the initial manufacturer device report number 1220648-2025-26872. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10906 | 1220648-2025-26874 | 21664130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21664130&pc=OZD | 02/25/2025 | 03/27/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607737 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 78 year old female patient with a cp placed for the patient undergoing a high risk percutaneous cardiac intervention (hrpci). the pump was placed and allowed for the intervention and then the pump was explanted. the explant was complicated by groin closure devices failing to achieve hemostasis. the team had to consult with vascular surgery due to the loss of flow and repair of the access site. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was most likely use issue since the perclose failed. the root cause of the limb ischemia cannot be determined since the product was not returned and insufficient clinical details were provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26874 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26874. g4 pma/510(k)number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26874. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10907 | 1220648-2025-26875 | 21664457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21664457&pc=OZD | 01/10/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 01/21/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575775 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 37 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient who experienced endured hemolysis with a "possible" relationship to the impella device used. the patient was a heart transplant rejection and would subsequently expired. the patient was on impella cp mechanical circulatory support prior for heart kidney transplant evaluation. that impella was successfully weaned. however, the patient has since decompensated requiring increased inotropes/pressors. the patient had a pulseless electrical activity arrest at the outside hospital requiring venoarterial extracorporeal membrane oxygenation (va ecmo) and impella cp placement (ecpella). the patient was also started on continuous renal replacement therapy. on day 8 of ecpella support, the patient was transitioned to bipella support and va ecmo was removed. additionally, on this day per the registry, approximately eight days after the start of the impella cp support, during the night, the patient had impella rp suction alarm and was decreased to performance level p-7. per notes, the impella cp also decreased to p-7 for hemolysis. the patient was transfused with packed red blood cells due to increasing lactate and escalating norepinephrine. there was no bleeding noted but monitored for signs. per the registry, it was reported that the patient/event did not resolve. the patient would subsequently expire five days later. it was noted the patient was not improving in heart and renal failure and the patient asked for withdrawal of life-sustaining treatment. although care was withdrawn, there is not enough information to exclude the impella cp device as an associated factor given the unresolved hemolysis event. | [5 empty] | The pump was not returned for investigation. however, data logs were received and analyzed. hemolysis was reported during impella cp support. clinical information mentions brief suction on the impella cp and impella rp flex pumps which ultimately was resolved. blood was given. data log showed brief suction alarm. no positioning issue occurred. no placement signal trend observed. no motor current spike outside of p-level changes observed. the cause of the hemolysis issue could not be determined since the complaint device not returned for analysis and insufficient clinical data was provided. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10908 | 1220648-2025-26873 | 21664469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21664469&pc=OZD | 02/21/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599291 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Frayed | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | GUIDEWIRE 3000067 | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant was prepping to insert a impella cp pump when the tip of the guide wire was found to be frayed. a new guide wire was used. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10909 | 1220648-2025-26877 | 21664888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21664888&pc=OZD | 02/25/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025526612 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 44 year old male patient was on impella 5.5 support with a direct impella insertion access site. after the implant, the patient had bleeding from chest tube drainage. three units of packed red blood cells were given to the patient. additionally, the patient had hemolysis. the doctor believed the red urine was a combination of traumatic foley insertion and hemolysis. two units of blood products were given to the patient. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10910 | 1220648-2025-26876 | 21665103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21665103&pc=OZD | 02/24/2025 | 04/07/2025 |  | 03/21/2025 |  |  | 02/24/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607238 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | Failure to Advance | Patient-Device Incompatibility |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placement going on for a patient admitted in with cardiogenic shock and cardiomyopathy. the pump delivery was complicated. the heart team took the impella in and out of the body twice and were unable to advance. the impella sleeve was found to have not been pulled back prior to insertion and was preventing advancement. the patient was very tortous which the heart team thought might have also contributed. the team was able to re-advance and upon starting the pump, high motor current and pump stop occurred. they were unable to restart after two attempts so the decision was made to remove and implant a new impella pump. the new pump started. there was no harm to the patient. | [5 empty] | The investigation of delivery issues and pump stop has been completed. the pump was returned for investigation. the data log was not provided for review. the returned pump passed electrical testing. no physical abnormalities were observed on the returned pump. some biomaterial was observed in the purge gap. the pump was then tested in a bucket of water, and saturated flow was observed. the pump was primed successfully, and no issues related to purge pressure were noted. as per the clinical details, pump was inserted and removed twice per the nurse but could not be advanced due to the impella sleeve not being pulled back, along with the patient's tortuous anatomy. after re-advancing, the pump was started but experienced high motor currents and a pump stop. the pump stop failure mode was changed to 'pump did not start' failure mode based on the investigated clinical details. the cause of the pump did not start issue was not determined since it was not reproduce during testing and data logs were not provided. the cause of the delivery issue was most likely patient condition-related, as the clinical report indicated that the patient had tortuous anatomy. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system . section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention . section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10911 | 1220648-2025-26878 | 21665275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21665275&pc=OZD | 02/19/2025 | 04/24/2025 |  | 03/21/2025 |  |  | 02/22/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616122 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support, and approximately three days into support, bleeding developed from artificial vascular anastomosis. a blood transfusion of an unknown quantity was performed, and anticoagulation was put on hold in response to the bleed. this resulted in an increase in purge pressure, and due to concerns of a potential pump stop, the decision was made to replace the pump which was successfully completed. however, the patient would expire 23 days later. it was noted heart failure persisted and the patient expired. the relationship between the bleeding and impella 5.5 device was noted to be unknown. medical safety review noted there is not enough evidence to exclude impella 5.5 pump 1 as an associated factor in the patient's outcome (demise). bleeding developed from the artificial vascular anastomosis. a blood transfusion of unknown quantity was performed and anticoagulation withheld. | [5 empty] | The investigation of high purge pressure and vascular damage - peripheral vessel has been completed. high purge pressure: clinical details report that on (b)(6) 2025, an increase in purge pressure was noted, though not high enough to trigger an alarm. earlier the same day, anticoagulation was stopped to troubleshoot bleeding. data log review confirmed the rise in purge pressure. the purge pressure rose over ~2 minutes to 850 mmhg. there were no motor current spikes leading up to the event, but motor current modulation did dampen. the purge pressure remained elevated through (b)(6) 2025, and during this time there was an increased jagged amplitude in the purge pressure. on the (b)(6) 2025, it dropped to previous levels within a few minutes. the only other pump metric that showed any change at the same time was a slight step up in the ps. returned product was investigated and there were no kinks in the catheter, biomaterial in the purge gap, or any other visual abnormalities. the pump and purge cassette were connected to the console, successfully primed and run for over an hour without high purge pressure. that being said, pump flows were saturated for the first half of the hour, before dropping within specification on their own. the patterns noted from data log review and returned product could be evidence of several causes for high purge pressure. therefore, based on the inconclusive evidence from data log and product review, the root cause of the high purge pressure could not be determined. vascular damage: it was reported that on the 4th day of support bleeding was noted from the vascular graft. the bleed had to be treated surgically by open hemostasis and transfusions. data logs are not relevant to the complication. returned product was investigated and there were no visual abnormalities noted. since limited clinical details were provided and returned product did not exhibit any abnormalities, the root cause of the vascular damage could not be determined. h6 health effect - impact code 1802 death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26878. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10912 | 1220648-2025-26879 | 21665765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21665765&pc=OZD | 02/25/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A literature article found that a patient was on impella support and during support, when the patient attempted to wean from cardiopulmonary bypass (cpb) after resuming percutaneous cardiopulmonary support, a left ventricular free wall rupture was observed. the patient continued to bleed afterwards and underwent surgical treatment (left ventricular reconstruction under cardiac arrest). the details of the left ventricular rupture and its relationship to impella were unknown. the patient expired. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1, a2, a.5 are unknown. d.4 without exact model number, the catalog number, serial number, lot number, expiration date and unique identifier (udi) #, d.6a and d.6b implant and explant dates, and h.4 device manufacture date is unknown. e.1 name prefix/title, first name and last name are unknown.
10913 | 1220648-2025-26880 | 21665992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21665992&pc=OZD | 02/25/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599068 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had a gastrointestinal bleed (gib). the patient was on heparin drip prior to the bleed. two units of blood products were given to the patient. the patient was deemed stable for an explant and the pump was explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10914 | 1220648-2025-26855 | 21666010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21666010&pc=OZD | 02/20/2025 | 04/14/2025 |  | 03/21/2025 |  |  | 02/24/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the impella guidewire was bent and kinked which made access to be difficult. the impella procedure was aborted. there was no patient harm reported and the procedure will be rescheduled for a different time. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1, a.4 and a.5 are unknown. d4-exact model number, catalog number, serial number, lot number, expiration date, and udi number are unknown. d6a/d6b-unknown implant and explant date. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. | The investigation of product damage (guidewire damage - peripheral vessel) has been completed. the product was not returned for investigation. data log review was not required for this case run. guide wire was reported bent and kinked during procedure. the cause of the guidewire damage issue was not determined since product was not returned and sufficient clinical information was not provided. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-26855 e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26855 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10915 | 1220648-2025-26881 | 21666163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21666163&pc=OZD | 02/25/2025 | 03/25/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610576 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 59 year old female patient was on impella cp support and during support, there was a shockwave to the left anterior descending artery and the console alarmed placement signal not reliable. it was assumed that the shockwave damaged the optical sensor. the impella was escalated to an impella 5.5. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. the clinical details mention that the placement signal went out after shockwave was utilized. according to the data logs, the parameters resemble that of shockwave. no product was returned. the minimum distance between shockwave and the sensor is unknown. shockwave should not be used less than 20mm within impella distance. the root cause of the optical signal issue is most likely due to a damaged sensor face from shockwave. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26881 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26881.
10916 | 1220648-2025-26882 | 21666589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21666589&pc=OZD | 11/19/2024 | 04/01/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560870 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported via abiomed's long-term outcome and quality indicator (loqi) impella registry that a 44-year-old male patient endured an occlusive thrombus of the right innominate vein during impella rp flex support. the thrombus was stated to have a definite relationship to the impella device. per the registry, the acute partially occlusive thrombus was in the right innominate vein with an additional abnormality in the right jugular vein. heparin was given to the patient to help manage the condition. patient bridged to transplant. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-26882. d4 model number was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2025-26882. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26882 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26882. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26882. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella rp with smartassist system states the following: section: using the automated impella controller with the impella catheter "achieving an act: 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup."
10917 | 1220648-2025-26883 | 21667683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21667683&pc=OZD | 02/25/2025 | 03/21/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568651 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a 57 year old male patient endured hemolysis during impella 5.5 support. the condition was believed to have a possible relationship to the impella device. hemoglobin levels were dropping and three units of packed red blood cells were given to treat the condition. after thirty six days of support, the patient's family made the decision to withdraw care and the patient passed away. there is not enough evidence to disassociate the use of the impella with the patient's passing. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions ¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes ¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10918 | 1220648-2025-26884 | 21668205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21668205&pc=OZD | 02/25/2025 | 04/02/2025 |  | 03/21/2025 |  |  | 02/25/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592330 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support a patient admitted in with cardiogenic shock and awaiting transplant. the pump was placed but came out of position and was within the aorta. the team sent the patient back to the operating room for replacement. the team additionally had heparin paused due to concerns of hematoma. the patient had a new pump placed and support continued. the pump had supported for 19 days before replacement. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of positioning issues and access site bleeding has been completed. the data logs confirmed pump was in improper position corresponded with clinical description (about 386 hours into the case) that triggered alarms impella position in aorta. pump then ran for more than 72 hours and stop manually. no other issues were found on the logs. product was returned in damaged condition (missing pump plug). visual inspection found catheter was stretched from mark 47-49. there was no issue found on the cathater anchor. the anchor go back and forth easily. no other issues were found on the rest of the pump. the root cause of positioning issue was unable to be determined as limited clinical details related to pump came out of left ventricle were provided. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no problem might related to the bleeding was found on the pump. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26884 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10919 | 1220648-2025-26885 | 21668255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21668255&pc=OZD | 02/26/2025 | 12/02/2025 |  | 03/21/2025 |  |  | 02/26/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during inspection of the console (aic), the selection knob was not working. there was no patient involvement. | [5 empty] | The device has been received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. corrections to the initial mfr report: d2 device name has been updated. h6 type of investigation, findings, and conclusion codes added. | The investigation for the reported issue has been completed. data analysis: no log evidence was found to confirm the reported failure mode. after the complaint date, the logs recorded abnormal optical signal parameters, which is unrelated to the reported failure mode. device analysis: the rotary knob issue was identified during service at fs. the issue was reproducible, upon pushing the knob, it does not return to its home position. there was no evidence of contamination on the rotary knob, also there was no fluid ingress inside the aic or behind the impellatronic board. regarding the abnormal optical parameters, the front panel optical connector was found to be contaminated, still unable to reproduce this issue (unreported and unrelated to the reported fm). no hardware issue was found related to the abnormal optical parameters. root cause: the root cause for the knob not returning to its home position was most likely a rotary encoder malfunction. h6 investigation type, findings and conclusions d9 date device returned to manufacturer. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10920 | 1220648-2025-26886 | 21668259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21668259&pc=OZD | 02/26/2025 | 10/16/2025 |  | 03/21/2025 |  |  | 02/26/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1889477 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the aortic (ao) and left ventricle (lv) curves were out of range on the impella cp. after impella was weaned and disconnected, the console (aic) was still alarming retrograde flow. unable to turn off aic. a hard reset was done. there was no patient harm reported. | [5 empty] | Based on the clinical data, data and device analysis the root cause of the automated impella controller (aic) shutdown issue was the impellatronic malfunction. the root cause of the ao and lv curves being out of range was most likely due to optical bench malfunction. d4 lot number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26886. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26886 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26886. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
10921 | 1220648-2025-26887 | 21668260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21668260&pc=OZD | 02/24/2025 | 09/22/2025 |  | 03/21/2025 |  |  | 02/26/2025 | 03/21/2025 |  |  |  |  | 03/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589720 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 55-year-old male patient developed a hematoma at their axillary site during impella 5.5 support. the patient mentioned experiencing pain at the site and a pressure dressing was placed. heparin was stopped the following day. support continued uninterrupted. there was no patient harm reported. | The complainant also reported a placement signal not reliable alarm, with absence of aorta and left ventricle waveforms. a second instance of the placement signal going out overnight and occurring before was reported. the pump position was verified with echocardiography. |  |  |  |  | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-26887. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26887 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26887. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26887. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. data logs were investigated and there was a placement signal not reliable (psnr) alarm. upon review of data logs, it is apparent that the console is the potential cause of the issue. therefore, as per the data log review, the root cause of the optical signal issue was sensor silicone wear. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue descriptions. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-26887. d6b updated. d10 concomitant medical products and therapy dates updated.
10922 | 1220648-2025-26834 | 21650964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21650964&pc=OZD | 02/24/2025 | 05/07/2025 |  | 03/20/2025 |  |  | 02/24/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617552 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patent had hemolysis and there were positioning issues. the impella cp was measuring 3.1-3.3 centimeters from the aortic valve, but the inlet was close to the septal wall. the doctor did not want to reposition. additionally, the patient had hemolysis. the patient had bright red urine. labs were hemolyzed. the p level was reduced to attempt to reduce the hemolysis, but the hemolysis was never confirmed to be resolved. the impella cp was explanted after interventions were performed and intra-aortic balloon pump was placed. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. the data logs showed pump ran without issue during support. there was a positioning alarm triggered in the end of the case but resolved quickly. no motor current spike was seen. placement signal noisy initially and then rose after p-level decrease. product was not returned for review. as per clinical, high afterload and low cvp was confirmed indicating likely poor patient condition. the root cause of the hemolysis issue was most likely patient condition related per clinical and log review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review.
10923 | 1220648-2025-26835 | 21651052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21651052&pc=OZD | 02/23/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/24/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1727322 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that a purge failure occurred involving an automated impella controller (aic). the aic was swapped and the issue resolved. there was no patient harm. | [5 empty] | The investigation for the automated impella controller (aic) purge issue has been completed. in the ¿case start¿ wizard, before connecting the pump, the console recorded 'piston blocked,' which resulted in repeated pud reset. although the cassette was removed and reinserted several times, the piston block issue was never resolved, and the console displayed 'purge system failure (piston blocked) ¿ alarm,' which remained unresolved on this console per the data logs. this confirmed the reported failure mode. this console was tested. dextrose residue was confirmed behind the purge cassette insert and the purge motor gasket. after cleaning the dextrose, purge system failure alarm did not occur. the root cause of the purge system failure alarm was dextrose deposited between the purge motor gasket and behind the purge cassette insert.
10924 | 1220648-2025-26836 | 21651239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21651239&pc=OZD | 02/23/2025 | 04/09/2025 |  | 03/20/2025 |  |  | 02/23/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025575896 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of an impella 5.5 device in a patient with cardiomyopathy was unsuccessful. the physician was unable to advance the impella 5.5 device past the axillary artery despite several efforts and manipulation of the artery. the impella 5.5 device was removed and a 0.035 buddy wire over the pigtail was attempted. the impella 5.5 device once again was attempted but failed to advance further. at this point, the physician opted to abort further efforts of axillary insertion in favor of carotid approach. the physician expressed concern for damage to optical sensor related to the attempted implant and, therefore, requested a new pump for carotid approach. it is unknown whether a new pump was implanted. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation of delivery issues and optical signal issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issues was most likely patient-condition related since the clinical details note the patient's poor vasculature. the root cause of the optical signal issue was not determined since no abnormalities were found in the data logs, product was not returned, and insufficient clinical details were provided. h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26836. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10925 | 1220648-2025-26837 | 21651333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21651333&pc=OZD | 02/21/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/21/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580648 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and one hour after the impella cp device was implant, the purge pressure began to increase coinciding with a drop in purge flow. tissue plasma activator-lysis was added to the purge line, and the purge values improved. impella mcs continued uninterrupted with the implanted pump until successfully weaned and explanted as planned. there was no patient harm to the patient as a result of the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10926 | 1220648-2025-26838 | 21651410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21651410&pc=OZD | 02/22/2025 | 03/31/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603501 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, there was bleeding at the access site. a jp drain to the chest was placed and emptied multiple times. another suture was placed at the drain site. bleeding then subsided around at the access site. systemic anticoagulation was started later that night. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of vascular damage was most likely use issue since hemostasis was achieved by adding additional suture.
10927 | 1220648-2025-26839 | 21651489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21651489&pc=OZD | 02/22/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025609797 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that roughly six days into impella 5.5 support, flows became lower than expected. the pump was subsequently replaced to resolve the noted issue. it was additionally documented that biomaterial was found at the inlet cage of the removed pump. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10928 | 1220648-2025-26840 | 21651618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21651618&pc=OZD | 02/22/2025 | 04/04/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600769 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 84-year-old female on a cp pump for 5 days. the pump was explanted with loss of pulses in the cp limb. the explant took place and thrombus burden was observed which prompted surgical repair. patient now supported by the axillary 5.5 pump. | [5 empty] | The investigation of limb ischemia and thrombus has been completed. the pump was returned for investigation and no abnormalities were found upon investigation. the data logs were reviewed and it was noted that there is a spike in purge pressure and an alarm for high purge pressure was noted at that time. the spike corresponds to the presence of biomaterial and was reported for failure mode thrombus during the case. as the necessary information was not provided, the root cause of the thrombus and limb ischemia was not determined g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26840. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26840. | The device was received and an evaluation will be done. upon completion of the investigation, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation; section: impella cp with smart assist catheter insertion; section: axillary insertion of the impella cp with smart assist catheter; section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10929 | 1220648-2025-26841 | 21651643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21651643&pc=OZD | 02/22/2025 | 04/16/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599945 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support when the pump stopped unexpectedly. the pump was exchanged as a result of the failure. there was no patient harm. | [5 empty] | The investigation of pump stop has been completed. the failure mode (fm) high purge pressure was added as it was claimed that there was increasing purge pressure and a decrease in purge flows. there was no product returned. the data logs were analyzed and on the 22nd there was a gradual build in purge pressure over the course of 3 hours with no prior motor current spikes or long bag/cassette changes near the instance of of this gradual climb. at 7pm on the 22nd there was a rapid rise in purge pressure ~ 45 minutes that then triggered hpp alarms. purge pressure remains high for about two hours until 9:40pm the pump is disconnected. based on data log analysis, the root cause of the high purge pressure is most likely due to biomaterial buildup but the cause is unknown. the clinical details state that there was a pump stop at 8:40pm but there is no sign of a pump stop in the logs. there is no upward trend in motor current. due to lack of product returned and no pump stop alarm seen in the logs, the root cause of the pump stop cannot be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10930 | 1220648-2025-26842 | 21651736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21651736&pc=OZD | 02/22/2025 | 04/29/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548402 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient's condition post impella was noted as critical and approximately one day later the patient underwent a femoral crossover bypass due to limb ischemia that was noted to be related to impella access. the patient would later expire on support that evening. no further details regarding the event were noted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10931 | 1220648-2025-26844 | 21652054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652054&pc=OZD | 02/22/2025 | 10/01/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508818 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had a impella 5.5 pump placed as care escalation from the impella cp. it was reported that the patients had on and off concentrated amber/red urine output. it was further reported that care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant also reported that the implanted impella 5.5 pump experienced intermittent placement signal issues during patient support. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | H1 type of report erroneously selected on the initial manufacturer device report number 1220648-2025-26844. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue and hemolysis. no product was returned for analysis. the data logs were evaluated and the 5s parameters were analyzed and it was suspected that there was possibly the presence of an additional device causing an effect of the placement signal. the root cause of the optical signal issue was not determined as no product was returned and limited clinical information related to failure was provided. the data logs showed no motor current spikes. clinical mention that good pump position was confirmed. it is unknown if hemolysis resolved with pump repositioning or removal. the root cause of the hemolysis was not determined as no product was returned and insufficient clinical information related to failure was observed. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding.
10932 | 1220648-2025-26843 | 21652055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652055&pc=PYX | 02/21/2025 | 04/29/2025 |  | 03/20/2025 |  |  | 02/21/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551160 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (post cardiothoracic surgery) was on the unit, decompensated and was brought to the catheterization lab for an impella rp flex implant. the right ventricle decompensation was noted on echocardiogram. the arterial line was placed with mean arterial pressure in the 10's. cardiopulmonary resuscitation was initiated. venous access was obtained. the impella rp flex was primed and connected. however, insertion was aborted due to patient anatomy reason. the patient was pronounced as deceased. no further details surrounding the event were noted, and there is not enough information to exclude the impella rp flex as an associated factor. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through right ventricle and lack of forward flow.
10933 | 1220648-2025-26845 | 21652152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652152&pc=OZD | 02/22/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  | 07/31/2025 |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443606 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a patient. it was reported that while on support there was issues with limb ischemia. the team placed fem-fem bypass externally. after 22 hours the pump was explanted for a new axillary delivered 5.5 pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10934 | 1220648-2025-26846 | 21652267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652267&pc=OZD | 02/22/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504003 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a computed tomography scan revealed that a patient suffered from an acute bilateral cerebral infarction during impella cp support. support continued uninterrupted with the implanted pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10935 | 1220648-2025-26847 | 21652452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652452&pc=OZD | 02/23/2025 | 04/25/2025 |  | 03/20/2025 |  |  | 02/23/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603506 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during placement of impella 5.5, the impella 5.5 was unable to pass through the axillary artery. the impella 5.5 was removed and then loaded back on the 0.018 wire. the impella 5.5 was advanced across the aortic valve. transesophageal echocardiogram showed inlet facing the mitral valve and the impella 5.5 was unable to be repositioned due to vascular anatomy. the decision was made to remove the impella 5.5 and move forward with an impella cp insertion. the patient was stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since it was unable to reposition impella due to vascular anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10936 | 1220648-2025-26848 | 21652705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652705&pc=OZD | 02/23/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/23/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558905 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | Swelling/ Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient (unknown ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced thrombus leading to explant of the device. the patient decompensated and was placed emergently on extracorporeal membrane oxygenation (ecmo) after presumed myocarditis as culprit. ejection fraction was found to be 5%. the decision was made to implant an impella cp device for decompression of the left ventricle, and the patient received distal perfusion catheter on both lower extremities. approximately one day after start of the support, the distal perfusion catheter (dpc) had multiple clots. a thrombectomy and dpc cutdown was performed. the patient had increased swelling in their leg after the intervention; however, the doppler showed distal flow. the team made plans to explant the impella cp device which was successfully completed the following day. the patient's status following the event was noted as unknown. however, the patient was noted to have survived the impella explant. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10937 | 1220648-2025-26849 | 21652817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652817&pc=OZD | 02/21/2025 | 10/01/2025 |  | 03/20/2025 |  |  | 02/21/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579082 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy implanted with an impella 5.5 device for mechanical circulatory support experienced hemolysis that led to the device being explanted. the patient was being worked up for heart transplant and was transferred in with an impella cp device in place. the patient was escalated to an impella 5.5 device which was successfully implanted via the axillary artery using double barrel technique. the impella cp device was removed without complication. approximately 11 days after start of the support on the impella 5.5 device, there was an impella in aorta alarm with pulsatility to motor current; however, correct position verified with echocardiogram per the report. the patient continued on support, and the next day it was noted there were no active alarms present. now day 26 on support, the patient was noted to have developed hemolysis. consequently, the device was explanted and ultimately replaced with venoarterial extracorporeal membrane oxygenation after unsuccessful attempts at a new impella 5.5 (case aborted) and then unsuccessfully attempting an intra-aorta balloon pump placement. | The customer reported that the device exhibited an optical signal issue. the pump remained on for support despite the reported problem. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two devices associated with the patient's implant experience. refer to manufacturing report number 1220648-2025-26812 for the second used impella 5.5 serial number (b)(6). instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. optical signal issue data logs show several impella in aorta alarms throughout the case. the snr is relatively low throughout the case and there are abnormal spikes in the raw optical signal. the position was verified to be correct in the field. upon inspection of the data logs, it is apparent that the console is the potential cause of the issue. please refer to 25-41340-2 for an investigation into the console. hemolysis data logs show no abnormalities in motor current. minimal clinical details were given about the hemolysis. the root cause of the hemolysis was not determined since product was not returned and insufficient clinical details were provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10938 | 1220648-2025-26850 | 21652860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652860&pc=OZD | 02/21/2025 | 04/08/2025 |  | 03/20/2025 |  |  | 02/21/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582867 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful. the guidewire was advanced with difficulty into the left ventricle. once the impella was inserted the physician, however, was unable to remove the guidewire. both devices were removed as a united and another impella cp device was implanted. there was no report of adverse effects to the patient. | [5 empty] | The impella guidewire and pump were returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. | The pump was returned for inspection and no damages were found with the pump, catheter and the guide wire upon investigation. guidewire able to be removed as usual. the root cause of delivery issue is determined to be most likely use issue because the second pump was successfully implanted without any issues and there were no damages to be noted or any issues found with the returned product and the guidewire was able to be removed. d6a and d6b were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26850. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26850 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10939 | 1220648-2025-26852 | 21652935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652935&pc=OZD | 02/22/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536798 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that following impella 5.5 removal, the patient developed moderate to severe aortic insufficiency post impella removal. the physician believed the condition could potentially be related to the impella device. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10940 | 1220648-2025-26853 | 21652965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652965&pc=OZD | 02/23/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/23/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582871 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute decompensated heart failure, atrial flutter with rapid ventricular response and diuresis continued to decompensate was taken to the catheterization lab for an impella 5.5 device implant for support and left ventricular assist device/transplant work up. approximately five days after start of the impella mechanical circulatory support (mcs), the patient had clotting in the right arm; however, the arm still had pulses and was warm. heparin treatment was initiated to manage the thrombus, and the impella 5.5 mcs continued uninterrupted. approximately 16 days later the patient was bridged to heart transplant. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smart assist system instructions for use for the related event are as follows: section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10941 | 1220648-2025-26854 | 21652992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21652992&pc=OZD | 02/12/2025 | 04/15/2025 |  | 03/20/2025 |  |  | 03/11/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1170637 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Upon review of the impella cp log, it was noted that a pump stop occurred during patient support. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - pump stop has been completed. based on clinical details and data logs, the root cause of the aic issue was a defective impellatronics board. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26854 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26854 and were added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10942 | 1220648-2025-26857 | 21653311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21653311&pc=OZD | 02/24/2025 | 04/03/2025 |  | 03/20/2025 |  |  | 02/24/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601586 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 placement, the impella 5.5 could not be torqued enough to pass the innominate. the procedure had to be aborted. the patient's outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the device could not be torqued enough to pass the innominate. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26857 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10943 | 1220648-2025-26856 | 21653448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21653448&pc=OZD | 02/24/2025 | 04/02/2025 |  | 03/20/2025 |  |  | 02/24/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621511 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old female with an impella cp for acute myocardial infarction. during the bipella procedure, once both pumps were in and flow balance achieved, the physician used the cp side port to verify flow. no flow noted below the impella sheath, so 7fr antegrade sheath was placed to the lfa. this was connected externally to the 7fr retrograde sheath in the right femoral artery (rfa). no pulses were dopplered, so the physician performed contrast shot through the antegrade sheath. there was occlusion noted at the level of the left knee. the physician then performed balloon angioplasty until flow returned to the anterior tib. the patient was sent to icu. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26856 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26856. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26856.
10944 | 1220648-2025-26859 | 21653774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21653774&pc=OZD | 02/22/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/22/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636279 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, pulses became intermittent to the left leg, with no flow. days later, an above knee amputation was performed on the impella side leg. support was continued until successful explant. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10945 | 1220648-2025-26860 | 21654634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21654634&pc=OZD | 02/23/2025 | 04/23/2025 |  | 03/20/2025 |  |  | 02/23/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610287 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the staff was unable to get optimal positioning for the pump. despite multiple attempts to reposition, the staff had to settle for a depth of 3.4 centimeters with the device pointing towards the mitral valve. as support continued, the patient coded and passed away. it is unknown if the positioning issue contributed to the patient's passing. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the positioning issue most likely was use related since the impella stacked against mitral valve. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10946 | 1220648-2025-26861 | 21654639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21654639&pc=OZD | 02/23/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/23/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600845 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a patient admitted in with a st-segment elevation myocardial infarction. the pump was supporting when there were signs of hemolysis. the patient had urine color change. position and fluid status were evaluated. the pump remained on for support. the hematuria was improving the following day. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10947 | 1220648-2025-26862 | 21658059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21658059&pc=OZD | 02/19/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/19/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603992 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced low pump flow which resulted in replacement of the device with a demise outcome. the patient cardiac arrested on the golf course. cardiopulmonary resuscitation was performed, and the patient was shocked twice in field, arrived at the emergency department and was intubated. electrocardiogram showed anterior st-segment elevation myocardial infarction. the patient was taken to the catheterization lab for percutaneous coronary intervention. post left anterior descending artery stent for 95% lesions, the diagnostic sheath was exchanged for 14fr x 13 cm peel away abiomed sheath. the pigtail catheter was advanced to left ventricle with 0.035j wire. the guidewire was exchanged for 0.018 abiomed wire, pigtail removed and impella advanced to left ventricle. cardiopulmonary resuscitation was being administered during this time. the pump was turned on but was only able get flows of 2l and was getting suction at performance level p-4. the patient regained blood pressure; on levophed at max dose. hung epinephrine, patient pressure increased to around 60 mean arterial pressure. troubleshooting low flows and suction, the position was checked under fluoroscopy, repositioned the device with no resolution of flow. echocardiogram was completed; images were challenging. however, the physician did not feel the right heart was an issue. there was a small effusion. the physician felt all options were exhausted and the mean arterial pressure dropped to 40. therefore, the physician decided to explant and replace the impella cp device due to low flows. the second impella cp was inserted. the pigtail catheter was advanced to the left ventricle, abiomed 0.018 wire advanced, the pigtail removed, and the new impella cp was advanced and placed. the pump was turned on and the same low pump flow issue occurred with flow. flows was unable to go above 1.5l, cpr was started due to mean arterial pressure dropping into the 30¿s. doctor went to speak with family. code protocol followed. the pump was turned on and off multiple times to see the underlying pressure. after speaking with family, the decision was made to turn on pump and the patient remained in pulseless electrical activity and expired. the treating physician commented that he suspected the patient might have had a pulmonary embolism and that the team was unable to get enough flow from the right heart. medical safety reviewed the event and noted there is not enough evidence, however, to exclude impella cp device pump 1 as an associated factor in the patient's outcome. there was low pump flow for reason that is unknown and delay in needed therapy initiated, resulted with impella cp device pump 1. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two device manufacturing reports associated with this event. refer to initial medical device report for the impella cp device (pump 2) serial number (b)(6) with date of event 19-feb-2025 and identifier ending in (b)(4).
10948 | 1220648-2025-26863 | 21658279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21658279&pc=OZD | 02/19/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/19/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501262 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with the second impella cp device for mechanical circulatory support experienced low pump flow as well. the patient cardiac arrested on the golf course. cardiopulmonary resuscitation was performed, and the patient was shocked twice in field, arrived at the emergency department and was intubated. electrocardiogram showed anterior st-segment elevation myocardial infarction. the patient was taken to the catheterization lab for percutaneous coronary intervention. post left anterior descending artery stent for 95% lesions, the diagnostic sheath was exchanged for 14fr x 13 cm peel away abiomed sheath. the pigtail catheter was advanced to left ventricle with 0.035j wire. the guidewire was exchanged for 0.018 abiomed wire, pigtail removed and impella advanced to left ventricle. cardiopulmonary resuscitation was being administered during this time. the pump was turned on but was only able get flows of 2l and was getting suction at performance level p-4. the patient regained blood pressure; on levophed at max dose. hung epinephrine, patient pressure increased to around 60 mean arterial pressure. troubleshooting low flows and suction, the position was checked under fluoroscopy, repositioned the device with no resolution of flow. echocardiogram was completed; images were challenging. however, the physician did not feel the right heart was an issue. there was a small effusion. the physician felt all options were exhausted and the mean arterial pressure dropped to 40. therefore, the physician decided to explant and replace the impella cp device due to low flows. the second impella cp was inserted. the pigtail catheter was advanced to the left ventricle, abiomed 0.018 wire advanced, the pigtail removed, and the new impella cp was advanced and placed. the pump was turned on and the same low pump flow issue occurred with flow. flows was unable to go above 1.5l, cpr was started due to mean arterial pressure dropping into the 30¿s. doctor went to speak with family. code protocol followed. the pump was turned on and off multiple times to see the underlying pressure. after speaking with family, the decision was made to turn on pump and the patient remained in pulseless electrical activity and expired. the treating physician commented that he suspected the patient might have had a pulmonary embolism and that the team was unable to get enough flow from the right heart. medical safety reviewed the event and noted there is not enough evidence, however, to exclude impella cp device pump 1 as an associated factor in the patient's outcome (demise) which is captured as a related report (noted in section h10 of this report). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two device manufacturing reports associated with this event. refer to initial medical device report for the impella cp device (pump 1) serial number (b)(6) with date of event (b)(6) 2025 and identifier ending in (b)(6).
10949 | 1220648-2025-26864 | 21658354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21658354&pc=OZD | 02/23/2025 | 03/20/2025 |  | 03/20/2025 |  |  | 02/23/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529870 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had oozing at the access site. angle matching was performed, leg immobilizer was in place. heparin was changed to bicarb in the purge. the oozing slowed and support continued. the patient also had hemolysis. it was further reported that the patient expired while on support however no other information was provided. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10950 | 1220648-2025-26865 | 21658488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21658488&pc=OZD | 02/23/2025 | 04/15/2025 |  | 03/20/2025 |  |  | 02/23/2025 | 03/20/2025 |  |  |  |  | 03/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589795 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Additional information was received that the impella 5.5 had purge pressure and purge system blocked alarm while in use with the patient. the patient has been on impella support since (b)(6) 2025 as no long-term advanced therapy was planned at this point. the purge cassette and tubing were changed without resolution. tissue plasminogen activator (tpa) was added to the purge. | The complainant reported that the patient on impella 5.5 support developed a hematoma at the access site noted upon arrival to the intensive care unit. sandbag was applied and blood products were administered. additionally, during initiation of gastrointestinal (gi) capsule study, the patient began showing signs cerebrovascular accident (cva). right sided deficits. the patient is off systemic anticoagulation for four days due to gi bleed. successful stent of left main coronary artery was done. neurological symptoms slowly were improving. it was further reported that the patient previous embolic cva from (b)(6) developed into a hemorrhagic cva evening of (b)(6). the patient remains off systemic anticoagulation. serial computed tomography scans will be done to monitor development. no intervention available at this time per neurology. the patient remains on support. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  | B.1 revised to reflect both adverse event and product problem due to additional information received from the clinical site. b.5 additional information was received from the clinical site. h.6 added due to medical device problem code additional information received from the clinical site. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26865 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of access site bleeding, stroke/cva, vascular damage - peripheral vessel, and high purge pressure has been completed. the root cause of the high purge pressure issue was most likely biomaterial per log analysis. as this clinical information was not provided, the true root cause of access site bleeding, stroke/cva, and vascular damage - peripheral vessel, could not be determined. b2 revised to reflect the patient outcome of death. b5 revised to reflect the patient outcome of death. d6b explant date added. h1 type of reportable event revised to reflect the patient outcome of death. h6 health effect - impact code 1802 death added.
10951 | 1220648-2025-26818 | 21641019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21641019&pc=OZD | 02/21/2025 | 04/08/2025 |  | 03/19/2025 |  |  | 02/21/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600780 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had been attempted to be implanted with an impella cp and during bleeding and a hematoma were noted at the impella site. femoral artery extravasation was noted under fluoroscopy. the decision was made to abort the procedure and heal the groin. the patient remained stable. | [5 empty] | The investigation of delivery issues and vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. as per the clinical information provided, at the time of insertion the pump did not pass through sheath due to patients anatomy. therefore, the root cause of the delivery issue was most likely patient condition related to patient anatomy. as per the clinical information provided, the delivery was canceled when the pump was unsuccessful to be advanced and caused bleeding during delivery. the clinical mentions that the insertion was cancelled due to patients anatomy issues. therefore, the root cause of the vascular damage issue was most likely patient condition related to patient anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10952 | 1220648-2025-26819 | 21641567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21641567&pc=OZD | 02/20/2025 | 04/08/2025 |  | 03/19/2025 |  |  | 02/20/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103168 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller (aic) would not recognize the purge disk to initiate purge. two different purge cassettes were attempted to no avail. the aic was replaced. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. h.6 investigation conclusions code 4315 was inadvertently selected on the initial manufacturer device report number 1220648-2025-26819. code 11 and 3233 should have been selectd as the device was received and the investigation was pending.
10953 | 1220648-2025-26820 | 21641718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21641718&pc=OZD | 02/24/2025 | 04/07/2025 |  | 03/19/2025 |  |  | 02/24/2025 | 03/19/2025 |  |  | 07/31/2025 |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392413 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp pump to support a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was placed, however, there were flow issues that prompted the team to troubleshoot with red blood cells infusion. the patient had a hemoglobin and hematocrit drop prior but no bleed was noted. the pump position was also an issue for this patient. the pump was shallow and repositioned only to be planned to be explanted due to the pump causing aortic insufficiency. no valve surgery or repair was noted and the plan was to explant but no other information was provided. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues, low pump flow, and heart valve damage has been completed. the complaint device was not returned for analyzing. data log was reviewed and no motor current spike was observed outside of p-level changes. no positioning issue alarms were seen. many suction alarms occurred on the first day of impella support and resolved by the end of the first day. clinical team said the patient's cvp was low. there were flow issues after insertion that prompted troubleshooting with red blood cell (rbc) infusion and the clinical team reported that the patient responded well to volume. the cause of the low pump flow issue most likely was patient condition related since data logs showed suction alarms improved after clinical mention of rbc infusion and that thepatient responded well to volume. clinical information stated: the pump was shallow and repositioned only to be explanted due to the pump causing aortic insufficiency. explant happened after 32 hours. no valve surgery or repair was noted. limited clinical details on how it became shallow were given. the cause of the positioning issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. not enough clinical data provided regarding heart valve damage issue. the cause of the heart valve damage issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. d.6b explant date was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26820.
10954 | 1220648-2025-26821 | 21642942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21642942&pc=OZD | 02/20/2025 | 03/19/2025 |  | 03/19/2025 |  |  | 02/20/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569308 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 device insertion to patient with postcardiotomy cardiogenic shock for mechanical circulatory support was unsuccessful. due to the patient's complex congenital anatomy, the impella 5.5 pump was unable to advance past the vessel on the right axillary site during attempted delivery. the surgeon attempted insertion on the left axillary side, however, again, the vessel anatomy was too small and the impella 5.5 device could not be inserted. the implant was aborted, and plans were made to place an impella cp pump for patient support. there was no resulting patient harm, and the patient was noted to be stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance." impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10955 | 1220648-2025-26822 | 21643764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21643764&pc=OZD | 02/17/2025 | 03/19/2025 |  | 03/19/2025 |  |  | 02/17/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 96 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team was forwarded a medwatch report stating, "patient experienced a significant pull of the impella 5.5 driveline causing a separation of the power wires within the driveline and subsequent device failure. patient was emergently taken to the operating room for pump exchange. concern is that the device is not manufactured to be as durable externally as anticipated and what should be expected compared to similar devices with external leads. due to the nature of the event and subsequent finding, the exact location of the break within iha driveline may point to a specific area of concern with current manufacturing process and durability testing." there was no other information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10956 | 1220648-2025-26823 | 21643768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21643768&pc=OZD | 02/20/2025 | 03/19/2025 |  | 03/19/2025 |  |  | 02/20/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605359 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support (mcs) experienced bleeding with heparin discontinuation and limb ischemia. however, the patient would subsequently expire. the patient had a positive electrocardiogram for inferior lateral st-segment elevation myocardial infarction. and was electively intubated before being taken to the catheterization lab. the left dominant circumflex proximally occluded and there was severe mitral regurgitation noted on ventriculography. the decision was made to place patient on impella cp mcs and consult cardiothoracic surgery for surgical intervention. the impella was successfully implanted, and the patient was transported to the unit. three days later the patient underwent mitral valve replacement and coronary artery bypass grafting surgery. the patient was at this time place on venoarterial extracorporeal membrane oxygenation (va ecmo). now day six of impella mcs the patient was started on continuous renal replacement therapy (crrt) with goal removal. additionally, it was noted the patient developed a gastrointestinal bleed and the purge solution was switched to sodium bicarbonate. three units of platelets were administered. two days later the patient was administered two units of packed red blood cells for low hemoglobin. per the report it was stated the staff does not feel it is impella related (however unconfirmed). the patient was having persistent melena and also now net negative with crrt removal, unable to wean flows on ecmo. the following day, the right leg had loss of pulses. a plan was noted to remove the impella cp with hopes to improve flow. persistent low hemoglobin noted with units transfused was noted. however, the patient would expire later in the day; care was withdrawn. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10957 | 1220648-2025-26824 | 21643797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21643797&pc=OZD | 02/21/2025 | 03/19/2025 |  | 03/19/2025 |  |  | 02/21/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612297 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and after the impella removal, patient moved arm breaking out of restraint while peel away sheath was still placed. bleeding was noticed, angio showed axillary dissection, covered stent placed and was soon noted dissection proximal and distally from covered stent, additional self expanding stents placed proximal and distal with great results. upon angio no extravasation noted post stents. blood transfusions were given to the patient and the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10958 | 1220648-2025-26826 | 21643910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21643910&pc=OZD | 01/17/2025 | 03/19/2025 |  | 03/19/2025 |  |  | 02/20/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594160 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding the patient who experienced femoral artery occlusion with a "possible" relationship to the impella device used. during the catheterization procedure, the team was unable to follow the pulses in the left lower extremity and completed an angiogram, which found a left superficial femoral artery occlusion. vascular surgery was consulted. cardiac and vascular surgery teams confirmed the impella cp device should be explanted, which was completed after a total of approximately two days of support. the following day, vascular surgery confirmed pulses on to the left foot via posterior tibial artery and right foot pulses. the patient was restarted on apixaban three days later. the status outcome of the patient was unknown; however, per the report the patient survived the explant was treated medically for clot prevention. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10959 | 1220648-2025-26825 | 21644026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21644026&pc=OZD | 02/06/2025 | 03/19/2025 |  | 03/19/2025 |  |  | 02/21/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594763 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Insufficiency | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured respiratory insufficiency post operative with a "unknown (undetermined)" relationship to the impella device used: ¿(b)(6) 2025 post-op respiratory insufficiency requiring heated high flow oxygen, multiple doses of lasix to diurese, prone positioning, and aggressive pulmonary toilet. (b)(6) 2025 transitioned to nasal canula." the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10960 | 1220648-2025-26827 | 21644847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21644847&pc=OZD | 01/26/2025 | 03/19/2025 |  | 03/19/2025 |  |  | 02/21/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592338 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient experienced a gastrointestinal bleed (gi) and expired on impella 5.5 support. both the gi bleed and death were believed to be possibly related to the use of the impella. no additional details were provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10961 | 1220648-2025-26828 | 21645962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21645962&pc=OZD | 02/19/2025 | 04/22/2025 |  | 03/19/2025 |  |  | 02/21/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603142 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced access site bleeding requiring blood transfusion. the patient was initially hospitalized with influenza, chronic depressed left ventricular ejection fraction. an impella cp device was implanted via the right femoral artery. approximately 15 days after start of the support, "increased" bleeding from the access site was noted requiring seven units of blood transfusion. no surgical intervention was required; the insertion site was put under pressure with a sandbag. the patient was noted to be in stable condition following the event. three days later it was noted the impella cp was surgically removed, and the patient was escalated in care to an impella 5.5 for continued mcs. | [5 empty] | The investigation of access site bleeding has been completed. clinical reports bleeding from access site on (b)(6) 2025, 15 days into the case. the patient received 7 units and the insertion site was put under pressure with a sand bag. the pump was returned and no abnormalities or damage was observed. the root cause of the access site bleeding was not determined due to insufficient clinical detail. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
10962 | 1220648-2025-26829 | 21647244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21647244&pc=OZD | 02/22/2025 | 04/01/2025 |  | 03/19/2025 |  |  | 02/22/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597155 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there were high purge pressures. tissue plasma activator was added to the purge which improved the pressure and support continued. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump was returned for investigation. no physical abnormalities were observed on the returned product. the cannula, impeller, and purge gap were clean. the pump was successfully primed, and purge pressure was within the specification range during testing. the data log shows that the motor current spiked slightly before the purge pressure gradually reached 1100 in 1.5 minutes, triggering a high purge pressure alarm. this trend returned to normal after 2 hours, and clinical observations suggest that tissue plasma activator (tpa) helped restore normal limits. ¿the cause of the high purge pressure issue was most likely biomaterial, based on data log trend and returned product investigation. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26829 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10963 | 1220648-2025-26830 | 21647311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21647311&pc=OZD | 02/22/2025 | 04/11/2025 |  | 03/19/2025 |  |  | 02/22/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607844 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 had suction alarms that did not resolved with repositioning. on transesophageal echocardiogram 3d image noticed debris on inlet of impella. the impella was exchanged and the patient remained stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of low pump flow and thrombus has been completed. the data logs confirmed the failure mode & clinical narrative. logs showed pump started and ran without issue for about 516 hours, flow suddenly dropped that triggered auto suction response & suction alarms. this event remained to the rest of the case. pump was returned in damaged condition (missing pump plug). visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of the low flow was biomaterial ingestion. as the necessary information was not provided, the root cause of the thrombus was not determined e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26830 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10964 | 1220648-2025-26831 | 21647349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21647349&pc=OZD | 02/24/2025 | 05/16/2025 |  | 03/19/2025 |  |  | 02/24/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025598492 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the pump had dislocated, and since the pigtail was already positioned above the aortic valve, he did not want to advance pump any further into the left ventricle. furthermore, the patient did not require as much hemodynamic support, so the team decided to completely withdraw the pump. unfortunately, no pump was returned for complaint. thus, the case was closed without any negative outcome or any harm to patient. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. clinical reported successfully implanted the pump. after transporting the patient from the catheterization lab to the icu, the pump was positioned in the aorta. position was verified by echo. repositioning was not attempted and pump was explanted since the patient did not require as much hemodynamic support and the pigtail was already positioned above the aortic valve. the root cause of the positioning issue was most likely patient movement since clinical reported the pump was positioned in the aorta after transport. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26831 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source; foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26831. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
10965 | 1220648-2025-26832 | 21649139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21649139&pc=OZD | 02/21/2025 | 03/19/2025 |  | 03/19/2025 |  |  | 02/21/2025 | 03/19/2025 |  |  |  |  | 03/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604708 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Contamination | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and experienced positioning issues. the patient presented to the emergency department with worsening shortness of breath and was sent for a right heart catheterization (rhc) due to pressor requirements. the patient became combative and severely hypoxic requiring intubation. the rhc revealed significant congestion. an impella cp was placed in left femoral artery independently achieving flows of 3.5 liters. the patient still looked poor with severe mitral regurgitation. without ability to cannulate the right femoral artery, the decision was made with vascular surgery to transition to an impella 5.5 device via the left axillary artery which was successfully completed. the impella cp was removed to cannulate the left femoral artery and the right femoral vein. three days later, all puncture sites were noted to be bleeding (action to treat was not noted/is unknown). however, hemoglobin and hematocrit were stable and holding. the patient was thereafter noted being transferred to an outside facility for higher level of care. now four days later, day seven of impella 5.5 support, the patient's delirium had increased with thrashing in bed despite nurses holding the patient down. the patient ripped the impella 5.5 repositioning unit from the suture hub and the catheter became contaminated with impella position in aorta alarms. the impella 5.5 support level was decreased to p-2, and the patient was transferred to the operating room for emergent impella 5.5 removal. it was noted, thereafter, overall, the patient remained hemodynamically stable. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10966 | 1220648-2025-26785 | 21630289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21630289&pc=OZD | 02/19/2025 | 03/19/2025 |  | 03/18/2025 |  |  | 02/19/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603955 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support experienced an embolic stroke. prior to the stroke diagnosis, impella in aorta and impella in position unknown alarms occurred. an echocardiogram was obtained which demonstrated correct impella positioning. a spike in motor current was noticed before the motor current dampened, and then another spike in motor current was noticed before returning to normal. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke has been completed. the root cause of the stroke could not be determined since the product was not returned and insufficient clinical details were provided. e.1 revised facility name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26785. h.10 investigation results were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26785.
10967 | 1220648-2025-26786 | 21630414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21630414&pc=OZD | 02/19/2025 | 05/01/2025 |  | 03/18/2025 |  |  | 02/19/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592640 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Also, the patient experienced torsades, the patient was shocked 64 times. the patient continued to have tachycardia during support. | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient experienced low pump flow, access site bleeding and hemolysis and subsequently expired. five days prior to the impella implant, the patient underwent three-vessel coronary artery disease complicated by cardiac arrest requiring venoarterial extracorporeal membrane oxygenation (va ecmo). the impella cp was emergently placed successfully. thereafter, the patient was transported to the unit on ecmo and impella support. two days later it was noted the patient underwent percutaneous coronary artery intervention of drug-eluting stent placement to the left anterior descending (lad) artery. the next day, low pump flow was noted. the impella cp pump was unable to go above performance level p-1 without suction. the physicians noted blood in chest creating heart constriction. chest tube was placed at bedside. echocardiogram showed little filling in left ventricle and right ventricle. right lung whited out on x-ray. ecmo was turned down to 3.0lpm due to chugging the night prior. the patient would receive four units of packed red blood cells and two units of fresh frozen plasma (unknown if for the chest bleed or low pump volume). urine output was around 30ml/hr. with pink tinge. five days later bleeding at the access site was noted requiring a dressing change. there was uncertainty if ecmo cannula or impella was the site of bleeding. the physician placed a suture at the impella access site and ecmo cannulas. the next day patient was experienced hemolysis. the patient being weaned off of va ecmo but was unable to. the patient was accepted to an outside facility for ventricular assist device evaluation. however, the patient's family member decided against this transition, opted for comfort care, and the patient would expire later this same day. |  |  |  |  | The investigation of access site bleeding, hemolysis, low pump flow, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. data log reviewed: many suctions during support occurred. no positioning issue seen. many impella position unknown observed. on (b)(6) 2019 a dip in placement signal (ps) and concurrent drop in pump flow was observed. echocardiogram shows little filling in left ventricle and right ventricle on (b)(6) 2019. soon after the dips, a motor current spike is observed. clinical stated: on (b)(6) 2025 bleeding from every tube including extracorporeal membrane oxygenation (ecmo) cannulas, impella, feeding tube etc. and blood given with adding suture of support. the cause of the access site bleeding - major most likely was patient condition related since bleeding observed from multisite of support. the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. the cause of the low pump flow was most likely patient condition related since echocardiogram shows little filling in left ventricle and right ventricle at the time of the issue. as the necessary information was not provided, the root cause of the vt/vf was not determined. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26786. h6 health effect - clinical code 2095 and health effect - impact code 4647 were inadvertently omitted on the initial manufacturer device report number. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10968 | 1220648-2025-26787 | 21630465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21630465&pc=OZD | 02/19/2025 | 04/07/2025 |  | 03/18/2025 |  |  | 02/19/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595529 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced low pump flow. the patient was supported for six days with an intra-aortic balloon pump and began to have ventricular tachycardia events. the decision was made to place an impella cp device which was completed. upon advancing the repositioning sheath, the implanted impella cp pump began to experience persistent suction. both the volume and positioning were verified at the time of the event. upon removing and reinserting the pump, a kinked cannula was noted on angiogram. the pump was subsequently replaced as a result of the low pump flow and kinked cannula. there was no resulting patient harm. | [5 empty] | The investigation has been completed. the clinical information stated upon advancing the reposition sheath, constant suction occurred. volume was assessed and the position was verified with echocardiogram. the pump was removed, and re-inserted, kink and cannula evidenced on angiogram. the returned complaint pump was visually inspected. cannula kink was observed near the outlet. no biomaterial or broken blade was observed. no other abnormality were found. data log was reviewed and no motor current spike was observed outside p-level change. a drop in flow was observed from case start followed by suction alarms. no positioning issue was observed. the cause of the low pump flow issue was due to cannula kink. lack of clinical details on the delivery of the pump or any patient conditions/anatomy. the cause of the product damage (cannula kink) was not determined due to lack of clinical details. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10969 | 1220648-2025-26788 | 21630599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21630599&pc=OZD | 02/20/2025 | 03/25/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1850881 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support ongoing when on day 10 of the support there was an automated impella controller (aic) malfunction. the aic controller failure alarm was observed by the perfusionist and the console was removed/replaced. the support continued without any noted harm or injury from interruption. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console log confirmed the reported failure mode. the console was tested and after power cycling the console 200 times, it was unable to reproduce the reported failure mode - controller error. the cause of the issue was not determined due to the inability to reproduce on the returned aic d4 unique identifier (udi) # was erroneously entered on the initial manufacturer device report. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10970 | 1220648-2025-26789 | 21630675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21630675&pc=OZD | 02/19/2025 | 04/07/2025 |  | 03/18/2025 |  |  | 02/19/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612246 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced low pump flow followed by a clot. heart catheterization showed extensive multi-vessel disease. after consulting with the cardiothoracic surgery, the decision was made to place an impella cp. the impella cp device was successfully implanted. there were "excellent" waveforms and flows of 3.7 liters per minute on support level p-9 were obtained. the peel away sheath was removed and repositioning sheath was inserted. an angiogram was performed thru the side port that showed good flow distally. forward tension sutures and an angle match dressing were then applied. the patient was transferred to the unit on impella mcs. two days after start of the support, there was low flow and suction thought to be due to biomaterial ingestion. the impella cp device was explanted, and at the time of explant, the team observed a clot in the femoral angiogram. penumbra thrombectomy was considered. however, the decision was made to start integrilin drip due to the thrombus not being flow limiting. the patient was noted to be doing well after the event. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump produced saturated flows (4.1 liters per minute at p9) during bench test. the root cause of low flow was unable to be determined as limited clinical details were provided and the low pump flow was unable to be reproduced. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10971 | 1220648-2025-26790 | 21630850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21630850&pc=OZD | 11/09/2024 | 03/18/2025 |  | 03/18/2025 |  |  | 02/19/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563732 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Cardiac Arrest | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed's long-term outcome and quality indicated impella registry and noted a patient experienced atrial fibrillation with rapid ventricular response. the relationship to impella cp device is listed as "possibly related." the patient would ultimately expire. the patient was alternated between normal sinus rhythm and atrial fibrillation during index stay prior to impella cp device procedure. in the afternoon, post impella cp device placement, the patient went into atrial fibrillation with ventricular rates of 110-130 beats per minute. the patient was cardioverted twice, amiodarone intravenous bolus was given and amiodarone drip was initiated. the patient remained in atrial fibrillation until death. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10972 | 1220648-2025-26791 | 21631496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21631496&pc=OZD | 01/09/2025 | 04/24/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164186 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced placement signal difficulties. while investigating the impella cp complaint, it was noted there was an issue also with the automated impella controller, and a request for the return of the product from the user facility was made. | [5 empty] | The investigation of automated impella controller (aic) - loss of opm data has been completed. during the device and data analysis was unable to reproduce the reported issue. the root cause was unable to be determined. | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-25755 for the report on the impella pump.
10973 | 1220648-2025-26792 | 21631666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21631666&pc=OZD | 02/18/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/18/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612786 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation | Pericardial Effusion | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced left ventricle perforation followed by limb ischemia both requiring intervention. the impella cp device was successfully place and following, the patient went into a violent coughing and vomiting spree causing the impella pump to dive deep into the left ventricle. shortly after repositioning the impella cp device, the impella wave forms diminished, and all went flat along with the pressure and flows. an echocardiogram was urgently performed, and a large and growing pericardial effusion was noted. at this time, a pericardiocentesis was performed. all pressures and waveforms returned to normal but quickly diminished each time blood was drawn from the pericardium. cardiovascular surgery was called in and the patient's chest was opened on the table and an open left ventricle perforation repair was performed. the patient was stabilized and sent to the unit on impella mcs. the following day in the morning, it was noted pulses were diminished and no doppler pulses were noted in the impella leg. the decision was made to explant the impella cp device and institute venoarterial extracorporeal membrane oxygenation (va ecmo) as the patient was still on levophed epinephrine and vasopressin. the patient's status after the event was indicated as unknown. however, it was noted that the patient survived the explant with transition to va ecmo. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10974 | 1220648-2025-26793 | 21631690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21631690&pc=OZD | 02/19/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/19/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594139 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced bleeding and arrhythmias. the patient would subsequently expired. the impella was placed, and percutaneous coronary intervention was completed to the left main artery. the patient remained hemodynamically stable throughout case and thereafter was sent to the unit on support. overnight, there was some ectopy noted, and amiodarone was started. additionally, it was noted there was bleeding from the contralateral sheath site. the sheath was removed, heparin was withheld, and the physician monitored the situation. the following day the patient was free of arrhythmias through the night and amiodarone drip was discontinued. the patient was following some commands and distal pulses were present on doppler. following day, the patient paced at 80bpm, and nurse reports asystole underling rhythm. the patient remained intubated and on vasoactive medications. heparin, restarted, was withheld the prior night due to blood-tinged sputum. distal pulses present on doppler. now day 7 on support the patient was in atrial fibrillation with rapid ventricular response. the following day the patient expired; care was withdrawn. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10975 | 1220648-2025-26794 | 21631699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21631699&pc=OZD | 02/19/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/19/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573404 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support and the following day experienced bleeding requiring surgical intervention. the patient was identified as high risk, came off bypass onto an impella 5.5 device. following the impella 5.5 implant, the patient had residual surgical bleeding of the access site. the patient was taken back to surgery for a washout and the bleeding subsided. two units of packed red blood cells were given, and the patient was reported to be stable following the event. support continued uninterrupted for an additional six days. the patient was successfully weaned, and the pump removed at bedside without difficulty per the plan of care. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10976 | 1220648-2025-26795 | 21631708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21631708&pc=OZD | 01/06/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592752 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Contamination | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device and upon removing the purge cassette from its packaging the leur lock that attaches the cassette to the spike was not screwed on and the spike fell out of the packaging into the sterile field. it was stated the staff was able to use the purge cassette; a new cassette was not obtained. the patient underwent the pci and the impella was successfully weaned and explanted the same day. there was no report of patient harm as a result of this event. after investigation of this complaint, it was found that this issue compromised sterility of the product. | [5 empty] | The impella device was not received from the customer. however, the investigation has been completed and noted the following: the product was not returned for investigation, and a data log review was not required for this case run. according to the clinical details, the purge cassette spike was not attached to the luer lock inside the packing, causing the spike to fall out onto the sterile field. the doctor was able to use the purge cassette. the reported purge cassette failure mode was changed to packaging issues - sterility since clinical details confirm that the poor connection between the luer and spike caused the spike to fall off and could impact the sterility of the product. no packaging/product image was provided. the cause of the packaging issue was not determined since product was not returned.
10977 | 1220648-2025-26796 | 21631733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21631733&pc=OZD | 02/20/2025 | 04/07/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612239 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, there were air bubbles in the patient's left ventricle (lv). the bubbles were noted during sternal closure. p levels were reduced and the lv was dilated. additionally, there was condensation/fluid inside the purge side arm. the purge cassette was changed and support was continued. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of embolism and purge leak - pump has been completed. the data logs were reviewed and there were no drops in purge pressure out of specification noted. there is only 1 purge pressure low alarm, and it was on (b)(6) and resolves after 5 seconds. the purge cassette change referenced in the clinical details was seen on (b)(6). at this time there was a slight drop in the purge pressure but never lower than ~500 mmhg. at this time data logs also show an increase of heparin in the purge fluid. product was returned with a cut in the catheter and a cut in the purge cassette line. to be able to run purge though what was returned, a cut was made in the purge cassette line just distal to the pressure reservoir, and a syringe was connected. the entire system was flushed with colored purge fluid and left two hours to see if any leaks would reproduce. there were no leaks of the colored fluid noted, particularly in the sidearm, after this test. embolism: clinical details report that air was noted in the left ventricle (lv). the patient had an existing patent foramen ovale (pfo), and also had intravenous (iv) and centrally placed venoarterial extracorporeal membrane oxygenation (va ecmo). the complication resolved the following day, but it is unclear which troubleshooting method was successful. data logs were reviewed and there were no signs of a purge leak noted during the case. returned product did not leak, nor were there any cracks noted in the filter. clinical details report that the patient was put on vasopressors and filters were added to iv lines, and the complication resolved the following day. however, it is not stated that these methods were the fix. since clinical details provided are limited and returned product did not leak from the filter, the root cause of the embolism could not be determined. purge leak: clinical details report that condensation/fluid was noted in the purge sidearm. no troubleshooting methods were reported, and the sidearm was stable/unchanged the following day. data logs do not show any signs of a purge cassette leak during the case. returned product did not reproduce a leak. since the full pump and purge cassette were not returned for investigation, the root cause of the purge leak could not be determined. d9 date device returned to manufacturer added.
10978 | 1220648-2025-26797 | 21631824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21631824&pc=OZD | 02/20/2025 | 04/02/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Field service reported that during preventive maintenance, the console (aic) was noted with corrosion in pbm. there was no patient involvement. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation for the console (aic) system self check error has been completed. console logs from the reported day of event are consistent with complaint and show persistent communication errors of sau_powermod_1. there were no such issues the last time console was used in the field, prior to its pm. inspection of the pbm revealed corrosion on both sides of the board due to electrolyte leak from supercaps c69, c70, which has impacted a neighboring rs232 communication channel. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
10979 | 1220648-2025-26798 | 21631832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21631832&pc=OZD | 02/20/2025 | 04/08/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025590022 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support, and four days after start of the support an echocardiogram was ordered to check position as the patient experienced frequent ventricular tachycardia (vt) runs. however, this was non-sustained. the electrophysiologist was consulted and noted the vt runs may be impella related. the device was repositioned and the following day it was noted there were no issues since repositioning. the patient continued on support and was eventually bridged to a left ventricular assist device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10980 | 1220648-2025-26799 | 21632054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21632054&pc=OZD | 02/20/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503022 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that upon removing the peel-away sheath following impella 5.5 implant, an impella in aorta alarm appeared. it was verified on transesophageal echocardiogram and fluoro that the pump had been inadvertently pulled back across the aortic valve. multiple attempt to re-advance the pump were unsuccessful and the sterile sleeve subsequently tore. the pump was removed and a new impella 5.5 device was placed for ongoing support. there was no resulting patient harm. | [5 empty] | The investigation has been completed. the returned pump was inspected and the anticontamination sleeve was confirmed to be separated from housing. clinical mentioned that on securing device, impella pulled back into aorta. multiple attempts to advance back across the av. unsuccessful. the root cause of the positioning issue was determined to be use related as evidenced by clinical details. the root cause of the product damage (sterile sleeve damage) was most likely due to a use-related issue causing separation.
10981 | 1220648-2025-26800 | 21632127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21632127&pc=OZD | 02/18/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/18/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611466 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a polysubstance abuse patient who developed vegetative endocarditis was taken to the operating room (or) for mitral valve repair which turned into a mitral and aortic valve replacements. the patient came to the unit on extracorporeal membrane oxygenation and intra-aortic balloon pump. an echocardiogram showed valves not opening and the decision was made for impella 5.5 placement. transesophageal echocardiogram in the or showed a left atrial clot and more clots possible on/around the mitral valve. the decision was made to reopen the chest and place back on bypass. the impella 5.5 was opened and primed in preparation for placement before coming off bypass. however, the heart was not beating. with contraction aortic valve not opening, the impella 5.5 was not able to be placed. therefore, there was no harm to the patient caused by the impella. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10982 | 1220648-2025-26801 | 21632383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21632383&pc=OZD | 02/20/2025 | 04/07/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576526 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 pump to support cardiogenic shock/cardiomyopathy. the pump was placed and the access site was intervened upon. extra sutures, thrombin and gelfoam were utilized to achieve hemostasis. the pump was supporting when after six days there was purge blocked/high purge pressure alarms. tissue plasminogen activator was started. the pump remains on despite these issues. | [5 empty] | The impella pump, returned, purge cassette and data stick were returned and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure and access site bleeding has been completed. information was reported having slight resistance at anastomosis site during implantation. the pump was placed and run at p-2 and p-5 with no issues. bleeding was controlled with more repair sutures, thrombin, gelfoam and evarrest. no blood products were given. the root cause of the access site bleeding - major was not determined due to insufficient clinical details. information reported that tissue plasminogen activator (tpa) was starting to resolve high purge pressure and purge blocked alarms and continued running tpa into the next day. biomaterial was found on around the inflow cage of the returned pump. data logs showed high purge pressure alarms that started and resolved the next day. no motor current spikes but motor current and flow starts trending upwards during high purge pressure alarms and trends back down when alarms resolve.¿ the root cause of the high purge pressure was most likely biomaterial buildup since tpa resolved the high purge pressure and there were no motor current spikes.
10983 | 1220648-2025-26802 | 21632454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21632454&pc=OZD | 02/19/2025 | 04/22/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485514 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was attempted to be implanted with an impella 5.5 pump for support of myocarditis. the patient's native anatomy did not allow for the impella 5.5 to be delivered. the team removed and placed an impella cp pump instead. no harm or injury was noted to the patient. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. clinical reported being unable to deliver the 5.5 as the axillary artery was too small and the procedure was aborted. during the removal of the 5.5, it was broken from pulling on it. it was replaced with a cp and there were no delivery issues with the cp. the pump was returned and the catheter was kinked and stretched beginning at the motor housing and the cannula was deformed, potentially from removal as stated in clinical details. the root cause of the delivery issue was most likely patient condition related since clinical reported the axillary artery was too small and the replacement cp was delivered with no issues. the root cause of the product damage was most likely a cannula kink due to trying to deliver a pump through a small axillary artery. the root cause of the product damage was most likely a catheter kink due to attempting to remove the pump through a small axillary artery and facing resistance.
10984 | 1220648-2025-26803 | 21633099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21633099&pc=OZD | 02/20/2025 | 04/22/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  | 02/28/2025 |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310509 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp and was escalated to an impella 5.5. while on the impella 5.5 the motor flat initially when started. pulse pressure was 5 when impella first turned on. motor remained flat as pulse pressure widened to 20 however later on unit motor still flat. unable to flow above p4 even with reduction on extracorporeal membrane oxygenation flows. the patient remained stable and the plan was to remove the pump however no other information provided. | [5 empty] | The investigation of low pump flow has been completed. the returned product was inspected and there were no physical abnormalities. there was dried dextrose around the impeller. the pump was run in the lab after soaking and low pump flow was not reproduced. the data logs show continuous suction alarms while the pump was on p4. there were also continuous 'impella position unknown" alarms. the clinical details mention patient was on as on 5lpm flow venoarterial extracorporeal membrane oxygenation (va ecmo) support and transesophageal echocardiogram (tee) confirmed inlet deviated toward apex and 5.1cm from aortic valve annulus. the logs indicate that pump was unable to flow higher than p4. the root cause of the low pump flow was most likely due to patient condition as evidenced by continuous suction and ecmo support. d9 date device returned to manufacturer added. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26803. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26803 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10985 | 1220648-2025-26804 | 21633648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21633648&pc=OZD | 02/20/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589788 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Pump | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on impella cp for acute myocardial infarction. it was noted that the patient had diminished bilateral pedal pulses. the impella side was the worse of the two. vascular surgery was following, however, there was no plan to remove the impella or surgical intervention planned. additionally, the unit nurse stated that the pump stopped spontaneous. attempts were made to restart. however, they were unsuccessful. the automated impella controller was swap out and was also unsuccessful. they noticed a motor current spike on connect. ultimately, the patient coded and eventually was made do not resuscitate and expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10986 | 1220648-2025-26805 | 21633759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21633759&pc=PYX | 02/20/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  | 02/28/2026 |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478737 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella rp flex support due to right heart failure. during support, the placement signal disappeared and flows represented zero. the pump was explanted. there was no patient harm reported. | [5 empty] | The investigation has been completed. the returned product was inspected and there were no physical abnormalities. the pump passed electrical testing and there were no sensor abnormalities. the data logs from the case showed that the pump was running on p6 with normal placement signal and flow readings until 12-feb-2025. the placement signal started drifting down and flow started drifting up until placement signal went out of range and flow showed 0 milliliters per hours. cvp was also out of range. the alarm dp signal out of range was triggered. this lasted for about three hours before placement signal recovered. upon recovery drifting resumed with placement signal drifting up and flows drifting down. the returned pump was run in the flow loop test and sensor drift was reproduced with placement signal trending up and no pump flow readings while cvp was negative. the root cause of the placement signal issue was determined to be due to a sensor drift as issue was reproduced on the returned product.
10987 | 1220648-2025-26806 | 21633817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21633817&pc=OZD | 02/21/2025 | 05/06/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614203 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that roughly one day into impella 5.5 support, the purge flows began to drop while the purge pressure rapidly increased. tissue plasminogen activator was added to the purge line and the flow rate and pressure began to normalize. there was no patient harm. | [5 empty] | The investigation of high purge pressure (hpp) has been completed. product wasn't returned for investigation. data log review confirmed the high purge pressure event. on 2/21, the purge pressure trended up gradually over the course of 1.25 hours until the first alarm triggered. there were no motor current spikes leading up to the event. purge pressure remained elevated for 7 hours before it resolved for the rest of the case. during the hpp, purge flow data shows sodium bicarb being removed from the purge fluid and then reintroduced after it had resolved. clinical details report there were no kinks noted in the purge line and that tissue plasminogen activator (tpa) was added to the purge. based on the gradual purge pressure rise, tpa being added to the purge, and the complication resolving, the root cause of the high purge pressure was determined to most likely be biomaterial buildup. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10988 | 1220648-2025-26807 | 21633949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21633949&pc=OZD | 02/21/2025 | 04/08/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609848 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | High Blood Pressure/ Hypertension | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was placed, a percutaneous coronary intervention (pci) procedure was started. during the pci procedure while utilizing shockwave on the left main, the automated impella controller placement signal disappeared with a placement signal not reliable alarm. the patient¿s blood pressure and flows were affected for about 30 seconds and then they both recovered simultaneously. the impella cp functioned without issue for the remainder of the case, but the placement signal never resolved. after completion of the case, impella cp support was rapidly weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue has been completed. the pump was returned and evaluated. the pump passed the laser continuity test, but its 5s parameters indicated a hard failure of the sensor with signal not reliable (snr) at 0 and the cavity length maxed out above 32000. the data logs show the placement signal (ps) signal go out of range about 2 hours into the case with snr dropping to 0, gain increasing, contrast decreasing, and lamp increasing. the clinical details mention the ps goes out when using shockwave; however, the distance between shockwave and the sensor is unknown. the distance between shockwave and the sensor should be greater than 20 mm to avoid interfering with and/or damaging the optical sensor per the instructions for use. the root cause of the optical signal issue was most likely a damaged sensor from shockwave interaction. d9 date device returned to manufacturer added.
10989 | 1220648-2025-26808 | 21634156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21634156&pc=OZD | 02/21/2025 | 04/02/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610835 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp to support acute myocardial infarction. the impella cp remained in placed during the percutaneous coronary intervention (pci) but the access site continued to bleed. the patient received multiple units of blood for the access site bleed. after 62 hours, the pump was explanted. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient condition related, bled from both femoral sites overnight with large hematoma at non impella site.
10990 | 1220648-2025-26809 | 21634206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21634206&pc=OZD | 02/20/2025 | 03/24/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611468 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65 year old male patient with an impella 5.5 pump began to experience saturated flow. the pump was removed and a new impella 5.5 device was implanted for ongoing support. there was no resulting patient harm. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed, a supplemental mdr will be filed. | The investigation of saturated flow has been completed. the product was returned and evaluated. there was a small amount of potential biomaterial with a clear appearance on the purge gap. no other abnormalities were found during inspection. the pump was run for 20 minutes and saturated flow was reproduced. the motor was torn down and there was no evidence of biomaterial within the motor housing. the data logs showed a large motor current spike followed by saturated pump flow and a subsequent rise in purge pressure. the root cause of the saturated flow was most likely biomaterial since data logs indicate biomaterial ingestion and saturated flow was reproduced with the returned product. g.1 added reporting contact facility name as it was inadvertently omitted on initial manufacturer device report number 1220648-2025-26809.
10991 | 1220648-2025-26810 | 21634441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21634441&pc=OZD | 02/20/2025 | 04/21/2025 |  | 03/18/2025 |  |  | 02/20/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607195 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 56-year-old male patient on an impella 5.5 for cardiomyopathy. it was reported that the patient developed high purge pressure/purge blocked alarm shortly after placement. tissue plasminogen activator (tpa) protocol was started but the pump did not improve. additionally, through impella connect it was noted that the pump stopped twice and restarted. the 5.5 was removed. the patient survived the procedure. | [5 empty] | The investigation of pump stop, high purge pressure, and saturated flow has been completed. the pump was not returned for investigation. the data log indicates a temporary pump stop accompanied by a motor current spike. pump was able to restart in two attempts. following this, the purge pressure increased within six minutes, leading to a high purge pressure (hpp) alarm. additionally, the pump flow became saturated after the pump stop event. the cause of the temporary pump stop issue was most likely biomaterial, based on data log review. the cause of the high purge pressure issue was most likely biomaterial, based on data log review. the cause of the saturated pump flow issue was most likely biomaterial, based on data log review. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26810 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26810. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
10992 | 1220648-2025-26811 | 21634634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21634634&pc=OZD | 02/21/2025 | 05/15/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1850911 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic | Black Or African American | Malfunction | No | Yes | Break | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62 year old female patient with an impella cp. during support, the automated impella controller (aic) displayed intermittent controller error alarms during support. the aic was replaced in response to the failure. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed upon return of the aic, it was noted that the console was returned in an unprotected box and was subsequently the rear housing was broken. aic - damage before therapy was added as a failure mode (fm). the data analysis confirmed the reported failure mode. this console was tested and upon running the optical test pump for 3 hours and 30 minutes, unable to reproduce the reported fm ¿ controller error or the communication issues with any of the subcomponents. the root cause of the controller error was not determined due to the inability to reproduce it. the root cause of the broken rear housing was most likely a use issue. d9 date device returned to manufacturer added. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10993 | 1220648-2025-26812 | 21634686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21634686&pc=OZD | 02/21/2025 | 04/29/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607018 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was advanced into the left ventricle, but the surgical team decided to abort due to the patient having a small left axillary vasculature. intra-aortic balloon pump was then placed via left fem artery but was not enough support per surgical team. decision was made to cannulate for veno-arterial extracorporeal membrane oxygenation due to increasing inotrope/pressor requirements. the right axillary impella was then explanted after cannulation with no issues. there was no patient harm reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to small left axillary vasculature and the insertion was aborted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10994 | 1220648-2025-26813 | 21635432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21635432&pc=OZD | 02/21/2025 | 03/18/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582270 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Detachment of Device or Device Component | [5 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp pump to support a patient admitted in with multiple cardiac comorbidities and suffering from acute myocardial infarction and cardiogenic shock. the patient had been supported by an intra-aortic balloon pump (iabp) on the right side prior to the impella cp. the impella cp was inserted and positioned in left ventricle successfully. there was difficulty removed iabp ultimately ending with applying femstop. it was noted that iabp explant took approx 20 minutes. shortly after femstop was applied, it was noted that impella was encountering more frequent suction. impella position under fluro was noted to be unchanged however support decreased to the point where impella was able to achieve flows higher than p3. echocardiogram (echo)was requested by local support to assess right ventricle (rv) function. pressure was noted to be dropping. the doctor initially wanted to transfer patient back to intensive care unit however patient coded before transfer could be done. echo arrived and limited study was performed. rv showed little to no function at time of study. return of spontaneous circulation was achieved and the doctor opted to transfer patient to cardiovascular intensive care unit (icu). upon arrival to icu, automated impella controller waveforms indicated device in aorta. impella access was inspected by local support showing impella sterile sleeve had become disconnected and device appeared to have been pulled back. patient¿s pressure dropped and patient ultimately coded again. echo arrived and limited study again performed showing device no longer across atrioventricular (av). the doctor attempted to recross av but was not successful. patient continued to be coded however, despite aggressive efforts the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10995 | 1220648-2025-26815 | 21636587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21636587&pc=OZD | 02/21/2025 | 04/08/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606090 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 66 year old male patient with a cp for support of acute myocardial infarction. after placement of the pump, the pump failed to start and was removed and replaced. new pump successfully placed and patient survived without harm or injury. | [5 empty] | The investigation of pump did not start has been completed. no product was returned for analysis. the data logs show impella stopped alarm electrical alarm. there were no clinical details related to excessive force. the root cause of the pump did not start was most likely due to an electrical open as evidenced from data logs but the reason for electrical open could not be determined as no pump was returned. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26815 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26815. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26815. | D.9 date device returned/information was added. the investigation have been updated as the device was returned for evaluation. the returned pump was inspected and there were no visual abnormalities. the electrical readings were within specification and there were no issues. the pump was able to run in the lab without issues. the data logs showed the impella stopped electrical alarm. there were no clinical details related to excessive force. the root cause of the pump did not start was not determined as the issue did not reproduce in the lab and no pump abnormalities were observed and limited information related to the failure was provided e.1 added us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26815. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
10996 | 1220648-2025-26816 | 21636639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21636639&pc=OZD | 02/21/2025 | 04/30/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614950 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant was trying to deliver a impella 5.5 pump for a patient admitted in with mitral regurgitation and other systemic comorbidities like cancer. it was reported that the 10mmx30 vascutek gelweave graft attached to innominate and tunneled out supraclavicular, but unable to maintain anastomosis due to friable arterial tissue therefore the doctor decided to abort the impella 5.5 pump. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely use issue non-abiomed product was used which is not recommended per instructions for use. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26816 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10997 | 1220648-2025-26817 | 21636759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21636759&pc=OZD | 02/21/2025 | 05/06/2025 |  | 03/18/2025 |  |  | 02/21/2025 | 03/18/2025 |  |  |  |  | 03/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603391 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during repositioning of an implanted impella cp pump, the pigtail wrapped around the motor housing. the automated impella controller subsequently displayed elevated left ventricle pressures and the motor current began to rise. a purge system blocked alarm appeared and the decision was made to replace the pump. there was no resulting patient harm. | [5 empty] | The investigation of high purge pressure has been completed. data logs showed that flows were slightly saturated from the beginning of the case. there were no abnormal motor current (mc) spikes or speed deviations at any time in the case. halfway through the case, the mc stepped up while motor speed remained constant. at this same time, the purge pressure began trending up for 8 minutes until the purge pressure high alarm triggered. the purge pressure and flow calculations remained elevated for the rest of the case. additionally, the ps began to trend up this time, particularly the lv ps. high purge pressure: product was not returned for investigation. data logs show that the motor current was slightly elevated from case start, but there was a distinct step up in motor current at the same time that the purge pressure began to trend up gradually. since there were no motor current spikes during the case, biomaterial likely started to build - adding resistance to the motor - from case start, and then started to impact the purge gap when purge pressure began to rise. clinical details report that a swan-ganz catheter was being manipulated and appeared to interact with the motor housing of the pump at the time of the issue. however, it's reported that visual inspection revealed no problems, and there weren't motor current spikes to indicate an interaction. based on data log review, the root cause of the high purge pressure was most likely biomaterial buildup. saturated pump flow: this failure mode had to be added once data log review revealed that pump flows were saturated during the entirety of the case. there were no abnormal motor current spikes, however, based on the fact that a high purge pressure event occurred shortly after case start, biomaterial most likely caused the saturated flows. from case start, the biomaterial likely added resistance to the motor's rotation - elevating the mc - and then began to affect the purge gap as well when purge pressure began trending up at the same time as the abnormal step up in motor current. product was not returned for investigation. based on data log review the root cause of the saturated pump flow was most likely biomaterial build up. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10998 | 1220648-2025-26740 | 21617361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21617361&pc=OZD | 02/18/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556271 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and shortly following the impella cp implant, distal pulses in the impella limb became absent. the pump was explanted in response to the condition. the patient was reported to be stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
10999 | 1220648-2025-26741 | 21617432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21617432&pc=OZD | 02/18/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567488 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device and was escalated to an impella 5.5 device for continued mechanical circulatory support. the following day after start of the impella 5.5 support, the patient had a ventricular tachycardia arrest followed by a pulseless electrical activity for 30 minutes. the patient was taken to the catheterization lab for a repeat left heart catheterization, which showed a patent left anterior descending stent. an echocardiogram was completed, and the impella 5.5 pump was repositioned. veno-arterial extracorporeal membrane oxygenation (va ecmo) was placed. the patient continued on continuous renal replacement, va ecmo and the impella 5.5. the patient survived and was noted to be in stable condition following the event. additional information was received. the physician believes the patient experienced a vasovagal episode, which led to the event. following the episode, an echocardiogram was performed, and the pump was repositioned as needed. prior to the event, a trans-esophageal echocardiogram was conducted in the operating room, confirming that the pump was appropriately positioned in the mid-left ventricle at that time. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11000 | 1220648-2025-26742 | 21617499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21617499&pc=OZD | 02/09/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584970 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced bleeding during impella cp support. actions taken are unknown. the patient expired while on impella support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11001 | 1220648-2025-26743 | 21617937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21617937&pc=OZD | 02/18/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568581 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the insertion of impella 5.5 device to a patient with an acute myocardial infarction/cardiogenic shock was unsuccessful due to the anatomy. the device was unable to pass the vascular anatomy and the case was aborted. there was no report of adverse effects to the patient. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11002 | 1220648-2025-26745 | 21618178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618178&pc=OZD | 02/14/2025 | 04/02/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601982 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Liver Failure | [11 empty] | Other | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced renal and liver failure during impella cp support. actions taken are unknown. the issue was not resolved. additional information noted that care was withdrawn. the patient was placed on hospice and subsequently expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist: section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | Corrected information provided in b2, h1 and h6.
11003 | 1220648-2025-26744 | 21618188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618188&pc=OZD | 02/18/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610872 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs)and approximately 8 days after start of the support the patient developed a hematoma. post implant of the impella 5.5 device, the patient was transferred to the tertiary center. subsequently, the team observed a hematoma at the site, performed a washout, titrated down the heparin and replaced the purge with sodium bicarbonate. the status of the patient following the event was indicated as unknown. however, the patient continued on the impella 5.5 mcs, which appears to be still implanted. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11004 | 1220648-2025-26746 | 21618278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618278&pc=OZD | 01/17/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573846 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient had shunting caused by ventricular septal perforation (vsp). although impella may have increased the shunt, the vsp itself had expanded over time, making it difficult to suppress the shunt. the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ section: warnings & cautions: warnings: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11005 | 1220648-2025-26747 | 21618432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618432&pc=OZD | 02/19/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261097 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the automated impella controller (aic) had a preventative maintenance (pm) performed and during the pm, the aic kept rebooting on start up, and later it displayed: "unexpected controller shutdown". there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. on february 18th, console logs indicate an unexpected shutdown at 19:08:05, followed by another shutdown shortly after at 19:11:53. both incidents triggered "unexpected shutdown" alarms upon system reboot. after each shutdown, the console's software automatically restarted. additionally, on february 19th, the console experienced another unexpected shutdown, resulting in "unexpected shutdown" alarm after rebooting. there were no warning signs or precursor events leading up to the shutdown, and no indications of any software failures. the console was tested, but the issue could not be reproduced during testing with the pump. no unexpected shutdowns or performance issues were observed. a thorough visual inspection of the internal components was conducted, but no physical problems were found. there was no interruption in the supply detected. the root cause of the intermittent, unexpected controller shutdown could not be determined, as there was no clear indication of the cause, and the issue could not be reproduced.
11006 | 1220648-2025-26748 | 21618689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618689&pc=OZD | 02/19/2025 | 04/02/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263212 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a console found to have damage to the port dust cover. the damage was thought to be broken in patient transfer. no patient was involved. | [5 empty] | The investigation for the automated impella controller (aic) damage before therapy has been completed. no log evidence was found because the reported failure mode was a broken gray cover of the 'front panel optical connector. this console was returned for evaluation and during inspection it was confirmed that the gray cover was missing. the root cause for the missing gray cover was most likely a use issue. a.1 added information as it was inadvertently omitted from the initial manufacturer device report number. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11007 | 1220648-2025-26749 | 21618807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618807&pc=OZD | 02/17/2025 | 05/07/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  | 12/31/2025 |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441831 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a surgical patient supported by an impella cp and then escalated to an impella 5.5 pump. the impella 5.5 was placed but while adjusting its position it was pulled too far and ended up being dislodged into the aorta. after checking with the doctor and medical staff, it turned out that this had happened because the positioning had been adjusted without using an ultrasound. the patient survived the loss of position/support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump was dislodged into aorta while repositioning of the pump without echo guidance. g2 report source; foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26749.
11008 | 1220648-2025-26750 | 21618823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618823&pc=OZD | 02/19/2025 | 04/03/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124467 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 42 year old male who while on impella 5.5 support, the impella console (aic) experienced a controller error-yellow alarm upon initializing. the aic was exchanged. the patient survived the procedure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the data logs were returned. the console was returned for testing but failure mode was not reproduced. the first initial placement signal issue was most likely an air gap given some of the symptoms disappeared after the pump reconnection. second, the periodic nature of the optical signal oscillations could not be determined because the failure mode was not reproduced but could be due to a malfunction of the power battery manager¿s optical bench supply or optical bench. the controller error and optical bench data loss is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26750. g.1 added reporting contact facility name as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26750.
11009 | 1220648-2025-26751 | 21618958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618958&pc=OZD | 02/19/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 110504 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 59 year old male patient on a impella 5.5 support for heart failure. while patient was being transported, the automated impella controller (aic) had an impella battery low alarm occur after only 30 minutes unplugged. alarm condition resolved after plugging device into outlet on bed. no further alarms reported. there was no patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
11010 | 1220648-2025-26752 | 21618999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21618999&pc=OZD | 02/21/2025 | 04/11/2025 |  | 03/17/2025 |  |  | 02/21/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594309 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 pump to support cardiogenic shock. the pump was placed via the axillary and support was ongoing when there were concerns of a purge leak. the purge cassette was assessed and no leak was noted. the crack was observe in the purge reservoir and the support was continued. the crack was thought to be due to ambulation. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. no product was returned. data logs show purge pressure steady around ~600 mmhg for the majority of the case and purge flow between ~4-10 ml/hr. clinical details state that there was a crack in the reservoir that lead to a slow leak. the reservoir crack was covered in tegaderm. the root cause of the purge leak was most likely a damaged sidearm but the exact location is unknown since no product/images were returned.
11011 | 1220648-2025-26753 | 21619154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619154&pc=OZD | 02/19/2025 | 04/22/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  | 02/28/2026 |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478504 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received that after 78 days of support, high purge pressure alarms occurred. the patient still remains on the pump for support. | The user facility reported a 63-year-old male patient with an impella 5.5 and ecmo for the support of a patient admitted in with cardiogenic shock (cgs) and cardiomyopathy. the pump was supporting while patient awaited decision on left ventricular assist device (lvad) or other advanced therapies. after 28 days of support the pump had to be troubleshot with tissue plasminogen activator (tpa) for high purge/purge blocked alarms. the pump remains on and there was no patient harm reported. |  |  |  |  | B.5 additional information was received. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26753. g.1 added reporting contact facility name as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26753. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11012 | 1220648-2025-26754 | 21619246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619246&pc=OZD | 02/21/2025 | 04/02/2025 |  | 03/17/2025 |  |  | 02/21/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607919 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient developed hemolysis during impella cp support. the pump was repositioned, but the condition persisted. support was continued for two more days until the decision was made to withdraw care due to the patient's significant comorbidities and poor prognosis. the patient subsequently passed away. it is unknown if the condition contributed to the patient's passing. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. no product was returned for analysis. the clinical details mentioned the pump was pulled 10 centimeters back. clinical information mentioned the hemolysis did not resolve with repositioning and with pump removal. clinical details also mentioned the patient had poor prognosis and significant comorbidities. the data logs returned did not span the entire case duration. logs only show 20-feb-2025 and no motor current spikes. the root cause of the hemolysis issue could not be definitively determined as no product was returned for analysis and insufficient data logs were provided. g.1 reporting contact us phone number should of been left blank on the initial manufacturer device report 1220648-2025-26754. added manufacturing site fax number as it was inadvertently omitted from initial manufacturer device report 1220648-2025-26754. h.1 type of reportable event was previously entered incorrectly on the initial manufacturer device report 1220648-2025-26754.
11013 | 1220648-2025-26755 | 21619380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619380&pc=OZD | 02/19/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616504 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a patient on support with an impella cp. after exchanging to the repositioning sheath, the blue suture hub was secured at the access site with forward tension sutures and a hematoma rapidly developed. manual pressure held and angle matching was done as well. femoral angiogram was done via wire re-access port. no noticeable extravasation was noted. a decision was made to remove the impella cp due to the hematoma. the impella cp was removed and hemostasis was achieved. the patient remained stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11014 | 1220648-2025-26756 | 21619393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619393&pc=OZD | 02/17/2025 | 04/23/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612703 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the impella cp insertion the cannula kinked and got stuck in the brachiocephalic truncus. the cp was pulled back and was exchanged. there was no patient harm. | [5 empty] | Pump returned and visually inspected. multiple minor kinks found along cannula. no other damages or abnormalities found. delivery issues reported unable to pass through vasculature with resistance met at aortic arch. clinical reported noticing the cannula was kinked and stuck in the brachiocephalic truncus after pushing up the pump. pump was removed and replaced. delivery issue: the root cause of the delivery issue was most likely patient condition related since the pump was unable to pass through vasculature and got stuck in the brachiocephalic truncus. product damage: the root cause of the product damage cannula kink was most likely patient condition since the pump got stuck in the brachiocephalic truncus during challenging delivery. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26756 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 component code 925 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26756. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11015 | 1220648-2025-26757 | 21619435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619435&pc=OZD | 02/19/2025 | 04/07/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571234 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70 year old male who after the impella cp placement, the patient developed a hematoma and bleeding from the access site. sutures and angle matching was done but there was still some oozing. seven unit of blood was administered. the impella was removed, the hematoma was evacuated, and the right femoral artery was repaired. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - peripheral vessel has been completed. the data logs were not applicable. the pump was returned and evaluated. there were no abnormalities such as tip damage found with the repositioning sheath. the 14 french sheath was not returned. based on the clinical details, there was oozing and a small hematoma noted at the access site. hemostasis was said to be achieved after pump removal with right femoral artery repair and hematoma evacuation. it was believed by the physician that angle of stick and entry caused impella sheath to kink and caused uncontrollable blood loss. the root cause of the vascular damage was most likely use-related since the physician noted a poor angle of stick and entry causing sheath to kink followed by uncontrollable blood loss. e.1 revised phone number as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26757. g.1 added reporting contact facility name as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26757. h.6 added health effect - clinical code 1884 and health effect - impact code 4627 as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26757. code 4315 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26757 as the device was returned and was investigation was pending. codes 3233 and 11 should of been on the initial manufacturer device report number 1220648-2025-26757. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26757 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
11016 | 1220648-2025-26758 | 21619477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619477&pc=OZD | 01/04/2025 | 05/19/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124584 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and while on support, greater than two months later, placement signal not reliable alarm triggered, and aorta and left ventricle placement signal were unavailable. there was no harm to the patient as a result of the event; the patient was eventually bridged to a left ventricular assist device per the report. abiomed's investigation engineer has now requested the return of the automated impella controller (aic) to further investigation if the issue could have been a result of the aic. | [5 empty] | The investigation of automated impella controller (aic) - loss of opm data has been completed. based on the data analysis and the device analysis, the cause of the aic issue was a defective optical bench. d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26758 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The automated impella controller device was not received from the customer at the time of this mdr writing; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11017 | 1220648-2025-26759 | 21619497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619497&pc=OZD | 02/18/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579638 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and approximately 12 days after start of the support, the patient developed bleeding from the access site. the patient lost approximately 1.2 l of blood as a result of the condition. manual pressure was applied for 45 minutes, and one unit of packed red blood cells was transfused as a result of the event. support continued following the intervention, and the patient was reported to be in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11018 | 1220648-2025-26760 | 21619534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619534&pc=OZD | 02/18/2025 | 05/01/2025 |  | 03/17/2025 |  |  | 02/18/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support (mcs) and experience access site bleeding same day of implant. the patient was transported to unit with femstop in place due to continued oozing around repositioning sheath. distal flow remained adequate with palpable pedal pulse. the following day, the physician tightened the perclose sutures which stopped the oozing. the groin site was unremarkable. the patient's outcome after the event was indicated as unknown. however, the patient was reported to have remained on the impella mcs for an additional three days before successful weaning and explant of the impella cp device at bedside. no bleeding or hematoma was noted after explant. | [5 empty] | Product-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was most likely use issue since the bleeding stopped after the physician tightened the sutures. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26760 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11019 | 1220648-2025-26761 | 21619600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619600&pc=OZD | 02/22/2025 | 05/28/2025 |  | 03/17/2025 |  |  | 02/22/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1168284 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | White | Death | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 pump support for 14 days while the patient was awaiting transplant/left ventricular assist device (lvad) decisions. the automated impella controller (aic) exhibitied a controller alarm yellow for a failure. the aic was removed/replaced and support was continued without harm or injury for just under 21 days until the decision was made to withdraw care and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude the impella 5.5 as an associated factor in the patient's outcome. although the console error occurred, the patient's lvad, critical cardiac nature, and multiple other issues contributed to the patient's demise. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. console logs from the reported day of the event are consistent with the complaint. the console was sent back to field service for further investigation and was tested after arriving in fs. voltage from the power battery manager (pbm) at connector (j3) was 4.9v. the voltage on the optical bench lamp connector (p4) was 4.8v. the lamp power was 2.65uw, which is below the specified value stated in the pm procedure, but did not cause the reported issues. the analog output from the optical bench (envelope/fringe) was observed without any distortions. the controller error was identified as a result of an anomaly in the optical bench firmware communication protocol, leading to misalignment of the data communication packets received by the epc. meanwhile, the aic software functioned as intended. further examination of the aic revealed that the 250 hz placement signal was experiencing sample loss, resulting in negative spikes. root cause: the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)" and the "controller error (optical bench ambient pressure out of range) [optical pump connected] - alarm issued" was due to a firmware issue (optical bench fw anomaly). the root cause for the distortion on the placement signal was most likely due to the optical bench malfunction.
11020 | 1220648-2025-26762 | 21619678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619678&pc=OZD | 02/19/2025 | 04/08/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609825 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 84 year old male patient on a impella cp pump support developed hemolysis despite of repositioning the cp. the impella remained on overnight and was successfully weaned the next day. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the data logs showed the pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. no motor current spikes in the data logs to indicate ingestion event. downward trending placement signal throughout the data log run. the pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. the pump passed hemolysis test. mih was 6.31. the root cause of hemolysis was most likely patient condition related since the patient had poor left ventricle function and no problem was found on the pump. e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26762. add fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26762. g.1 added reporting contact facility name as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26762. h.6 investigation conclusion code 4315 was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26762. codes 3233 and 11 should have been on the initial manufacturer device report number 1220648-2025-26762 as the device had been received and the investigation was pending. h.10 additional manufacturer narrative on the initial manufacturer device report number 1220648-2025-26762 stated that the device was not returned but the pump was returned.
11021 | 1220648-2025-26763 | 21619686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619686&pc=OZD | 02/22/2025 | 04/23/2025 |  | 03/17/2025 |  |  | 02/22/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593681 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that following impella cp implant for a patient on support for acute myocardial infarction, the motor current, placement signal, and left ventricular tracings became irregular. as the pump continued to function, the planned surgery was successfully performed with support of the implanted pump. the pump was then replaced as a result of the noted issue. there was no noted patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Optical signal issue: the pump was returned. data log review confirmed high mc and pump flow saturation. the returned pump showed no optical issue: 5s parameters were nominal, resistance testing passed, laser continuity testing passed, and the pump did not trigger psnr alarms. the cause of the optical signal issue was not determined since issue could not be recreated during investigation as pump ran at nominal 5s parameters while passing electrical and wiring testing. saturated flow: the pump was returned. data log confirmed high mc and saturated pump flow . the returned device was investigated and although no visual abnormalities were noted, saturated flow reproduced upon pump startup. pump was torn down and biomaterial was noted on the components of the motor housing. the cause of the saturated flow was most likely biomaterial due to saturated flow reproducing in lab and pump teardown showing biomaterial within the components of the motor housing. it is unknown how the biomaterial ingressed into the motor housing.
11022 | 1220648-2025-26764 | 21619987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619987&pc=OZD | 02/23/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/23/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025602061 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp for acute myocardial infarction and during support, the impella cp became back into the aorta. the pump was explanted a day earlier than planned. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11023 | 1220648-2025-26765 | 21619988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21619988&pc=OZD | 05/12/2021 | 03/17/2025 |  | 03/17/2025 |  |  | 02/17/2025 | 03/17/2025 |  |  | 03/31/2023 |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060403 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the medical team was having delivery issues with the impella 5.5 pump because the axillary artery was too small so the pump was aborted. there was no harm to the patient reported. | [5 empty] | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of the delivery issue - anatomy was most likely due to patient condition. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11024 | 1220648-2025-26766 | 21620452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21620452&pc=OZD | 05/14/2021 | 03/17/2025 |  | 03/17/2025 |  |  | 02/17/2025 | 03/17/2025 |  |  | 10/31/2022 |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021010193 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was unable to pass through the patient's tortuous vasculature during attempted delivery. as a result of the noted resistance caused by the patient's anatomy, the decision was made to abort the implant. there was no reported patient harm. | [5 empty] | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of the delivery issue - anatomy was most likely due to patient condition. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11025 | 1220648-2025-26767 | 21620669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21620669&pc=OZD | 02/24/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/24/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025604657 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 42 year old male patient with a impella cp for support of acute myocardial infarction. during support, the impella cp stopped without any prior signs or alarms. attempts to restart the pump were unsuccessful. the impella was exchanged. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
11026 | 1220648-2025-26768 | 21620978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21620978&pc=OZD | 02/24/2025 | 04/29/2025 |  | 03/17/2025 |  |  | 02/24/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553904 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump experienced high purge pressure during support. lysis therapy was initiated and the purge levels stabilized. the patient passed away on support 15 days later. there is not enough information to disassociate the death with the use of the impella. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. clinical states that purge pressure of 1070mmhg and purge flow of 1.6ml/hr was observed. the pump and data logs were not returned. therefore, the root cause of the high purge pressure issue was not determined since product and data logs were not returned. (only the cassette was returned) b1 revised from the initial manufacturer device report number 1220648-2025-26768 to reflect both adverse event and product problem. d9 should have been left blank on the initial manufacturer device report number 1220648-2025-26768, as the pump was not received for investigation, only the cassette. g1 reporting contact address and country were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26768. g1 manufacturing site name, address, and fax number were erroneously on the initial manufacturer device report number 1220648-2025-26768. h3 should have reflected the pump was not returned for investigation on the initial manufacturer device report number 1220648-2025-26768.
11027 | 1220648-2025-26769 | 21621176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21621176&pc=OZD | 02/24/2025 | 05/16/2025 |  | 03/17/2025 |  |  | 02/25/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026614362 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unsealed Device Packaging | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp pump was delivered with a broken seal. the customer would like a replacement. there was no patient involvement. | [5 empty] | The investigation of packaging issues- sterility has been completed. the seal on the packaging of a delivered pump was reported to be broken when the parcel arrived at the hospital and they required a replacement. logs are not applicable. product was returned and evaluated. the sticker label on one side of the pump kit packaging was broken. all other labels were intact. there were no other abnormalities found. based on the returned product, the broken seal was on the outer packaging, meaning that the fm should be updated to "packaging issue" since sterility of the pump is not associated with the outer packaging. the root cause of the packaging issues was most likely a damaged label as confirmed on the returned product. d9 date device returned to manufacturer added. e3 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26769. g1 reporting contact name, address, and country were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26769. h6 medical device problem code 4008 was erroneously entered on the initial manufacturer device report number 1220648-2025-26769. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
11028 | 1220648-2025-26770 | 21621681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21621681&pc=OZD | 02/19/2025 | 04/08/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585447 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp after a patient was admitted in with a cardiac history. the pump was placed but the team noted signs of hemolysis. the urine color changed and the pump position was deep. the team repositioned the pump and the urine cleared. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the data was not returned for review. visual inspection found no issue on the pump. pump ran without issue under bench test. per clinical description, the cause of hemolysis was due to pump was in improper position and resolved after repositioning of the pump. pump analysis revealed no issue and ran without issue under bench test. the root cause of hemolysis was most likely due to pump was not in the proper position.
11029 | 1220648-2025-26771 | 21623635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21623635&pc=OZD | 05/04/2021 | 03/17/2025 |  | 03/17/2025 |  |  | 02/17/2025 | 03/17/2025 |  |  | 07/31/2022 |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1494728 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was unable to pass through the patient's vasculature during attempted delivery. as a result of the resistance caused by the patient's anatomy, the decision was made to abort the implant. an intra-aortic balloon pump was subsequently implanted for patient support. there was no reported patient harm. | [5 empty] | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of the delivery issue - anatomy was most likely due to patient condition. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11030 | 1220648-2025-26772 | 21623895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21623895&pc=OZD | 02/18/2025 | 04/16/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump support placed in a patient. the support was ongoing for a day when the team removed the impella cp and escalated to a 5.5 pump due to limb ischemia and access site bleeding. the patient outcome was noted to be critical. | [5 empty] | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11031 | 1220648-2025-26773 | 21626177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21626177&pc=OZD | 02/19/2025 | 04/22/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558669 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had optical sensor problems. the physician didn¿t want to implant the impella with this issue. troubleshooting was done but unsuccessful. the impella cp was exchanged and the new impella worked without any problems. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of optical signal issue has been completed. during pump implantation procedure, placement signal did not work out of box. the pump was returned. data logs were not recovered. investigation of the pump revealed a damaged fiber line on the patient side of the red pump plug during laser continuity test. the optical fiber sat at a strenuous angle in order to route around the other wires and the post. the patient cable was unable to rotate freely in the kink protector. the cause of the optical signal issue was a damaged optical fiber line due to the damaged optical fiber line found during laser continuity inspection. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26773 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact address line 1, reporting contact us city, reporting contact state, reporting contact us zip code and reporting contact country as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26773. g.2 added foreign as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26773.
11032 | 1220648-2025-26774 | 21626182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21626182&pc=OZD | 02/19/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563970 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support had pink/red urine in the foley and noted hematuria. the urine clear but the following day, it when back to pink. a decision was made to decannulate from extracorporeal membrane oxygenation and escalate the impella cp to impella 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11033 | 1220648-2025-26776 | 21626269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21626269&pc=OZD | 02/19/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540589 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient (unknown race and ethnicity) undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. when trying to remove the dilator from the peel away sheath difficulty was encountered. only slight force could be used or the ring for attaching the dilator comes loose. the implanted pump ran without issue. the pump was successfully weaned and explant after approximately four hours of support. there was no patient harm as a result of the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11034 | 1220648-2025-26777 | 21627204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21627204&pc=OZD | 02/20/2025 | 04/09/2025 |  | 03/17/2025 |  |  | 02/20/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608248 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a fluid leak was seen on the yellow luer lock during impella cp support on a 53 year old male patient. a purge system open alarm appeared on the console (aic) at the time of the leak. the purge cassette was replaced but the issue persisted. a purge bypass was subsequently performed to resolve the leak and support continued with the implanted pump. there was no resulting patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the product was not returned for investigation. a data log review was not required for this case. according to the clinical report, the doctor performed a sidearm bypass after a leak was reported from the yellow luer, which successfully resolved the issue. the cause of the purge leak issue was a leak from the sidearm but the exact location of the leak was unknown. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-26777. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26777 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name, address, and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26777.
11035 | 1220648-2025-26778 | 21627225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21627225&pc=OZD | 02/20/2025 | 04/07/2025 |  | 03/17/2025 |  |  | 02/20/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025609731 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male patient was on impella 5.5 support and during support, the pump was in suboptimal position. there were decreased flows. repositioning was attempted without success. the pump was pulled back into the aorta without increase in flows. the pump was exchanged. upon explant a thrombus was noted in the inflow. the patient survived the procedure. | [5 empty] | The investigation of positioning issues, low pump flow, and thrombus has been completed. pump was not returned for investigation. as per the clinical information provided, the pump was pulled into aorta by the doctor while repositioning. therefore, the root cause of the positioning issue was most likely use related to improper technique. pump was not returned for investigation. data logs were not returned as well. without the product or the data logs, the failure cannot be further investigated. therefore, the root cause of the low flow issue was not determined since product was not returned and data logs were not returned as well. as the necessary information was not provided, the root cause of the thrombus was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26778 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26778. h6 code 4582 was reported incorrectly on manufacturer device report 1220648-2025-26778. new code was added to health effect - clinical code. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock entering cardiac output (impella cp with smart assist) section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
11036 | 1220648-2025-26779 | 21627253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21627253&pc=OZD | 02/21/2025 | 04/02/2025 |  | 03/17/2025 |  |  | 02/21/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported 76 year old female patient was on impella cp support and during support, the patient had limb ischemia. the patient had no pulses present in right leg distal to impella access. the impella was removed. additionally, further details state that care was withdrawn and the patient subsequently expired. there were no details provided as to the cause of the death. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.7 revised and added information since the initial manufacturer device report number 1220648-2025-26779 was submitted. e.1 added us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26779 was submitted. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26779 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact facility name and fax number as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26779. h.6 component code 925 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26779. a new component code was added. code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26779.
11037 | 1220648-2025-26780 | 21627262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21627262&pc=OZD | 02/21/2025 | 04/07/2025 |  | 03/17/2025 |  |  | 02/21/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2025572689 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 63 year old male patient was on impella support and during support, the purge disc was not recognized. staff attempted to change the purge cassette however, the issue was with the automated impella controller (aic). the aic was replaced. there was no patient harm reported. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. there was no evidence found in the logs to confirm the reported failure. the console was and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26780 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name, address, and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26780.
11038 | 1220648-2025-26781 | 21627276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21627276&pc=OZD | 02/19/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542542 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs). the patient was a transfer from an outside hospital on an intra-aorta balloon pump for a left ventricular assist device workup. approximately nine days after start of impella mcs, the pump flow saturation was noted and the pump stopped. pump stopped. the automated impella controller (aic) was exchanged; however, the pump did not restart. following, the decision was made to pull impella across the aortic valve and start comfort care. the patient would subsequently expire. | [5 empty] | The impella device was not received from the customer at the time of this mdr writing; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11039 | 1220648-2025-26782 | 21627278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21627278&pc=OZD | 02/19/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630079 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced positioning issue. the patient presented to catheterization lab on bipap for a high-risk percutaneous coronary intervention (pci). the patient became unresponsive as wire was attempting to cross valve. pressures in the 40's transitioning to pulselessness. cardiopulmonary resuscitation was initiated, and the patient was intubated. return of spontaneous circulation was obtained, and the impella cp was placed with flows of 3.2 l/min and mean arterial pressure of 138. single access pci was attempted with alternate site right femoral artery. pci was completed to the left internal mammary and left anterior descending (lad) arteries with one stent placed to ostial lad to proximal lad. the impella cp pump position appeared on echocardiogram to have a depth of 3.5cm with the pigtail in the mitral valve apparatus. an attempt was made to reposition the pump but left as there was concern for worsening mitral valve regurgitation. the patient was transported to the unit on support. an attempt was made to reposition the pump under echocardiogram at bedside. the patient had existing mitral regurgitation and per the physician the mitral regurgitation was presently consistent with baseline. the plan was to attempt to wean patient as tolerated. the patient remained intubated and sedated. the following day the impella position remained unchanged. the next day thereafter the patient was successfully weaned, and the pump was explanted. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11040 | 1220648-2025-26783 | 21627288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21627288&pc=OZD | 02/19/2025 | 04/25/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572022 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. following the implant access site hematoma developed around the repositioning sheath. manual compression was applied, and forward tension suture was added. two units of pack red blood cells were administered. the patient continued on support. however, approximately three days later, there were no flows to impella extremity and leg was severely mottled/dusky. the impella cp device was, therefore, weaned and explanted. the patient survived the explant and was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding and limb ischemia has been completed. the root cause of the access site bleeding issue was most likely use related, manual compression and forward tension suture added to stop the bleeding. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26783 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26783.
11041 | 1220648-2025-26784 | 21627765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21627765&pc=OZD | 02/05/2025 | 03/17/2025 |  | 03/17/2025 |  |  | 02/19/2025 | 03/17/2025 |  |  |  |  | 03/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594775 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | A notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured tachyarrhythmias with a "definite" relationship to the impella device used. following, the patient continued to have runs of tachyarrhythmias overnight and into the day requiring increased amiodarone and cardioversion (arrhythmia). the patient recovered with no sequelae. per the report the patient continues on the impella 5.5 mechanical circulatory support. | [5 empty] | The impella device was not received from the customer as it last indicates the device remains implanted supporting the patient. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11042 | 1220648-2025-26719 | 21604447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21604447&pc=OZD | 02/17/2025 | 03/26/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578424 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of an impella 5.5 device to a patient with cardiomyopathy was unsuccessful due to the patient's anatomy. the impella 5.5 device was advance to innominate and after several attempts was unable to get down the ascending aorta due to difficult acute angle of vessel. the physician decided to place an intra-aortic balloon pump for support. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely use issue since a non-abiomed product was used and the doctor lost the access of wire.
11043 | 1220648-2025-26720 | 21604576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21604576&pc=OZD | 02/17/2025 | 03/14/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576975 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced limb ischemia. the next day after start of the support, the patient's right lower extremity became mottled. a plan was made for the removal of the impella cp device with potential escalation to an impella 5.5 device pending vasopressor requirements with the impella cp removal. the patient was noted to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11044 | 1220648-2025-26721 | 21604852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21604852&pc=OZD | 02/17/2025 | 03/14/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605411 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced access site bleeding requiring intervention and led to device removal. the patient was prepped and draped in usual fashion. the right common femoral artery was accessed and direct placement of 4f x 13 cm peel away sheath. aortic valve was crossed with pigtail. upon impella cp catheter insertion through hemostatic valve, the peel away sheath began to leak blood. it was unclear from where the blood was leaking from; however, physician felt he may have inserted impella catheter not through center of hemostatic valve. impella was advanced into left ventricle under fluoroscopy into good position and started. the peel away sheath was quickly removed due to bleeding. after this, a hematoma was noted in groin and bleeding from insertion noted whenever manual pressure was removed. the patient was transferred to unit after given proteins with hopes that bleeding would stop, which did not occur and led to the decision to remove impella cp device (after approximately one hour of support). the patient was hemodynamically stable pre impella-cp insertion; therefore, the physician did not feel need to place second impella. the patient was noted to be stable following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation).
11045 | 1220648-2025-26722 | 21605522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21605522&pc=OZD | 02/17/2025 | 04/22/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593572 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 53 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Cardiac Tamponade | Pseudoaneurysm | Vascular Dissection | [7 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was electively implanted with an impella 5.5 device for mechanical circulatory support and experienced a vascular dissection with a demised outcome. a computed tomographic angiography of the chest was completed after the development of spontaneous tamponade approximately 17 days after impella implant. findings revealed a 2.2 centimeter contained leak arising from the anterior wall of the distal ascending aorta immediately superior to the entry point of the impella device with a surrounding hematoma. the decision made to take the patient to the operating room for chest exploration and explant of the impella 5.5 device. upon retracting the chest open, a leak was noted in the aorta. the patient was emergently cannulated for cardiopulmonary bypass and was noted to have an extensive aortic dissection. the patient expired on the operating room table. the physician's comments afterwards noted unknown cause why this event occurred. the physician noted the possibility of infection at the site though no obvious signs of infection were noted. extensive repair was required when the patient was placed on cardiopulmonary bypass. the dissection extended to where no flow was getting to head vessels while attempting to put on circulatory arrest. | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular tachycardia with a "possible" relationship to the impella device used: ¿post cabg and impella implant, subject had multiple runs of vt requiring cardioversion x2.¿ the patient recovered with no sequelae. additional information received: loqi reported that a cta chest showed a leak from the impella's graft's entry point. the patient taken to or for impella removal. ¿during impella removal, subject's aorta began bleeding profusely. impella removed and a large pseudoaneurysm was found in aorta. this dissected down towards the root of the aorta. subject expired in the or.¿ the patient endured cardiac pseudoaneurysm with a "possible" relationship to the impella device used. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B.5 additional information was received. g.2 added report source due to new information received. h.6 additional coded were added due to new information received. h.10 added investigation results for ventricular tachycardia issue. the investigation for ventricular tachycardia issue has been completed. the root cause of the ventricular tachycardia issue was not determined since the product was not returned and insufficient clinical information was provided. | The investigation of vascular damage - great vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11046 | 1220648-2025-26723 | 21606047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21606047&pc=OZD | 02/16/2025 | 04/01/2025 |  | 03/14/2025 |  |  | 02/16/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123072 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during preparation of the automated impella controller (aic) for impella 5.5 mechanical circulatory support for a patient with an acute myocardial infarction/cardiogenic shock, the aic experienced system self-check error at start up. when the aic was turned on, the screen was blue and displayed ¿system self-check failed change console immediately." the battery was turned off with a five-minute wait before turning it back on. however, the error was not resolved. the aic was exchanged and connected to the impella 5.5 for mechanical circulatory support. the patient would expire approximately 1.5 days later for reasons determined unrelated to the aic. however, events related to the impella pumps were considered and notation of related reports are captured in h.10. medical safety review of the event was completed and noted the aic failed self-check at start up and was exchanged with no indication or report of interference mid support or inadequate support. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs from the reported day of event are consistent with complaint and show system self-check failures due to persistent loss of communication with power battery manager (pbm)1. the inspection revealed corrosion on the pbm, due to leaking electrolyte from supercaps also, console¿s batteries were fully depleted, most likely due to pbm failing to charge them. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of three device manufacturing reports associated with the event and/or implant experience. refer to the following: medical device report for the impella cp serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). medical device report for the impella 5.5 serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). medical device report for the automated impella controller serial number (b)(6) with date of the event 16-feb-2025 and patient identifier (b)(6) (this report).
11047 | 1220648-2025-26724 | 21606235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21606235&pc=OZD | 02/18/2025 | 03/14/2025 |  | 03/14/2025 |  |  | 02/18/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581475 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the device was removed because it had become dislodged into the aorta. a new impella cp was placed. the issue occurred upon moving the bed from the catheter table. it was noted that the fixing ring may have been forgotten to be closed. the patient later expired after eight days of support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11048 | 1220648-2025-26725 | 21606356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21606356&pc=OZD | 02/18/2025 | 04/04/2025 |  | 03/14/2025 |  |  | 02/18/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594277 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced distal leg ischemia below the femoral access site. a contralateral femoral artery to impella leg bypass was required. the next day there was improvement in right leg color and temperature, but it was still very faint distal pulses. additionally, it was reported that the physician accidentally ripped the sterile sleeve when manipulating the repo sheath. tegaderm was used to seal it. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of limb ischemia and product damage (sterile sleeve damage) has been completed. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of the product damage issue was not determined since product was not returned and due to insufficient clinical imaging/wording. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26725. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11049 | 1220648-2025-26726 | 21606851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21606851&pc=OZD | 02/17/2025 | 03/14/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601149 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced a stroke. care team is unclear when the stroke occurred, however. the patient was transferred in and had worsening respiratory status requiring intubation and increasing pressor requirement. the decision was made to for impella cp insertion and left and right heart catheterization with possible intervention. the impella was implanted and in auto mode flowing 3.5lpm. single access technique utilized with companion sheath for intervention. percutaneous transluminal coronary angioplasty (ptca) and intravascular ultrasound (ivus) to left main, opened up left circumflex artery for ptca and ivus. synergy drug-eluting stent was placed to left circumflex. the physician was pleased with final results and decided to leave the impella cp in place at least overnight for evaluation. companion sheath and wires removed. peel-away sheath was removed and repositioning sheath advanced in place. impella slack removed, position confirmed and locked down at 91cm. forward tension sutures and angle matching provided for site hemostasis. patient to recover in the intensive care unit. the following day, the impella cp device was removed without incident; access site hemostasis achieved. five days after explant of the device it was reported the patient had a stroke. computed tomography scan and magnetic resonance imaging results showed embolic stroke to left frontal, parietal, and occipital lobes. patient experienced right-sided deficits. care team is unaware of exactly when the stroke occurred as patient was sedated heavily since admission. of note, the patient on heparin drip prior to, during and post impella support. the physician did ask if the impella can cause a stroke and was educated on the risk of stroke with anticoagulation, impella support, balloon angioplasty in coronary intervention, and known patient history of peripheral vascular disease and coronary artery disease. the physician shared worries of the impella cp pigtail making her "nervous with explant" as the physician does not feel comfortable removing any pigtail without wire support. the physician provided feedback to be able to reinsert wire through pigtail to be able to alleviate the stress of removing impella in regard to risk of clot shearing. | [5 empty] | As the actual day of the stroke is unknown, the aware date has been captured as the provisional date of event. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11050 | 1220648-2025-26727 | 21606897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21606897&pc=OZD | 02/18/2025 | 03/28/2025 |  | 03/14/2025 |  |  | 02/18/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592762 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. due to difficult anatomy, the physician was unable to fully insert the 14fr peel away sheath. the physician made the decision to abort impella placement. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. upon insertion due to difficult anatomy, the physician was unable to fully insert the 14 french peel away sheath into the arteriotomy. b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26727. e.1 added us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26727. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26727 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4611 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26727.
11051 | 1220648-2025-26728 | 21606908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21606908&pc=OZD | 02/18/2025 | 03/27/2025 |  | 03/14/2025 |  |  | 02/18/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607733 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bruising and swelling to impella site on right chest wall as well as bruising to inner aspect of right upper arm. two units of packed red blood cells and one unit platelets were administered to the patient. sandbag was applied. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 was inadvertently left blank on the initial manufacturer device report number 1220648-2025-26728. d6a was erroneously entered on the initial manufacturer device report number 1220648-2025-26728. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26728 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11052 | 1220648-2025-26729 | 21606928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21606928&pc=OZD | 02/17/2025 | 05/06/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612485 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support experienced a pump stop seven days into support possibly caused by an automated impella controller failure (according to the controller alarm). the impella 5.5 device was rapidly weaned and explanted. therapy, if any, after explant of the impella device is unknown. there was report or indication of harm to the patient as a result of the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-26666 for the report on the automated impella controller. | The investigation of pump stop has been completed. no product was returned. log review showed an impella stopped alarm 119 about 164 hours into support, indicating an electrical failure. the pump was unable to be restarted successfully. data logs showed pump use from 10-feb-2025 to 17-feb-2025. no relevant additional details were provided from the clinical. the root cause of the pump stop was most likely an electrical open based on the alarm in the data logs but it is unknown how the failure occurred due to insufficient clinical details. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26729.
11053 | 1220648-2025-26730 | 21607093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21607093&pc=OZD | 02/16/2025 | 03/14/2025 |  | 03/14/2025 |  |  | 02/16/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590655 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced limb ischemia requiring intervention. the patient would later expire. the patient presented with st-segment elevation myocardial infarction. percutaneous coronary intervention was completed to the left anterior descending artery with one stent. left ventricular diastolic pressure was 30 and left ventricular ejection fraction was 30%. post pci the impella cp device was implanted, and the patient was sent to the unit on support. approximately twelve hours later into the next day the patient developed limb ischemia. creatine level was rising, and lactate level was increasing. there were no pulses distal to the impella cp. the decision was made to explant and replace with escalation to an impella 5.5. additional information received indicating that sutures were used when they did a femoral vascular cut down to remove the impella cp and then did a fasciotomy to the leg. the patient was placed on the impella 5.5. this same the patient was preemptively started on continuous renal replacement therapy and the patient had washout lactic from the ischemic leg. the patient remained critically ill and was planned for dialysis for lactic acidosis. bleeding issues from all sites (impella site, femoral site, femoral fasciotomy, and orally) was observed. impella 5.5 flows and waveforms were noted within normal limits. heparin was held post procedure and blood products were administered. later in the night it was noted that the patient was unable to tolerate prolonged intermittent renal replacement therapy. urine output was poor, and the nurse reported unchanged bleeding from line sites and incisions. the patient remained critically ill and would expire the following day. a transesophageal echocardiogram revealed a large ventricular septal defect and care was withdrawn. medical safety review of the event noted there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome (demise). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of three device manufacturing reports associated with the event and/or implant experience. refer to the following: medical device report for the impella cp serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). (this report). medical device report for the impella 5.5 serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). medical device report for the automated impella controller serial number (b)(6) with date of the event 16-feb-2025 and patient identifier (b)(6). instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation).
11054 | 1220648-2025-26731 | 21607404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21607404&pc=OZD | 02/16/2025 | 04/03/2025 |  | 03/14/2025 |  |  | 02/16/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507967 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device and was escalated to an impella 5.5 device for mechanical circulatory support. the patient developed limb ischemia while on impella cp mcs (on day after start of this support). the decision was made to escalate to an impella 5.5 which was successfully completed. on the same day of the impella 5.5 implant, the patient was preemptively started on continuous renal replacement therapy and the patient had washout lactic from the ischemic leg. the patient remained critically ill and was planned for dialysis for lactic acidosis. bleeding issues from all sites (impella site, femoral site, femoral fasciotomy, and orally) were observed. impella flows and waveforms were noted within normal limits. heparin was held post procedure and blood products were administered. later in the night it was noted that the patient was unable to tolerate prolonged intermittent renal replacement therapy. urine output was poor, and the nurse reported unchanged bleeding from line sites and incisions. the patient remained critically ill and would expire the next day. a transesophageal echocardiogram revealed a large ventricular septal defect and care was withdrawn. medical safety review of the event noted there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome (demise). related reports are noted in h.10 accordingly. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of three device manufacturing reports associated with the event and/or implant experience. refer to the following: medical device report for the impella cp serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). medical device report for the impella 5.5 serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). (this report). medical device report for the automated impella controller serial number (b)(6) with date of the event 16-feb-2025 and patient identifier (b)(6). | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 selection of life-threatening was erroneously selected on the initial manufacturer device report number 1220648-2025-26731. h1 type of reportable event was erroneously selected on the initial manufacturer device report number 1220648-2025-26731. h6 health effect - impact code 1802 death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26731.
11055 | 1220648-2025-26732 | 21607703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21607703&pc=OZD | 02/17/2025 | 04/08/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129861 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported shortly after start of the impella 5.5 support to a patient with an acute myocardial infarction/cardiogenic shock, the placement and left ventricle signals both would intermittently go out in the absence of any alarms such as placement unreliable. the decision was made to switch out the automated impella controller (aic) mid-support with report of no harm to the patient. the patient remained on the impella mechanical circulatory support for nine days before support was removed. the patient was not a candidate for any other support, expressed to discontinue the support, and subsequently expired. the direct cause of death is unknown. abiomed's position is the aic is not a main factor in the demise. however, there is not enough information to exclude the device as an associated factor to the event per medical safety review. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of optical signal issue has been completed. data analysis: this failure mode reproduced on the complaint date with the optical bench constantly resetting as optical communication was malfunctioning. this resulted in sometimes a one second drop to -1638.4 mmhg immediately followed by rise to +1638.4mmhg for two seconds before returning to previous range of signal. additionally, the placement signal spiked to 3000 mmhg during this time which is what the user would have seen. pump 570404 was moved over to ic1755 and the failure mode did not reproduce on that console. device analysis: the failure mode was consistently reproduced during testing whenever a pump was connected to the console. the optical bench was replaced with a known good bench and the malfunction did not reproduce. root cause: the loss of placement signal was caused by a consistently resetting optical bench that was malfunctioning. b1 revised from the initial manufacturer device report number 1220648-2025-26732 to reflect both adverse event and product problem.
11056 | 1220648-2025-26733 | 21607876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21607876&pc=OZD | 02/18/2025 | 04/02/2025 |  | 03/14/2025 |  |  | 02/18/2025 | 03/14/2025 |  |  | 12/31/2025 |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438156 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced non-sustained ventricular tachycardia (vt). the p-level was adjusted and staff confirmed the position. upon removal of the impella, severe aortic insufficiency noted on transesophageal echocardiography. the left coronary cusp appeared to be damaged. patient had aortic valve repair. the patient outcome was noted to be critical. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (unites states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter: ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation of heart valve damage and vf/vt/af/at has been completed. the device was implanted on (b)(6) 2025. mitral valve regurgitation with tissue valve and tricuspid ring intervention performed without complication. on 21feb25, device explanted. upon removal of impella (pulled out in one motion with no resistance noted), severe aortic insufficiency noted on transesophageal echocardiography. the left coronary cusp damaged. the pump was returned. the returned product showed no cannula kinks, no catheter kinks, and no other abnormalities. the cause of the heart valve damage and vt/vf was not determined due to lack of clinical details.
11057 | 1220648-2025-26734 | 21608403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21608403&pc=OZD | 02/16/2025 | 04/04/2025 |  | 03/14/2025 |  |  | 02/16/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496345 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced low pump flow and became hemodynamically unstable. the patient was admitted and had mitral valve replacement and was hemodynamically unstable since the surgery. the patient coded the following day after the surgery at night with ventricular fibrillation arrest. the patient was transported to the catheterization lab for a left heart catheterization and impella cp placement. noted was left ventricular ejection fraction was 70% preoperative and 45% postoperative mitral valve replacement surgery with severe mitral regurgitation. the impella was implanted and was running on auto with flows at 3.6. the patient abruptly became hypotensive with systolic blood pressure was in the 60¿s and mean in the low 50¿s. the left ventricle waveform was above the aorta waveform, motor current over 1000, and the flow was high at 4.3 liters per minute with maximum and minimum flows of 4.3/4.0. performance level was reduced to p-1 for crossing the valve and p-0 for explant of the device from the patient. following, the patient was implanted with an impella 5.5 as escalation in therapy and impella rp device for right heart support. the patient was reported to be in stable condition following the event and now on bipella support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of three device manufacturing reports associated with the event and/or implant experience. refer to the following: medical device report for the impella cp serial number (b)(6) with date of the event (b)(6) 2025 and patient identifier ending in (b)(6). (this report) medical device report for the impella 5.5 serial number (b)(6) with date of the event (b)(6) 2025 and patient identifier ending in (b)(6). medical device report for the impella rp serial number (b)(6) with date of the event (b)(6) 2025 and patient identifier ending in (b)(6). | The investigation of low pump flow has been completed. the data was not returned for review. visual inspection found biomaterial inside the pump housing. no other issues were found on the rest of the pump. the root cause of saturated flow was most likely biomaterial. note: changed failure mode fm to saturated flow as the flow increasing to 4 l/min not dropping. d9 date device returned to manufacturer added.
11058 | 1220648-2024-17106 | 21608435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21608435&pc=OZD | 05/23/2021 | 05/21/2025 |  | 03/14/2025 |  |  | 05/23/2021 | 03/14/2025 |  |  | 02/28/2023 |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071129 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the physician discovered a wire perforation to the left ventricle after the impella was placed. this perforation was successfully repaired. the patient was transfused with multiple blood products. additionally, the patient experienced a gastrointestinal (gi) bleed; however, the actions taken to resolve the issue are unknown. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter." ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
11059 | 1220648-2025-26735 | 21608453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21608453&pc=OZD | 02/18/2025 | 04/09/2025 |  | 03/14/2025 |  |  | 02/18/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129883 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the automated impella controller (aic) had a controller failure mode and placement signal was lost. support was continued. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. on the complaint date, before the complaint, impella cp was run for 16 minutes, unplugged then run for 9 minutes and had no malfunctions with this pump. three hours later, the impella 5.5 was used for two minutes before the pump was disconnected and reconnected, after which there were four instances of the logs of placement signal crc errors, twice lasting long enough before resolving to trigger controller error alarm. the pump was switched over to another console and the error did not reproduce. real time logs confirmed that all signals received from optical bench (placement, signal-to-noise ratio, contrast, lamp, gain) were represented as -16384 values due to the crc error. the console was sent back for investigation but the failure mode was not reproduced. the controller error and loss of placement signal is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the automated impella controller software operated as designed.
11060 | 1220648-2025-26736 | 21608706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21608706&pc=OZD | 02/17/2025 | 03/14/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512861 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in postcardiotomy cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support (mcs) and experienced low pump flow and became hemodynamically unstable. the patient was escalated to an impella 5.5 device and also placed on an impella rp device for bipella mcs. the patient would experience hemolysis, low pump flow and subsequently expire. the following day after bipella placement, continuous renal replacement therapy was started with the plan to increase flow as patient tolerates and not removing volume at such time. the patient was noted to be pink warm and dry. impella rp dressing was dry and intact in right femoral vein and the left subclavian impella 5.5 dressing serosanguineous. the patient's urine was noted to be purple and pink tinged. patient did receive methylene blue the prior night it was noted. the possible hemolysis was discussed and suspected. renal failure was noted and also the patient was reported to have systemic organ failure. two units of red blood cells were given. the following day it was noted the patient required greater vasopressor support overnight after crrt started pulling ~100/hr. the patient was required one shock due to abnormal rhythm. chest x-ray showed the impella rp shallow but appeared to be across pulmonic valve. the impella rp was at performance level p-9 and flowing 3.4 l/min, impella 5.5 was at p4 and flowing 2.5. low purge pressure alarms from the impella rp device occurred and were resolved. discussion was made with family, and decision was made to move the patient to comfort care. the patient would expire thereafter. medical safety review of the event noted there is not enough evidence to exclude impella rp as an associated factor in the patient's outcome (demise). the impella rp flows dropped; hemodynamic consequence with the impella rp. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of three device manufacturing reports associated with the event and/or implant experience. refer to the following: medical device report for the impella cp serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). medical device report for the impella 5.5 serial number (b)(6) with date of the event 17-feb-2025 and patient identifier ending in (b)(6). (this report). medical device report for the impella rp serial number (b)(6) with date of the event 17-feb-2025 and patient identifier ending in (b)(6). instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11061 | 1220648-2025-26737 | 21610030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21610030&pc=OZD | 02/17/2025 | 03/14/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025486974 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | Ischemia Stroke | [11 empty] | Required Intervention | Disability | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced a stroke. the patient had his impella discontinued and decompensated overnight. lactate the next morning was 8.96. the physician decided to place the patient back on impella mcs via an impella 5.5. the pump was placed, and team considered left ventricular assist device and reviewed the right ventricle function. two days after start of this support the patient was taken for a head scan and results showed large left middle cerebral artery (mca) stroke. however, since the initial lmca stroke the patient converted to a hemorrhage stroke/subarachnoid hemorrhage on three days later. systemic heparin was withheld. the team and family were discussing potential decisions depending upon the neurological status. approximately 15 days later the device was explanted, and the patient is still noted to have survived the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11062 | 1220648-2025-26738 | 21610055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21610055&pc=PYX | 02/17/2025 | 05/02/2025 |  | 03/14/2025 |  |  | 02/17/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025515406 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Renal Failure | Tachycardia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in postcardiotomy cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support (mcs) and experienced low pump flow and became hemodynamically unstable. the patient was escalated to an impella 5.5 device and also placed on an impella rp device for bipella mcs support. the patient would experience hemolysis, low pump flow and subsequently expire. the following day after bipella placement, continuous renal replacement therapy was started with the plan to increase flow as patient tolerates and not removing volume at such time. the patient was noted to be pink warm and dry. impella rp dressing was dry and intact in right femoral vein and the left subclavian impella 5.5 dressing serosanguineous. the patient's urine was noted to be purple and pink tinged. patient did receive methylene blue the prior night it was noted. the possible hemolysis was discussed and suspected. renal failure was noted and also the patient was reported to have systemic organ failure. two units of red blood cells were given. the following day it was noted the patient required greater vasopressor support overnight after crrt started pulling ~100/hr. the patient was required one shock due to abnormal rhythm. chest x-ray showed the impella rp shallow but appeared to be across pulmonic valve. the impella rp was at performance level p-9 and flowing 3.4 l/min, impella 5.5 was at p4 and flowing 2.5. low purge pressure alarms from the impella rp device occurred and were resolved. discussion was made with family, and decision was made to move the patient to comfort care. the patient would expire thereafter. medical safety review of the event noted there is not enough evidence to exclude impella rp as an associated factor in the patient's outcome (demise). the impella rp flows dropped; hemodynamic consequence with the impella rp. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of three device manufacturing reports associated with the event and/or implant experience. refer to the following: medical device report for the impella cp serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). medical device report for the impella 5.5 serial number (b)(6) with date of the event 17-feb-2025 and patient identifier ending in (b)(6). medical device report for the impella rp serial number (b)(6) with date of the event 17-feb-2025 and patient identifier ending in (b)(6). (this report) | The investigation of hemolysis has been completed. the product and data were not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position as the rp pump was unable to advance into left pulmonary artery. on 2/18, patient went into ventricular tachycardia at 1003, impellas turned to p2 and time of death called at 1013. the cause of the vt storm was unable to be determined due to insufficient clinical details. b1 product problem was erroneously selected on the initial manufacturer device report number 1220648-2025-26738. h6 health effect - clinical code 2095 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26738.
11063 | 1220648-2025-26739 | 21610205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21610205&pc=OZD | 02/18/2025 | 03/14/2025 |  | 03/14/2025 |  |  | 02/18/2025 | 03/14/2025 |  |  |  |  | 03/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600508 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced arrythmia and renal failure during impella cp support. actions taken are unknown. the patient continued on support and eventually was escalated to an impella 5.5 pump. during follow-up it was noted that the patient was in an arrhythmia and coded multiple times support was initiated. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
11064 | 1220648-2025-26686 | 21592708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21592708&pc=OZD | 02/17/2025 | 04/02/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146152 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during this support, the automated impella controller (aic) had an ongoing controller failure alarm and the purge system stopped. the team was able to switch to a back up aic. there was no resulting harm to the patient. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis: cases previous to this complaint had no evidence with relation to this complaint. on (b)(6) 2025, cp pump 561747 was inserted and did not have purge failures while the pump was running. after the pump was put to p-0 but the pump was still connected, console had a controller alarm for battery level below 50% but the console was performing as expected since the console had been running a pump on battery power for 4 hours so there was no power related malfunction. after the pump was disconnected, the console was shut off. 30 minutes later, the console was booted back up and had the same battery level low alarms and also a 418 purge pressure sensor failure for current or voltage out of range. the console was power cycled and plugged into ac power so the battery level low alarm resolved though the purge pressure 418 alarm returned, intermittently resolving and re-triggering. device analysis: the failure mode reproduced intermittently during testing and was able to be isolated to the purge pressure board. that board was swapped out with a known-good and the failure mode would not reproduce. root cause: the cause of the purge pressure sensor alarms was a defective purge pressure sensor board. h8 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26686.
11065 | 1220648-2025-26687 | 21592824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21592824&pc=PYX | 02/17/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025578280 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. it was noted that the physician made several attempts to access the pulmonary artery (pa), however difficulty was noted to be caused by venous cannula from perfusion. left pulmonary artery access was eventually obtained and impella 0.027 wire was positioned. an impella rp flex was then inserted over the wire but was unable to pass the cannula. access to the pa continued to be difficult but was obtained and impella 0.027 wire was positioned in right pa. the impella rp flex was inserted and positioned in right pa under fluoroscopy. the wire was removed, and the impella was slowly started. flows were approximately 3.5 liters at p7 for approximately four to five minutes before it was noted that flows dropped to 1.0 liter. no alarms were noted. motor current was noted to be flat and placement signal was severely dampened. additionally, it was noted that the right ventricle showed a perforation. attention was switched to attempt to close perforation. the impella flow noted to remain at 1.0 liter despite increasing p level. the physician attempted to close the perforation but ultimately was not successful and the patient expired. additional information was received during follow-up that indicated that the physician does not blame the impella for the right ventricle perforation. the pump was unable to be sent back due to the patient passing. of note, there was a suspected volume issue with the patient resulting in potential suction and suboptimal pump flows. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist system with automated impella controller. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
11066 | 1220648-2025-26688 | 21592950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21592950&pc=OZD | 02/18/2025 | 03/23/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1671391 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that due to the issue being identified when turning off the device, there is no information that reasonably suggests interruption and/or delay in impella therapy. no report or indication that the automated impella controller was exchanged or that this issue impacted patient's therapy in any way. use of the device cannot conclusively be associated with the outcome. | The user facility reported that upon turning off the console following support, the selector knob was noted to not function properly. there was no patient involvement. |  |  |  |  | B.5 additional information was received, and abiomed's medical safety officer review was completed. | The investigation for the console (aic) rotary knob issue has been completed. data logs were reviewed finding no evidence to confirm the reported failure mode. the console was returned for evaluation and upon testing, was unable to reproduce the reported failure mode. upon testing, the rotary knob - scrolling up, scrolling down, and selecting functions - worked as intended. the root cause of the selector knob issue was not determined due to the inability to reproduce the issue. corrections to the initial mfr report: g! mfr site name and address updated. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11067 | 1220648-2025-26689 | 21593050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21593050&pc=OZD | 02/18/2025 | 03/19/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597065 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella 5.5 pump and extracorporeal membrane oxygenation placed to support due to cardiogenic shock. after a month of support, the pump had high purge pressures and purge flow alarms. the team consulted with medical affairs and tissue plasminogen activator was added to the purge. the pump remains on for support while decisions are being made of either a left ventricular assist device or a transplant. | [5 empty] | The investigation for the high purge pressure has been completed. no product was returned for analysis. the clinical details mention tissue plasminogen activator (tpa) was added and that resolved the issue. clinical details mention patient also has lower ptt levels and runs of v-tach due to heart failure requiring cardioversion. the data logs confirm purge pressures in the 1000s and align with clinical details of tpa addition and purge pressure trend. there were no motor current spikes or long bag changes. the logs show a gradual rise in purge pressure and corresponding decrease in purge flows starting on (b)(6) 2024. these trends continued until complaint date. the tpa likely controlled the issue but purge pressures remained in the 800s until explant. the root cause of the high purge pressure was most likely due to gradual biomaterial build up as evidenced by data logs. added h6 impact code 4644. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11068 | 1220648-2025-26690 | 21593282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21593282&pc=OZD | 02/17/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571584 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the following day, the patient experienced hemolysis concerns. the labs hemolyzed. the device was explanted approximately two days after the event. no therapy after explant of the device was indicated, and it was noted the patient's heart recovered after cardiogenic shock. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11069 | 1220648-2025-26691 | 21593451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21593451&pc=OZD | 02/17/2025 | 04/07/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597946 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Sense | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. it was noted that the impella 5.5 was backloaded and advanced the vessel was tight and hit was difficult making the turn from anastomosis. wire access was lost so the pump was removed. when pump came out, placement signal was off by greater than 10. the pump was exchanged. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. the product was returned and evaluated. there were no abnormalities found with the 5s parameters. no significant kinks or biomaterial was found. the pump was plugged into the console and it read 6/6 mmhg ao ps, so no abnormalities found. data log review showed the pump was inserted into the body and then removed shortly after as clinical details mentioned. about three minutes after the pump was removed, the placement signal dropped to about -10 mmhg from 0 mmhg. the 5s parameters were unaffected. the root cause of the optical signal issue was not determined since the issue was not reproduced in the lab and no abnormalities were found with the pump. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26691. e.1 fax number and us phone number were added as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26691. h.6 code 4627 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26691 and a new code was added.
11070 | 1220648-2025-26692 | 21593541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21593541&pc=OZD | 02/16/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584448 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced ischemia in the lower limb. although perfusion of the lower limb was attempted from the contralateral side using a sheath-to-sheath technique, this was ineffective and the impella was removed and replaced with an intra-aortic balloon pump. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
11071 | 1220648-2025-26693 | 21593747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21593747&pc=OZD | 02/17/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  | 02/28/2026 |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474572 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump experienced saturated flow coinciding with a spike in motor current during patient support. despite troubleshooting steps being taken, the pump stopped a short time later. the pump was replaced in response to the failure. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk cpi. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11072 | 1220648-2025-26694 | 21594046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21594046&pc=OZD | 02/17/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595521 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported delivery difficulty was encountered during an attempted delivery of an impella cp pump for mechanical circulatory support to a patient with an acute myocardial infarction/cardiogenic shock. upon insertion, the 14fr introducer kinked. the impella was then inserted successfully, however, it then kinked within the sheath. both the sheath and pump were replaced for the planned implant. it was noted there was no patient harm as a result the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11073 | 1220648-2025-26695 | 21594137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21594137&pc=OZD | 02/16/2025 | 04/08/2025 |  | 03/13/2025 |  |  | 02/16/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610440 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and a bridge to transplant. six days after start of the support, the impella was weaned to performance level p-6 due to consistent ventricular tachycardia episodes. no suction alarms were noted with the ventricular tachycardia. echocardiogram confirmed the impella 5.5 placement. nine days later the pump stopped and was unable to be restarted. consequently, the impella 5.5 device was explanted and replaced with a new impella 5.5 device. the patient was reported to be stable following the event. | [5 empty] | The investigation of pump stop and vf/vt/af/at has been completed. the data logs were reviewed and on (b)(6), a sudden alarm 119 triggered a pump stop without a motor current spike. subsequent attempts to restart the pump continued to trigger alarm 119 with mc spikes to 30,000 ma. the pump metrics were all stable leading up to the event. returned product was investigated and all three phases of the motor cables were open lines. there was also a kink in the catheter just distal to the kink protector/red pump plug. the red pump plug was opened, and the breaks in all three lines were found with clear necking just inside of the kink protector. based on data log review and returned product, the root cause of the pump stop was determined to be electrical open. as the necessary information was not provided, the root cause of the vt/vf was not determined." e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26695 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed.
11074 | 1220648-2025-26696 | 21594149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21594149&pc=OZD | 02/16/2025 | 04/21/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548947 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support and five days after start of the support, a fluid leak was observed from inside the sidearm retainer of an impella 5.5 pump during support. the leak was observed after the purge fluid was replaced. there was no patient harm and support continued uninterrupted. the device was successfully weaned and explanted the following day. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. | The investigation of purge leak ¿ pump has been completed. clinical reports leaking from the side arm retainer with no alarms after replacing the purge fluid on (b)(6) 2025, 5 days into the case. leak was found on returned pump at the radial seal. the root cause of the purge leak was most likely sidearm reservoir damage based on leak found at the radial seal on the returned pump. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ d7a was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26696. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26696 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26696.
11075 | 1220648-2025-26697 | 21594228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21594228&pc=OZD | 02/16/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/16/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574337 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support. now day two of support, the impella cp was explanted for escalation in care to an impella 5.5 device which resulted in femoral dissection requiring surgical intervention. however, attempts to implant an impella 5.5 pump were unsuccessful. although the physician used multiple maneuvering techniques, the impella 5.5 device was unable to advance through the calcified area. the patient was eventually implanted with another impella cp device via axillary access and continued on impella mcs therapy for an additional five days until the device was successfully weaned and explanted. ejection fraction fully recovered. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this is one of two device manufacturing reports related to the patient's implant experience. refer to the following: medical device report for the impella 5.5 serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). (this report) medical device report for the impella cp serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6) (under manufacturing report number: 1220648-2025-26681).
11076 | 1220648-2025-26698 | 21594491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21594491&pc=OZD | 02/17/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575783 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump presented high purge pressure alarm during patient support. troubleshooting was completed, but the issue remained. the decision was ultimately made to replace the pump to resolve the failure. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11077 | 1220648-2025-26699 | 21594693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21594693&pc=OZD | 02/18/2025 | 03/23/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1626081 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Durng prep, the automated impella controller (aic) presented an issue with the blue/grey purge disc component. it was falling out and the aic did not detach the disc correctly and the part fell out completely. the aic was replaced. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) purge disc/flag issues has been completed. data logs were reviewed but did not find anything relevant. the console was returned for evaluation finding that the reported broken purge disc retainer was confirmed. the purge disc retainer was confirmed to be broken. purge retainer was completely broken off. corrections to the initial mfr report: h5 should have been no. h8-should have been reuse.
11078 | 1220648-2025-26700 | 21595055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595055&pc=OZD | 02/18/2025 | 04/23/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610281 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding from the access site. the dressing was changed, and one unit of packed red blood cells was administered. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11079 | 1220648-2025-26701 | 21595207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595207&pc=OZD | 02/17/2025 | 08/12/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597110 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [5 empty] | AUTOMTED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was difficulty during the impella ramp up. it kept ¿flipping into mitral valve¿ per the physician. impella was ramped down to p2 and inevitably placement was optimized. additionally, bedside echo measurement was >5.5cm and was giving low flows with continuous suction on lv placement screen. the device was explanted due to the positioning issues. harsh coughing and ambulation were noted contributing factors. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1-updated to adverse event. b2-selected required intervention. h1-updated to serious injury. h6 added clinical code 1888. abiomed received information via an impact study stating that bleeding periprocedural from the surgery as well as from the impella graft site. upon initial assessment, there was a definite relationship to the impella device and procedure. | The investigation of positioning issues and low pump flow has been completed. the failure mode positioning issue was removed as it is the root cause of low flow. the data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in aorta/impella position unknown. no motor current spikes observed in the logs. no other issues were found on the logs. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely due to pump was not in the proper position. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26701as it is unknown if the initial reporter also sent the report to the food and drug administration. as it is unknown if the initial reporter also sent the report to the food and drug administration.
11080 | 1220648-2025-26702 | 21595242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595242&pc=OZD | 02/17/2025 | 03/27/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589237 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support there was high purge pressure. tissue plasma activator was added to the purge. support continued. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of high purge pressure has been completed. the pump was not returned for investigation. the data log indicates a gradual increase in purge pressure over four days, accompanied by a high alarm that was resolved by a decreasing trend over the following two days. additionally, on the explant day, a purge system block alarm was observed, with a three-minute rise in purge pressure and a small spike in motor current. this trend also resolved with a gradual decrease over the course of ten hours. the clinical team also mentioned that the tissue plasminogen activator (tpa) was administered while a high purge pressure alarm was observed, along with monitoring of the motor current trend. the cause of the high purge pressure issue was most likely biomaterial buildup, based on data log review. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26702 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11081 | 1220648-2025-26703 | 21595359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595359&pc=OZD | 02/18/2025 | 03/31/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1409579 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported at the conclusion of cp support on a male patient, the team tried to stop/power down the console and this failed. they were unable to turn off the power. no harm or injury. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation for the console (aic) shutdown/restart issue has been completed. console logs from the reported day of event are consistent with complaint and show that after p-level was set to 0, and pump 490027 was unplugged, a single shutdown request was recorded (operator confirmed request by pressing the button), but the console did not shut down (shutdown script was not executed). while imc log data shows a single shutdown request, ltlog data shows multiple such requests, each unsuccessful. imc logs prior to shutdown showed multiple and frequent interruptions of data streaming, and attempts to reestablish handshake, as well as inability of rtlogger process to compress log files. rlm journal from corresponding time frame was not available for analysis. console was returned for evaluation but issue was not reproduced. based upon previous investigations with similar details, the inability to shut down the console was identified as a software anomaly, that was triggered by erratic behavior of the rlm: frequent unexpected restarts (most likely due to microsd card corruption) resulting in interruptions in data streaming and repeated handshakes.
11082 | 1220648-2025-26704 | 21595519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595519&pc=OZD | 02/18/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601559 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed a hematoma at the access site that required surgical intervention; the hematoma was evacuated. additionally, heparin was discontinued. the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11083 | 1220648-2025-26705 | 21595551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595551&pc=OZD | 02/18/2025 | 04/30/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597645 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Peripheral Nervous Injury | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient had been very slow to wean with impella 5.5 support. they were able to successfully wean the patient for explant. after the explant, the patient was experiencing weakness and loss of sensation in the right arm. this was the same side as the cutdown for the impella. the patient was able to move the arm and pulses were present by doppler. of note, vascular was monitoring the patient in the event that further intervention was required. | [5 empty] | The investigation of nerve compression has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11084 | 1220648-2025-26706 | 21595700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595700&pc=OZD | 02/17/2025 | 04/04/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600033 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. low pump flow occurred. the cp was removed and no clot was observed. the impella was exchanged. no patient harm was reported. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the data logs confirmed the failure mode and clinical narrative. the logs showed after the pump started, during ramp up, multiple motor current spike events occur with a drop in pump flows post the motor current spike. this low flow event remained to the rest of the case. based on clinical description and data review, the root cause of low flow was most likely due to biomaterial ingestion. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26706. e.1 added facility name as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26706. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26706 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11085 | 1220648-2025-26707 | 21595719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595719&pc=OZD | 02/17/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578369 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient underwent percutaneous transluminal coronary angiography to the left circumflex, left anterior descending and ramus arteries and had regain of some coronary flow. however, the patient remained a surgical case. three days following, the patient was successfully weaned and taken to catheterization lab for explant. following successful impella cp explant, a right common femoral artery angiography revealed a thrombus burden at the access site. a mechanical thrombectomy was performed to treat the condition. the patient was reported to be stable following the event and planned for coronary artery bypass graft surgery the next day. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11086 | 1220648-2025-26708 | 21595750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595750&pc=OZD | 02/17/2025 | 04/02/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263046 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The automated impella controller presented button related issues as it was noted that buttons were sticking. there was no patient involved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the automated impella controller rotary knob issue has been completed. according to the logs, on the complaint date, the console was on for 12 minutes, and there was no evidence related to the impella connect issue. according to the rlm logs, on the complaint date, after the initial boot-up of the rlm, there were continuous 'button pressed blinking' logs for more than 15 seconds, indicating that the flight button was pressed, either due to contamination or a manual button press. the continuous button presses likely caused the rlm to enter and exit maintenance mode multiple times. by design, each exit from maintenance mode caused the rlm to reboot. this console was tested. heavy contamination was confirmed on the rlm - flight button side. they were unable to enter maintenance mode by pressing the flight button for 15 seconds due to contamination. upon connecting the rlm to the ic-test network, it remained connected for three hours without any issues, and unable to reproduce the automatic button press, as the contamination had dried out. the root cause of the rlm reboot was contamination. the cause of the contamination was most likely a use issue. a.1 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26708. d.2 revised common device name since the initial manufacturer device report number 1220648-2025-26708 was submitted in accordance with updated procedures. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26708 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6. revised component code as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26708.
11087 | 1220648-2025-26709 | 21595767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595767&pc=OZD | 02/17/2025 | 05/07/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588370 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the insertion/delivery was aborted due to native anatomy. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, insertion aborted due to patient anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26709 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11088 | 1220648-2025-26710 | 21595776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595776&pc=OZD | 02/18/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557320 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced an embolic stroke during support. a thrombectomy was performed to treat the condition and the patient remained off heparin due to concerns of potential hemorrhagic cardioversion. support with the implanted pump continued uninterrupted. it was noted that the physician suspected that the patient was likely under anticoagulated. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11089 | 1220648-2025-26711 | 21595884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595884&pc=OZD | 02/18/2025 | 03/21/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609922 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the tuohy borst failed to lock, allowing catheter to move. the impella cp was removed using side access port and new impella cp was placed. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of positioning issues has been completed. clinical information reported that after placement of the impella cp, they experienced a tuohy borst failure to lock allowing the catheter to move. the returned complaint pump was visually inspected. no cannula kink or catheter kink was observed. no issue or damage observed on the repo unit and tuohy borst. the tuohy borst moved along the catheter normally. the sterile sleeved was noted to be detached. the cause of the positioning issue most likely was use related as the returned pump repositioning unit functioned properly. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26711. d.4 revised serial number, lot number and unique identifier (udi) # as they were entered incorrectly on the initial manufacturer device report number 1220648-2025-26711. e.1 added us phone number and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26711. h.5 revised selection as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26711.
11090 | 1220648-2025-26712 | 21595931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21595931&pc=OZD | 02/18/2025 | 03/21/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the pump would not start after placement. the pump was removed and replaced with another impella cp. no patient harm was reported and the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of pump did not start has been completed. pump and data log were not returned for investigation. as per the given clinical, pump did not start after placed into patient. the cause of the pump did not start issue was not determined since product was not returned and sufficient clinical information was not provided.
11091 | 1220648-2025-26713 | 21596468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21596468&pc=OZD | 02/18/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583086 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the staff inadvertently tripped and fell over the impella cable causing the proximal end of the red plug to come out. in response to the resulting stoppage, the patient was brought to the operating room where the pump was replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11092 | 1220648-2025-26714 | 21597684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21597684&pc=OZD | 02/18/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561720 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. per best practices, the physician utilized reaccess port to assess distal blood flow. no distal blood flow was present down the patients leg. the physician attempted to place antegrade and contralateral distal perfusion but was not successful. cut down was performed; however, distal flow were still not restored. the physician made decision to remove impella. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
11093 | 1220648-2025-26715 | 21599446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21599446&pc=OZD | 02/17/2025 | 03/31/2025 |  | 03/13/2025 |  |  | 02/17/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571474 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and experienced placement signal issues resulting in replacement of the device. during the catheterization it was discovered that the graft to the ostial main and graft to the right posterior descending artery were down. upon insertion of the impella, the pressure measurements appeared distorted, as well as the waveforms. the physician attempted to try to reposition the impella inside of the left ventricle to see if that would help. however, the numbers became more distorted, and the left ventricle waveform disappeared. despite troubleshooting efforts, the decision was made to remove the impella cp and to place a new one. as pumps were being exchanged, it was observed that the circulator had accidentally hung normal saline instead of dextrose 5 percent in water (d5w). a new d5w bag was hung and the impella was primed with a successful insertion of a new device without issues. there was no harm to the patient as a result of the event. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of placement signal issue has been completed. the root cause of the optical signal issue was not determined since no issue was found on returned product and data log were not provided to review the case run activities.
11094 | 1220648-2025-26716 | 21599545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21599545&pc=OZD | 02/18/2025 | 03/31/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612862 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the impella 5.5 was unable to pass through the right subclavian artery. the implant was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to delivered due to calcification in the vessels.
11095 | 1220648-2025-26717 | 21601327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21601327&pc=OZD | 02/18/2025 | 04/01/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601232 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. bedside echocardiogram revealed large pericardial effusion with collapse of right atrium and right ventricle. decision was made to take to operating room for sternotomy and evacuation of pericardial effusion. hemodynamics improved with removal of effusion. the patient continued on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the ventricle perforation has been completed. product was not returned for investigation and data logs were not available for review. the clinical details provided report that once the patient arrived at a second hospital, a baseline echo showed pericardial effusion. there were no signs of active bleeding. no further details were provided about the potential timing or cause of the issue. based on limited clinical details and no logs or product return, the root cause of the ventricle perforation could not be determined.
11096 | 1220648-2025-26718 | 21601697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21601697&pc=OZD | 02/18/2025 | 03/13/2025 |  | 03/13/2025 |  |  | 02/18/2025 | 03/13/2025 |  |  |  |  | 03/13/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1738105 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complaintant reported the automated impella controller (aic) had a "running on batter power" error even while the console was plugged in. this issue occurred during commercial use. there was no patient harm. | [5 empty] | The investigation into the reported issue has been completed. the complaint could not be confirmed in the logs due to the nature of the complaint but there was no ac power connected input on the complaint date, which does not refute the complaint .the failure mode was not reproduced during testing as the console recognized ac power and charged every time it was plugged in. additionally, the batteries were depleted down to 95% and charged back to 100%. the cause of the batteries clinically not charging was not determined since failure mode was not reproduced.
11097 | 1220648-2025-26654 | 21582243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21582243&pc=OZD | 02/13/2025 | 03/25/2025 |  | 03/12/2025 |  |  | 02/13/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610375 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an impella cp implant experienced oozing from their access sites. anticoagulant was stopped, dressing was changed, site was re-sutured, and manual pressure was applied to manage the condition. support continued uninterrupted. the patient was eventually weaned and explanted. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient condition related as the patient was oozing from all invasive lines.
11098 | 1220648-2025-26655 | 21582310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21582310&pc=OZD | 02/13/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/13/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606194 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that after rolling a patient who was implanted with an impella cp pump, an impella in aorta alarm triggered. the patient was hemodynamically stable and the decision was made to explant the pump. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11099 | 1220648-2025-26656 | 21582524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21582524&pc=OZD | 02/14/2025 | 05/13/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1790445 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The automated impella controller (aic) presented a blinking light-emitting diode (led) on the display. both the yelow and blue lights were blinking all the time, non-stop. the aic was removed from use. there was no patient involved. | [5 empty] | The investigation of automated impella controller (aic) - impella connect - usd card corruption has been completed. the console logs did not contain any information relevant to this failure mode. console was tested in and the issue was reproduced. initial boot-up sequence (u-boot) captured over serial terminal connection, revealed continuous reboots of the rlm that was booting up in ¿emergency mode¿. after replacing micro sd card, the issue was resolved and rl11552 was able to fully boot-up. the cause of the issue was not determined since issue could not be reproduced. d9 date device returned to manufacturer added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11100 | 1220648-2025-26657 | 21582661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21582661&pc=OZD | 02/13/2025 | 04/09/2025 |  | 03/12/2025 |  |  | 02/13/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599849 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. a few days into support, the patient experienced persistent ectopy for which the impella was repositioned. a few days later, the patient experienced hemolysis coinciding with an increase cardiac resynchronization therapy and lactate dehydrogenase levels. the pump was bridged to apical protek for ongoing support. the patient outcome was noted to be stable. | [5 empty] | The investigation has been completed. the data logs show that there were some suction alarms on (b)(6) 2025 and impella position unknown alarms on (b)(6) 2025. there was no trend in placement signal or motor current spikes seen. there were no damages seen with the device or biomaterial found. the pump was hemolysis tested and the mih was 11.93. due to lack of clinical information and no problem found with the returned product, the root cause of the hemolysis cannot be determined. as the necessary information was not provided, the root cause of the ectopy was not determined. d.7a added selection as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26657. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11101 | 1220648-2025-26658 | 21582979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21582979&pc=OZD | 02/13/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600521 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis. the urinary output was dark red in foley. additionally, the patient was oliguric. the pump was explanted and an additional impella device was needed. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ section: suction, ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11102 | 1220648-2025-26659 | 21583082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583082&pc=OZD | 02/14/2025 | 03/20/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558092 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp for acute myocardial infarction. it was reported that after successful percutaneous coronary intervention, the sterile drape was removed and the patient's right leg was mottled with absent of distal pulses. the physician decided to remove the impella cp and replaced it with an intra-aortic balloon pump via left leg. the patient survived the procedure. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11103 | 1220648-2025-26660 | 21583288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583288&pc=OZD | 02/13/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/13/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597601 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 60 year old male stemi patient presented to the er and taken to cath were revealed severe mitral valve disease. decision was made to proceed with impella cp placement to unload then reevaluate right ventricle. it was noted that the patient began to decompensate and required 1 round cpr requiring intubation during impella placement. after impella placement the patient became more hemodynamically stable. during support it was reported that the patient went into ventricular tachycardia requiring two shocks and returned to sinus tach. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11104 | 1220648-2025-26661 | 21583340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583340&pc=OZD | 02/16/2025 | 04/03/2025 |  | 03/12/2025 |  |  | 02/16/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592765 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the patient had a hematoma at the access site. heparin was withheld and the hematoma was expressed. support was continued. the patient was eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient movement since hematoma noted after arrival to the unit.
11105 | 1220648-2025-26662 | 21583405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583405&pc=OZD | 02/15/2025 | 05/07/2025 |  | 03/12/2025 |  |  | 02/15/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595825 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Infusion or Flow Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy implanted with an impella cp device for mechanical circulatory support experienced limb ischemia followed by saturated pump flows. the patient would subsequently expire. the patient transferred from an outside hospital for acute chronic heart failure exacerbation. the patient was started on dobutamine and nitroprusside. the following day a right heart catheterization revealed a reduced cardiac index. after discussion between physicians, the decision was made to place impella cp to unload and bridge to a decision. after the right heart catheterization, the right femoral artery was accessed for impella cp implant. it was noted heparin 5000 units was given. activated clotting time was 210 seconds and more heparin 3000 units was given. the pump was started on auto 3.5lpm. peel-away was removed and repositioning sheath was advanced. sleeve and tuohy bost valve was locked into placed. blue suture hub anchored with forward tension sutures. after the case was completed, the patient complained right lower extremity (rle) pain. the extremity was cool to touch and paler than the other leg. doppler to the rle was negative. the re-access port was used to inject contrast into the artery and imaging showed the artery was occluded. at this point, the team decided to attempt an external fem-fem bypass. wiring the right superficial artery proved unsuccessful. the patient still did not have dopplered pulses, but the leg was warmer than before. decision was made to take the patient to the unit and discuss with the team about next steps. vascular surgery was consulted about the patient's potential ischemic leg. the next day, the patient was taken back to the catheterization lab to assess rle flow. aortogram showed sluggish flow in the distal rle. discussion with intensive care unit, cardiology and vascular teams was made with the decision to leave the impella cp in place as there is some flow distal from collaterals. if patient' s condition worsens, then the plan was noted to remove the impella cp and place percutaneous subclavian impella cp. later in the day it was noted there were no pulses in rle, but leg was relatively warm. the patient was going to be worked up for a left ventricular assist device, and possible impella 5.5 implant in a couple days after a decision was made. two days later, the rle was now with palpable pulse and three days thereafter, it was noted patient would go for durable lvad. the patient was noted coughing up bloody sputum/clots and did have some blood in the bowel movement. days later, now eight days after start of the support, heparin was stopped at 1700 due to concerns of heparin-induced thrombocytopenia. argatroban was started. later in the evening, the impella cp motor current spiked and left ventricle waveform became abnormal. min/max and mean flows all 4.2; flow saturation concerns were highlighted. the following day, discussion was made regarding pump failure and the need to urgently exchange the impella cp device. the decision was made to plan for device escalation to an impella 5.5 considering durability and benefit of axillary placement. it was additionally noted there was concerns for infection- probable pneumonia and ruling out aspiration as cause. the patient was pan cultured to cover all bases. hemoglobin continued to be low, and patient continued to have bloody stools (although less frequent per the nurse). patient was considered for gastrointestinal (gi) scope procedure, but the gi physician was reluctant as the patient was on argatroban. however, the following day, the patient was noted to have expired; it was indicated that care was withdrawn. further details around the patient's demise were not provided and it does not appear the patient was able to be escalated to the impella 5.5 and the flow saturation remained unresolved. details surrounding the patient's demise are unclear other than care was withdrawn. however, given the impella-related events ([it appears] partially resolved limb ischemia and saturated flow) there is not enough information to exclude the impella cp as an associated factor to the event. | [5 empty] | The investigation of limb ischemia and saturated flow has been completed. no product returned. reported spike in motor current with saturated flows around 2/23 19:30. heparin was stopped prior at 2/23 17:00 due to concerns of hit and argatroban started shortly after. data logs show dips in purge pressure and spikes in pump and purge flow beginning around 2/23 2:00am and continue till a spike in motor current at 2/23 19:30. the motor current spike is followed by an upward shift in flow, motor current, and purge pressure with a downward shift in purge flow. the root cause of the saturated flow was most likely biomaterial since there was a motor current spike followed by an upward shift in flow, motor current, and purge pressure with a downward shift in purge flow in data logs. the root cause of the limb ischemia was not determined since product was not returned and insufficient clinical information. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-26662. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11106 | 1220648-2025-26663 | 21583428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583428&pc=OZD | 02/16/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/16/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560142 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and two days after the start of the support, the impella cp device was accidentally pulled into the aorta. the patient tolerated the minimal support. ejection fraction was up to 40%, and the decision was made to explant the impella cp device. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11107 | 1220648-2025-26664 | 21583531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583531&pc=OZD | 02/16/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/16/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610416 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support a bleed at the access site was noted requiring blood transfusion. the patient presented to the emergency room with chest pains, was admitted and diagnosed with st-segment elevation myocardial infarction and was taken to catheterization lab. the decision was made for impella cp placement, which was successfully completed. the patient was transferred to the unit on support. later this day, bleeding at impella site and right internal jugular site was noted. the patient was given two units of packed red blood cells, albumin bolus and vitamin k. heparin dripped was turned off. it was noted the impella site was soft with no evidence of hematoma. the patient continued on impella mcs. three days later the patient underwent percutaneous coronary intervention (coronary angioplasty and stent placement) and was noted to have tolerated the procedure well. the impella access site was clean, dry and intact with no evidence of bleeding or hematoma. two days thereafter, the impella cp device was successfully weaned to performance level p-2 without any hemodynamic compromise. heparin was stopped, the device was explanted without incident, and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11108 | 1220648-2025-26665 | 21583545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583545&pc=OZD | 02/17/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/17/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610777 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp for acute myocardial infarction. it was reported that roughly an hour after impella cp explant, the patient had a change in mental status and a stroke call was performed. the patient was taken for an emergent thrombectomy to treat the condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11109 | 1220648-2025-26666 | 21583677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583677&pc=OZD | 02/17/2025 | 04/11/2025 |  | 03/12/2025 |  |  | 02/17/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1287971 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller failure error while in use on a patient with an impella 5.5 pump for cardiogenic shock. the pump was rapidly weaned and explanted. no harm or injury was noted. the patient survived the procedure. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console logs from the reported day of the event are consistent with the complaint. the logs showed multiple instances of the following error "impella stopped (currents mismatch) - alarm issued event #119", "impella stopped. retrograde flow. - alarm issued event #18, and "pmd has detected a pump speed deviation - alarm issued event #138". long term and real time logs showed a spike in motor current after it drops to zero along with the pump flow and motor speed drops to zero. the console was returned. a visual inspection of the internal circuitry of the console did not reveal any issues. the console was tested with a known good pump and purge system, and no anomalies were observed during testing. the root cause for the pump stop was not related to any problems with the aic. no issues were found within the aic.
11110 | 1220648-2025-26667 | 21583807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583807&pc=OZD | 02/15/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/15/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582865 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMTED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced red urine and hemolyzed labs. the pump was explanted and escalated to an impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11111 | 1220648-2025-26668 | 21583894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583894&pc=OZD | 02/16/2025 | 04/01/2025 |  | 03/12/2025 |  |  | 02/16/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599801 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and following day, the impella 5.5 pump came back across the aortic valve during patient support. at performance level p-7, the placement signal showed aortic wave, the left ventricle (lv) signal showed lv wave. the pump alarmed, placement signal unknown at p-7 and the mean motor current was around 350mah. a transesophageal echocardiogram revealed the pump was completely in the aorta. when the performance level was dropped to p-2, the alarmed triggered, pump completely in the aorta. the pump was removed and replaced as a result of the event. the patient was noted to be stable following the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. | The investigation of positioning issues has been completed. the reported placement signal showing aortic waveforms and confirmed pump was in aorta with imaging. pump was visually inspected and no damages or abnormalities were found. the root cause of the positioning issue was not determined due to insufficient clinical detail.
11112 | 1220648-2025-26669 | 21583910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21583910&pc=OZD | 02/16/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/16/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025598930 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Renal Failure | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support it was noted that the right lower extremity was cold and pale. later that day, it was noted that pulses were still absent but the skin was warm and normal in color. additionally, the nurse reported hemolyzed labs and lactic was not clearing so continuous renal replacement therapy was started. ultimately the patient was escalated to an impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11113 | 1220648-2025-26670 | 21584031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21584031&pc=OZD | 02/17/2025 | 03/18/2025 |  | 03/12/2025 |  |  | 02/17/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575783 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after initiating support, high purge pressure alarm was triggered. it was not possible to correct the issue with troubleshooting measures. the impella was exchanged. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of high purge pressure has been completed. the returned product was inspected and there was biomaterial was observed in the purge gap. the pump had no other physical abnormalities. the pump was run in the lab after soaking and clearing the biomaterial and the high purge pressure did not reproduce. the data logs from the field show a gradual rise in purge pressure upon pump start for about 12 minutes before alarm was triggered. there was also a decrease in purge flows. the high purge pressure alarm resolved. purge pressure values remained elevated. later during the case there was a motor current spike with speed deviation and a shift in pump flows. the issue did not resolve. the root cause of the high purge pressure was determined to be biomaterial as issue did not reproduce upon clearing and running in the lab and as evidenced by data log. h5 selection was inadvertently entered incorrectly on the initial manufacturer device report number: 1220648-2025-26670.
11114 | 1220648-2025-26671 | 21586301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21586301&pc=OZD | 02/14/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610320 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 53 year old female with a impella cp for acute myocardial infarction. it was reported that after gaining access for impella placement, the staff was unable to pass the introducer through the iliac. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿
11115 | 1220648-2025-26672 | 21586968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21586968&pc=OZD | 02/14/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601560 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 56 year old male patient was on impella 5.5 support and during support, there were positioning issues. the pump was eventually explanted due to positioning issues. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
11116 | 1220648-2025-26673 | 21587521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21587521&pc=OZD | 02/17/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/18/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2026613164 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 25 YR |  | Male |  | Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 25 year old male patient with impella 5.5 pump support placed for over a month for a due to cardiogenic shock. this 2nd pump was placed as a replacement and the patient suffered a stroke on the pump. the patient had sodium bicarb in the purge and historically heparin in the systemic iv drips. the pump remains on for support despite the stroke and patient has survived to date. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11117 | 1220648-2025-26674 | 21587663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21587663&pc=OZD | 02/13/2025 | 03/17/2025 |  | 03/12/2025 |  |  | 02/13/2025 | 03/12/2025 |  |  | 12/31/2025 |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454535 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported on a 86 year old male on an impella cp to support a stemi patient. when they explanted the pump due to bleeding and being supratherapeutic with ptt >160. the team explanted, held manual pressure and femstop. the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely anticoagulation management since they explanted the pump due to bleeding and being supratherapeutic with ptt >160 corrections to the initial mfr report: h5 should have been yes | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11118 | 1220648-2025-26675 | 21587688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21587688&pc=OZD | 02/14/2025 | 03/31/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432077 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during day 2 of support, the automated impella controller presented a controller failure alarm. the controller was swapped. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) controller failure-red alarm has been completed. data logs were reviewed finding the purge communication fails and never recovers. this triggers a pud communication loss controller failure as noticed in the complaint and the pud is power cycled multiple times. during future boot ups, the console is not able to complete boot up due to the lack of purge communication and continuously triggers the sw watchdog. the aic was returned for evaluation. this failure was reproduced repeatedly during testing. replacing the pud pcb, the failure mode resolved. the original pcb was examined closer outside of the console and capacitor c9.1 was found to be broken and a fluid from an unknown source had leaked extensively over the board including the communication trace which is what caused the loss of communication. the controller failure was triggered by a loss of communication to the pud which was caused by capacitor c9.1 breaking and an unknown capacitive liquid leaking onto and corroding the pcb. corrections to the initial mfr report: d2 common device name updated h5 should have been no.
11119 | 1220648-2025-26676 | 21587861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21587861&pc=OZD | 02/14/2025 | 04/11/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2026613189 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 60 year old male on ecmo support who was having an impella 5.5 pump placed. the pump was unable to be delivered. the team was unable to deliver across the valve. a new 5.5 pump was placed. the patient survived the procedure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. reported having difficulty getting the pump past the aortic valve. a guidewire was used. implant was aborted due to vascular anatomy and replaced. there were no delivery issues with the replacement pump. catheter kink found between motor housing and 20 mark on the returned pump. no other damages or abnormalities found. the root cause of the delivery issues was most likely patient condition related since delivery was aborted due to vascular anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26676 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name, address, country and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26676. h6 investigation conclusions code 4315 was erroneously on the initial manufacturer device report number 1220648-2025-26676.
11120 | 1220648-2025-26677 | 21587872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21587872&pc=OZD | 02/14/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590631 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) the impella cp device implant was uneventful; however, during transition of the patient off the catheterization lab table, the sterile sleeve broke (no clear reason could be noted per the reporter) and impella pulled back across aortic valve into descending aorta (likely related to getting caught on table during move). the patient started to decompensate. the physician returned to the catheterization lab and made a quick decision to replace the impella cp with a new device instead of re-advancing due to lack of sterility of catheter after sterile sleeve broke. a new impella cp device was eventually implanted and the patient transferred to the unit on impella continued mcs. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11121 | 1220648-2025-26678 | 21587993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21587993&pc=OZD | 02/14/2025 | 03/12/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610798 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump came out of position within the left ventricle and this malposition caused the team to remove and replace the pump. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
11122 | 1220648-2025-26679 | 21588116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21588116&pc=OZD | 02/14/2025 | 03/28/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605291 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 72 year old female patient on an impella cp pump for support due to acute myocardial infarction. the support was on day 2 when the pump site was seen to be oozing at the groin. the team redressed and applied angle matching. the team additionally removed the heparin and gave instead sodium bicarb in the purge. the pump remained on for 6 days until escalated to the 5.5 pump. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time (act) values, heparin was stopped to achieve hemostasis as per the clinical description. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26679 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26679. h5 selection was erroneously selected on the initial manufacturer device report number 1220648-2025-26679. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11123 | 1220648-2025-26680 | 21588156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21588156&pc=OZD | 02/15/2025 | 04/15/2025 |  | 03/12/2025 |  |  | 02/15/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025603662 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 58 year old female patient was on impella 5.5 support and during support, there was high purge pressure. purge flow decreased and the pressure increased. staff added tissue plasma activator to the purge and support continued. | [5 empty] | The impella device will not be received from the customer.. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. clinical details reported that a high purge pressure event occurred. the purge line showed no abnormalities, and replacing the purge cassette did not resolve the issue. tissue plasminogen activator was added to the purge and the issue was resolved. the patient was on support for five more days without any further reported incidents. data logs were not available for review and the pump was not returned for investigation. since neither data logs nor product were able to be investigated, the root cause of the high purge pressure could not be determined. d.6b revised explant date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26680. e.1 revised first name and address line 1 as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26680. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26680 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name, address, country, and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26680.
11124 | 1220648-2025-26682 | 21588318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21588318&pc=OZD | 02/16/2025 | 04/01/2025 |  | 03/12/2025 |  |  | 02/16/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580387 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 52 year old female on a impella 5.5 pump experienced high purge pressure and triggered purge pressure blocked alarms during patient support. a tissue plasminogen activator (tpa) infusion was performed but the alarm persisted. the decision was ultimately made to replace the device with another impella 5.5 pump. there was no patient harm. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the cut pump was returned for investigation. biomaterial was observed on the purge gap and outflow cage. since the pump was cut, priming was tested on both sides of the cut. no issues were noted during priming on the side opposite the motor. the motor side purge line was tested, and no purge was seen coming out, which indicates a blockage inside the motor housing. the data log showed a small motor current spike followed by a gradual increase in purge pressure within eight minutes. the purge pressure remained high with a partially resolving trend during the rest of the case. according to the clinical information, tissue plasminogen activator (tpa) did not resolve the high purge pressure issue. the cause of the high purge pressure issue was biomaterial injection, based on data log review and returned product investigation. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26682. d.6b explant date was revised as it was previously entered incorrectly on the initial manufacturer device report 1220648-2025-26682. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26682. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26682 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26682. h.6 codes 3233 and 11 were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26682. they were added to this report but since the initial report has been submitted these codes have changed as the investigation has been completed now. code 4315 was previously entered incorrectly on the initial manufacturer device report 1220648-2025-26682.
11125 | 1220648-2025-26681 | 21588323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21588323&pc=OZD | 02/16/2025 | 04/04/2025 |  | 03/12/2025 |  |  | 02/16/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597756 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during explant for escalation in care to an impella 5, a femoral dissection was noted to have occurred and required surgical intervention. a decision was made to perform a left heart catheterization and upon injection to the left system, the patient's blood pressure dropped. angiogram showed "severe" left main disease. the decision was made to place the impella cp for stabilization. the impella cp device was implanted and percutaneous transluminal coronary angioplasty to the left main artery was completed to restore flow. the patient stabilized after restoring of flow and was sent to the until for monitoring until transport to an outside hospital for urgent coronary artery bypass graft surgery. the patient was subsequently placed on extracorporeal membrane oxygenation therapy. now day two of support, the pump was explanted for escalation in care to an impella 5.5 device. upon explant, perclose x2 failed. bleeding at the site was observed. when the artery was exposed, the physician noted a femoral artery dissection had occurred. a surgical cutdown with vessel repair was completed. contributing factor to the event was noted as patient obesity and high stick. the impella 5.5 would not be implant as the case was aborted due to the inability to pass a calcified area. the patient was eventually implanted with another impella cp via axillary access and continued on impella mcs therapy for an additional five days until the device was successfully weaned and explanted. ejection fraction fully recovered. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the pump was returned for investigation. there were no abnormalities observed on the pump. as per clinical information patient vessel was dissected likely due to failed perclose device closure. therefore, the root cause of the vascular damage issue was most likely use related due to failed perclose device closure post explant. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26681 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. this is one of two device manufacturing reports related to the patient's implant experience. refer to the following: medical device report for the impella cp serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6). (this report) medical device report for the impella 5.5 serial number (b)(6) with date of the event 16-feb-2025 and patient identifier ending in (b)(6).
11126 | 1220648-2025-26683 | 21588374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21588374&pc=OZD | 02/18/2025 | 04/08/2025 |  | 03/12/2025 |  |  | 02/18/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606204 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 65 year old male patient on impella cp support and during the implant, the cp came up and over the arch the cannula was noted to be buckling. the device was placed into mid left ventricle space and slack removed but the cannula remained kinked. wire was removed and impella was started but flows of 1.0lpm noted. at this point in time it was recommended to pull the impella and replace it. the pump was explanted and replaced. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the data was not returned for review. visual inspection found a kink on cannula about 15 mm from of cage and cannula bonding site. no other issues were found on the rest of the pump. the root cause of low flow was most likely kinked cannula due to patient condition. the cause of the product damage (cannula kink) issue was a kink likely due to patient condition since the patient had aortic stenosis causing resistance during/post insertion. product damage (cannula kink) failure mode (fm) added per new sop updates to ha coding guide and based on reported information on the event/failure. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26683 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26683.
11127 | 1220648-2025-26684 | 21588405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21588405&pc=OZD | 02/18/2025 | 03/18/2025 |  | 03/12/2025 |  |  | 02/18/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607530 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 46 year old male patient was on impella cp support and during support, the patient had access site bleeding. the peel away sheath was left in place and this led to oozing at the access site. heparin was withheld, blood products were given to the patient and the pump was explanted. the patient survived the procedure | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was most likely use issue since the peel away sheath was left in from transferring facility leading to continuous oozing around the site. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26684 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26684. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
11128 | 1220648-2025-26685 | 21588954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21588954&pc=OZD | 02/14/2025 | 04/25/2025 |  | 03/12/2025 |  |  | 02/14/2025 | 03/12/2025 |  |  |  |  | 03/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600671 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a patient in cardiogenic shock from an acute myocardial infarction implanted with an impella cp device for mechanical circulatory support (mcs) experienced positioning issues that resulted in explant with replacement of the device. the patient would subsequently expire. the patient presented to the emergency department with chest pain, shortness of breath and hypotensive. the patient was on two pressors and had an ejection fraction of 15%. of note, the patient declined cardiothoracic surgery three years prior. the patient was transferred to the catheterization lab for an impella cp device placement and possible percutaneous coronary intervention. the impella cp device was implanted via the right femoral artery without incident. catheterization revealed severe left main artery disease. the patient now on impella mcs was chest pain free and normal electrocardiogram. the physician made the decision not to intervene and consult with cardiothoracic surgery. the patient was sent to the unit on support and vomited on arrival. it was noted approximately two days after start of the support, the pump was pulled back into aorta and repositioned at bedside. the position was questionable, however. the patient was transported out to a heart failure hospital and on arrival, the pump positioning was still questionable. multiple attempts, therefore, were made to reposition, and all were unsuccessful. the patient was brought to the catheterization lab for repositioning under fluoroscopy; however, this attempt too was unsuccessful. the decision was made to explant and replace with a new impella cp device which was completed for continuation of therapy. however approximately 1.5 days later the patient expired on support. no further details around the demise were provided. although the impella cp device was explant and replaced prior to the patient's demise association cannot be excluded given the positioning difficulty encountered and potential compromise to circulatory support. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. reported the patient vomited which moved the pump out of position with impella in aorta alarms. the pump was pushed back into the ventricle without the guidewire but unable to reposition pump using echo and under fluoro. pump was removed. the root cause of the positioning issue was most likely patient movement since the improper positioning occurred when the patient vomited. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26685 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11129 | 1220648-2025-26611 | 21571371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21571371&pc=OZD | 10/05/2024 | 03/11/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 61 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A publication was received titled "left atrial appendage cannulation for left ventricular unloading in a patient with ventricular thrombus on extracorporeal life support," which indicated an impella cp device was implanted for mechanical circulatory support and was prematurely explanted due to insufficient flows during patient support. the patient presented with normotensive cardiogenic shock after subacute anterior wall myocardial infarction deteriorated with pulmonary edema and ventricular fibrillation, requiring venoarterial extracorporeal life support under ongoing cardiopulmonary resuscitation. an impella cp device was placed for left ventricular unloading. however, it was unable to generate flow and was removed. a large left ventricle thrombus was detected as the cause for insufficient impella flow. for urgent left ventricular unloading, a vent cannula was placed via the left atrial appendage through a thoracotomy to bridge the patient to total artificial heart implantation. however, intraoperative transesophageal echocardiogram showed resolution of the left ventricle thrombus, enabling to change the strategy to left ventricular assist device implantation only, which was performed successfully. the patient made a full recovery and was noted to have done well post the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." literature citation: schaefer, a.-k., wiedemann, d., heinz, g., riebandt, j., & zilberszac, r. (2024). left atrial appendage cannulation for left ventricular unloading in a patient with ventricular thrombus on extracorporeal life support. https://doi.org/10.21203/rs.3.rs-5210610/v1.
11130 | 1220648-2025-26612 | 21571451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21571451&pc=OZD | 02/12/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580349 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the insertion of an impella 5.5 device for escalation to a patient initially implanted with an impella cp was unsuccessful. the device via the axillary artery was unable to pass or advance due to the patient's anatomy. the implant of the impella 5.5 device was, therefore, aborted. the patient's status outcome following the event was noted as unknown. however, there was no report or indication of adverse effects to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11131 | 1220648-2025-26613 | 21571497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21571497&pc=OZD | 02/12/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599682 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient suffered from a left ventricle free wall perforation during impella 5.5 support when the heart was lifted for ventricular septal defect repair. the perforation was repaired and support was continued with the implanted pump. the patient's outcome at time of explanted was survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11132 | 1220648-2025-26614 | 21571660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21571660&pc=OZD | 10/17/2024 | 03/11/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576537 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient suffered ventricular tachycardia during impella cp support. the condition was believed to have a possible relationship with the use of the impella. the patient survived the condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11133 | 1220648-2025-26615 | 21571772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21571772&pc=OZD | 02/13/2025 | 03/19/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585982 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on an impella 5.5 pump for support. the patient was admitted in for high risk surgery. the pump was placed but after two days was found to have high purge pressure which the team troubleshot with tissue plasminogen activator. the pump remained on for support for almost 20 days. the patient was not a candidate for transplant and so the care was withdrawn and the patient expired. | [5 empty] | The investigation for the high purge pressure has been completed. no product was returned for analysis. the data logs do not show any motor current spikes but there was a gradual rise in purge pressure before alarm was triggered. there were continuous suction alarms and an effect (decrease) in purge flow suggesting biomaterial build up. the high purge pressure alarms persisted for about an hour before purge pressure values returned to normal levels. this aligns with clinical detail of tissue plasminogen activator (tpa) being added to the purge and alarms resolving after addition. there were no other high purge pressure alarms for the remaining duration of the case. the root cause of the high purge pressure was most likely biomaterial as evidenced by log analysis and tpa resolution. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist, section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11134 | 1220648-2025-26616 | 21571845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21571845&pc=OZD | 02/13/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1850881 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported an automated impella controller (aic) would not charge while in use on a patient. the battery was not charging and the aic was running on the battery while plugged in. the aic was replaced and no harm was noted. | [5 empty] | The investigation for the automated impella controller battery charging issues has been completed. the console logs from the reported complaint date were partially consistent with the complaint description. the only slight indicators of aleternating current (ac) power detection issues observed in the logs were some unusual toggling of ac power detection. the console was tested. the reported charging issue was able to be reproduced. the console would not charge and detect ac power. the power cord was inspected but no issues were observed. the fuses were intact. the voltage output from the power supply to the power battery manager was measured and was getting 0 volt. after swapping the power supply with a known good one, the charging issue resolved. the console was able to charged and detected ac power as normal with the known good power supply. the root cause of the reported charging issue was determined to be a defective power supply.
11135 | 1220648-2025-26617 | 21571895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21571895&pc=OZD | 02/12/2025 | 04/23/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559845 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during the impella cp insertion on a 60-year-old male, the physician noted clot within the impella sheath. the physician declined to swap the sheath for a new one and decided to aspirate it all out. the impella was then inserted successfully. it was further reported that the patient started to developed hemolysis. plasma free was over 100. the impella was repositioned with improvement of hemolysis. the patient survived the procedure. | [5 empty] | The investigation of hemolysis and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26617. h6 code 4582 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26617. new codes were added to health effect -clinical code. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿
11136 | 1220648-2025-26618 | 21572024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572024&pc=OZD | 02/12/2025 | 09/30/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597929 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMTED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after the impella was implanted, the patient went into ventricular tachycardia (vt). the patient was defibrillated. the pump was pulled back from the septal wall and the vt resolved. support continued. additional information noted that the patient's family decided to withdraw care. the patient subsequently expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant also reported that there were placement signal issues. pump position was confirmed. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue and cardiac arrythmia. the pump was not returned for investigation. the data log shows the signal-to-noise ratio (snr) trending down after the pump was remapped to p-9, with several improper position alarms during the start of the case. no known optical failure pattern was confirmed from the data log review. the cause of the optical signal issue was not determined since product was not returned. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. b7 other relevant history, including preexisting medical conditions updated.
11137 | 1220648-2025-26619 | 21572026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572026&pc=OZD | 02/14/2025 | 04/03/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585507 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and had dark colored urine and there was a concern for hemolysis. placement on echocardiogram appeared to be optimal, at 3.5 centimeters. repositioning was attempted to resolve the hemolysis. no plans on attempting to escalate due to the patient having severe axillary stenosis. the patient was successfully weaned and explanted. | [5 empty] | The investigation of hemolysis has been completed. the impella device and data logs were not returned for evaluation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11138 | 1220648-2025-26620 | 21572120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572120&pc=OZD | 02/14/2025 | 05/02/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  | 01/31/2026 |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464453 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of an impella cp device in a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful. the device was unable to be delivered due tortuous anatomy and the case was aborted. the patient was unable to receive impella mechanical circulatory support. the patient was reported to be stable following the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because of hyper tortuous fem arteries of the patient and the insertion was aborted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11139 | 1220648-2025-26621 | 21572172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572172&pc=OZD | 02/12/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025605778 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 18 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER, 0.035 EXCHANGE WIRE | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. 0.035 exchange wire was used despite it not being recommended for direct insertion. the guidewire splintered the aortic valve, leading to severe ventricular deficiency. the implant was aborted. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11140 | 1220648-2025-26622 | 21572274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572274&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  | 02/28/2026 |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494886 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced limb ischemia which resulted in the pump being explanted. the patient went into cardiac arrest at home and was transported via emergency medical services to the emergency department where the patient coded four additional times. return of spontaneous circulation was achieved and thrombolytic therapy was administered. the decision was made to wait three hours before completing a catheterization due to thrombolytic therapy. the right heart catheterization was eventually completed, and the impella cp was implanted pre-percutaneous coronary intervention to the right coronary artery where two drug-eluting stents were implanted. dobutamine was weaned, and the patient decompensated requiring cardiopulmonary resuscitation for a short time until drips were restarted. the patient was made stable for transfer to an outside facility and underwent extracorporeal membrane oxygenation placement the following day after start of impella mcs, however, it was reported the patient lost pulses in both feet due to being clamped down for a long period of time. consequently, the impella cp device was explanted. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
11141 | 1220648-2025-26624 | 21572410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572410&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608494 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient's lower extremities were cold to the touch during impella cp support. a doppler study showed significantly decreased blood flow to the bilateral lower extremities. it was documented that both legs had shutdown and a bilateral amputation was planned for the patient. fasciotomies on both legs were performed, but the decision was made by the family to withdraw care. an amputation was not performed and the patient passed away. there is not enough evidence to disassociate the patient's passing with the use of the impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11142 | 1220648-2025-26625 | 21572564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572564&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608882 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced a hematoma requiring blood products. the patient would subsequently expire. st-segment elevation myocardial infarction was confirmed, and the patient had two stents placed. on the way to the unit, the patient decompensated and was brought back to the catheterization lab for impella cp device placement. the impella cp device was placed via the right common femoral artery without incident. the patient was transferred to the unit on impella mcs. following the implant, a hematoma at the access site was noted. blood products were required and manual pressure and pressure dressing were applied. care goals were discussed with the family and the patient was made do not resuscitate with care withdrawn. the patient would expire, thereafter. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11143 | 1220648-2025-26626 | 21572639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572639&pc=OZD | 02/11/2025 | 09/22/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025598360 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 25 YR |  | Male |  | Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported that the placement signal was unstable during support. positioning was checked via echocardiography, and support continued uninterrupted. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. a purge system blocked alarm was encountered, for which a lysis protocol was initiated. there was no patient harm. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement optical issue description. corrections that were made from the initial manufacturer device report number 1220648-2025-26626. b7 other relevant history, including preexisting medical conditions updated. d10 therapy dates updated. g2 report source foreign updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11144 | 1220648-2025-26627 | 21572678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572678&pc=OZD | 02/15/2025 | 04/03/2025 |  | 03/11/2025 |  |  | 02/15/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610455 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and subsequently developed a hematoma of the access site. heparin was held, pressure was held for 30 minutes, and blood products were given. a computed tomography angiogram showed the bleeding stopped. the patient was noted to be in stable condition following the event. the device was successfully weaned and explanted the following day. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time values. heparin was stopped to achieve hemostasis as per the clinical description.
11145 | 1220648-2025-26628 | 21572741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572741&pc=OZD | 02/13/2025 | 03/27/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572090 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the pump stopped on day nine. the pump was explanted and anticoagulation was stopped. the physician commented that the pump was likely clotting off. motor current doubled and the left ventricle placement screen appeared incorrect. there were plans to escalate the patient to impella 5.5. | [5 empty] | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. the data logs were reviewed, and there was a temporary pump stop. leading up to the stop, there was a slight trend up in the motor current. then, the motor current (mc) spikes to 1500 ma along with a speed deviation. it also caused a subsequent rise in the purge pressure to ~600's without triggering any alarms. the pump is able to start back up on its own with saturated flows, but several other pump stops follow. the pump is ultimately able to continue running for 16 hours before the case is ended. for this reason, the failure mode was changed to temporary pump stop. returned product was investigated and the pump passed electrical testing. visual inspection revealed a large purge gap, indicating bearing wear. the pump stop occurred on the 9th day of use. per design trace matrix the impella cp's design life is 8 days. clinical details report the patient was not on anticoagulation for several days, however, there were no signs of biomaterial on returned product and the gradual rise in mc followed by a spike is more likely to be due to bearing wear. based on returned product and data log review, the root cause of the temporary pump stop was determined to most likely be bearing wear.
11146 | 1220648-2025-26629 | 21572772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572772&pc=OZD | 02/12/2025 | 04/30/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595451 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Perforation | [12 empty] | Required Intervention | Life Threatening | Disability |  |  |  | COREVALVE | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient was admitted to hospital for nausea, diarrhea, mental status changes, indigestion. while on the unit, the patient had chest pain and ekg changes and was taken to the cardiovascular lab where in-stent restenosis was revealed. during this intervention the patient became unstable requiring intubation and the decision made to place an impella cp device. the impella cp was implanted without incident. however, following the implant, the pump stopped multiple times due to interference with the struts of the corevalve. the pump was restarted successfully each time with both motor current and flow rates within normal limits. the device was manipulated to resolve the issue. as support continued, suction alarms appeared, and the patient's hemoglobin levels dropped. an echocardiogram was performed and showed a pericardial effusion. a pericardiocentesis was performed and removed 250ml of blood. it was unknown if the pump caused a perforation. the physician believed the pump was causing too much interference, and the impella cp pump was removed. the patient coded shortly after but was documented to have survived the event. no further updates on the patient's status post the event were provided. | [5 empty] | The investigation of temporary pump stop, ventricle perforation, and low pump flow has been completed. product was returned and evaluated. no biomaterial was observed in the cannula or near the impeller. the corners of the impeller blades were damaged. the edges of the blades were rough in appearance and one blade was bent. there was also evidence of a kink on the pump cannula. the pump was run in a bucket and low flow was not reproduced. clinical details mention interference with the patient's transcatheter aortic valve replacement (tavr) to be a possible cause of the temporary pump stops. after pump startup, there are multiple motor current spikes up to 1500 ma and 4 temporary pump stops due to high motor current within 15 minutes of pump startup seen in the data logs. the pump was able to be restarted each time. the returned product had visible damage to the impeller. the root cause of the temporary pump stops was most likely interaction with another device since the impeller was damaged on the returned product, there are motor current spikes seen in the data logs, and the clinical details mention the potential interaction between the temporary pump stops and the patient's tavr. clinical details state that patient's blood pressure (bp) dropped while suction alarms occurred. correct position was verified with fluoroscopy, and echocardiogram revealed pericardial effusion. angiogram was not repeated to confirm possible perforation. there was evidence of a cannula kink on the returned product, but the timing/cause of the possible perforation is unknown. the root cause of the ventricle perforation was not determined since insufficient clinical details were provided. the returned product had a damaged impeller and evidence of a cannula kink, but no biomaterial found and low flows not reproduced in the lab. the root cause of the low pump flow was not determined since low flow was not reproduced in the lab and lack of clinical information. d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26629 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-26629. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: precaution impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11147 | 1220648-2025-26630 | 21572812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572812&pc=OZD | 02/13/2025 | 09/30/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591479 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | AUTOMATED IMPELLA CONTROLLER | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it following 19 hours of impella support, the patient decompensated overnight and went into ventricular tachycardia and arrested. team was unable to resuscitate the patient. the patient expired. | The complainant also reported a placement signal not reliable alarm after the impella initiation. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the pump was not returned for investigation. data logs were provided. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-26630. d10 concomitant medical products and therapy dates updated.
11148 | 1220648-2025-26631 | 21572830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572830&pc=OZD | 11/14/2024 | 03/17/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568574 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported that a study documented a report of a 64-year-old male patient who experienced short runs of ventricular tachycardia during impella 5.5 support. the condition was believed to have a definite relationship to the impella device, impella procedure, and surgical procedure. no treatment was required and support continued uninterrupted. | [5 empty] | The investigation of ventricular tachycardia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia was not determined. b.2 revised as the wrong selection was reported on the initial manufacturer device report number 1220648-2025-26631. b.3 revised date of event as it was inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26631. e.1 added us phone number as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26631. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26631 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name and reporting contact fax number as they were inadvertently omitted from initial manufacturer device report number 1220648-2025-26631. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26631 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
11149 | 1220648-2025-26632 | 21572848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572848&pc=PYX | 02/13/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  | 06/30/2025 |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024360935 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was optical sensor failure and flows were inaccurate. the pump was explanted and replaced. no patient harm was reported. there was no visible damage or clot after removal upon inspection by the physician. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11150 | 1220648-2025-26633 | 21572868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572868&pc=OZD | 02/10/2025 | 04/16/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570135 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump was not detected by the automated impella controller (aic), the pump was successfully replaced. there was no patient harm reported. | [5 empty] | The investigation of pump not detected has been completed. pump was returned. data logs were recovered. data log review showed troubleshooting measures taken in the field to disconnect the pump and restart the console. although attempted multiple times, impella defective alarms reappeared upon every attempted case start. investigation of the pump showed no abnormalities. impella defective and pump not detected alarms were unable to be reproduced during investigation. there was no issue found during the case. the device functioned/operated to specification due to impella defective alarms not being reproduced during investigation. b3 date of event was erroneously on the initial manufacturer device report number 1220648-2025-26633. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26633 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect - impact code 2645 was erroneously on the initial manufacturer device report number 1220648-2025-26633. new code was entered. h6 type of investigation code 4114 was erroneously on the initial manufacturer device report number 1220648-2025-26633, the device had been received for investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11151 | 1220648-2025-26634 | 21572880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572880&pc=OZD | 02/13/2025 | 03/24/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. as groin was being dressed for icu, impella alarmed in aorta. after multiple attempts to advance the impella cp back into the left ventricle, a decision was made to exchange the impella cp. no patient harm was reported and the patient outcome was noted as stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the pump was not returned for investigation. the data log shows the impella position in the aorta alarmed after 21 hours of case run. also, clinical observations confirm that the pump was not able to reposition after it was pulled into the aorta. the cause of the positioning issue was most likely use related based on given clinical detail. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26634 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11152 | 1220648-2025-26635 | 21572885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572885&pc=OZD | 02/13/2025 | 10/01/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578419 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient had long runs of ventricular tachycardia (vt). position was checked and the pump was found to be in the mitral valve. the pump was repositioned. the patient continued on support until the device was successfully weaned. the patient outcome was noted to be stable. | Us complainant reported the patient was on impella 5.5 support and during support, there were placement signal issues. position was checked and the pump was found to be in the mitral valve. the pump was repositioned. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal: no product was returned. data logs were reviewed, and no psnr alarms observed. several impella position unknown alarms were found during the case. optical 5s parameters snr and contrast were periodic low during whole case and pulsatile. gain, lamp and light were found to be stable and no major issues, but ps was spiking to 3000 during the whole case. no similar pattern was identified with before pump on the console. the root cause of optical signal issue is not determined due to inconclusive evidence per log review and with product not being returned. please refer to complaint (b)(4) for an investigation into the console. vf/vt/af/at: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue(b)(6) as part of a retrospective review of optical signal issues in accordance with fda recommendation.
11153 | 1220648-2025-26636 | 21572909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572909&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a 39 year old female with impella 5.5 device when the purge pressure was out of the normal range. the purge cassette was replaced but the problem continued. the impella aic was exchanged, and the issue was resolved. there was no patient harm reported. | [5 empty] | The investigation for the console (aic) purge issues has been completed. console logs were not consistent with what was reported in the complaint description. the reported complaint date was 2025.02.14 but no logs exist from that date. the actual complaint date was likely 2025.02.09 since it was reported that the issue did not resolve after a purge cassette change and that was the most recent day that the purge cassette was replaced. no evident purge pressure issues were observed in the logs. however, after reviewing the logs, numerous imc-ahp errors were observed which are indicative of communication issues between the purge cassette and the rfid. the console was returned for evaluation. the reported purge pressure fluctuation issue could not be reproduced. the console operated as expected. however, isolated testing was performed to identify the root cause of the imc-ahp error issues observed in the logs. the imc-ahp error issues were isolated to the rfid pcb. after swapping the rfid pcb for a known good one and testing the console with a pump and purge cassette kit, the imc-ahp error issues subsided. the reported fluctuating purge pressure issue could not be determined due to the inability to reproduce. the root cause of the observed imc-ahp error issues was isolated to the rfid pcb. the rfid issues did not have a clear correlation to the reported but not confirmed purge pressure fluctuation issue.
11154 | 1220648-2025-26637 | 21572910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572910&pc=OZD | 02/13/2025 | 04/24/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540315 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the impella presented positioning issues; the impella was pulled back into the descending aorta during the case and decision was made to explant the device post-operatively. the patient outcome was noted to be stable. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of positioning issue was most likely use related since the pump was intentionally pulled back into the descending aorta during coronary artery bypass grafting (cabg) and was explanted after the case. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11155 | 1220648-2025-26638 | 21572924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572924&pc=OZD | 02/14/2025 | 05/06/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593246 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, and post implant, same day, suction was encountered on all performance levels. the flow was only 0.2l/min. it was further noted the patient was being moved when suddenly the impella cp lost its position. alarm triggered with displayed message "impella in ventricle." however, the physician stated in echocardiogram, the impella, neither in the ventricle, nor in the aorta could be seen. repositioning attempted and now, at this time, the impella appeared to be stuck in the aortic valve. there was a suspected kink of the canula. the issue could not be resolved, and as the patient was noted as stable, the physician opted to explant the impella cp device. the patient was noted to be stable following the event, | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (cannula kink) and positioning issues has been completed. the data was not returned for review. visual inspection found a kink on can about 15 mm from if cage and can bonding site. no other issues were found on the rest of the pump. pump ran without issue under bench test. the root cause of low flow was most likely kinked cannula since clinical team confirmed cannula kink under echocardiogram and kinked cannula found under pump analysis. the root cause of positioning issue was most likely patient movement related since physician reported that impella lost its position while they were moving the patient d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26638 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11156 | 1220648-2025-26639 | 21572940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572940&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603150 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the procedure, the impella cp device was explanted and post closure technique perclose failed twice. the team decided to use a new prostyle system. however, this system also failed along with other actions. the bleeding persisted. manual pressure was maintained, and the patient was immediately taken to the vascular operating room for emergency closure. the patient required and was given one unit of blood. after the surgical closure, hemostasis was achieved. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11157 | 1220648-2025-26640 | 21572954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21572954&pc=OZD | 02/15/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/15/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577118 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced limb ischemia requiring intervention. the patient was taken to the catheterization lab and cardiac arrested upon arrival. return of spontaneous circulation was achieved and the impella cp device was placed using best practices. the left anterior descending artery was opened with a drug-eluting stent placement. following, the peel-away sheath was removed, and repositioning sheath was inserted. valve wire marked and distal runoff completed. there was no distal flow visualized; therefore, the physician placed antegrade and retrograde sheaths for fem-fem bypass to perfuse limb. the patient was reported to be stable following the event. three days later the patient was successfully weaned of the impella mcs and all drips. closure of impella access site as well as fem-fem bypass were successful. the patient remained intubated until end of best rest. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11158 | 1220648-2025-26641 | 21573218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21573218&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594813 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Perforation | [12 empty] | Required Intervention | Life Threatening | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support experienced low pump flow and pericardial effusion. the patient was successfully treated noted to have survived the event. the patient arrived via emergency medical services and became unconscious and received two rounds of cardiopulmonary resuscitation. the patient was in ventricular tachycardia and was shocked twice. the patient was taken to the cardiovascular lab for percutaneous coronary intervention (pci) and an impella cp device placement. the impella cp device was successfully implanted without incident. the impella was started on auto and, however, began to have low flow and motor current. echocardiogram was completed and showed pericardial effusion. it was unclear if the pericardial effusion was caused by the impella cp or as a result of the pci procedure. however, pericardiocentesis was completed, and the impella cp pump was explanted. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: precaution impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11159 | 1220648-2025-26642 | 21573543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21573543&pc=OZD | 02/14/2025 | 04/15/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139985 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon starting the automated impella controller and using demonstration pump for a conference teaching, a controller error alarm was received multiple times. a hard reset was completed; however, the controller error alarm issue remained unresolved. | [5 empty] | The investigation of automated impella controller (aic) ¿ controller error (yellow alarm) has been completed. on the compliance date, the console displayed the message ¿controller error (defective optical sensor lamp) ¿ the optical parameters were abnormal, particularly the light: 0.69 and the snr: 1.4. although the console was rebooted several times, it continued to display event each time after connecting the demo pump. this confirms the reported failure mode. after the fourth reboot, it is most likely that the console was shut down due to the battery switch being disengaged without an ac connection. this aligns with the claim that a ¿hard reset was completed.¿ the console was then manually rebooted. after the reboot, the console displayed ¿unexpected controller shutdown device analysis: this console was tested and upon power cycling the console 10 times and running the optical test pump for two minutes each time, unable to reproduce the reported failure mode. the envelope from the optical bench was good. the 5s parameters with the optical test cavity are within specification. found the front panel optical connector contaminated, unable to remove the debris upon cleaning. root cause: the root cause of the controller error was most likely due to contamination of the front panel optical connector. g2 report source; health professional and user facility was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26642. | The automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11160 | MW5167480 | 21573633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21573633&pc=OZD | 02/18/2025 | 03/05/2025 |  | 03/11/2025 |  |  |  | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 74 YR |  | Female | 69 KG | Non Hispanic | White | Injury | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | On removal of the impella device the metal portion of the lmpella intake and outflow, separated from the impella cord.
11161 | 1220648-2025-26643 | 21573849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21573849&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599142 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Contamination | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support when part of the device fell off below the sterile field on the operating table (contaminating the device). the decision was made to remove and replace with a new impella cp device. there was no harm to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11162 | 1220648-2025-26644 | 21574821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21574821&pc=OZD | 02/13/2025 | 03/31/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594019 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported on 73 year old male with impella 5.5 pump due to acute myocardial infarction/cardiogenic shock. it was reported that a leak in the area of the yellow-to-yellow connection or in the covered area of the filter of the impella 5.5 was observed. the nurse could not see where it was coming from. the patient continued on support until he was successfully weaned and explanted. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the data logs were reviewed and the purge pressure remained stable at around 550 mmhg through the case and there were no significant changes in purge flow. the returned product was purged for 10 minutes on a lab console and a cassette from the lab. no leaks were reproduced. there was some biomaterial around the impeller and the purge gap was seen. a syringe was taken and connected to the side arm to add pressure to try to exacerbate if there was a leak. no leaks were reproduced. there were no damages seen of the pump. the root cause of the purge leak - pump could not be determined as it was not able to be reproduced and unknown what occurred in the field. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26644. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26644. g.1 added contact facility name as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26644.
11163 | 1220648-2025-26645 | 21574976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21574976&pc=OZD | 02/14/2025 | 09/23/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557434 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [15 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported a placement signal not reliable alarm which was addressed, and an arterial line was being used for hemodynamic monitoring. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support while awaiting a heart transplant. during support the patient coded multiple times and required a nerve block at bedside. additionally, the patient presented with an infection so the swan was removed. the patient also experienced a drop in hemoglobin and bleeding from esophagus. a unit of blood was transfused. anticoagulants were paused. the patient continues on support while awaiting a heart transplant. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. infection/sepsis: the root cause could not be determined due to insufficient information. optical signal issue: data logs were not recovered, and the product was not returned. due to the lack of clinical details and absence of data, the root cause could not be determined. ventricular tachycardia/fibrillation, atrial fibrillation/tachycardia: the device was not returned. to determine the root cause of the vt/vf, clinical information and analysis of the returned device would be necessary. as this information was not available, the root cause could not be established. vascular damage ¿ peripheral vessel: the product was not returned, and there was a lack of clinical details. as such, the root cause could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-26645. b7 other relevant history, including preexisting medical conditions updated. d4 serial number corrected d6b explant date was corrected. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11164 | 1220648-2025-26646 | 21575353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21575353&pc=OZD | 02/13/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570929 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support the patient presented with duct-dependent pulmonary circulation and loss of pulses to the right limb. treatment indicated was vascular surgery. actions taken resolved the ischemia. additional information received indicated that this patient was on maximum amount of chemical support that caused significant vasoconstriction that limited flow to all extremities (not just the impella leg). subsequently, it was reported that the patient decompensated and veno-arterial extracorporeal membrane oxygenation (ecpella) was required, with emergent permanent pacemaker (ppm). the ejection fraction remained decreased, so the patient was escalated to impella 5.5. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
11165 | 1220648-2025-26623 | 21575393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21575393&pc=OZD | 02/12/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/12/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119121 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a automated impella controller (aic) turned on for a swap but, there was an issue reported with unresponsive buttons. they were unable to mute, change the flow or initiate other functions. another device was used. no patient was involved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11166 | 1220648-2025-26647 | 21575547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21575547&pc=OZD | 02/12/2025 | 04/16/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599663 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp pump was being placed on a patient for the support of an acute myocardial infarction/cardiogenic shock when there was a guidewire perforation of the left ventricle. a thoracotomy hemostasis was performed and circulation continued. the patient continued on support. circulation became difficult to maintain, and the device was removed due to death. the cause of death was stated by the facility to be acute myocardial infarction/cardiogenic shock and the left ventricle perforation. | [5 empty] | The investigation of ventricle perforation has been completed. no product was returned for evaluation. clinical details noted that perforation occurred during use of the guidewire during impella insertion steps. additionally, physician noted that the perforation of the ventricle was sharp and resembled a puncture from a guidewire and not from fragile tissue and pigtail perforation. the root cause of the ventricle perforation was most likely due to a guidewire puncture during impella insertion. b2 date of death was erroneously entered on the initial manufacturer device report number 1220648-2025-26647. d6b explant date was erroneously entered on the initial manufacturer device report number 1220648-2025-26647. h6 code 2135 was erroneously entered on the initial manufacturer device report number 1220648-2025-26647. new code was added to health effect - clinical code. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11167 | 1220648-2025-26648 | 21575759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21575759&pc=OZD | 02/13/2025 | 04/16/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  | 04/30/2025 |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344456 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was a kinked cannula observed after three hours of support. the kink was restricting flow and despite efforts to advance and twist the cannula, the issue was not resolved and the device was explanted. the device was replaced with another device, not another impella. there was no patient harm. the patient was noted to be stable. | [5 empty] | The investigation of product damage (cannula kink) has been completed. pump was returned for investigation. there was a cannula kink observed near the outlet of the pump. there is no clinical information provided about the reason of the kink. therefore, the root cause of the product damage issue was a kinked cannula but the reason of the kink is unknown. pump was investigated and there was a cannula kink observed near the outlet of the pump. data logs were not returned. clinical information provided is very limited and states kink observed that was restricting flow. therefore, the root cause of the low pump flow issue was likely a kinked cannula but the reason of the kink is unknown. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11168 | 1220648-2025-26649 | 21575983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21575983&pc=OZD | 02/13/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556225 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient presented for shortness of breath and was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that heparin was started. the patient coughed and experienced bleeding at the access site. dressing was changed and pressure was held. heparin was turned off to obtain hemostasis. no hematoma was present. support continued. subsequently it was noted that the patient decompensated after rolling for cleaning. he went into a wide complex rhythm, then into ventricular tachycardia and then ventricular fibrillation. advanced cardiovascular life support (acls) protocol was initiated and patient coded for 20 minutes, until it was called and he was pronounced expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11169 | 1220648-2025-26650 | 21576021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21576021&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607202 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and as a bridge to transplant. the impella 5.5 device was secured with 3-point fixation attached to abdomen. the patient was sent to the unit on support. approximately seven days after start of the support the pump stopped. the patient stood up to use restroom and accidentally stepped on impella cord. when the patient realized the cord was being stepped on, the patient sat down. within seconds the impella pump stopped and was unable to be restarted even after connecting to a different automated impella controller. the patient's mean arterial pressure dropped into the 30s and was very unstable. the patient coded and cardiopulmonary resuscitation was initiated, maintained for approximately ten minutes, and the patient was shocked multiple times. emergent extracorporeal membrane oxygenation placement was completed at bedside, and the patient was then urgently transported for a new impella 5.5 device placement which was indicated as successfully completed for continuation in mcs therapy. the patient returned to the unit in critical condition. post replacement of the impella 5.5 (timeframe unknown) the patient experienced runs of ventricular fibrillation and was shocked by the implantable cardioverter defibrillator. the patient was noted to have survived the event; however, the patient was noted as critical. | [5 empty] | The impella device and data logs were received from the customer, and the evaluation/analysis and investigation were completed with details captured below. product malfunction: pump stop: according to clinical, the patient stood on the impella and then sat down while standing on the device cord. the pump and the data logs received for evaluation and analysis indicated the following: the data logs were reviewed and there is an alarm 115 pump stop with motor current spiking to 30000 on the day of the event. the returned product showed that there was sufficient glue on the kink protector to red handle bond as well kink protector to catheter. the catheter was not able to freely rotate inside the kink protector. there was a large kink in the catheter seen distal to the kink protector and then the motor cable lines were all open was the red handle was opened. they were also frayed. product history review was completed, and the pump passed all sterile inspection checks. in conclusion, the root cause of the pump stop was determined due to open electrical lines coming from the excessive force from the user. clinical adverse event: cardiac arrythmia: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. in conclusion, to determine the true root cause of the ventricular tachycardia/ventricular fibrillation, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the ventricular tachycardia/ventricular fibrillation was not determined.
11170 | 1220648-2025-26651 | 21576056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21576056&pc=OZD | 02/14/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/14/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571930 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the impella unable to advance through the patient's vasculature during delivery due to calcification in the innominate artery. the implant was aborted and an impella cp pump was placed for planned support. there was no patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11171 | 1220648-2025-26653 | 21577963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21577963&pc=OZD | 02/13/2025 | 04/02/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577788 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported on a 85 year old female patient on impella cp pump support placed for cardiogeni shock/acute mycocardial infarction. the pump was placed but it had low flows and was unable to be troubleshot. the pump was explanted and replaced. no harm or injury was reported. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. a cannula kink was found near the outflow cage. no other damages or abnormalities. reported immediate suction alarms and low flows around 1.2 liters per minute when starting the pump. data logs confirmed suction alarms and consistent flows around 1.2 liters per minute at p-9 for 10 minutes. the root cause of the low pump flow was most likely a kinked cannula since a cannula kink was found. d.9 device return date/information was added. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26653. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact fax number and reporting contact facility name as they were inadvertently omitted from the initial manufacturer device report number.
11172 | 1220648-2025-26652 | 21579546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21579546&pc=OZD | 01/08/2025 | 03/11/2025 |  | 03/11/2025 |  |  | 02/13/2025 | 03/11/2025 |  |  |  |  | 03/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | It was reported by abiomed's field service representative that a console (aic) that had been in use since on (b)(6) 2024, was unable to download data onto the usb drive. there was no patient involvement. | [5 empty] | The investigation for the reported aic issue downloading to the usb has been completed. the console logs were successfully downloaded, revealing multiple instances of the "algs suspended opm signal invalid" error, followed by several placement signal low (event#: 439) alarms, triggered by the signal-to-noise ratio (snr) falling below the minimum threshold. the lt_log data revealed that the unreliable placement signal persisted throughout most of the case, with varying intensity, which corresponded to the low snr. this caused repeated placement signal low alarms (#: 439) due to the low snr. additionally, the rt_log data showed low contrast and snr values. further data analysis revealed a controller error alarm (opm comm crc error ¿ event#: 1045). the rt logs confirmed that all signals from the optical bench (placement, snr, contrast, lamp, gain) were recorded as 16384 values due to the crc error. the imc logs also show multiple "impella stopped (#: 13, motor current high) [detected by epc] - alarm" and "impella stopped (#: 18) retrograde flow alarm". for further information on the impella stop alarm. the console ic8602 was returned to field service for further investigation. the aic data download issues were not reproduced, as aic data logs were successfully downloaded. testing of the optical system revealed that the snr value (one of the 5s parameters) was only marginally passing (~3000), but this didn't cause optical signal issues when testing with a known good pump. the analog output of the ccd from the optical bench no distortions in the signal (envelope/fringe) was observed. the issues with aic data download could not be reproduced/ cause not determined, as no issues were found with the initial data log download process. the controller error due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the root cause of the optical signal issue was most likely due marginally low snr optical bench.
11173 | 1220648-2025-26585 | 21558216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21558216&pc=OZD | 02/16/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/16/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610826 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella 5.5 support had frequent runs of ventricular tachycardia and was shocked multiple times. the patient was started on a lidocaine and amino dose was increased. impella position was confirmed via echocardiogram. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11174 | 1220648-2025-26586 | 21558268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21558268&pc=OZD | 02/17/2025 | 04/21/2025 |  | 03/10/2025 |  |  | 02/17/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592671 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 pump was supporting a patient admitted in for high risk surgery. the pump was placed but it was reported that there was negative aorta pressure readings, suction alarms over p5 and therapy was unable to be escalated. the patient remained stable and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Product and data were not returned for review. the root cause of low flow was unable to be determined as limited clinical details, no data logs were provided and no product was returned for review. the root cause of optical signal issue was unable to be determined as limited clinical details, no data logs were provided and no product was returned for review.
11175 | 1220648-2025-26587 | 21558402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21558402&pc=OZD | 02/08/2025 | 04/15/2025 |  | 03/10/2025 |  |  | 02/11/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025514480 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and developed high purge pressure and would expire the same day noted to be from multi-organ failure. the impella was implanted via a femoral artery prior to percutaneous coronary intervention. vessel treatment information was noted as unknown. the patient was eventually sent to the unit on impella mcs. it was noted there were no other therapies prior to impella mcs. approximately 11 days after the start of the support, the impella cp pump experienced a rise in purge pressure coinciding with a decrease in purge flow. troubleshooting was completed without success as the purge cassette was replaced, but there was no improvement with the high pressures. no further details were provided regarding the purge pressure issue were made available. however, it was noted the patient would expire this same day from multi-organ failure as the cause of death. no details surrounding and cause of the multi-organ failure were provided other than this clinical adverse event is unrelated to the impella device. although the reporter indicates the cause of death as multiple organ failure and unrelated to the impella device used, it is unknown if the unresolved high purge pressure device malfunction may have impacted the patient's support. the high purge pressure issue was unresolved with the cassette change and other device function factors cannot be ruled out with only "not related." thus, there is not enough evidence to exclude the impella as an associated factor in overall outcome. | [5 empty] | The impella cp pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of high purge pressure has been completed. impella cp was returned for evaluation. upon visual inspection, biomaterial was present in purge gap but not fully covering all of the purge gap. the returned pump and returned purge cassette were connected to the console and purged in the lab for approximately 30 minutes and no high purge pressure alarms reproduced. the pump was turned up to p-9 and again no high purge pressure alarms were triggered. data logs were reviewed and long-term log of full case showed two instances of high purge pressure occurring during the case. no ¿high purge pressure¿ alarms were triggered during either instance of purge pressure increases. both instances of increases in purge pressure occurred over a period of approximately 12 hours for each instance and no motor current spikes were present prior to purge pressure rise. first instance of high purge pressure begins to decrease prior to a purge cassette change. clinical details noted that high purge pressure alarms occurred on day 12 of support. the clinical team attempted to exchange the purge cassette but did not resolve the issue. the purge composition at time of alarms was not noted. the root cause of the high purge pressure could not be definitively determined as high purge pressure was unable to be reproduced in the lab. b.2 required intervention was reported incorrectly on initial manufacturer device report number 1220648-2025-26587.
11176 | 1220648-2025-26588 | 21558973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21558973&pc=OZD | 02/12/2025 | 03/31/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621756 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Perforation | Cardiac Tamponade | Hemodynamic instability | [10 empty] | Death | Required Intervention | Life Threatening |  |  |  | JR4 GUIDE CATHETER AND 0.035 J GUIDEWIRE | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) undergoing a high-risk percutaneous coronary intervention (pci) and implanted with an impella cp device for mechanical circulatory support (mcs) experienced left ventricle perforation with a demise outcome. the patient presented to the hospital with chest pain. a catheterization revealed multi-vessel disease including a heavily calcified left main and an occlusion of the right coronary artery. cardiothoracic surgery was consulted. after further work up, patient was turned down for surgery due to having a porcelain aorta. the plan was noted for cardiology to perform protective pci with impella support to rotablation to the left main/left anterior descending arteries and stent the right coronary artery. the impella cp device was inserted right femoral artery, utilizing 14fr. peel-away sheath. jr4 and 0.035 j guidewire utilized for ventricular access. the guidewire exchanged for abiomed 0.018, of note 0.018 wire was curled prior to insertion. impella primed and backloaded onto 0.018 and inserted into left ventricle. the guidewire removed and the impella cp started in auto mode and encountered reduced flows. patient became more hemodynamically unstable, correlating with poor pump performance. patient required cardiopulmonary resuscitation, and an emergent echo was completed revealing a large cardiac tamponade. emergent pericardiocentesis and cardiothoracic surgery was summoned to the bedside. patient continued to require resuscitation including two units of packed red blood cells. the decision made to emergently transport the patient to the operating room for pericardial window. in operating room, the patient had median sternotomy, cardiac tamponade was relieved, 500 cc of blood evacuated, and the patient was found to have left ventricle perforation. at this point impella was weaned to performance level p-2 and withdrawn into the descending aorta to allow for left ventricle repair. the decision was made to reestablish impella mcs with a new impella cp. the first impella cp was removed through the peel-away sheath, and new impella cp was inserted without issue. the replacement impella cp was started in auto mode. however, the pump and patient were dependent on cardiac massage. the patient had no signs of native activity; therefore, the decision was made by cardiothoracic surgery and cardiology that patient was declared expired. following, the physician commented a second impella cp was utilized due to concerns that the left ventricle perforation was caused by the initial impella cp. | [5 empty] | The investigation for the ventricle perforation and low pump flow has been completed. the product was returned and evaluated. evidence of cannula kink was found. there was no biomaterial found near the impeller or in the cannula. the pump was run in a bucket at p-6, p-9, and p-2. low flow was not reproduced. in the beginning of the case, reduced flows were reported with the patient becoming more hemodynamically unstable. data log review confirms reduced flow in the first 45 minutes of the case. during this time, there are several suction, impella position unknown, and impella in ventricle alarms. the pump is then turned up to p-6 from p-2 and eventually reaches adequate flows (~2.7 l/min). the pump is then turned down to p-2 and the calculated pump flow stays high around 2.7 l/min until the end of impella use. the root cause of the ventricle perforation was not determined since insufficient clinical details were provided related to timing/cause of the perforation. the root cause of the low pump flow was most likely a cannula kink since the returned product showed evidence of a kink and data logs confirm positioning issues. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: precaution impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11177 | 1220648-2025-26589 | 21559050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559050&pc=OZD | 01/27/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562112 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) experienced a stroke some days after explant of the impella cp device. the patient with disease in the right coronary artery, left anterior descending, diagonal arteries and chronic total occlusion of the circumflex with an ejection fraction of 55% was offered and refused surgery. the patient opted for high-risk pci of the right coronary artery, diagonal and left anterior descending arteries. the impella was placed, and the pci being completed resulted in a large vessel rupture perforation and subsequent pulseless electrical activity (pea) arrest for approximately 30 minutes until the covered stent and pericardiocentesis were performed. the patient recovered - blood flow maintained during the pea. at the end of the case, the patient was hemodynamically stable with the impella cp device in place and sent to unit with norepinephrine, aggrastat, and heparin drips running. the following day the decision was made to transfer patient to an outside hospital for advanced management. the patient's lactic acid and pressor requirements had lowered, and the decision was made to remove impella cp after approximately 27 hours of total support. upon removal (impella staff not present) it was reported clots were found on the pump. during the pci and during the pumps indwelling time patient was on heparin and aggrastat continuous drips with activated clotting time greater than 300. once the patient was transferred to the outside facility where the pump was removed, abiomed's data documentation of drips and anticoagulation is unavailable. therefore, it is unclear what anticoagulation strategy was implemented on their management until removal later the same day. however, (approximately 4) days later per the physicians at the hospital, a head scan to assess the patient's cerebral blood flow was completed. the scan showed multiple embolic strokes. it was commented, since the impella cp had a clot on the pump when it was removed, the team stated the clots came from the impella cp. the patient was noted to have survived the event. however, the patient's outcome status following the event is indicated as "unknown." | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: precaution impella cp with smartassist system. section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11178 | 1220648-2025-26590 | 21559152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559152&pc=OZD | 02/12/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555066 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump was placed but due to lack of distal flow down the limb, the team made attempts to place a distal perfusion line. the placement failed and so the pump was weaned and replaced by an intraaortic balloon pump. the patients status was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
11179 | 1220648-2025-26591 | 21559381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559381&pc=OZD | 01/30/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605360 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via protect iv study regarding a patient who experienced access site during impella cp support. after removal, there was noted oozing of the femoral site. manual pressure was held. protamine was also administered to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11180 | 1220648-2025-26592 | 21559573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559573&pc=OZD | 02/12/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597235 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported on a 65 year old female patient with an impella cp for acute myocardial infarction/cardiogenic shock. the cp was supporting when there was an access site bleed. the bleed was treated by suturing, femostop, stopping of the iib/iiia aggrastat and heparin was paused. the pump was weaned and explanted after 35 hours. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11181 | 1220648-2025-26593 | 21559818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559818&pc=OZD | 02/12/2025 | 04/11/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584368 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER, IMPELLA RP FLEX | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device was unable to go above p2 without triggering continuous suction alarms. the device was repositioned, volume was provided, and an implanted impella rp flex device was adjusted, but the low flow persisted. the impella was removed. no patient harm was reported. | [5 empty] | The investigation of low pump flow has been completed. the product was returned and there were no kinks in the cannula seen, biomaterial or damages to the impeller blades. the field data logs show that it was disconnected and reconnected twice to see if lower flows were still present. there were instances of suction alarms #14 and #73. there were no motor current spikes seen or a trend in the placement signal. however, the motor current was lower than the expected range for the pump while running on the respected p-levels as in ~less than 150ma at p-2. when the returned device was ran on a different console in the lab, lower pump flows were not reproduced at p-2 and the motor current was ~375ma. this is the expected range for the motor current for the pump. there was no problem found with the pump. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26593 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 investigation conclusions 4315 was erroneously entered on the initial manufacturer device report number 1220648-2025-26593. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11182 | MW5167441 | 21559824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559824&pc=OZD | 02/17/2025 | 02/28/2025 |  | 03/10/2025 |  |  |  | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED | [8 empty] | Yes | [5 empty] | Y | Health Professional |  |  | Voluntary report | Yes | Other Health Care Professional | [6 empty] | 65 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [9 empty] | Patient experienced a significant pull of the impella 5.5 driveline causing a separation of the power wires within the driveline and subsequent device failure. patient was emergently taken to operating room for pump exchange. concern is that the device is not manufactured to be as durable externally as anticipated and what should be expected compared to similar devices with external leads. due to the nature of the event and subsequent findings, the exact location of the break within iha driveline may point to a specific area of concern with current manufacturing process and durability testing.
11183 | 1220648-2025-26595 | 21559860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559860&pc=OZD | 02/12/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599977 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced limb ischemia at explant and was ultimately resolved. the patient presented with grafts of a prior coronary artery bypass graft (cabg) occluded and chronic coronary disease; no interventions were completed. the patient was also in complete heart block and transvenous pacer was placed. the patient was brought to electrophysiology lab for placement of leadless pacemaker. however, upon induction of anesthesia, patient¿s blood pressure dropped into the 50s requiring max dose of levophed and stopping sedation. stat echocardiogram showed stunning and decreased right ventricular and left ventricular contractility with dilated right ventricle and decreased left ventricular ejection fraction. decision made to place impella cp and complete a right heart catheterization to evaluate if an impella rp support was also needed. impella cp advanced over wire into left ventricle without difficulty and wire removed. impella started and ramped up on auto to 3.5 l/min flow and then placed on performance level p-7 with 3.2 l/min flow. the physician proceeded with the right heart catheterization. another echocardiogram was completed, and contractility seemed improved; right ventricle no longer distended and also had improved contractility. therefore, decision was made not to place an impella rp. the patient was more stable, therefore, the physician decided to proceed with leadless pacemaker placement. at end of procedure, impella remained secured, and the patient was returned to the unit on support. the patient did well overnight, and the impella cp was weaned and explanted. manual compression was held for 40 minutes, and hemostasis was achieved. the patient had palpable pulse following explant. however a few hours following explant, the staff was unable to obtain distal pulses with doppler. arterial doppler was completed and there was no distal flow. the patient was returned to catheterization lab for lower extremity angiogram and intervention. a clot was noted in the right femoral artery access site. a thrombectomy was performed using penumbra and then a stent was placed. the physician believed there was an underlying lesion and possible dissection. distal flow was restored and patient tolerated procedure well. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: precaution impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11184 | 1220648-2025-26596 | 21559871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559871&pc=OZD | 02/12/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610850 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump. upon attempting via right axillary access, it was noted that the axillary kit was unable to advance. the left axillary artery was accessed successfully, but upon insertion of the pump, the cannula kinked and further advancement was not possible. the implant was aborted. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11185 | 1220648-2025-26597 | 21559876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559876&pc=OZD | 02/12/2025 | 03/19/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592781 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient had ventricular tachycardia (vt). the pump position was adjusted. this occurred twice during support and after the second repositioning the staff stated the patient's vt resolved. the patient continued on support until the device was successfully weaned. the patients condition was stable. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11186 | 1220648-2025-26599 | 21559921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559921&pc=OZD | 02/16/2025 | 04/29/2025 |  | 03/10/2025 |  |  | 02/16/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609813 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed hematuria due to impella positioning issues. impella was repositioned three times since implant. due to persistent hematuria and request from the doctor the patient was taken to the cardiac catheterization laboratory for removal. the patient remained stable and there was successful removal with no complaints or allegations. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. note: removed failure mode positioning as it was the root cause of hemolysis. product and data were not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11187 | 1220648-2025-26598 | 21559922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559922&pc=OZD | 02/12/2025 | 03/24/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601344 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the 71-year-old female patient experienced multiple complications during impella cp support. following implant, it was observed that the device was obstructing flow to the patient's leg. a fem-fem bypass was placed coinciding with an antegrade sheath to reestablish flow to the impella leg. it was additionally noted that the patient experienced steady oozing from their insertion site. manual pressure was held but the oozing continued. an additional suture was placed, the angle of insertion was matched, protamine was given, and femostop was applied to achieve hemostasis. blood products were given to counteract the blood loss and drop in hemoglobin levels. support continued with the implanted pump following the intervention. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. e.1 added fax number and us phone number as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26598. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26598 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name and fax number as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26598.
11188 | 1220648-2025-26600 | 21559936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21559936&pc=OZD | 02/16/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/16/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592559 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant had a impella cp pump placed via the 14fr sheath to support a patient admitted in with a st-segment elevation myocardial infarction. it was reported that the right leg was mottled and no distal pulses. plan to remove impella cp and upgrade to impella 5.5. during removal of the cp there was vascular damage and a repair of the right femoral artery was done. the 5.5 was then successfully placed. the procedural outcome was the patient survived surgery. the device was discarded and not available for investigation. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
11189 | 1220648-2025-26601 | 21560080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21560080&pc=OZD | 11/25/2024 | 03/10/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured femoral access site bleeding with a "possible" relationship to the impella device used: ¿on arrival to p5, patient was intubated. he developed hypotension and bradycardia. he was found to have profuse bleeding at the femoral sheath site. ¿. the patient outcome is unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11190 | 1220648-2025-26602 | 21560822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21560822&pc=OZD | 02/13/2025 | 04/21/2025 |  | 03/10/2025 |  |  | 02/13/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571439 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 33 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured renal failure with a "possible" relationship to the impella device used: ¿aki secondary to hemolysis. eventually transitioned to hemodialysis via permacath.¿ the patient not recovered. | The user facility reported a 33 year old female patient on a cp pump placed to support a patient admitted in with cardiogenic shock and myocarditis. the pump was placed but there were signs of hemolysis. the free hemoglobin (pfhg) was rising, pump repositioning did not remedy, and when the patient stopped making urine and creatinine rose, the team made decision to explant pump. the patient was successfully weaned and explanted. |  |  |  |  | B5 updated with additional information received from the clinical site. h6 updated to reflect the additional failure mode reported from the clinical site. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26602. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26602. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11191 | 1220648-2025-26603 | 21561521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21561521&pc=OZD | 02/13/2025 | 03/10/2025 |  | 03/10/2025 |  |  | 02/13/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626745 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 73 year old male patient on impella cp pump for mechanical circulatory support. upon insertion of the impella 14fr long sheath, the component kinked on the curve from the common femoral to the iliac. a long dilator was advanced, the sheath was pulled back, and an impella 14fr sheath was placed without complication. the impella cp implant was successful upon swapping the introducer. there was no patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
11192 | 1220648-2025-26604 | 21562213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21562213&pc=OZD | 02/11/2025 | 04/25/2025 |  | 03/10/2025 |  |  | 02/11/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579640 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of the impella 5.5 device to a patient with postcardiotomy cardiogenic shock was unsuccessful due to the anatomy. the patient was taken to cardiovascular operating room for the impella 5.5 placement in the setting of post operative cardiogenic shock requiring extracorporeal membrane oxygenation support with failure to liberate from support. cutdown of right axillary artery was completed with graft anastomosis of 8mm gelsoft graft. it was noted there was "extreme" difficulty with wiring aortic valve due to acute angle of new aortic graft repair. after successfully passing wire past aortic valve, the impella 5.5 was backloaded and advanced. however, the impella 5.5 device was unable to cross aortic valve and guidewire access was lost. the impella 5.5 was removed and found to have "significant" clot obstructing inflow/outflow with unsuccessfully attempts to clear. the physician opted to repeat insertion attempt with a new impella 5.5 device which was noted to be successful. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: precaution impella 5.5 with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of delivery issues and thrombus has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely use related,5.5 insertion pump was unable to be advanced past atrioventricular (av) with resulting loss of wire access to left ventricle (lv). as the necessary information was not provided, the root cause of the thrombus was not determined.
11193 | 1220648-2025-26605 | 21562450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21562450&pc=OZD | 02/05/2025 | 04/03/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565435 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with systemic circulatory failure was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump stopped on the 10th day of support. the pump was explanted and there were no plans to escalate care. the patient had an impella dependent condition and expired shortly after the pump was explanted. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. the impella cp was returned for evaluation. no physical abnormalities were found on the returned product, and the pump passed electrical testing. upon further inspection, a large purge gap was found which is due to bearing wear. data logs were not provided for this case run. as per the clinical information, the pump stop happed on the tenth day of use. the root cause of the pump stop was due to bearing wear as the pump was run past the indicated design life.
11194 | 1220648-2025-26606 | 21562489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21562489&pc=OZD | 02/12/2025 | 03/22/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554053 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced oozing and a hematoma at the access site. the team remedied the issue with femstop placement, blood products, and consulted with vascular surgery. the patient continued on support until the device was successfully weaned. the groin was clean and dry before the patient was transferred. additional information was requested. however, no further information is available. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of has access site bleeding been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11195 | 1220648-2025-26607 | 21562735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21562735&pc=OZD | 02/12/2025 | 03/26/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604945 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. after the implant, a small pin hole was found at the posterior left ventricle. it was quickly repaired but it is unclear if it was caused by the impella placement. additionally, upon explanting the impella, the pump detached into two pieces at the nitinol cannula/outlet seam. the nitinol cannula remained in the graft and required surgical removal. the remaining portion was removed successfully. two units of blood products were administered to the patient. the patient outcome was noted to be critical. | [5 empty] | The investigation has been completed. according to the additional information provided by the representatives, the left ventricle was pretty dry and the physician may have placed the pump too deep. in the parts of the cannula that were not returned, there was no kink seen. due to a lack of conclusive evidence, the root cause of the ventricle perforation could not be determined. according to the representatives, it was unclear of what fully happened during the procedure. the product was returned and there was a piece of the cannula missing where the bond is to the motor housing. the cannula was seen clamped, cut and there were holes in it. there was some clot on the impeller as well as the graft was returned with a large clot in it as well. due to inconclusive evidence for clinical details as well as the missing bond on the returned product, the root cause of the product damage (detached inflow/outflow - great vessel) could not be determined. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11196 | 1220648-2025-26608 | 21566586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21566586&pc=OZD | 02/12/2025 | 04/03/2025 |  | 03/10/2025 |  |  | 02/12/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025590879 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 18 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 had difficulty advancing through the axillary artery. the impella was removed and clot was observed at the impella inlet area. the team made the decision to open a new impella 5.5. | [5 empty] | The investigation of delivery issues and thrombus has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related as the doctor aborted insertion due to patient's anatomy. as the necessary information was not provided, the root cause of the thrombus was not determined. b.2 revised selections as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26608. e.1 revised facility name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26608. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26608 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added address line 1, city, state, zip code and country as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26608. h.6 code 4114 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26608 and was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: iimpella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
11197 | 1220648-2025-26610 | 21566757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21566757&pc=OZD | 02/13/2025 | 03/12/2025 |  | 03/10/2025 |  |  | 02/13/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609773 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 72 year old female patient was on impella cp support and during support, the pump was pulled into the aorta and staff were unable to advance the pump across the aortic valve. the pump was replaced. the patient survived the procedure. | [5 empty] | The investigation of positioning issues has been completed. no product was returned for analysis. the clinical details mention pump was pulled back across the valve. the root cause of the positioning issue was most likely use-related as evidenced by clinical detail of pump being pulled back across the valve e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name was missing on the initial manufacturer device report number. g1 reporting contact fax number was missing on the initial manufacturer device report number. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
11198 | 1220648-2025-26609 | 21566832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21566832&pc=OZD | 02/12/2025 | 05/05/2025 |  | 03/10/2025 |  |  | 02/13/2025 | 03/10/2025 |  |  |  |  | 03/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026612496 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during patient transport impella 5.5 pump, a purge pressure high / purge system blocked alarm appeared on the console (aic). the patient was taken to the operating room where a ventricle rupture was observed. the physician attempted to change the pump, but the patient died due to severe bleeding. | [5 empty] | A.2 additional information was received confirming that the patient was an adult over the age of 50. a.3 revised as selection was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26609. b.3 revised date of event as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26609. e.1 revised name prefix/title, first name and last name as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26609. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26609 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact facility name, address, and fax number as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26609. | The investigation of high purge pressure and ventricle perforation has been completed. the pump was not returned for investigation. the data log shows a sudden rise in purge pressure, accompanied by a motor current spike and saturated pump flow. the purge pressure and pump flow values returned to normal after 1 hour. according to the clinical report, a high purge pressure alarm was observed during transport. additionally, a ventricular rupture was observed in the operating room. the cause of the high purge pressure issue was most likely biomaterial, based on data log review. the cause of the ventricle perforation issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
11199 | 1220648-2025-26565 | 21546159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546159&pc=OZD | 02/12/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609818 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced bleeding from their impella cp access site during support. manual pressure was applied and the sheath was manipulated to achieve hemostasis. three units packed red blood cells, two units' fresh frozen plasma, and two unit's platelets were transfused to treat the bleed. support was set to continue, however, due to the severity of the patient's disease, the family opted to move to comfort care. the pump was removed and the patient later passed away. a relationship between the bleed and the patient's passing was not defined. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11200 | 1220648-2025-26566 | 21546214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546214&pc=OZD | 02/07/2025 | 09/16/2025 |  | 03/07/2025 |  |  | 02/10/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583207 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 86 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Other | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The study months later again reported upon another ae of hemolysis which was attributed to be a "possible" relationship to the impella device. | A notification was received via protect iv study notification regarding a patient who experienced renal failure during impella cp support. actions taken are unknown. the patient continued on support and the patient condition was noted to be stable. no further information was provided. | Later in (b)(6) 2025 the study reported upon yet another ae of the access site bleed and new intervention information was shared. to achieve hemostasis the study noted, "pressed on the bleeding site and renewed corsafix." |  |  |  | Additional information received from the clinical site and the following fields have been updated. b5 updated to reflect additional information h6 health effect - clinical code updated. pump was not returned for investigation. clinical information provided is very limited and not sufficient. therefore, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. pump was not returned for investigation. data logs were not returned as well. without the data logs and the pump the failure cannot be further investigated. clinical information provided is very limited. therefore, the root cause of the renal failure issue was not determined since product and data logs were not returned. g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26566. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B5 updated with additional information received from the clinical site. h6 updated with additional adverse event reported by the clinical site.
11201 | 1220648-2025-26567 | 21546267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546267&pc=OZD | 02/12/2025 | 04/22/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1170637 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that two impella pumps were plugged into the automated impella controller (aic) outside of the patient¿s body and the aic did not work. the aic was exchanged. there was no patient harm reported. | [5 empty] | The investigation of automated impella controller (aic) - a/c power issues has been completed. based on the data and device analysis the cause of low flow issues was a defective impellatronics board. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11202 | 1220648-2025-26568 | 21546383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546383&pc=OZD | 02/12/2025 | 03/28/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604559 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was placed along with extracorporeal membrane oxygenation in a patient when purge alarms occurred. purge pressure was low and purge flow was high. they attempted to troubleshoot with purge cassette replacement and best practice for bypass was communicated. luer lock was possibly missing. the patient was in very good hemodynamic condition. however, the pump was weaned to p2 and explanted. the patient survived the procedure. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the purge pressure low and purge flow high near the end of the case was noted. replacement of the purge cassette could not resolved the issue. the pump then was explanted as the patient's hemodynamic was stable. data was not returned for review. visual inspection found a crack on the yellow luer of the purge side arm. leak was reproduced. no other issues were found on the rest of the pump. the root cause of the purge leak - pump was determined to be due to a damaged sidearm yellow luer due to high tension in combination with disinfectant and certain drug ingredients such as oil, alcohol, lipids etc. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
11203 | 1220648-2025-26569 | 21546418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546418&pc=OZD | 02/10/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533056 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the team placed the cp via a non-manufacturer approved or provided 16fr sheath. the limb became ischemic and the ecmo circuit was utilized along with a perfusion catheter line to perfuse the limb. the team eventually made decision to explant the cp and limb resolved. the patient remained on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿.
11204 | 1220648-2025-26570 | 21546729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546729&pc=OZD | 02/12/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605735 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. following implant, fluid was seen leaking from the purge filter and purge reservoir system. a purge filter bypass walk through was given to the site the manage the issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
11205 | 1220648-2025-26572 | 21546804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546804&pc=OZD | 02/09/2025 | 09/30/2025 |  | 03/07/2025 |  |  | 02/09/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572810 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Death | Yes | Yes | Malposition of Device | Deformation Due to Compressive Stress | Device Sensing Problem |  |  |  |  | Hemolysis | [12 empty] | Death | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). however, during insertion, a kink in the catheter was noted to have damaged the cable and aortic monitoring was dysfunctional. the impella cp device continued to support the patient. two days later, hemolysis was noted when blood draws from a-line, but no hemolysis when blood drawn from v-line. echocardiogram was completed and showed the impella catheter in too far. the physician decided not to reposition the pump at such time. two days following, the patient was noted on a very high dose of levophed and vasopressin. labs are all hemolyzing. therefore, a new set was initiated. good flow on the impella placement was observed on echocardiogram. the outcome of the patient as a result of the event is unknown. however, the patient continues impella mcs. | Additional information received the patient expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the data logs confirm placement signal not reliable alarm. at the instant of failure, the signal-to-noise ratio (snr) decreased to zero, contrast decreased. gain and lamp increased. clinical mentioned that there might have been a kinked catheter. the 5s parameters were characteristic of a possible fiber break due to a kink. the root cause of the optical signal issue could not be definitively determined as no product was returned for analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-26572. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The investigation of product damage (cannula kink), positioning issues, and hemolysis has been completed. the root cause of the hemolysis was most likely biomaterial as evidenced by ingestion pattern in the logs. the root cause of the positioning issues and product damage (cannula kink) was not determined as no product was returned for analysis and limited information related to time of failure occurrence was provided. | B2 revised to reflect the additional information received regarding the patient outcome death. b5 revised to reflect the additional information received regarding the patient outcome death. d6b explant date added. h1 revised to reflect the additional information received regarding the patient outcome death. h6 health effect - impact code 1802 added reflecting patient outcome.
11206 | 1220648-2025-26571 | 21546836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546836&pc=OZD | 02/11/2025 | 03/14/2025 |  | 03/07/2025 |  |  | 02/11/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601903 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella 5.5 support developed a hematoma at the insertion site. the physician evacuated the site and heparin was paused. the patient was eventually weaned and explanted. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11207 | 1220648-2025-26573 | 21546965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21546965&pc=OZD | 02/12/2025 | 04/01/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531349 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was leaking from the sidearm. visible damage was noted to the sidearm. the impella was exchanged and there was no reported patient harm. | [5 empty] | The investigation for the purge leak-pump has been completed. the product was returned and evaluated. the pump and purge cassette were inspected for any visible damage and a scratch mark/potential crack was found on the face of the reservoir. the pump was purged for 6 total hours and no external leaking was observed from the sidearm. upon closer inspection, it appeared that the purge fluid had leaked past the seal of the reservoir "bubble," but had not leaked external to the sidearm. the leaked purge fluid inside the reservoir had not seeped through the potential crack found on the face of the reservoir. the field data logs show the purge pressure stable throughout the case around 450 mmhg and no evidence of leaking. the clinical details mention that the damage occurred during the case and the purge sidearm was cracked/damaged during transport or check in to icu. it is unknown if the leaking into the radial seal was due to an out-of-box abnormality or a result of the damage/crack. the root cause of the pump purge leak was most likely a damaged sidearm reservoir but the cause of the damage is unknown. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿.
11208 | 1220648-2025-26574 | 21547226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21547226&pc=OZD | 02/05/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/05/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1338418 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Failure to Read Input Signal | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient turned down for coronary artery bypass graft surgery presented for a high-risk percutaneous coronary intervention (pci) with impella cp mechanical circulatory support (mcs). during initial purge setup, the automated impella controller displayed a purge disc not detected message. the aic was exchanged and no further issues with the purge system were encountered. the patient, however, would eventually pass away. the impella cp had been placed and started in auto with decision to place performance level p-5 and eventually increasing to p-7 - 3.1l for the procedure. contralateral access was obtained for pci and extensive effort and time for team to confidently wire the left anterior descending (lad) and left circumflex (lcx) arteries; ultimately access was successful. the left main and lad were then ballooned and stented from mid-lad to proximal left main with "great" result. after the pci was completed, repositioning sheath was then tipped into the 16fr. cook sheath and per the plan it was left in for the following 24-48 hours for the patient to rest on support. the following day after start of the impella mcs, continuous renal replacement therapy (crrt) was started to help with fluid removal and function. the patient continued on impella mcs and crrt for two additional days. last day, overnight, the patient was intubated, maxed out on pressors and later this day the patient would expire. family withdrew care. the cause of death is unknown at this time, and the temporary period to address the purge disc issue compromising support is an unclear association with the event outcome and therefore cannot be dissociated. after exchange of the aic, the patient would continue with impella mcs for an additional three day until his/her demise. medical safety reviewed the event noted there is not enough evidence to exclude aic as an associated factor in the patient's outcome. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the console logs from the reported day of the event revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. a product history review was completed, and no action was required as a result of the review. in conclusion, the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. applicable repairs and full functional check on console and optical system were completed.
11209 | 1220648-2025-26575 | 21547412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21547412&pc=OZD | 02/15/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/15/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589901 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump being placed to support a patient. at the end of the procedure an angiogram of the right leg and left leg showed a dissection and compromised flow to left femoral artery. the team noted the cp was causing occlusion to the superficial femoral artery. the decision was made to consult vascular surgery and send patient directly to cardiovascular operating room for intervention. the patient was escalated to impella 5.5 via axillary cutdown by cardiothoracic surgery, a double barrel technique was utilized. the cp was explanted from right femoral artery (rfa) and cutdown to rfa performed to address limb ischemia. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11210 | 1220648-2025-26576 | 21548057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21548057&pc=OZD | 02/07/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/07/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567762 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the impella cp was explanted due to bleeding and hematoma at the access site area. the patient was resuscitated during implantation. the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11211 | 1220648-2025-26577 | 21548569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21548569&pc=OZD | 02/04/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611410 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the day after upgrading to impella 5.5, the patient expired while still receiving ecpella support. a computed tomography scan showed superior mesenteric artery (sma) occlusion due to a thrombus. it is unknown what cause this occlusion. as a result of the sma occlusion the patient subsequently died. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
11212 | 1220648-2025-26579 | 21548664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21548664&pc=PYX | 02/16/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/16/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025548103 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient on impella rp flex support. while one support the doctor decided to place another figure of 8 stitch due to some slight oozing noted on one side of the first stitch and the sheath was forward tension sutured and angled match dressing was applied. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿
11213 | 1220648-2025-26580 | 21548868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21548868&pc=OZD | 02/12/2025 | 04/11/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609428 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER, IMPELLA RP FLEX. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the device was unable to go above p2 without triggering continuous suction alarms. the device was repositioned, volume was provided, and an implanted impella rp flex device was also adjusted, but the issue persisted. the pump was ultimately removed as a result of the low pump flow. upon an initial data evaluation, it was recognized that the pump had also experienced a pump stop. no patient harm was reported. | [5 empty] | The investigation of low pump flow and pump stop has been completed. the device was returned to the lab and there were no kinks in the cannula seen, biomaterial or damages to the impeller blades. in the lab run the motor current was approximately 350 ma at p-2 which is the expected range for the motor current for that pump. there was slight saturation of flows towards the beginning but that shortly resolved. the field data logs show that there is variability in motor current and flows at the beginning. when the pump is started again, motor current is still very low and begins to trend down reaching as low as 50ma with at p-2 with no change in p-level. flows are almost 0. about 20 seconds after it reaches 50ma the motor current spikes to 32767 and continues to have those spikes for 4 minutes with (no changes in the p-level). it was changed to p-1 and then a high motor current pump stop occurred and then the pump was disconnected approximately 9 seconds later. there was no problem found with the pump. h5 selection was erroneously entered on the initial manufacturer device report number 1220648-2025-26580. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11214 | 1220648-2025-26581 | 21549092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21549092&pc=OZD | 02/12/2025 | 10/01/2025 |  | 03/07/2025 |  |  | 02/12/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610528 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the device exhibited a placement signal issue. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella cp was removed for escalation to impella 5.5 and upon removing the pump and 14fr, bleeding was noted at the access site. preclose sutures failed. manual pressure was applied and blood products were administered. access was obtained on the left femoral artery and covered stent and balloon tamponade were applied to the right femoral artery. the patient condition was noted to be critical. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. only pump was returned for investigation. no physical abnormalities were observed, and all the 5s optical parameters were within specifications. during testing, the pump was receiving a psnr alarm. pump was then sent for pressure testing for optical sensor. the optical sensor maintained pressure, and ps was available when the pressure was set to 100 mmhg and higher. the pressure test data log was reviewed and noted that the -110 ps adjustment was automatically applied, resulting in a negative reading in the ps value and causing the psnr alarm when operating at pressures less than 100 mmhg. field data log: case start data log and imc logs were not provided for investigation. the transfer console ic3234 data log was reviewed, and no abnormalities were observed in the 5s optical signal parameters. the log suggests that there was a negative placement signal from the start of the case, with intermittent losses in the placement signal during the case. a -110 ps adjustment was noted from the start of the case on the transfer console ic3234. the cause of the optical signal issue was not determined since sufficient clinical information was not provided such as case start data log. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed has been completed. only pump was returned for investigation and not the introducer. no physical abnormalities were observed, and all the 5s optical parameters were within specifications. as per the provided clinical information, the perclose suture failed after the removal of the 14fr sheath. the access site bleeding was managed with a covered stent and balloon tamponade. vascular repair was required to control breeding. the cause of the vascular damage issue was not determined since introducer was not returned and sufficient clinical information was not provided.
11215 | 1220648-2025-26582 | 21549392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21549392&pc=OZD | 02/09/2025 | 03/07/2025 |  | 03/07/2025 |  |  | 02/09/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609837 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hemorrhagic Stroke | [10 empty] | Disability | Required Intervention | Death | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock (ami/cgs) implanted with an impella cp device for mechanical circulatory support (mcs) limb ischemia, access site bleeding and intracranial hemorrhage with a demised outcome. the patient was taken for emergent percutaneous coronary intervention (pci) to left anterior descending (lad) artery. during the stay, patient began to experience respiratory distress, progressively became worse the following morning and was emergently intubated and required inotrope/pressor support. troponin level was greater than 200, and the physician was concerned that lad stent was occluded. the decision was made to place an impella cp impella in the presence of ami/cgs, which was successfully implanted without incident. following the implant, however, both of the patient's legs were cold to the touch with no palpable pulses. distal flows were optimized and the issue resolved by the following day. a computed tomography (ct) scan then showed a small subarachnoid hemorrhage with no retroperitoneal bleed or any other source of bleeding. the clinical consultant went back to the ct scanner to check again, and it noted intraparenchymal hemorrhage. the patient's neurological responses declined. the following day another ct scan was completed, and it showed stable hemorrhage. the decision was made to hold plavix and start heparin. a bleed from the impella site developed, and one unit of fresh frozen plasma and one unit of packed red blood cells were given as treatment. support continued for approximately another four days with the implanted impella cp pump despite the noted conditions. it was noted the patient also remained on mechanical ventilator support, veno-arterial-venous extracorporeal membrane oxygenations, and continuous renal replacement therapy. after five days of total support, the patient's family made the decision to withdraw care due to the patient's lack of neurological progress. the patient expired, and the impella related complications may have possibly added to or compounded an already complex situation. medical safety review of the event noted there is not enough information to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: precaution impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11216 | 1220648-2025-26583 | 21552804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21552804&pc=OZD | 02/17/2025 | 04/07/2025 |  | 03/07/2025 |  |  | 02/21/2025 | 03/07/2025 |  |  | 10/31/2025 |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2024417469 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cerebral Edema | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 64 year old female patient on an impella 5.5 pump for escalation of therapy and the requirement for long term therapy. the patient's heart function improved some, however, the patient died of cerebral edema. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 selection requiring intervention was erroneously selected on the initial manufacturer device report number 1220648-2025-26583 b3 date of event was erroneously on the initial manufacturer device report number 1220648-2025-26583. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26583 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26583. g2 report source; foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26583. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this complaint is for escalation to the impella 5.5 device from impella cp device used in the 1220648-2025-26536.
11217 | 1220648-2025-26584 | 21553418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21553418&pc=OZD | 02/08/2025 | 07/15/2025 |  | 03/07/2025 |  |  | 02/08/2025 | 03/07/2025 |  |  |  |  | 03/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087214 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown ethnicity) with cardiomyopathy was implanted with an impella device for mechanical circulatory support. while on support, approximately 38 days later, the patient was ambulating when the air in purge system alarm triggered and was unable to be resolved until exchange of the automated impella controller (aic). it was noted a placement signal not reliable alarm triggered as well and resolved with the aic exchange. there was no resulting harm to the patient as a result of this aic event. the patient was noted to be in stable condition and remained on impella mcs until bridged to a left ventricular assist device 20 days later. | [5 empty] | D9 device returned to manufacturer date added. h3 revised to reflect the device was received for investigation. data review: console logs are partially consistent with what was reported in the complaint. there were multiple instances of air in purge system alarms [event set (b)(4)] observed in the imc logs. rtlog data from the time of the air in purge system alarm show that the purge pressure momentarily did not increase per forward stroke of the cassette¿s piston. attempts were made to de-air the purge system but the air in purge system alarm would not resolve. no evidence of placement signal issue was observed in the imc logs from this case. device analysis: console was tested after arriving in field service. the reported air in purge system and placement signal issues could not be reproduced when running a known-good pump and purge cassette. the console operated as intended while testing on an engineering lab bench. conclusion: despite the issue resolving when swapping to the backup controller, there are other possible causes of an air in purge system alarm that were not confirmed, which are unrelated to the console: air in the purge system. kinked or open purge lines. empty purge fluid bag. the cause of the air in the purge system alarm was not determined due to the inability to reproduce. no placement signal issues were found during this case (no problem found). repair: before sending this console back to the customer, field service was instructed to perform sections 10-12 of pm procedure [0042-7309] as a precaution and perform a full functional check on the console and optical system [0042-7316]. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26584 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The automated impella controller device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11218 | 1220648-2025-26534 | 21535301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535301&pc=PYX | 02/10/2025 | 03/27/2025 |  | 03/06/2025 |  |  | 02/10/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025571400 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in postcardiotomy cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support (mcs). following a difficult delivery, (the impella was unable to advance through the sheath initially), a critical alarm appeared on the automated impella controller when the impella rp flex device was started for patient support. after several failed troubleshooting attempts, the pump ultimately would not start and had to be replaced, which was successfully completed from a different access point. the following day, the patient passed away noted from the ("very poor prognosis") condition. impella as a possible associated factor cannot be discounted; it is unclear how the impella related event may have affected the outcome. | [5 empty] | The investigation of pump did not start has been completed. the data logs are unavailable. the product was returned and evaluated. two large kinks were found in the catheter. no other abnormalities were found. the pump passed all electrical testing and was run in a bucket of water for 15 minutes without issue. the root cause of the pump not starting was not determined since the issue was not reproduced and data logs are unavailable. g1 manufacturing site name and address was erroneously entered on the initial manufacturer device report number 1220648-2025-26534. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
11219 | 1220648-2025-26535 | 21535360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535360&pc=OZD | 01/25/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/10/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612185 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and approximately three day later, a fluid leak was observed from the purge sidearm while on impella mcs. fluid was noted to be leaking from the connection between the filter and the pressure reservoir. as the parameters were still within the allowable range, the decision was made to maintain support while monitoring the situation. there was no harm to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
11220 | 1220648-2025-26536 | 21535513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535513&pc=OZD | 02/10/2025 | 03/13/2025 |  | 03/06/2025 |  |  | 02/10/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547577 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | After further review, the patient was reported as successfully escalated from the impella cp to the impella 5.5. the patient remained on the impella 5.5 for an additional seven days before expiring. | The user facility reported a patient with cardiomyopathy implanted with an impella cp device for mechanical circulatory support (mcs) experienced a stoke and subsequently expired after explant of the device for escalation in care. the patient was admitted in with fulminant myocarditis and initially managed with an intra-aortic balloon pump (abp) and extracorporeal membrane oxygenation (ecmo). the patient's condition did not improve and therefore the patient was escalated from an iabp to an impella cp device for mcs. as it was determined the patient would require long-term support, escalation to an impella 5.5 was considered. a computed tomography scan for preliminary evaluation was completed, and a cerebral hemorrhage was observed. event noted to have occurred approximately 8 days after starting impella cp mcs. it was noted heparin in the purging fluid was 10u/ml and activated clotting time was controlled at 160-180 seconds. the patient was escalated to the impella 5.5 and would remain on impella mcs for an additional 7 days before expiring. additional information noted the cerebral hemorrhage was in the occipital lobe and was determined that there was no risk to life. the cause of death was noted as cerebral edema. the physician commented the cause of the edema, however, is unknown. the patient had myocarditis, and therefore, may have been caused by a virus, etc. there is no relationship between impella and the edema per the report. the cerebral hemorrhage, however, may have been caused by the anticoagulant associated with the use of the impella. although it is still unclear whether the cerebral hemorrhage was caused by the impella cp. although the cause of death cerebral edema with an unknown cause was noted to be unrelated to the impella, there is not enough information to discount the impella cp device given the unclear association with the hemorrhagic stroke. |  |  |  |  | After further review of manufacturer device report 1220648-2025-26536, the patient was reported as successfully escalated from the impella cp to the impella 5.5. the patient remained on the impella 5.5 for an additional seven days before expiring. the death while on impella 5.5 has been captured in manufacturer device report 1220648-2025-26583. please see the following revisions to manufacturer device report 1220648-2025-26536 to reflect these updates. b.2 death and date of death omitted. b.5 revised information . h.1 type of reportable event revised to serious injury . h.6 health effect impact code 1802 and 4614 reported in error - new code added for escalated device to impella 5.5. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11221 | 1220648-2025-26537 | 21535599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535599&pc=OZD | 02/11/2025 | 03/26/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597742 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the upon explant of impella 5.5 on a 61-year-old male patient, a clot was noted on the tip of the impella. there were no device flow issues or problems with functionality. there was no patient harm reported. the patient survived the procedure. | [5 empty] | The investigation has been completed. biomaterial was found around the impeller in the outflow cage. no other damages or abnormalities were found. as the necessary information was not provided, the root cause of the thrombus was not determined. b.2 revised as selection reported on the initial manufacturer device report number 1220648-2025-26537 was incorrect. e.1 added fax number and us phone number as they were inadvertently omitted from initial manufacturer device report number 1220648-2025-26537. g.1 added contact facility name as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26537. h.6 code 4582 and 2199 were reported incorrectly on the initial manufacturer device report 1220648-2025-26537. new codes were added to health effect - clinical code and health effect - impact code. | The impella device has been received from the customer but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11222 | 1220648-2025-26538 | 21535664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535664&pc=OZD | 02/11/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608879 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 46 year old on impella cp to support a high risk cardiac procedure. it was reported that after the pump was placed, it had to be explanted due to a kink in the cannula which was giving poor flows. the case continued without impella support. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11223 | 1220648-2025-26484 | 21535676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535676&pc=OZD | 02/06/2025 | 03/14/2025 |  | 03/06/2025 |  |  | 02/06/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1111989 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and approximately two days after start of the support, there was an unexpected shut down of the automated impella controller (aic). the loading screen displayed and then the aic automatically restarted. there was no harm to the patient as a result of the event. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. the console logs from the reported day of the event are consistent with the complaint, showing an unexpected shutdown, followed by "unexpected shutdown" alarms after the boot-up. there were no warning signs or precursors prior to the unexpected shutdown, and no software processes showed signs of failure. the pump operation was interrupted and restarted after 30 sec of reboot. the console was tested, but the issue could not be reproduced during testing with the pump. no unexpected shutdowns or performance issues were observed. a thorough visual inspection of the internal components was conducted, but no physical problems were found. there was no interruption in the supply to the 4-in-1 (epc). however, the aic unexpected shutdown can most likely caused due to a undetermined malfunction of the 4-in-1 assembly and/or power battery manager (pbm) assembly. the root cause of the unexpected shutdown was not determined due to the inability to reproduce. however it can most likely caused due to a undetermined malfunction of the 4-in-1 assembly and/or pbm assembly. there is no evidence of a design or manufacturing defect associated with the failure modes above. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26484. d.4 revised serial number, lot number and unique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26484. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26484.
11224 | 1220648-2025-26539 | 21535888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535888&pc=OZD | 02/11/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540754 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that multiple complications were encountered during impella cp support. it was initially reported that the pump experienced suction events shortly after implant. repositioning was attempted, but the catheter was inadvertently pulled out across the aortic valve. the support level was reduced and the decision was made to escalate to an impella 5.5 for support. the pump was thrown away. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11225 | 1220648-2025-26540 | 21535912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535912&pc=OZD | 02/11/2025 | 04/01/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025596993 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during placement of an impella 5.5, the placement was aborted due to difficulty in advancing catheter into the aorta. there was no reported patient harm. | [5 empty] | The impella device was received from the customer but has not yet been evaluated. when the investigation has been completed or any new information is received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to the patient's anatomy. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26540.
11226 | 1220648-2025-26541 | 21535983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21535983&pc=OZD | 02/07/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/07/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551253 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and prior to starting the inserted impella cp pump, an impella defective alarm appeared on the automated impella controller (aic). the device was unplugged and plugged back in; however, the issue persisted. the pump was subsequently removed and a new impella pump was placed for patient support. there was no harm to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11227 | 1220648-2025-26542 | 21536041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536041&pc=OZD | 02/07/2025 | 03/17/2025 |  | 03/06/2025 |  |  | 02/07/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559939 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced oozing at the access site. heparin was removed. sutures and femostop were placed. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was most likely use issue since impella was not sutured in and was secured via statlock instead. also repositioning sheath was 1 inch outside the body. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26542 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11228 | 1220648-2025-26543 | 21536048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536048&pc=OZD | 02/08/2025 | 03/13/2025 |  | 03/06/2025 |  |  | 02/08/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301039 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. purge disc not detected alarmed during cassette change. there was no damage to the blue disc hold observed. the nurse reported that the same alarm presented after cassette replacement twice. the aic was replaced and the alarm resolved. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from after the reported complaint date revealed that the console issued purge disc not detected alarm. the device analysis revealed a broken purge disc flag was identified. the root cause of this issue was a broken purge disc flag. d2 common device name was revised on manufacturer device report in accordance with updated procedures on manufacturer device report 1220648-2025-26543. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26543 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11229 | 1220648-2025-26544 | 21536055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536055&pc=OZD | 02/08/2025 | 03/17/2025 |  | 03/06/2025 |  |  | 02/08/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507953 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMTED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced episodes of ventricular tachycardia during impella 5.5 support. the patient was shocked over 20 times in the period of a twenty-four hours to manage the condition. the pump was repositioned in response to the episodes and a ganglion block was utilized to resolve the issue. support continued with the implanted pump. twenty-four days into support, care was withdrawn and the patient passed away. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of vf/vt/af/at has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26544 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11230 | 1220648-2025-26545 | 21536072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536072&pc=OZD | 02/09/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/09/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515056 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Thrombosis/Thrombus | [10 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella 5.5 device for mechanical circulatory support (mcs) experienced a clot on the catheter inlet and an access site hematoma requiring one unit of red blood cells. the patient would subsequently expire after choosing palliative care. the patient underwent an emergent left heart catheterization for st-segment elevation myocardial infarction; thrombectomy to the left anterior descending artery was completed, and an iatrogenic lad perforation occurred. while on the unit, the patient decompensated and the decision was made to place an impella 5.5, which was successfully completed via the right axillary artery. the patient was transferred back to the unit in stable condition. approximately eight days after starting impella mcs, an echocardiogram was completed and a clot on the impella inlet was observed. the patient was on systemic heparin; plans were to monitor motor trends. the following day, performance level was reduced due to a suspected clot on the left ventricle. two days thereafter, a hematoma at the access site was noted requiring one unit of packed red blood cells. the patient had increasing pressor and dobutamine requirements with continuous renal replacement therapy fluid removal overnight. central venous pressure was on the higher end despite aggressive fluid removal. the plan was discussed with staff to continue with current therapies; however, potentially switch to comfort care. the following day the patient went into atrial fibrillation and an amiodarone drip was started. the patient stated being tired and no longer wanted to be on support. therapies would continue until the next day. the patient with family decided on palliative care. performance level was decreased to p-2 (from p-4) to prepare for impella explant by the surgical team, which was successfully completed. the patient was hemodynamically stable until the demise. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11231 | 1220648-2025-26546 | 21536079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536079&pc=OZD | 02/11/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578938 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 69-year-old male patient had an impella cp placed due to a stemi. after the cp was placed, there was suction that was treated by red blood cells (rbcs) and other blood products/volume. the team then observed there was a hematoma at the access site. patient condition continued to deteriorate and the team observed in the or that the lv was perforated. the perforation was repaired and the pump remains on while patient being sent to tertiary center. further details state that the patient subsequently expired while on support as the bleeding could not be controlled. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11232 | 1220648-2025-26547 | 21536084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536084&pc=OZD | 02/10/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/10/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490183 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) implanted with an impella cp device for mechanical circulatory support (mcs) experienced access site bleeding. the impella cp was implanted without incident, and the pci was completed. the pump weaned at end of case and removed. perclose sutures failed resulting in a bleed from the access site. manual pressure was held, protamine was given, and femostop was placed. the patient was noted to be in "good" condition following the event, transferred to the unit for recovery. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11233 | 1220648-2025-26548 | 21536157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536157&pc=OZD | 02/11/2025 | 03/24/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577886 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp pump was placed to support a high risk percutaneous coronary intervention (hrpci) patient. the patient had poor vasculature anatomy with the right iliac artery being occluded. left femoral access was obtained with micropuncture. upon angiography of the distal vessel, the team found occlusions and elected to obtain antegrade access in the event that a bypass might be needed for flow. the antegrade access was obtained by the provider who left the wire in place with no sheath. the patient remained on the support for the hrpci til weaned and explanted after two hours. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
11234 | 1220648-2025-26550 | 21536300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536300&pc=OZD | 02/11/2025 | 03/28/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589291 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an 85 year old male patient on an impella 5.5 for support during a high risk surgery. it was reported that the impella 5.5 was not able to flow above 2l/min without suction alarms. the 5.5 was repositioned but it did not resolve the low flow. the physician opted to leave pump in place despite reduced flows. motor current stable. the patient was successfully weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the product was not returned for investigation. data log review confirmed the low flow issues early in the case. clinical details report that the first pump was explanted with a significant clot, which supports the likelihood that the complaint pump picked up some clot during implant as well. based on clinical details provided and data log review, the root cause of the low pump flow was determined to most likely be biomaterial. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26550 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26550.
11235 | 1220648-2025-26549 | 21536301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536301&pc=OZD | 02/10/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591219 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was placed on an impella cp pump for support. after the pump was placed, there was an observation made of a pericardial effusion that prompted the team to perform pericardiocentesis. the pump was explanted and the patient survived. the pump was discarded. additional information was obtained. the care team is not sure what caused the wire perfusion. the team was trying to cross a chronic total occlusion in the left circumflex artery and that is where the perfusion happened. the perfusion was noted the next day. the patient did go down to the catheterization lab for pericardiocentesis where they drained the fluid. the patient is in stable condition with a return of normal heart function. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11236 | 1220648-2025-26551 | 21536324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536324&pc=PYX | 02/11/2025 | 03/25/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025598981 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 54 year old female patient on impella rp. the pump was removed due to increase of placement signal, flows dropping, purge pressure and purge flows increasing. additionally, plasma free jumped from 10 to 52 and ldh was hemolyzed. the patient is stable. | [5 empty] | The investigation of low pump flow and hemolysis has been completed. the data logs confirmed the failure mode and clinical narrative. logs showed after pump started and run for 63 hours, motor current (mc) spiked suddenly spiked to 1600 ma that triggered auto suction response and suction alarms. product analysis: visual inspection found foreign material inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. conclusion: most likely, the foreign material was ingested into the cannula which indicated by mc spike, then biomaterial was building up inside the purge gap caused high pressure purge (pp) but resolved and low flow. no biomaterial was found during around the impeller under inspection visually. there was a small amount in the purge gap. no other issues were found on the rest of the pump. low pump flow was not reproduced in the lab. the root cause of the low flow was most likely biomaterial ingestion based on the ingestion pattern in the data logs. the root cause of hemolysis was most likely foreign material and biomaterial ingestion. d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26551 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26551. | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller states the following: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
11237 | 1220648-2025-26552 | 21536353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536353&pc=OZD | 02/11/2025 | 03/25/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568545 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support on a 84 year old female patient. following insertion, the physician was unable to advance the pump across the aortic valve. it was noted that the patient had severe aortic stenosis and that delivery of the device was expected to be difficult prior to insertion. the implant was aborted with no resulting patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to the patient's anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26552 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26552. | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11238 | 1220648-2025-26553 | 21536511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536511&pc=OZD | 02/11/2025 | 03/14/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591706 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed on an impella cp pump for ongoing support for acute myocardial infarction/cardiogenic shock. the patient had signs of hemolysis on day two of the support. the pump position was changed, as were p levels and the urine color cleared. standard troubleshooting resolved the signs of hemolysis. the patient was successfully weaned and explanted. | [5 empty] | The investigation of hemolysis has been completed. no product was returned for analysis. the clinical information mentions that the pump was too deep at 6.5cm past the valve. the data logs showed no abnormalities. the root cause of the hemolysis was most likely due to improper positioning of the pump as evidenced by resolution of hemolysis after pump reposition. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11239 | 1220648-2025-26554 | 21536624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536624&pc=OZD | 01/07/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/10/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590049 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received from the protect iv study regarding a patient who underwent high-risk percutaneous coronary intervention (pci) and was implanted with an impella cp device for mechanical circulatory support. the study reported after conclusion of impella mechanical circulatory support, the patient was seen in follow-up (approximately 41 days post device explant) and was noted to have a possible suture granuloma to the left groin access site, which was treated with antibiotics. the patient was free of signs and symptoms of an active infection. the event was determined to be not related to the pci but rather definitely related to th impella procedure and the impella device used. the outcome was noted as "ongoing;" however, no further report regarding this and other events has been received. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist, section: table 3.2 impella catheter components, ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings, ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11240 | 1220648-2025-26555 | 21536708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536708&pc=OZD | 02/07/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/08/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607922 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and approximately four days after start of support, the patient developed limb ischemia in the impella leg. the staff was unsure if the ischemia was due to the repositioning sheath or if the patient had a thrombus in the area. activated clotting time was between 150-160 seconds during impella therapy. the impella cp device was explanted, and the patient expired the same day as the ischemic event. the pump worked well so the user facility threw the pump in the trash. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11241 | 1220648-2025-26556 | 21536834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21536834&pc=OZD | 02/07/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/07/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605488 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the insertion of impella cp device to a patient was unsuccessful. a right axillary cutdown was performed and the graft was attached. the physician was unable to pass the impella, and the case was, therefore, aborted. other details were unknown. the patient was noted to be in stable condition following the event. the pump was not kept. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11242 | 1220648-2025-26557 | 21537941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21537941&pc=OZD | 02/12/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/12/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589805 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella cp implant, the access site developed a persistent bleed. manual pressure was held for 45 minutes, but the condition persisted. intravenous albumin was initiated, packed red blood cells were transfused, and vasopressors were utilized to support the patient's change in blood pressure. the decision was ultimately made to explant the pump. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11243 | 1220648-2025-26558 | 21540302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21540302&pc=OZD | 02/11/2025 | 03/27/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585980 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the impella was unable to pass through the patients vasculature due to calcification. it placed over .035 glidewire advantage for more pushability. viperslide utilized again without avail. after an additional 20min of attempted placement, decision made to use 10mm graft for arterial extracorporeal membrane oxygenation cannula. the procedure was aborted. the patient was noted to be stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the returned product was visually inspected and a cannula kink was observed. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature. calcification of the vessels was reported. h6 codes 4611 and 4627 were reported incorrectly on manufacturer device report 1220648-2025-26558. new code was added to health effect - impact code.
11244 | 1220648-2025-26559 | 21540594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21540594&pc=OZD | 02/04/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/04/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584964 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Perforation | Cardiovascular Dysfunction/Insufficiency | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella 5.5 pump to support a high-risk percutaneous coronary intervention patient. the team used impella and non-impella product. there is an allegation that the non impella wire caused a left ventricle perforation that was observed and was repaired. the patient survived. later in the month of the ongoing 5.5 pump support the loqi database reported upon 3 adverse events, 2 of which were of unknown relationship to the pump, these being right ventricle failure and atrial arrythmia/atrial fibrillation with rapid ventricular rate, and the 3rd was definite relationship, this being bleeding. care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11245 | 1220648-2025-26560 | 21542363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21542363&pc=PYX | 02/11/2025 | 04/01/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551046 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an impella rp flex pump was inadvertently pulled back during patient support. it was noted that the suture hub had broken off the repositioning sheath but was still sutured intact to the neck. the patient passed away later that day. it is unknown if the complication contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the product was not returned for review. it was noted in the complaint description that the suture pad was broken off the repo at the same time the impella rp flex was found to be dislodged into the ivc, indicating force was likely a factor. other details surrounding wat was happening at the time of the dislodgment are unknown. the root cause of positioning issue was unable to be determined as limited clinical details were provided e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26560 it is unknown if the initial reporter also sent the report to the food and drug administration.
11246 | 1220648-2025-26561 | 21542365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21542365&pc=OZD | 02/11/2025 | 03/06/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  | 02/28/2026 |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482015 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that an impella cp pump was in use for roughly 16 days when it stopped unexpectedly. attempts to restart the pump were unsuccessful and the decision was made to replace the device. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11247 | 1220648-2025-26562 | 21542370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21542370&pc=OZD | 01/22/2025 | 03/25/2025 |  | 03/06/2025 |  |  | 02/11/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573327 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced arrythmia during impella 5.5 support. actions taken are unknown. the patient continued on support until the device was successfully weaned. the patients condition was noted to be stable. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-26562. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26562 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26562. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11248 | 1220648-2025-26563 | 21542376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21542376&pc=OZD | 02/12/2025 | 04/01/2025 |  | 03/06/2025 |  |  | 02/12/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553371 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a cerebral hemorrhage and was noted to be in critical condition. the patients blood pressure was controlled by administering nicorandil and after cerebral hemorrhage was confirmed, the addition of heparin to the purge fluid was discontinued. the exact cause of the cva is unknown, and the causal relationship with impella is also unknown. the patient continued on support. additional information noted care was the patient was made do not resuscitate and the patient expired. the cause of death was acute myocardial infarction and cerebral hemorrhage. no issues were identified with impella operation. the causal relationship between cerebral hemorrhage, one of the causes of death, and impella is unknown. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cerebrovascular accident (cva) has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. e.1 added facility name as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26563. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26563 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.2 report source; foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26563.
11249 | 1220648-2025-26564 | 21542378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21542378&pc=OZD | 02/12/2025 | 04/01/2025 |  | 03/06/2025 |  |  | 02/12/2025 | 03/06/2025 |  |  |  |  | 03/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610683 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump had suction issues due to a kink and the pump was removed/replaced. no patient harm was reported. | [5 empty] | The investigation for the product damage-cannula kink has been completed. clinical details stated that the pump was exhibiting suction and a kink was noted prior to accessing with companion sheath. repositioning was attempted but unsuccessful and pump was removed. the pump was returned for evaluation and a cannula kink was found near the outflow cage. no other damages or abnormalities found. the root cause of the product damage (cannula kink) was not determined due to insufficient clinical detail. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26564. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26564. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26564 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11250 | 1220648-2025-26479 | 21524850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21524850&pc=PYX | 02/11/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590523 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following patient repositioning, an implanted impella rp flex pump experienced a drop in flow rate and spike in motor current. no clot was visible via echocardiogram. as a result of the noted issue, the decision was made to explant the device. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11251 | 1220648-2025-26480 | 21524917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21524917&pc=OZD | 02/10/2025 | 04/22/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604768 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella 5.5 axillary cutdown and anastomosis, there was initially some bleeding from the anastomosis that was then repaired. additionally, the patient had a hematoma and bleeding from the access site. the patient was taken to the operating room for a washout. two units of blood were administered, and heparin was held. the patient remains on support. | It was further reported that care was withdrawn, and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely patient condition since the patient had bleeding from multiple sites. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B2 outcome revised based on the additional information received from the clinical to site to include death and date of death b5 revised based on the additional information received from the clinical to site. d6b explant date added h1 type of report revised based on the additional information received from the clinical to site. h6 revised based on the additional information received from the clinical to site to include death.
11252 | 1220648-2025-26481 | 21524990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21524990&pc=OZD | 02/10/2025 | 03/21/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611474 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during placement of an impella 5.5 pump in a patient for cardiomyopathy support, the procedure was aborted due to patient anatomy. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the delivery issue was most likely patient condition related. the impella was unable to pass through vasculature.
11253 | 1220648-2025-26482 | 21524998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21524998&pc=OZD | 02/06/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582375 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and transplant work up. approximately seven days after starting the support the patient experienced runs of ventricular tachycardia that required cardioversions. the impella 5.5 device was "hitting" the septal wall, repositioned and shallowed to 4cm. however, the patient continued to experience random bouts of ventricular tachycardia but was noted in stable condition. impella 5.5 mcs continued for another 18 days before being explanted for a left ventricular assist device placement. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11254 | 1220648-2025-26483 | 21525071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525071&pc=OZD | 02/11/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597693 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the introducers utilized for impella cp implant both kinked and were partially peeled in the peel-away section due to arterial calcification and the patient's tortuous anatomy. a medtronic 16 fr sheath was subsequently utilized for successful implant. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11255 | 1220648-2025-26485 | 21525215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525215&pc=OZD | 02/08/2025 | 03/25/2025 |  | 03/05/2025 |  |  | 02/08/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584995 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that following implant, the flow rate of the impella cp pump dropped unexpectedly. it was documented that flow rates were not exceeding .2-.3 l/min on all p-levels. the pump was subsequently replaced for ongoing support. there was no noted patient harm. | [5 empty] | The investigation of low pump flow has been completed. the data logs confirmed the failure mode and clinical narrative. logs showed after pump started, motor current (mc) spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. the visual inspection found no issues on the pump. the returned pump ran with saturated flows for an impella cp (4.0 l/min) indicating a presence of biomaterial despite not reproducing the low pump flow. although no biomaterial was found under visual inspection but based on data and clinical description review, the root cause of low flow was most likely due to biomaterial ingestion. the biomaterial probably had been expelled during/after pump explant. a2 patient age was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26485. a3 patient gender was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26485. a5 patient race and ethnicity was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26485. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11256 | 1220648-2025-26486 | 21525534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525534&pc=OZD | 02/07/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1673731 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that a 74 year old male patient was on impella cp support and during support, there were placement signal issues. ao position signal values deviated significantly from the blood pressure on the vital signs monitor. the pump was repositioned to no avail. the automated impella controller (aic) was replaced and the issue remained. the patient remained on support without any harm until successfully weaned from the pump. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11257 | 1220648-2025-26487 | 21525539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525539&pc=OZD | 02/04/2025 | 04/07/2025 |  | 03/05/2025 |  |  | 02/04/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605098 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that a purge pressure high alarm presented. lysis protocol was sent to the site to manage the condition. additionally, the physician reported that the patient's right leg had to be amputated due to the impella. the pump was explanted a few days later and the patient expired. it was noted that the patient's treatment was palliative. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of high purge pressure and limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of the high purge pressure was most likely due to biomaterial as lysis protocols were able to resolve issue according to clinical description. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b1 selection of product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26487. b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26487. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26487 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2025-26487 in accordance with updated procedures. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation."
11258 | 1220648-2025-26488 | 21525545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525545&pc=OZD | 02/05/2025 | 05/06/2025 |  | 03/05/2025 |  |  | 02/05/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572621 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to an unknown age patient undergoing a high-risk percutaneous coronary intervention was unsuccessful due to the anatomy. impella cp device was could not reach the patient's left ventricle due to an elongated / twisted aorta. the implant was subsequently aborted. there was no patient harm as a result of this event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Based on the clinical details the root cause of the delivery issue was most likely patient condition related, left ventricle (lv) was not reached by impella, aorta elongated and twisted d6a and d6b were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26488. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26488 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact last name was erroneously entered on the initial manufacturer device report number 1220648-2025-26488.
11259 | 1220648-2025-26489 | 21525548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525548&pc=OZD | 02/05/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/05/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025605568 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device via right axillary artery for mechanical circulatory support (mcs) with possible bridge to transplant. approximately five days after start of the support, a hematoma at the right shoulder was noted. in response, heparin was held. the patient continued on impella mcs pending bridge to transplant. twelve days later some oozing of the shoulder was noted. physician was made aware with no orders given. it is unknown if heparin was still on hold during this time. the patient continues on impella mcs, and latest update noted the patient is doing well still pending bridge to heart and kidney transplant. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: potential adverse events: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation).
11260 | 1220648-2025-26490 | 21525573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525573&pc=OZD | 02/05/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/05/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539152 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant and encountered removal difficulty. the patient received a heart donor. the intention was to advance the impella and transect the impella catheter with the donor heart and pull out the remaining catheter. however, the physician was unable to advance the impella in order to transect the catheter and had to pull the impella back through the vessel causing vascular injury requiring surgical intervention. it was noted that the patient's vessel size was small (approximately 4.5mm) prior to impella implant, and the physician stated he does realize there is a risk of vascular injury in vessels less than 7mm as per the instructions for use. the patient was reported to be in stable condition following the event. | [5 empty] | The explant date of the device was not provided. therefore, the date of the event will be reflected as the explant provisional date. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11261 | 1220648-2025-26491 | 21525594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525594&pc=OZD | 02/06/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600661 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction implanted with an impella cp device for mechanical circulatory support (mcs) developed hemolysis and would later expire. the patient was brought in by ambulance after being down with ventricular fibrillation arrest for approximately ten minutes. cardiopulmonary resuscitation was initiated and patient shocked three times with return of spontaneous circulation achieved. the patient was intubated in field and three pushes of epinephrine was given; the patient was brought to cardiovascular lab on 1mcg amiodarone drip. angiography revealed multi-vessel disease including chronic total occlusion of the left anterior descending artery. ultrasound guided rfa access obtained and 14fr medtronic sentrant sheath placed. impella prepped and placed through medtronic sheath and started without complications, reaching mean flows of 3.6l/min. no percutaneous coronary intervention was performed; the patient was referred to surgical consult. repositioning sheath was advanced into medtronic sheath and sutured to leg. the patient was transported to the unit on impella mcs. once on the unit it was reported lactate was trending down and urine output was darkening. the physician acknowledged the pump was too deep and did not want to reposition at such time as the pump was adequately functioning. this day the patient was made a do not resuscitate (dnr) and care plans with family were discussed. the following day the patient was reported to be stable but labs x4 hemolyzed. performance level was decreased to p-7 from p-9 and position check under echocardiogram was recommended. the status of the patient following the event was noted to be unknown. however, later this day, the patient would expire. the family decided to withdraw care; all therapies were discontinued and the patient passed away after 16 total hours of impella mcs. the cause of death is unknown. the impella related event may have possibly added to an already complex situation. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11262 | 1220648-2025-26492 | 21525625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525625&pc=OZD | 02/06/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584321 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy implanted with an impella cp device for mechanical circulatory support experienced limb ischemia. approximately seven days prior, the patient had to have re-perfusion catheter placed due to no flow of blood down right leg. it was felt that pt was ready to come off extracorporeal membrane oxygenation (ecmo) for limb salvage, which was eventually amputated, and impella 5.5 would be placed for continuation of care. however, the impella 5.5 device was unable to be placed due to the patient's anatomy. as the artery was very small, the physician did not believe he could get either pump (impella 5.5 or cp) due to size of artery and calcified vessels; however, the physician wanted to give the patient a chance. the impella cp was implanted via the right axillary artery with much difficulty and ecmo had to be turned up from 2l to 5l for compensation due to low blood pressure. the following day after starting support on the impella cp, it was noted the patient's hand was turning cold. treatment and/or actions taken to manage the event is unknown. the patient was noted to be in stable condition following the event still on support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two medical device reports associated with the patient's implant experience. refer to manufacturing report number 1220648-2025-26409 for the impella 5.5 device report. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11263 | 1220648-2025-26493 | 21525662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525662&pc=OZD | 02/06/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  | 02/28/2026 |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512490 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported an unknown age patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was inserted via right femoral artery without resistance and advanced over wire into ventricle. the guidewire was removed and the impella was started. flows remained under approximately 2l/m. when considering positioning for reason of low flows, imaging was obtained, and cannula was noted to be kinked under the motor. decision made to place new impella cp device and place an impella rp flex. there was no resulting patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed.
11264 | 1220648-2025-26494 | 21525668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525668&pc=OZD | 02/06/2025 | 05/07/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  | 02/28/2025 |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307328 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump being placed to support a patient admitted in with cardiogenic shock and cardiomyopathy. it was reported that there was unknown vessel complication and pushing the device forward was not possible. the impella placement was aborted due to patient anatomy. there was no noted harm or injury. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to vessel complications and resistance in aorta. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26494 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source; foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26494. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11265 | 1220648-2025-26496 | 21525670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525670&pc=OZD | 02/06/2025 | 03/14/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577889 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that after impella cp was explanted because distal pulse diminished via doppler. the physician performed embolectomy and distal pulsed improved. the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26496 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11266 | 1220648-2025-26495 | 21525675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525675&pc=OZD | 02/08/2025 | 03/22/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025595612 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that four days into impella 5.5 support, the patient had to go to the operating room for a washout of their axillary artery cutdown site. it was documented that a hematoma had developed at the site and the patient's trip to the operating room was necessary to find the source of the bleed. the patient was stabilized and support was continued uninterrupted. the patient was eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be use issue because of the bleeding from the graft anastomosis and the best practices were not followed.
11267 | 1220648-2025-26497 | 21525689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525689&pc=OZD | 02/07/2025 | 03/24/2025 |  | 03/05/2025 |  |  | 02/07/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582914 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the device was pulled back accidentally post impella placement and unable to advance back across aortic valve. the device was removed and an impella cp was inserted instead. the patient outcome was noted to be stable. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issues most likely was use related since the impella pulled back accidentally post procedure and they were unable to advance back across aortic valve. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26497. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26497 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11268 | 1220648-2025-26498 | 21525749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525749&pc=OZD | 02/07/2025 | 03/19/2025 |  | 03/05/2025 |  |  | 02/07/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529761 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support the patient had questionable arterial flow in their right lower extremity distal to the impella implant site. the pump was removed in response to the condition. the patient was considered stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
11269 | 1220648-2025-26499 | 21525759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525759&pc=PYX | 02/10/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025600773 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility a 58 year old male patient who had placed a rp flex to support a post operative/open heart patient. the pump was supporting for 19 hours when pump flows suddenly dropped. after assessment, it was determined that the pump had likely ingested a clot. decision was made to remove and replace the pump. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual: section: potential adverse events (united states) states the following: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
11270 | 1220648-2025-26500 | 21525762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525762&pc=OZD | 02/10/2025 | 03/10/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133364 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon initiating the controller, there was a battery controller alarm. the battery of the automated impella controller (aic) was switched but the alarm did not resolved. the aic was removed from service. there was no patient involvement at the time as this occurred during prep. the procedure was completed with a different controller. | [5 empty] | H.6 code 2199 was reported incorrectly on manufacturer device report 1220648-2025-26500. a new code was added to health effect - impact code. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11271 | 1220648-2025-26501 | 21525773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525773&pc=OZD | 02/07/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/07/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584974 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who hemolysis experienced during impella 5.5 support. actions taken and patient outcome are unknown. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11272 | 1220648-2025-26502 | 21525779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525779&pc=PYX | 02/08/2025 | 03/22/2025 |  | 03/05/2025 |  |  | 02/08/2025 | 03/05/2025 |  |  | 02/28/2026 |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025476539 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump presented placement signal loss, motor current changes, and low flows. there was a concern for biomaterial ingestion. argatroban was noted to be sub-therapeutic there was noted troubleshooting through fluid delivery. after several minutes, flows, motor current, and placement signal all returned to baseline. the patient continue on support until the device was successfully weaned. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the cause of the low pump flow issue was most likely biomaterial due to ingestion trend observed based on log review. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h8 usage of device selection was inadvertently omitted on the initial manufacturer device report number. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11273 | 1220648-2025-26503 | 21525823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525823&pc=PYX | 02/08/2025 | 03/25/2025 |  | 03/05/2025 |  |  | 02/08/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025600713 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that flows dropped to 1.5 and there was a spike in motor current. the pulmonary artery waveform disappeared and motor current went flat. x-ray performed and position was confirmed to be acceptable. team elected to explant the pump. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the device was not returned for analysis. the data log was reviewed and there were many suctions that occurred at the time and date of the reported issue. motor current spike was observed at p-7 followed by a drop in flow to 1.5 liters per minute, at same time placement signal trended up. these changes indicate possible biomaterial ingestion. the root cause of the low pump flow issue most likely was biomaterial ingestion. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26503. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26503. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26503 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 added health effect - impact code as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26503.
11274 | 1220648-2025-26504 | 21525833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525833&pc=OZD | 02/08/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/08/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606356 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | A |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp via the right femoral artery for mechanical circulatory support (mcs). four days into the impella cp device support, the patient developed limb ischemia, and the patient was planned for escalation to an impella 5.5 device in response to the condition. the attempt with the impella 5.5 was ultimately unsuccessful and a new impella cp was implanted via the axillary artery. two days after the start of support on the axillary impella cp device, the patient was now in atrial fibrillation with rapid ventricular response at rate in 170¿s. it was noted the patient had been cardioverted for unstable ventricular tachycardia. echocardiogram was completed and physician had pulled back the impella cp device 2cm because it was measuring greater than 6 cm and was up against the wall. echocardiogram post repositioning noted the impella was measuring about 3.8 and was free floating in the left ventricle, and not up against any structure. the patient continued to go in and out of rapid atrial fibrillation, however. a couple of days following, the patient still required vasopressors and inotropes to maintain hemodynamic stability. lactic normalizing, however, the question remained how to address long term support for the patient as a bridge to transplant or other decision due to the inability to place impella 5.5. discussion made with the family and care was withdrawn after a total of six days on support. the cause of the death was not provided. medical safety review noted there is not enough evidence to exclude the axillary impella cp as a possible contributing factor to the outcome. although it should be noted long-term sustaining treatment options to bridge were limited, and the family withdrew care from the patient. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of three device manufacturing reports related to the patient's implant experience. refer to the following: medical device report for impella cp serial number (b)(6) with date of event 02-feb-2025 and patient identifier ending in (b)(6). medical device report for impella 5.5 serial number (b)(6) with date of event 06-feb-2025 and patient identifier ending in (b)(6). medical device report for impella cp serial number (b)(6) with date of event 08-feb-2025 and patient identifier ending in (b)(6) (this report).
11275 | 1220648-2025-26505 | 21525852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525852&pc=OZD | 02/10/2025 | 03/31/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559431 |  | 048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 32-year-old female patient who experienced multiple complications during impella cp support as a bridge to transplant. after one day of support, the patient lost doppler signal in their right leg. the pump was explanted in response to the ischemia, however a doppler signal was still unobtainable in the extremity. upon further examination, the patient was found to have a right iliac dissection and thrombus in her right femoral artery. the iliac was stented and a thrombectomy was performed to treat the thrombus. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia, vascular damage - peripheral vessel, and thrombus has been completed. the device was returned for investigation. the device was inserted via fight femoral access without issue. one day on support, the patient lost doppler signals to the lower leg on the implant side. lack of doppler signals continued post explant, with the patient having right iliac dissection and thrombus in the right femoral artery. the pump was returned. inspection of the returned pump showed no abnormalities. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. while on support, the patient was had right iliac dissection and thrombus in the right femoral artery. the pump was returned. inspection of the returned pump showed no abnormalities. the cause of the vascular damage - peripheral vessel was not determined due to lack of clinical information. while on support, the patient had right iliac dissection and thrombus in right femoral artery. the pump was returned. inspection of the returned pump showed no abnormalities and no biomaterial around impeller. as the necessary information was not provided, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.9 revised date device returned to the manufacturer was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26505. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26505 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name and fax number as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26505. h.3 selection was entered incorrectly as the device had been returned at the time of the initial manufacturer device report 1220648-2025-26505 was submitted. h.6 code 4114 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26505. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was returned. h.10 some of the instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2025-26505 and were added.
11276 | 1220648-2025-26506 | 21525867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525867&pc=OZD | 02/10/2025 | 03/18/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557312 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was purge fluid leaking out of the red plug. the fluid had the consistency of dextrose. there have been no purge alarms or high flow issues since the leak has been reported. the staff monitored and support continued. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. pump was not returned for investigation. data log review was not required for this case. as per the given clinical, purge was leaking from red handle without observed any purge related alarm. the leak location could not verified without product back. the cause of the purge leak issue was not determined since product was not returned and sufficient clinical information was not provided. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
11277 | 1220648-2025-26507 | 21525886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525886&pc=OZD | 02/10/2025 | 03/31/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 71 year old female patient who had an impella cp placed to support for cardiogenic shock. after 22 hours of support, the patient exhibited signs of hemolysis with urine color changes and increasing lab values. the patient received 1 unit of blood. the pump was then explanted. the urine cleared post-explant. the survived the procedure. | [5 empty] | The investigation for the hemolysis has been completed. no product was returned. reported signs of hemolysis, urine turned darker red and lactate dehydronase (ldh) increase morning of (b)(6) and patient received 1 unit of blood overnight. data logs show no motor current spike observed and stable flows. placement signal low and suction alarms that resolve night of (b)(6), 8 hours before reported hemolysis. unclear if alarms were resolved from receiving volume. the root cause of the hemolysis was not determined since product was not returned and insufficient clinical details. added g1 reporting contact fax number. | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11278 | 1220648-2025-26508 | 21525895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525895&pc=OZD | 02/10/2025 | 10/01/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553897 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the device exhibited optical sensor issues. | The complainant reported that the impella 5.5 pump became unreliable and persistent suction alarms were triggered during patient support. repositioning was attempted with echo guidance, but the suction alarm persisted. the alarms were disabled and support continued with the implanted pump. there was no documented patient harm. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was returned for investigation. the pump was cut open. no further investigation could be performed on the optical signal and low pump flow. tl low alarms during the entire pump run while running at the same p-level with several drops in snr was seen below 2000, and contrast levels dropped below 15. also, there were some drops in motor current, and the flow was noted on the reported third day, where map does not correlate with the a-line, and this trend continues at lower drag during the case without a change in p-level. no motor current spike was reported during the entire case run. the overall placement signal trend was decreasing, with a recovery seen later during the case run. the cause of the low pump flow issue was most likely patient condition related, based on data log review. it is noteworthy that the pump previously used on the same console exhibited a similar trend of low snr while running at a higher p-level, suggesting that console-related abnormalities may be causing the optical signal issues. the cause of the optical signal issue was not determined, as the returned product was not sufficient to derive any root cause. please refer to 25-41429-2 for aic-related issues. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: add optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11279 | 1220648-2025-26509 | 21525898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525898&pc=OZD | 02/10/2025 | 03/25/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610436 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported 65-year-old female patient implanted with a cp pump for support due to cardiogenic shock/cardiomyopathy. the patient had signs of hemolysis but the team made a choice to evaluate patient before repositioning or other troubleshooting measures. the pump remained on with the hemolysis until the family made a decision to withdraw care and patient expired. | [5 empty] | The investigation of hemolysis has been completed. the device and the data were not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26509 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26509. | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11280 | 1220648-2025-26511 | 21525918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525918&pc=OZD | 02/07/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584409 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella cp and extracorporeal membrane oxygenation (ecmo) support ongoing for cardiogenic shock. the pump was supporting for seven hours when the patient expired from bleeding that originated at the ecmo site. the contribution of the heparin in the impella purge and impella cp use and any contribution to the death is not known. the bleeding made stabilization of the patient difficult and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11281 | 1220648-2025-26513 | 21525955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21525955&pc=OZD | 02/09/2025 | 04/24/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593413 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported 63 year old male patient on impella cp for support for acute myocardial infarction. the procedure was completed but post-explant, the patient began to have complaints of pain in the groin area. ultrasound was ordered and a pseudoaneurysm was identified. interventional radiology was consulted, and plans are to inject thrombin into the pseudoaneurysm. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage issue has been completed. pump was not returned for investigation. clinical information provided is very limited. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-26513. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26513.
11282 | 1220648-2025-26514 | 21526023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21526023&pc=OZD | 02/08/2025 | 03/17/2025 |  | 03/05/2025 |  |  | 02/08/2025 | 03/05/2025 |  |  | 11/30/2025 |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424749 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed limb ischemia. the decision was made to consult vascular surgery to assist with a right superficial femoral artery distal perfusion sheath. left to right external femoro-femoral bypass was performed successfully with immediate improvement in pallor of the right lower extremity. support was continued. additional information noted that the decision was made to withdraw care and the patient expired. there is not enough information to exclude the impella cp device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11283 | 1220648-2025-26510 | 21526215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21526215&pc=OZD | 02/10/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585983 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump. the pump was inserted, but ultimately could not be advanced due to significant iliac disease and a tortuous anatomy. the implant was aborted in response to the noted resistance. there was no resulting patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11284 | 1220648-2025-26512 | 21526219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21526219&pc=OZD | 02/10/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/10/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573331 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, a left ventricle thrombus was noted. the thrombus was observed in the patient, but support continued and the patient remained stable. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
11285 | 1220648-2025-26516 | 21527247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21527247&pc=OZD | 02/11/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588761 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp pump to support a patient. it was reported that the sterile cover on the repositioning unit torn off. there was no harm to the patient and support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11286 | 1220648-2025-26517 | 21527467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21527467&pc=OZD | 02/11/2025 | 04/24/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580289 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and developed hemolysis due to the impella cp being too shallow. the impella cp was repositioned but to manage the hemolysis. after 12 hours of impella cp support, the patient was escalated to impella 5.5. the urine was clearing up post impella exchange. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and positioning issues has been completed. the data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved. no motor current (mc) spikes outside p-level changes. the product was not returned for review. clinical reports dark red urine and pump not in good position. repositioning the pump resolved the hemolysis. the root cause of hemolysis was most likely due to improper positioning as the clinical team stated hemolysis was resolved post repositioning of the pump. note: remove fm positioning issue as it is captured under the fm of hemolysis as the root cause.
11287 | 1220648-2025-26518 | 21527546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21527546&pc=OZD | 02/08/2025 | 04/08/2025 |  | 03/05/2025 |  |  | 02/08/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1724721 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The automated impella controller (aic) was priming the impella catheter and presented with an ¿optical sensor system error¿. reattempted but the same error code appeared. the aic was replaced with another console. no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic) - loss of opm data has been completed. the console logs from the reported day of event are consistent with the complaint and showed that during case start, the sensor calibration could not be performed due to absence of valid optical signal. the real time log data indicated very low signal-to-noise ratio and contrast and maxed-out gain and lamp, consistent with either a large air gap at the pump connector or loss of light due to optical bench/lamp/fiber issue. the reported issue was not reproduced in lab tests. analog output of the ccd detector in optical bench did not reveal any defects. the cause of the aic issue was a defective optical bench. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11288 | 1220648-2025-26519 | 21527688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21527688&pc=OZD | 02/08/2025 | 03/23/2025 |  | 03/05/2025 |  |  | 02/08/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025606980 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a small hematoma at the access site. blood products were given and a sandbag was applied. the patient was subsequently escalated from impella cp to impella 5.5. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added d5-operator of device should have been health care professional. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11289 | 1220648-2025-26520 | 21527767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21527767&pc=OZD | 09/15/2023 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  | 03/31/2025 |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346043 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular fibrillation with a "possible" relationship to the impella device used. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
11290 | 1220648-2025-26521 | 21527915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21527915&pc=OZD | 02/06/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518276 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support (mcs) and experienced access site bleeding. reportedly, the bleeding occurred the same day of the implant and was noticed at the insertion site after positioning the impella cp device. blood transfusion was given, and the bleeding was conservatively managed with compression and activated clotting time management. no surgery was required. the outcome of the patient as a result of the event was unknown. however, there was no report or indication of further harm to the patient. the patient continued on impella mcs for an additional five days until successfully weaned and explanted noting a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11291 | 1220648-2025-26522 | 21528137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21528137&pc=OZD | 02/06/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510307 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). following the implant, a dopplerable pulse was unobtainable in the right lower extremity. an angiogram through the re-access port showed that a slow flow was present past the impella sheath. mcs was continued for an additional two days, after which the impella cp device was explanted. upon explant, femstop was applied to the site to achieve hemostasis. after 30 minutes it was taken off, and the patient expressed discomfort. the patient's leg was cool and no pulse since shortly after manual pressure was started. the patient was noted to be in stable condition following the event and was transferred to an outside facility to treat the condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11292 | 1220648-2025-26523 | 21528402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21528402&pc=OZD | 02/09/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/09/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584527 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding at the access site. the site was redressed and sutured. the patient continued on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11293 | 1220648-2025-26524 | 21528435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21528435&pc=OZD | 02/11/2025 | 03/18/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583111 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support experienced a stroke. the patient was unresponsive, and pupils were unequal. the patient was on and off anticoagulation. it is unknown whether stroke was impella 5.5 related. care was withdrawn and the patient expired. the relationship between the impella 5.5 and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11294 | 1220648-2025-26525 | 21528495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21528495&pc=OZD | 02/06/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551817 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and unexpectedly stopped ten days into support. the pump was restarted but stopped again shortly after. the decision was made to explant the pump as a result of the failure. it was clarified the pump failure caused an explant, although the explant was likely planned for this same day. there was no harm to the patient as a result of the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen." impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11295 | 1220648-2025-26526 | 21528604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21528604&pc=OZD | 02/06/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542542 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the insertion of an impella cp device to a patient with acute myocardial infarction/cardiogenic shock was unsuccessful due to the anatomy. the device was unable pass due to calcification and additionally, the catheter kinked. consequently, the case was aborted. there was no report of adverse effects to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11296 | 1220648-2025-26527 | 21528716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21528716&pc=OZD | 02/11/2025 | 03/21/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528162 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a complication during impella 5.5 support. twelve days into support, the patient suffered from runs of supraventricular tachycardia. the pump was pulled back .5 cm to counteract the condition. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the supraventricular tachycardia (svt) has been completed. the purge cassette was returned for evaluation. the pump was not returned for evaluation. as the necessary information was not provided, the root cause of the svt was not determined. h.6 revised health effect - clinical code as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26527.
11297 | 1220648-2025-26528 | 21528873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21528873&pc=PYX | 02/06/2025 | 03/21/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025567876 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient postoperative cardiothoracic surgery was implanted with an impella rp flex for mechanical circulatory support and experienced low pump flow, thrombus and hemolysis requiring device removal. three days earlier the patient had an impella rp flex device placed and removed the following day after implant. it was decided that the patient required support again, and the patient was placed on another impella rp flex device. however, after approximately two days on support, the patient's motor currents and pulmonary artery pressures abruptly doubled. an x-ray confirmed correct position, and the team discussed likely hood of clot ingestion with presence of motor current and drop in flows. also noted was hemolyzing labs and drop in hemoglobin. removal of device was recommended as best option for patient, which was successfully completed. the patient was noted to be stable following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two devices associated with the patient's implant experience. refer to manufacturing report number 1220648-2025-26403 for the first rp flex device experience. | The investigation of hemolysis, low pump flow, and thrombus have been completed. the data logs confirmed low pump flow for the whole case with motor current (mc) spike and auto suction response. the product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. the root cause of hemolysis was most likely biomaterial ingestion. note: removed failure mode (fm) thrombus as it was a root cause for both hemolysis and low flow.
11298 | 1220648-2025-26529 | 21528943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21528943&pc=OZD | 02/06/2025 | 03/13/2025 |  | 03/05/2025 |  |  | 02/06/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132981 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Select Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was planned for an impella 5.5 device implant. however, during preparation, automated impella controller (aic) failed to recognize an installed purge cassette. the aic and purge cassette were replaced, and priming was successfully initiated. the patient was successfully implanted with the impella 5.5 device for impella mechanical circulatory support. there was no harm to the patient as a result of this event. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. a.5 revised ethnicity as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26529. b.7 added information as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26529.
11299 | 1220648-2025-26530 | 21529218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21529218&pc=OZD | 02/11/2025 | 03/05/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599710 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a plan to escalate from a impella cp to the impella 5.5 pump. upon attempting to insert the 5.5 via the right axillary artery the impella catheter kinked and was unable to advance into position. the case was aborted. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11300 | 1220648-2025-26531 | 21530701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21530701&pc=OZD | 02/11/2025 | 03/14/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551798 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that 62 year old male patient had an "impella in aorta" alarm appeared during impella cp support. the pump position was verified to be shallow in conjunction with the patient having high lactate dehydrogenase levels and rose urine. in response to the noted issues, the doctor opted to explant the pump. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation of positioning issues and hemolysis has been completed. no product was returned for analysis. the root cause of hemolysis was most likely due to improper positioning as evidenced by the clinical report of pump being too shallow confirmed by echocardigram. the root cause of the positioning issue was most likely due to patient condition as evidenced by clinical detail of small aortic arch. e.1 revised us state as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26531. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26531 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting facility name and reporting facility fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26531. g.2 revised report source as selections were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26531.
11301 | 1220648-2025-26532 | 21530789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21530789&pc=OZD | 02/11/2025 | 04/01/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604087 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiac Tamponade | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was placed for primary percutaneous coronary intervention. after the procedure was completed, an transthoracic echocardiogram (tte) was obtained showing no aortic (valve) insufficiency (ai) with impella, so decision made to leave cp in place. during sheath exchange impella went deep into ventricle and quickly withdrawn. during this time patient was losing pulsatility and was very hypotensive, patient started on vasoactive, labs drawn to rule out bleeding and tte showed cardiac tamponade. the team quickly performed pericardiocentesis. a massive blood transfusion protocol was followed. tte showed left ventricle (lv) hematomas and contrast administration with washout in the pericardium. unsure of lv perforation etiology. the patient was put on palliative care and after the device was withdrawn that patient expired shortly after. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation and positioning issues has been completed. the pump was not returned for investigation, and a data log review was not required for this case. according to clinical findings, the left ventricle (lv) is friable in nature, and during the perforation, the impella device was inserted deeply. the positioning issue failure mode was removed since it was addressed under the root cause of ventricular perforation. the cause of the ventricle perforation issue was most likely improper positioning of the device. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26532 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11302 | 1220648-2025-26533 | 21530848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21530848&pc=OZD | 02/11/2025 | 03/21/2025 |  | 03/05/2025 |  |  | 02/11/2025 | 03/05/2025 |  |  |  |  | 03/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576642 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was having lower flows than expected. after troubleshooting, assessing volume status, afterload, and position the issue was unresolved. the 5.5 was exchanged and there was no patient harm. | [5 empty] | The investigation of low pump flow has been completed. clinical states that low flow occurred for more than 48 hours. assessing volume status, afterload, and positioning did not resolve the issue so the pump was explanted. the data logs confirmed the failure mode and clinical narrative. logs showed after pump started and run for 193 hours, motor current (mc) suddenly spiked to 1250 ma followed by low flow. there is a slight increase in purge pressure, still in range, which resolves. auto suction response and suction alarms were also triggered. this event remained to the rest of the case. visual inspection found no biomaterial around the impeller. there was a small amount in the purge gap. no other issues were found on the rest of the pump. low pump flow was not reproduced in the lab. the root cause of the low flow was most likely biomaterial ingestion based on the ingestion pattern in the data logs and low flows not being reproduced in the lab. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26533 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11303 | 21513512 | 21513512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513512&pc=OZD | 10/10/2024 | 02/03/2025 | 02/03/2025 | 03/04/2025 | 02/03/2025 | 03/04/2025 |  | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided |  |  | Not Available | [6 empty] | User Facility report | No | Risk Manager | [6 empty] | NA |  | Unknown |  |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | At change of shift, i was made aware by primary nurse that the impella stopped running. i attempted to restart the impella with perfusion, and the impella rep at the bedside. after troubleshooting, the impella was not able to be restarted successfully. patient received mtp (massive transfusion protocol) overnight with multiple blood products and was hemodynamically stable at this time. decision was made to take the patient to the or (operating room) for impella replacement. impella was replaced, no clots found on previous device or in graft. replacement working as expected.
11304 | 1220648-2025-26457 | 21513582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513582&pc=OZD | 01/23/2025 | 03/04/2025 |  | 03/04/2025 |  |  | 02/04/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587463 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Disseminated Intravascular Coagulation (DIC) | Sepsis | [11 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was initially implanted with an impella cp device as preoperative placement and then escalated to an impella 5.5 device after coronary artery bypass graft surgery. the patient subsequently expired from sepsis and disseminated intravascular coagulation (dic) due to infection approximately 27 days after start of support. postoperative cardiac function did not recover well, and therefore the patient was escalated to an impella 5.5. report stated before the emergency bypass, the patient had developed "80% of infections," which was noted to be likely due to the patient's condition rather than the impella insertion site. however, no alternative cause for the event was identified as a likely cause of such event. additional information was subsequently received and noted the patient's demise was related to the impella 5.5 device as the patient expired from sepsis and dic (noted as the cause of death). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11305 | 1220648-2025-26458 | 21513610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513610&pc=OZD | 02/05/2025 | 04/02/2025 |  | 03/04/2025 |  |  | 02/05/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612488 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device and pump did not start at onset. the impella 5.5 device was explanted and successfully replaced to provide mechanical circulatory support to the patient. it was noted there was no harm to the patient as a result of this event. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. the pump was returned for investigation. the data logs were not returned for investigation. catheter kink was noted near the motor housing. 2 out of 3 motor line was found open on electrical test. additionally, a cannula kink was reported near the outlet cage. pump stop was reproduce with motor current high alarm. pump stop failure mode was replaced with pump did not start failure ode as investigated. the cause of the pump did not start issue was an open electrical line inside catheter due to excessive force. g1 reporting contact last name was erroneously entered on the initial manufacturer device report number 1220648-2025-26458 g1 manufacturing site name and address was erroneously entered on the initial manufacturer device report number.
11306 | 1220648-2025-26459 | 21513627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513627&pc=OZD | 02/09/2025 | 03/25/2025 |  | 03/04/2025 |  |  | 02/10/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591223 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp pump to support a patient. the patient was seen to have signs of hemolysis with elevated plasma free hemoglobin labs twice. an echocardiogram was performed and showed the impella cp was positioned in the papillary muscle. no attempts were made to reposition. the team also saw limb ischemia and consulted with the vascular team in which they said that the right leg was ¿dead¿ after assessing doppler studies and physical assessment. the patient did get blood products prior to the explant. the patient was placed on extracorporeal membrane oxygenation support and the impella cp was explanted. the day after explant, the patient expired. the relationship between the impella cp and cause of death is unknown. | [5 empty] | The investigation of hemolysis and limb ischemia has been completed. the product was returned and evaluated. there was biomaterial on the impeller, but no other abnormalities were found with the returned product. the data logs showed no motor current spikes and therefore no signs of biomaterial ingestion. clinical details stated that the pump was positioned in the papillary muscle upon arrival to the second hospital, but the pump was not repositioned. the root cause of the hemolysis was most likely improper positioning since the clinical details confirm the positioning issue and no signs of biomaterial ingestion were found in the data logs. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11307 | 1220648-2025-26460 | 21513704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513704&pc=OZD | 01/29/2025 | 03/27/2025 |  | 03/04/2025 |  |  | 02/07/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572134 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received regarding a protect 4 study patient who experienced hemolysis during impella cp support. actions taken are unknown. the procedure outcome was noted to be successful. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source; study was missing from initial manufacturer device report 1220648-2025-26460.
11308 | 1220648-2025-26461 | 21513757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513757&pc=OZD | 02/10/2025 | 03/23/2025 |  | 03/04/2025 |  |  | 02/10/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp placement, the physician attempted the right femoral artery, but lost wire access. micro puncture access was gained to left femoral artery (lfa) and access lost occurred again. the physician reattempted micro puncture access to lfa and was successful. the impella cp was inserted and started. the percutaneous coronary intervention was performed and once completed a significant hematoma was noted to the right groin. there was also a hematoma on the previous stick site to lfa. blood products were administered. the patient expired after four hours of support. the relationship between the impella cp and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ , renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, the doctor lost wire access multiple times this led to groin site hematoma¿s requiring blood administration.
11309 | 1220648-2025-26462 | 21513800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513800&pc=OZD | 02/10/2025 | 04/02/2025 |  | 03/04/2025 |  |  | 02/10/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599698 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, there was saturated flow. waveform started looking similar to pump in the aorta where the aorta and left ventricle waveform were completely overlapping. pump flow, systolic and diastolic flow were narrower, not exactly the same. the pump was explanted and replaced. the patient remained stable. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of saturated flow has been completed. the impella 5.5 was returned for evaluation. upon visual inspection, some biomaterial was present on returned pump within the purge gap. pump was purged and run for approximately 15 minutes. pump was run at p-9 and the saturated flow did reproduce. flows were fully saturated for approximately 15 seconds from pump start then motor current and pump flow were variable before becoming more stable in second half of pump run. after pump was run, small biomaterials were present near outflow cage struts and on impeller blade. data logs were reviewed lt logs from field showed no persistent suction alarms occurring throughout support. pump flow starts to become saturated on day 3 of support when pump was turned up to p-9. motor current at that time becomes less stable and slightly variable. pump was turned down to p-8 on day 5 of support but flows remain saturated above appropriate p-level range. the root cause of the saturated flow was most likely due to biomaterial.
11310 | 1220648-2025-26463 | 21513805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513805&pc=OZD | 02/08/2025 | 03/19/2025 |  | 03/04/2025 |  |  | 02/08/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578647 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient became agitated and pulled the pump out of the appropriate position. the pump was explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue most likely is patient movement since the patient became extremely agitated in the intensive care unit and kicked his legs. this motion dislodged the impella device and pulled the pump out of the appropriate left ventricle position.
11311 | 1220648-2025-26464 | 21513926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21513926&pc=OZD | 02/07/2025 | 03/04/2025 |  | 03/04/2025 |  |  | 02/07/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601499 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. a clot was noted around the insertion site. the decision was made to start dobutamine and treat the clot medically. the patient was weaned down and the device was explanted. the patient outcome was noted to be stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
11312 | 1220648-2025-26465 | 21514040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21514040&pc=OZD | 02/08/2025 | 03/25/2025 |  | 03/04/2025 |  |  | 02/08/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578976 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pumped presented persistent suction issues. volume was given and positioning was adjusted, but the suction alarms continued. after multiple failed attempts to troubleshoot the low flow, the decision was made to explant the pump. there was no patient harm. | [5 empty] | The investigation of low pump flow has been completed. the returned complaint device visually inspected. biomaterial around impeller and purge gap observed no broken blade or cannula kink observed. the cause of the low pump flow issue was biomaterial ingestion. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11313 | 1220648-2025-26466 | 21514109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21514109&pc=OZD | 02/05/2025 | 03/20/2025 |  | 03/04/2025 |  |  | 02/05/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596513 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support encountered a positioning issue that required device replacement. the patient was transported tothe catheterization lab to undergo emergent right and left heart catheterizations for investigation of ventricular fibrillation arrests as well as diffuse st depression noted on electrocardiogram. the left heart catheterization revealed no stenosis except for 40-50% left anterior descending (lad) stenosis. left ventricular end-diastolic pressure measured at 33. left ventriculography gram revealed ejection fraction of 20%. the decision was made to initiate impella cp support at this time. the impella cp was inserted in the right femoral artery in routine fashion and flows were initiated. after insertion placement was re-evaluated, the impella catheter was retracted and inadvertently pulled through the aortic valve. re-insertion was attempted with bowing of the cannula noted under fluoroscopy. the decision was made to remove the impella cp and reinsert a new impella cp device which was successfully completed without incident. flows were initiated. flows of 3.7 liters were achieved. the patient was transported to the unit to rest on impella cp support in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. reported inadvertently pulling the pump back across the aortic valve after pump was initially placed. attempted wireless re-access without success and noted a cannula kink under fluoro. the root cause of the positioning issue was most likely use related since the pump was inadvertently pulled back across the aortic valve during repositioning.
11314 | 1220648-2025-26467 | 21514296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21514296&pc=OZD | 02/04/2025 | 04/07/2025 |  | 03/04/2025 |  |  | 02/05/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567676 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support (mcs). prior to impella mcs, the patient was placed on venoarterial extracorporeal membrane oxygenation. four days after start of the impella mcs, the pump alarmed, unexpectedly stopped and was unable to restart. thrombus was observed in the cannula. activated clotting time was noted to be 183 seconds. consequently, the impella cp device was explanted and was successfully replaced without incident. the patient was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of thrombus and pump stop has been completed. the cut pump was sent back for investigation. the data log was not provided for this case run. no physical abnormalities were noted on the returned pump other than the cut catheter. biomaterial was observed in the purge gap. no blood was observed in purge line near me motor. motor housing was teardown and no biomaterial was observed inside it. as per the clinical details, thrombus was noted inside cannula after pump stop event. the cause of the pump stop issue was not determined since cut pump was sent back and data logs were not provided for field run. as the necessary information was not provided, the root cause of the thrombus was not determined. d9 date device returned to manufacturer added.
11315 | 1220648-2025-26468 | 21514507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21514507&pc=PYX | 02/04/2025 | 03/04/2025 |  | 03/04/2025 |  |  | 02/04/2025 | 03/04/2025 |  |  | 06/30/2025 |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024371162 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the insertion of an impella rp device in a patient with pulmonary embolism was unsuccessful. following the successful insertion of the 23 french introducer, the physician was unable to advance the impella rp pump into the pulmonary artery. the impella rp was unable to cross the pulmonic valve, and after multiple failed attempts, the case was aborted. the impella rp was not saved to be returned for investigation. the patient was noted to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
11316 | 1220648-2025-26469 | 21514655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21514655&pc=OZD | 02/07/2025 | 03/21/2025 |  | 03/04/2025 |  |  | 02/07/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025559479 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. the pump was unable to be delivered. the team removed the impella 5.5 and instead placed an impella cp pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the impella 5.5 was unable to be placed due to patient anatomy. b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26469. e.1 added us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26469. h.4 revised device manufacture date as it was reported incorrectly on the initial manufacturer device report number 1220648-2025-26469.
11317 | 1220648-2025-26470 | 21514805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21514805&pc=OZD | 02/05/2025 | 04/07/2025 |  | 03/04/2025 |  |  | 02/06/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114577 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) displayed a controller error during patient support. the alarm resolved with no noted harm to the patient. support was continued uninterrupted. the aic has been shipped for investigation but has yet to be received. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. based on the data analysis the root cause for the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. h5 erroneously selected on the initial manufacturer device report number 1220648-2025-26470. h8 erroneously selected on the initial manufacturer device report number 1220648-2025-26470.
11318 | 1220648-2025-26471 | 21514970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21514970&pc=OZD | 02/04/2025 | 03/22/2025 |  | 03/04/2025 |  |  | 02/04/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576648 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | Appropriate Term / Code Not Available | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella 5.5 device in a patient with severe multivessel disease undergoing a high-risk percutaneous coronary intervention was unsuccessful due to the anatomy. after backloading, the impella 5.5 device onto a .018 guidewire for implant, the device was unable to advance into the subclavian to make the turn into the ascending aorta. after multiple failed attempts, a significant kink was discovered on the guidewire. the devices were removed, and a new guidewire was inserted. however, once again, resistance was encountered in the same area. the physician then gained single access into the oscor sheath next to the impella and attempted to advance a glide advantage wire. however, the wire would not feed. at this time, the decision was made to abort the impella 5.5 implant and place an axillary impella cp device. with the wire remaining in the left ventricle space, the impella cp was primed and prepped. the physician loaded the impella cp onto the platinum plus wire and advanced the impella cp. some difficulty was met when the motor housing attempted to cross the same area of subclavian. however, the impella cp device was able to successfully cross and support the patient. after proper placement was identified and confirmed with transesophageal echocardiogram, platinum plus wire was removed and support was initiated on auto 3.7 liters per minute. there was no resulting harm, and the patient was transported to the unit in stable condition. the pump was returned for investigation but the guidewire is not available. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the product damage and delivery issues has been completed. during delivery, pump would not make the turn into the ascending aorta despite multiple attempts with 3 different guidewires, with resistance being met in the same location each time. the pump was returned. data logs are not applicable to this case. visual inspection revealed kink in the cannula at the bend. the cause of the delivery issue was patient condition related due to the patients noted severe carotid stenosis. during delivery, guidewire was found to be kinked. it was removed and 2 other wires were used to try to gain access with the 5.5 but to no success. one of the two additional guidewires used was a .018 wire from abiomed. resistance was met in the same area each time. no guidewires were returned. data logs are not applicable to this case. the cause of the delivery issue was patient condition related due to the patients noted severe carotid stenosis.
11319 | 1220648-2025-26472 | 21515123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21515123&pc=OZD | 02/10/2025 | 04/25/2025 |  | 03/04/2025 |  |  | 02/10/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611392 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had care being escalated from a impella cp to the 5.5 pump. the 5.5 pump placement took time and bleeding occurred at the site. the bleeding prompted the team to reduce/reverse the anticoagulation. additionally, during the delayed and long pump placement the team observed arrhythmias. the patient remained stable and on support. | Additional information received that the patient remained on the 5.5 pump for 69 hours until the decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding and arrythmia has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related as lengthy cut down/anastamosis due to bleeding was observed. as the necessary information was not provided, the root cause of the arrythmia was not determined. | B2 outcome revised to reflect the additional information received from the clinical site to include death and date of death. b5 revised to reflect the additional information received from the clinical site. d6b explant date added. h1 type of report revised to reflect the additional information received from the clinical site. h6 revised to reflect the additional information received from the clinical site to include death.
11320 | 1220648-2025-26473 | 21515397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21515397&pc=OZD | 02/06/2025 | 03/24/2025 |  | 03/04/2025 |  |  | 02/06/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576573 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support placed to support a patient admitted in with acute myocardial infarction and cardiogenic shock. the pump was supporting when there were positional alarms. the impella was shallow, impella in aorta alarm with inlet at the atrioventricular annulus. the device was repositioned under echocardiogram guidance. the team after 1 day of support made plan to escalate to the 5.5 pump but when taken to the operating room there was an observation made of left ventricle thrombosis. the pump remained on support for several days until successful escalation of the 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus and positioning issues has been completed. as the necessary information was not provided, the root cause of the thrombus was not determined. the cause of the positioning issues most likely was patient movement since it occurred upon moving the patient over to new bed and changing sheets, suction alarmed with decoupling of waveforms. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26473. h.6 added health effect - impact code as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26473.
11321 | 1220648-2025-26474 | 21516449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21516449&pc=OZD | 02/07/2025 | 03/24/2025 |  | 03/04/2025 |  |  | 02/07/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was also placed on extracorporeal membrane oxygenation (ecmo) and a temporary pacer. the patient experienced a gastrointestinal bleed(gi). the patient had been supported with heparin in the purge. the patient continued on support until the device was weaned. the patient outcome was noted to be stable. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. clinical information provided is very little and not sufficient to investigate the complaint. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11322 | 1220648-2025-26475 | 21517619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21517619&pc=OZD | 02/06/2025 | 09/26/2025 |  | 03/04/2025 |  |  | 02/06/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574096 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs). the next day after the start of the support, placement signal low alarm was triggered. it disappears when the performance level was reduced to p-7 from p-9. the physician noted, however, the device was correctly positioned. furthermore, on this same day, it was noted the patient was showing signs of hemolysis. treatment details or action taken as a result of the hemolysis event were not provided. the following day, it was reported that the patient was bleeding from the puncture site. entry angle and forward sutures were reviewed. a new dressing was applied and two units of red blood cells were given. the status of the patient as a result of the event is unknown. however, the patient remained on support for an additional four days before being successfully weaned and explant with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause of the optical signal issue was not determined since product was not returned and insufficient clinical details are available. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates
11323 | 1220648-2025-26476 | 21517636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21517636&pc=OZD | 02/05/2025 | 04/29/2025 |  | 03/04/2025 |  |  | 02/06/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582501 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 14 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support while awaiting "urgent" transplant. approximately 20 days after start of the support, the pump lost flow, left ventricle and placement signal decoupled, and repositioning attempt failed. it was further explained while intubating the patient, the patient arrested, and flows could not be improved for reasons unknown. consequently, impella 5.5 device was explanted, and the patient was moved to extracorporeal membrane oxygenation support. the status of the patient as a result of the event was noted to be unknown. however, the outcome at explant noted the patient survived the device removal. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow and positioning issues has been completed. the product and data were not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the optical signal issue could not be determined due to insufficient product or data logs returned.
11324 | 1220648-2025-26477 | 21517723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21517723&pc=OZD | 02/07/2025 | 03/04/2025 |  | 03/04/2025 |  |  | 02/07/2025 | 03/04/2025 |  |  | 11/30/2025 |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434139 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient had an attempted impella 5.5 implant and during the implant, the 0.018 wire kinked at initial insertion. when the doctor attempted to advance the 5.5 they lost pushability of the catheter. it appeared that the 5.5 also kinked behind the motor housing, which limited the pushability of the device. the doctor removed the pump from the graft and the wire came out of the left ventricle. the doctor was unable to re-access the ventricle despite using multiple wires. the implant was aborted and there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11325 | 1220648-2025-26478 | 21518127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21518127&pc=OZD | 02/08/2025 | 03/26/2025 |  | 03/04/2025 |  |  | 02/08/2025 | 03/04/2025 |  |  |  |  | 03/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607035 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding at the access site. manual pressure was held, medication was titrated and anticoagulation was withheld. femstop was applied and support continued. additionally, after the pump was implanted, the doctor had difficulty with the sterile sleeve so they removed the sleeve. support continued until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and product damage (sterile sleeve damage) has been completed. pump was not returned for investigation. as per the clinical information provided, the patient had arterial bleeding on access site with activated clotting time (act) greater than 420. therefore, the root cause of the access site bleeding issue was most likely anticoagulation management related due to high patient act values. as per the clinical information provided, the doctor removed the sterile sleeve since he was finding it difficult to secure the catheter after insertion. therefore, the root cause of the product damage - sterile sleeve damage issue was use related to improper technique. d3 manufacture address line 1 was erroneously entered on the initial manufacturer device report number 1220648-2025-26478.
11326 | 1220648-2025-26443 | 21503297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21503297&pc=OZD | 02/02/2025 | 03/20/2025 |  | 03/03/2025 |  |  | 02/02/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573976 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp via the right femoral artery for mechanical circulatory support (mcs). after the implant, the patient developed some oozing from their access site. forward sutures were placed, and dressings were applied to initially manage the condition. the following day, manual pressure was held on the resulting hematoma and femostop was applied on the site to achieve hemostasis. impella cp mcs continued for another four days and at this time distal pulse was loss in the impella limb and visible mottling of the right lower extremity was present. the patient was planned for escalation to an impella 5.5 device in response to the condition. the attempt with the impella 5.5 was ultimately unsuccessful and a new impella cp was implanted via an upper artery access. the section of graft was left external to accommodate for repositioning sheath. due to the patient's size and pre-existing issues with distal perfusion, the team did not want to risk insertion of the repositioning sheath and loss of flow to the right upper extremity (rue). there were positive pulses in the rue, color and temperature within the defined limits. explant site of femoral impella cp was intact. there was pre-existing bruising from the hematoma related to the initial implant. the team continued to monitor the right lower extremity for return of flow now that the impella cp pump was removed. long-term plan for the patient was uncertain as the team noted the patient would eventually require a device to provide long-term support as a bridge to transplant or other decision. the patient continued on impella mcs and was noted to be stable with no report or indication of unresolved limb ischemia. | [5 empty] | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was most likely use issue since the blue repositioning hub was implanted too far into tissue tract. the root cause of the limb ischemia cannot be determined since the product was not returned and insufficient clinical details were provided. h10 the investigation results inadvertently omitted on the initial manufacturer device report number 1220648-2025-26443. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of three device manufacturing reports related to the patient's implant experience. the other two reports are: medical device report for impella 5.5 serial number (b)(6) with date of event 06-feb-2025 and patient identifier ending in (b)(6). medical device report for impella cp serial number (b)(6) with date of event 08-feb-2025 and patient identifier ending in (b)(6). as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11327 | 1220648-2025-26444 | 21503335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21503335&pc=OZD | 02/09/2025 | 03/20/2025 |  | 03/03/2025 |  |  | 02/09/2025 | 03/03/2025 |  |  | 12/31/2025 |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446688 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was removed due to right groin bleeding in a patient. the patient remained stable and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. h.10 the investigation results was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26444. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11328 | 1220648-2025-26445 | 21504134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21504134&pc=OZD | 02/06/2025 | 03/20/2025 |  | 03/03/2025 |  |  | 02/06/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576559 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp via the right femoral artery for mechanical circulatory support (mcs). four days into the impella cp device support, the patient developed limb ischemia, and the patient was planned for escalation to an impella 5.5 device in response to the condition. the attempt with the impella 5.5 was ultimately unsuccessful and a new impella cp was implanted via an upper artery access. the impella 5.5 was inserted with the plan of utilizing double barrel technique to minimize time off support. the impella 5.5 was inserted to the aortic valve; however, despite repeated attempts, cardiologist with the assistance was unable to cross. the pump was pulled back, torqued, and re-advanced several times; it was pulled back into graft, rotated and then re-advanced, wire repositioned, all without success. the decision was then made to wean femoral impella cp to performance level p-1 and pull into descending aorta to provide more space to cross the patient's small aortic valve. the patient required high dose pressors and inotropes to support hemodynamics during this time. even with the impella cp no longer across the valve, and the impella 5.5 was in good position, it was still unable to cross the valve. the patient was deemed too unstable for a balloon aortic valvuloplasty. therefore, the physician decided to remove the impella 5.5 and was replaced with an axillary impella cp. the patient continued on mcs with the axillary impella cp for an additional six day before care was withdrawn for reasons unrelated to the impella 5.5 as per the review by medical safety. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of three device manufacturing reports related to the patient's implant experience. refer to the following: medical device report for impella cp serial number (b)(6) with date of event 02-feb-2025 and patient identifier ending in (b)(6). medical device report for impella 5.5 serial number (b)(6) with date of event: 06-feb-2025 and patient identifier ending in (b)(6). (this report). medical device report for impella cp serial number (b)(6) with date of event 08-feb-2025 and patient identifier ending in (b)(6). as noted in the impella ifu: impella 5.5 with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy and was aborted. d.6b revised explant date as it was reported incorrectly on the initial manufacturer device report number. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number. h.10 the investigation results were inadvertently omitted on the initial manufacturer device report number.
11329 | 1220648-2025-26446 | 21504194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21504194&pc=OZD | 02/03/2025 | 03/22/2025 |  | 03/03/2025 |  |  | 02/03/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599140 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced low pump flow. the decision was made to place an impella cp and send to the intensive care unit with the patient being on vasopressors. upon placing the repositioning sheath, it was noted that the physician grabbed the sterile sleeve and inserted, thus causing the cannula to kink followed by suction alarms and "impella in ventricle" alarm. the physician immediately pulled the impella cp back and placed into the proper position. however, suction alarms were unable to be resolved and they were unable to get flows above 1.5 liters per minute. the decision was made to remove the impella cp and replace with a new one, which was successfully completed. the patient was sent to the unit on impella mcs and was noted to be in stable condition. | [5 empty] | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. clinical details stated that physician grabbing the sterile sleeve and inserting when exchanging the repo sheath with the peel away which caused the cannula to kink and suction and "impella in ventricle" alarm. the pump was pulled back into place but unable to resolve suction and get flows above 1.5 l/min. product was visually inspected and a cannula kink was found. data logs show suction and "impella in ventricle" alarms followed by low flows. the root cause of the issue was most likely a kinked cannula since a cannula kink was found on the returned product.
11330 | 1220648-2025-26448 | 21504288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21504288&pc=OZD | 02/10/2025 | 03/28/2025 |  | 03/03/2025 |  |  | 02/10/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597734 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 had continuous suction alarms at all p-levels despite excellent positioning. the impella ran good and suction alarms were turned off. the medical doctor thinks there is an issues with the sensor that detects suction alarms. there was no patient harm. | [5 empty] | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the data logs confirmed low pump flow when pump started and remained to the rest of the case. the parameters were in normal range. the visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. the root cause of low flow was unable to be determined as no problem was found on the pump and failure mode could not be reproduced.
11331 | 1220648-2025-26447 | 21504289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21504289&pc=OZD | 02/09/2025 | 03/21/2025 |  | 03/03/2025 |  |  | 02/09/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579840 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 97 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. the impella implant was initiated after the patient began to rapidly decompensate following a pci attempt. following successful insertion, the pump was unable to cross the aortic valve. the implant was aborted and the patient passed away. it is unknown if the delivery issue contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related as impella cp inserted but unable to cross aortic valve. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26447.
11332 | 1220648-2025-26449 | 21504509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21504509&pc=OZD | 02/09/2025 | 03/31/2025 |  | 03/03/2025 |  |  | 02/09/2025 | 03/03/2025 |  |  | 08/31/2025 |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388553 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp to support a patient admitted in with st-segment elevation myocardial infarction and in need of coronary artery bypass graft. the cp was supporting when limb ischemia prompted the team to place a fem-fem bypass. the pump remains on for support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 life threatening was erroneously selected on the initial manufacturer device report number 1220648-2025-26449. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26449 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11333 | 1220648-2025-26450 | 21504604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21504604&pc=OZD | 02/10/2025 | 04/14/2025 |  | 03/03/2025 |  |  | 02/10/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1287215 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an automated impella controller (aic) being used when there were console alarms for "impella defective" which prompted pump replacement and when this failed, the aic was shutdown. after gaining support, the aic was restarted again and functioned. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) ¿ pump stop has been completed. the pump was not detected for the first attempt and case start wizard was cancelled. the case start wizard was reinitiated again and the pump was connected at 19:25:24 and detected without any issue the case proceeded smoothly and operated without complications for the following two days. a thorough visual inspection of the internal components was conducted and no other issues were found. the console was tested with a known good pump and purge cassette, and it functioned as expected. root cause: the root cause for the pump not detected "impella defective (error reading eeprom) - alarm issued" was likely a malfunction in the impellatronic board. based on the data and device analysis the root cause for the pump not detected "impella defective (error reading eeprom) - alarm issued" was likely a malfunction in the impellatronic board. d9 date device returned to manufacturer was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26450. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26450 as it is unknown if the initial reporter also sent the report to the food and drug administration. g4 pma/510(k) number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26450. h3 device not returned was erroneously selected on the initial manufacturer device report number 1220648-2025-26450. h6 type of investigation code 4114 was erroneously selected on the initial manufacturer device report number 1220648-2025-26450.
11334 | 1220648-2025-26451 | 21504626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21504626&pc=OZD | 02/10/2025 | 06/25/2025 |  | 03/03/2025 |  |  | 02/10/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594088 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was a leak at the white impella plug. purge flow and pressure were unaffected, tegaderm was wrapped around where the blue reservoir is. support continued and there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. due to lack of sufficient clinical details and no product returned, the root cause of the purge leak - pump could not be determined. e.1 revised first and last name as they were reported incorrectly on the initial manufacturer device report number 1220648-2025-26451. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26451 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 the investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26451 and were added. | The investigation of purge leak ¿ pump has been completed. the returned pump was inspected and there were no abnormalities to the purge sidearm on visual inspection. the leak test was performed and a leak was observed from the membrane radial seal of the purge reservoir. the root cause of the purge leak - pump was determined to be due to a damage sidearm reservoir as evidenced by the results of the leak test. d6b explant date added d9 device returned to manufacturer date added e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26451 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11335 | 1220648-2025-26452 | 21505007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21505007&pc=OZD | 02/10/2025 | 02/11/2026 |  | 03/03/2025 |  |  | 02/10/2025 | 03/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594779 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Inflammation | Nerve Damage | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported the patient was on impella cp support and during support, the patient had bleeding at the access site. a femstop was applied and when it was removed, the patient continued to bleed. a purse string suture was applied and a femstop was reapplied. heparin was withheld. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding is determined to be use issue because purse string and femstop applied to control the bleeding. the activated clotting time (act) values (146) are in range during the time of bleed. hence, the root cause is determined to be use issue. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26452. | Investigation summary no product returned. infection: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. nerve damage/inflammation: the cause of the injury was not determined due to insufficient clinical details. access site adverse event: patient had redness and yellow drainage at impella site. the cause of the access site adverse event was not determined due to insufficient clinical details.
11336 | 1220648-2025-26453 | 21505982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21505982&pc=OZD | 02/06/2025 | 03/20/2025 |  | 03/03/2025 |  |  | 02/06/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539048 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced ectopy on day one of support. adjusting the p-level was unsuccessful. the impella was repositioned and ectopy resolved. it was also reported that on day eight of impella support the patient had a stroke. the patient had two head computed tomography scans to follow, and both came back normal. the patient had full function of the right side of their body, but patient was still having difficulty moving their left arm and leg. the patient continued on support until the deice was successfully weaned. patient outcome was noted to be stable. | [5 empty] | The investigation of vf/vt/af/at and stroke/cva has been completed. the returned product was visually inspected and biomaterial was found on the returned product which doesn't affect the results of the investigation. as this clinical information was not provided, the true root cause of the stroke/cva and vt/vf could not be determined.¿ d9 date device returned to manufacturer was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26453. h3 selections on initial manufacturer device report number 1220648-2025-26453 were incorrect as the device was received for investigation h6 type of investigation code 4114 was inadvertently entered incorrectly on initial manufacturer device report number 1220648-2025-26453. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26453. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11337 | 1220648-2025-26454 | 21506175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21506175&pc=OZD | 02/05/2025 | 04/02/2025 |  | 03/03/2025 |  |  | 02/05/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608322 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. shortly after implant the flows on the pump dropped and the team made the decision to explant the replace the pump. the second pump was placed but, the patient expired later on that second pump while in the intensive care unit. the team noted the death was not due to the pump per the field rep. the second pump supported the patient for under 3 hours before the patient expired. | Medical safety review of the event noted that the patient did not receive support in critical need. there was a delay in treatment. a new impella cp was implanted and supported the patient about three hours before the patients demise. the noted device issue may have added to an already complex situation. there is not enough information to dissociate the delay in therapy. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety review of the event noted that the patient did not receive support in critical need. there was a delay in treatment. a new impella cp was implanted and supported the patient about three hours before the patients demise. the noted device issue may have added to an already complex situation. there is not enough information to dissociate the delay in therapy. b.1 revised as product problem was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26454. b.2 revised to add death and date of death as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26454. b.5 added medical safety officer review as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26454. e.1 revised city and address line 1 as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26454. added facility name as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26454. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26454 as it is unknown if the initial reporter also sent the report to the food and drug administration. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2025-26454 and should not have been. g.1 revised reporting contact first and last name and manufacturing site name and address section as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26454. g.2 added health professional as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26454. g.3 30-day should have been selected on the initial manufacturer device report number 1220648-2025-26454 and was not. h.1 revised as it was previously selected incorrectly on the initial manufacturer device report number 1220648-2025-26454 h.6 added death code as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26454. h.10 added medical safety officer review as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26454.
11338 | 1220648-2025-26455 | 21506464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21506464&pc=OZD | 02/10/2025 | 03/24/2025 |  | 03/03/2025 |  |  | 02/10/2025 | 03/03/2025 |  |  |  |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594845 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella 5.5 placement, the 5.5 was not able to successfully advance impella through the right axillary artery graft due to patient's anatomy. upon removal of impella, clots were found in inlet. there was no patient harm and the device was exchanged to successfully complete the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. | The investigation of delivery issues and thrombus has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition due to difficulty inserting through axillary artery to across aortic valve. as the necessary information was not provided, the root cause of the thrombus was not determined.
11339 | 1220648-2025-26456 | 21509812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21509812&pc=OZD | 02/07/2025 | 03/03/2025 |  | 03/03/2025 |  |  | 02/07/2025 | 03/03/2025 |  |  | 08/31/2025 |  | 03/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397663 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the team was unable to deliver the pump to the patient due to the anatomy. the team aborted the placement of the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11340 | 1220648-2025-26414 | 21490609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21490609&pc=OZD | 01/27/2025 | 03/06/2025 |  | 02/28/2025 |  |  | 01/29/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213121 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient, critical, with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs). of note, prior to implant of the impella, the patient was initially on extracorporeal membrane oxygenation (ecmo) for suspected biventricular disfunction, then rejected and decision was made to place the impella cp even with the first symptoms of multi-organ failure appearing. later, same day of impella implant, an optical sensor system error occurred with the automated impella controller (aic). it was further explained the light was visible; however, the console was unable to recognize the optical sensor. there was an aic-to-aic exchange, the issue resolved, and the patient continued on the same pump with impella mcs. following, three days later, the patient was reported to have developed hemolysis and was in systemic multi-organ failure. treatment for the hemolysis is unknown; however, it was shared through additional information, the hemolysis event was unresolved. the next three days would pass, and it was reported the patient expired from multi-organ failure after a total of 6 days on impella mcs. an aic optical sensor failure occurred; however, support continued with the same pump that continues to operate without placement monitoring. a standard aic-to-aic exchange appears to have occurred, no loss or interruption in support. the patient developed hemolysis that was not resolved and this, however, may have added to an already complex situation leading to a possible association to the outcome. medical safety review event noted there is not enough evidence to exclude impella cp pump as an associated factor in the patient's outcome although the likely cause of death appears to be, as reported, multi-organ failure already present before the impella cp implant. | [5 empty] | The investigation of optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show that upon each attempt to start case, after connecting the pump case start wizard presented with ¿optical sensor system error¿ message. operators were prompted first to disconnect and reconnect pump, then - since issue would not resolve - to replace controller. real time log data showed that while the pump was connected, signal-to-noise ratio was extremely low (less than 10), placement signal was invalid, and gain was at its maximum value (2.65) - which is consistent with either optical fiber break or an air gap within the optical path. at the same time, the light parameter around 2.7 volt indicated that optical bench ccd detector receives light. two days after the reported event, the console was field-tested with a demo/test pump, and the issue was reproduced and the optical bench test indicated failing ¿5s¿ parameters. the console was tested, but the issue was not reproduced, and the console¿s optical system operated within specification. considering that the same pump worked well on the replacement console, an intermittent malfunction of optical bench or optical pump connector was the most likely cause of the issue. g.2 revised as other was selected incorrectly on the initial manufacturer device report number 1220648-2025-26414. health professional was inadvertently omitted from the initial manufacturer device report 1220648-2025-26414. | The automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this is one of two device manufacturing reports associated with the event. refer to the following: medical device report for automated impella controller serial number (b)(6) with date of event (b)(6) 2025 and patient identifier ending in (b)(6). (this report) medical device report for impella cp serial number (b)(6) with date of event (b)(6) 2025 and patient identifier ending in (b)(6).
11341 | 1220648-2025-26415 | 21490825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21490825&pc=OZD | 01/30/2025 | 03/21/2025 |  | 02/28/2025 |  |  | 01/30/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025477039 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient, critical, with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs). of note, prior to implant of the impella, the patient was initially on extracorporeal membrane oxygenation (ecmo) for suspected biventricular disfunction, then rejected and decision was made to place the impella cp even with the first symptoms of multi-organ failure appearing. later, same day of impella implant, an optical sensor system error occurred with the automated impella controller (aic). it was further explained the light was visible; however, the console was unable to recognize the optical sensor. there was an aic-to-aic exchange, the issue resolved, and the patient continued on the same pump with impella mcs. following, three days later, the patient was reported to have developed hemolysis and was in systemic multi-organ failure. treatment for the hemolysis is unknown; however, it was shared through additional information, the hemolysis event was unresolved. the next three days would pass, and it was reported the patient expired from multi-organ failure after a total of 6 days on impella mcs. an aic optical sensor failure occurred; however, support continued with the same pump that continues to operate without placement monitoring. a standard aic-to-aic exchange appears to have occurred; no loss or interruption in support. the patient developed hemolysis that was not resolved and this, however, may have added to an already complex situation leading to a possible association to the outcome. medical safety review event noted there is not enough evidence to exclude impella cp pump as an associated factor in the patient's outcome although the likely cause of death appears to be, as reported, multi-organ failure already present before the impella cp implant. | [5 empty] | The investigation for the hemolysis has been completed. clinical details stated that placement signal not reliable alarms triggered during the case until console changed. insufficient data log provided. no positioning alarms observed. no suctions occurred. unclear if the pump repositioned or resolved at any point during impella support. the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this is one of two device manufacturing reports associated with the event. refer to the following: medical device report for automated impella controller serial number (b)(6) with date of event 27-jan-2025 and patient identifier ending in (b)(6). medical device report for impella cp serial number (b)(6) with date of event 30-jan-2025 and patient identifier ending in (b)(6). (this report).
11342 | 1220648-2025-26416 | 21490962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21490962&pc=OZD | 02/03/2025 | 03/18/2025 |  | 02/28/2025 |  |  | 02/03/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607754 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHARMACIST | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella for mechanical circulatory support and post implant, same day, access site bleeding occurred and was managed (steps to manage the bleed are unknown). heparin was paused due to high partial thromboplastin time (ptt) less than 120. the following day the patient was noted to be unstable and still bleeding from the groin site. it was noted heparin was reduced, ptt 71. discussion around escalation to extracorporeal membrane oxygenation was noted. however, two days later it was indicated that the patient expired on support. no further details were made known. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11343 | 1220648-2025-26417 | 21491447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21491447&pc=OZD | 02/04/2025 | 03/21/2025 |  | 02/28/2025 |  |  | 02/04/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1847282 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during setup there were noted purge issues with the console. the automated impella controller had to be changed out. the case then proceeded as usual. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) purge issues has been completed. the device nor the data logs were returned for evaluation. the root cause of this issue could not be determined because the console was not returned. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26417. e.1 added us phone number and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26417.
11344 | 1220648-2025-26419 | 21491724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21491724&pc=PYX | 02/04/2025 | 03/23/2025 |  | 02/28/2025 |  |  | 02/04/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560823 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. prior to insertion, there was manipulation of venous cannula and the patient developed a large tear in right atrium. extensive repair was attempted unsuccessfully. the implant was aborted and the impella was never inserted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the delivery issues has been completed. no product was returned and logs are not applicable. the rp flex placement was aborted due to the ra tear from the bypass venous cannula that led to the patient expiring. no abiomed products ever touched the patient. the root cause of the delivery issues was most likely patient condition related since the patient had expired before the impella was able to be placed due to an ra tear from the bypass venous cannula.
11345 | 1220648-2025-26420 | 21491808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21491808&pc=OZD | 02/06/2025 | 02/28/2025 |  | 02/28/2025 |  |  | 02/06/2025 | 02/28/2025 |  |  | 10/31/2025 |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409456 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that attempts to implant an impella 5.5 pump were unsuccessful due to the staff being unable to cross the subclavian/innominate junction with the guidewire. the implant was aborted and an intra aortic balloon pump was placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11346 | 1220648-2025-26418 | 21491865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21491865&pc=OZD | 02/07/2025 | 02/28/2025 |  | 02/28/2025 |  |  | 02/07/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608574 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during the impella 5.5 procedure, the introducer was leaking at the peel away point. the impella 5.5 was then advanced into the ventricle and support was initiated. the introducer was then removed. it was further reported that care was withdrew and the patient expired. the relationship between the impella 5.5 and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11347 | 1220648-2025-26421 | 21491893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21491893&pc=OZD | 02/04/2025 | 03/25/2025 |  | 02/28/2025 |  |  | 02/04/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592984 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that when plugged into the console, the pump presented a purge failure alarm. the purge cassette, purge bag, and console were replaced, but the issue persisted. the pump was subsequently replaced for planned patient support. there was no patient harm. | [5 empty] | The investigation of priming issue has been completed. the purge cassette and pump were returned for investigation. according to the clinical details, purge failure alarm occurred during priming. the doctor tried the de-airing steps several times without resolution and encountered the same purge failure alarm on the second console. the doctor replaced the pump. the data log shows a purge pressure low alarm at the start of the case. multiple case start cancellation alarms and de-air cancellation procedures were reported in the data log. the low purge pressure alarms were off during the de-airing procedure. no physical damage was noted on the returned product. the pump was plugged into the console, followed by the purge cassette. priming was initiated and completed without any reported abnormalities. the imc logs were reviewed during the test run, and no abnormalities were reported. the cause of the priming issue was not determined since issue could not reproduce on returned device and sufficient clinical information was not provided. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11348 | 1220648-2025-26423 | 21492132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492132&pc=OZD | 02/02/2025 | 03/14/2025 |  | 02/28/2025 |  |  | 02/02/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565599 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced continuous oozing at the access site. the device continued to migrate away from the patient¿s body so the doctor made the decision to suture down the suture pad using all four eyelets. the patient continued on support. additional information noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related md made the decision to suture down the suture pad using all four eyelets to prevent the bleeding. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11349 | 1220648-2025-26422 | 21492133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492133&pc=OZD | 02/07/2025 | 04/08/2025 |  | 02/28/2025 |  |  | 02/07/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595543 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon attempting to prime the setup, no purge fluid was observed flowing out of the pump assembly. it was noted that the purge line had been successfully primed prior to the pump being attached. the pump was replaced to resolve the issue. there was no patient harm. | [5 empty] | The investigation of priming issue has been completed. the pump was returned for evaluation. the pump was connected to console and purged for approximately 18 minutes without issue. purge fluid was present exiting from outflow cage. the data logs were reviewed and real time log of pump priming in field shows pump was able to prime and have a stable purge pressure and purge flow. purge flow ticks are consistent after priming of purge cassette was performed. placement signal is near 0, indicating that the pump was never inserted into the body. clinical details noted that when pump was being primed, no purge fluid was observed exiting from outflow cage in the field. the pump was exchange for a new impella cp for patient support. no problem found with pump in the field or on returned product. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11350 | 1220648-2025-26426 | 21492171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492171&pc=OZD | 01/12/2025 | 03/12/2025 |  | 02/28/2025 |  |  | 02/04/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584976 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced ventricular tachycardia during impella cp support. the pateint required antitachycardia pacing and mexitil was resumed. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event revised on manufacturer device report 1220648-2025-26426 to reflect correct event date. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26426 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source; study was missing from manufacturer device report 1220648-2025-26426. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11351 | 1220648-2025-26424 | 21492190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492190&pc=OZD | 02/04/2025 | 04/02/2025 |  | 02/28/2025 |  |  | 02/04/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605040 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp insertion was aborted due to the patient¿s anatomy. there was no patient harm reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump was unable to be implanted due to the patient's anatomy. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26424. d.3 added manufacturer site zip code as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26424. e.1 revised address line 1 as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26424. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26424 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact address line 1, city, state, zip code and country as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26424. h.6 code 4611 was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26424. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11352 | 1220648-2025-26425 | 21492205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492205&pc=OZD | 02/04/2025 | 03/20/2025 |  | 02/28/2025 |  |  | 02/04/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541730 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. during the insertion there was difficulty passing through the patient¿s abdominal aneurysm. the physician tried to remove the impella. the guidewire kinked due to tortuosity and the device was difficult to remove. the pump and guidewire were advanced to loosen it, and the pump was then able to be removed. the pump was then run in a bowl of saline to make sure there was no damage to the impeller. the physician gained left ventricle access with the pigtail and the pump was then advanced over a platinum plus since the 0.018 wire was damaged. impella was started in auto. the patient continued on support until the device was successfully weaned. the patient outcome was noted to be stable. | [5 empty] | The investigation for the product damage has been completed. the product was not returned for investigation. as per the clinical information provided, the guidewire kinked due to patients' tortuous vessels. the cause of the issue was most likely patient condition related since the kinked guidewire was reported to be due to tortuosity. the failure mode was changed from product damage to delivery issue since the guidewire was not fully/partially fractured. b.7 added information as it was inadvertently omittted from the initial manufacturer device report number. e.1 added us phone number as it was inadvertently omittted from the initial manufacturer device report number. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2889 reported on the initial manufacturer device report number was replaced with a new code due to investigation results. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11353 | 1220648-2025-26427 | 21492216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492216&pc=OZD | 02/01/2025 | 02/28/2025 |  | 02/28/2025 |  |  | 02/01/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590070 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with myocarditis had an intra-aortic balloon pump (iabp) and extracorporeal membrane oxygenation (ecmo) placed for initial treatment one day prior to the impella cp implant, which was an exchange from the iabp. the iabp was removed over-the-wire. a new sheath was inserted, and the same access (right femoral artery) was used to successfully implant the impella cp. approximately 10 days after start of the impella support, oozing at the access site occurred when the patient coughs. the event was treated with a stitch. the patient continued on impella mechanical circulatory support (and ecmo) for approximately four days until "successful" orthotopic heart transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11354 | 1220648-2025-26428 | 21492217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492217&pc=OZD | 02/06/2025 | 03/04/2025 |  | 02/28/2025 |  |  | 02/06/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1227632 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a console (aic) was observed to have console damage. the crack occurred at an unknown time/cause. the aic will be replaced and not used for a patient. | [5 empty] | The investigation for the console (aic) damage has been completed. console logs did not contain any data relevant to the complaint. console was returned for evaluation. the reported damage was confirmed. cracks were observed on the front housing near the glass keyboard as well as on the rear housing. persistent rebooting of the rlm was also observed while testing the console but was unrelated to the reported cracked console issue. the console functioned normally after testing with a pump and purge cassette. the damage on this console was likely use related. the unrelated rlm reboot issue was isolated to the som card and resolved after troubleshooting using another som card. | A.3 revised as selection was left blank on the initial manufacturer device report number 1220648-2025-26428. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number: 1220648-2025-26428. d.5 revised operator of device as it was previously entered incorrectly on the initial manufacturer device report number: 1220648-2025-26428. e.4 should have been left blank on the initial manufacturer device report number: 1220648-2025-26428 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact facility name address and phone number as they were previously entered incorrectly on the initial manufacturer device report number: 1220648-2025-26428. revised manufacturing site name and address as they were previously entered incorrectly on the initial manufacturer device report number: 1220648-2025-26428. h.6 codes were added as they were inadvertently omitted from the initial manufacturer device report number: 1220648-2025-26428.
11355 | 1220648-2025-26429 | 21492242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492242&pc=PYX | 02/06/2025 | 03/25/2025 |  | 02/28/2025 |  |  | 02/06/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025548247 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that following impella rp flex implant, the purge flow abruptly began to decrease. tissue plasminogen activator was run through the purge and the flow rate increased to a stable rate. there was no noted patient harm. | It was further reported that despite improvement with additional bipella support, the patient never recovered from the impact to his liver and kidneys. as a result, the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. the high purge flow did not interrupt support or cause a change in the therapy. the effect was to the user as the user had to take action (i.e. tpa to the purge system), user inconvenience. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Added b5-mso review statement. | The investigation for the high purge pressure has been completed. no product was returned for analysis. the data logs show a motor current spike with speed deviation, shift in placement signal and shift in pump flows characteristic of biomaterial ingestion. after motor current spike, purge pressure trended upwards and there was a decrease in purge flow. the purge pressure trended up but did not meet threshold for alarm trigger. as clinical mentioned, tissue plasminogen activator (tpa) protocol was initiated and purge pressure stabilized between 600mmhg to 800mmhg for the rest of the case. the root cause of the high purge pressure was most likely due to biomaterial ingestion as evidenced by data log analysis. the instructions for use for the impella rp with smartassist system states the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ added: h6 impact code 4644.
11356 | 1220648-2025-26430 | 21492337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492337&pc=OZD | 02/06/2025 | 03/26/2025 |  | 02/28/2025 |  |  | 02/07/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628426 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | Tachycardia | Vascular Dissection | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp to support a patient admitted in for a high-risk percutaneous coronary intervention. while on support it was reported that the patient became hypotensive and tachycardiac and required vasopressors. the decision was made to escalated to the impella 5.5 pump. upon cp removal a dissection of the right common femoral artery was noted and interventional cardiology performed intervention using two long balloon inflations. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. clinical information provided mentioned that the dissection was noted when during the angiography but the reason for the dissection is unknown. therefore, the root cause of the vascular damage issue was not determined since th product was not returned and insufficient clinical information was provided. d.4 revised catalog number, serial number, lot number, expiration date, and unique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26430. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26430. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26430. h.6 code 1721 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26430. a new code was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11357 | 1220648-2025-26432 | 21492826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492826&pc=OZD | 02/07/2025 | 03/31/2025 |  | 02/28/2025 |  |  | 02/08/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574769 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed along with extracorporeal membrane oxygenation for the support of a patient. the team did not peel away the 14 french sheath per the medical doctors preference. the team observed bleeding which prompted the team to stop the heparin and give two units of blood. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the pump was not received for investigation. based on the clinical details the root cause of access site bleeding is determined to be most likely patient condition because the patient is also oozing from the extracorporeal membrane oxygenation (ecmo) site. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11358 | 1220648-2025-26431 | 21492838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21492838&pc=OZD | 01/30/2025 | 04/01/2025 |  | 02/28/2025 |  |  | 01/30/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553372 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support and subsequently experienced intestinal ischemia. the report noted it is possible a clot developed during percutaneous coronary intervention or due to heparin-induced thrombocytopenia (hit). however, the cause is unclear. thus, the relationship to the impella device is unknown. the patient would expire as a result of the event. additional comments noted, the cause of death as intestinal ischemia, death relationship to the impella is unknown (however the pump was functioning normally) and meylon was added to the purging solution due to the suspected hit. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. an attempt to clarify the date of death, date of event and awareness date has been made. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
11359 | 1220648-2025-26433 | 21493384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21493384&pc=OZD | 02/07/2025 | 03/27/2025 |  | 02/28/2025 |  |  | 02/07/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568134 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that several attempts were made to insert impella 5.5 but was unsuccessful due to the patient¿s vessel tortuosity. while advancing the catheter, it was noted that the tortuosity was causing the wire to back out. the physician advanced the .018 wire, and it crossed without the catheter into the lv. at this time, rhythm noted to change to a very fast supraventricular tachycardia (svt). patient was given adenosine but did not convert and was shocked x 200 with successful cardioversion to pace rhythm. the physician removed .018 wire and impella 5.5 and noted bleeding from subclavian pocket. upon further assessment, a hole in the anastomosis of the graft and artery was noted and a repair stitch was placed which successfully controlled the bleed. three units prbc were administered. decision made to place patient on va ecmo with iabp placement. | [5 empty] | The investigation of access site bleeding, delivery issues, and vf/vt/af/at has been completed. the pump visually inspected and no damages or abnormalities found. reported multiple failed attempts gaining access to the ascending aorta due to tortuous anatomy and potential calcium deflecting the wire into the descending aorta. the .018 did gain eventually gain access to the lv but lost the position when advancing the 5.5. insertion was aborted. the root cause of the delivery issue was most likely patient condition related since reported tortuous anatomy. reported bleeding in subclavian pocket from a hole in the anastomosis of the graft and artery. a repair stitch controlled the bleed. the root cause of the access site bleeding - major was most likely use related since a repair stitch controlled the bleed. as the necessary information was not provided, the root cause of the vt/vf was not determined. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11360 | 1220648-2025-26435 | 21494003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21494003&pc=OZD | 02/05/2025 | 03/19/2025 |  | 02/28/2025 |  |  | 02/05/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596822 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 02/05/2025 |  |  |  |  | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced hematoma with a hemoglobin decrease. impella was removed and the puncture site was closed with a proglide and an angioseal and a pressure bandage was applied. blood units were ordered. patient was stable. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient movement since the hematoma occurred after the patient was transported from the catheterization laboratory to the intensive care unit. b.2 added death and date of death as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26435. b.7 added information as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26435. e.1 revised address line 1 as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26435. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26435 as it is unknown if the initial reporter also sent the report to the food and drug administration. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2025-26435 and should not have been. g.1 added reporting contact address line 1, us city, state, country and us zip code as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26435. h.6 health effect impact clinical code 1888 was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26435 and a new code was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11361 | 1220648-2025-26434 | 21494058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21494058&pc=OZD | 02/08/2025 | 05/08/2025 |  | 02/28/2025 |  |  | 02/08/2025 | 02/28/2025 |  |  | 02/28/2026 |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490998 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that three days into impella cp support, the patient's urine turned red and their hemoglobin levels dropped. three units of blood were given to initially treat the condition. the pump was later explanted in response to the hemolysis. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of hemolysis has been completed. returned complaint device visually inspected. no biomaterial observed. no broken blade or sharp edge found. no abnormality observed. no data log available to review. hemolysis test conducted successfully with mih value =5.69 hemolysis: the cause of the hemolysis issue cannot be determined due to insufficient clinical information, no data logs available, and the returned pump passed hemolysis testing. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11362 | 1220648-2025-26436 | 21494310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21494310&pc=OZD | 02/01/2025 | 04/02/2025 |  | 02/28/2025 |  |  | 02/01/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558117 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support (mcs) had a noted possible pump "failure" 4 days after start of support and noted pump "stop "12 days into support. the patient was cannulated peripheral venoarterial extracorporeal membrane oxygenation (va ecmo) in the emergency department before being transported to the catheterization lab. angiogram showed chronic total occlusion of the left anterior descending artery and 90% occlusion of the left circumflex and mid right coronary arteries. the impella was inserted and initiated; the patient was transported to another facility for care. four days into impella cp support, a report of possible pump failure was noted. no further information was provided. following now after 12 days of patient support the impella was dropped to performance level p-2 with an anticipated possible removal within the next couple of days. however, the pump stopped. the patient remained hemodynamically stable. the heart team followed protocol until moved to the catheterization lab for removal of the impella cp which was successfully completed. there was no harm to the patient as a result of this event. | [5 empty] | The investigation of pump stop has been completed. the data logs were returned and investigated. there was an alarm with a motor current spike to 30000 ma while the pump was undergoing weaning at p2 at the end of the case when the pump stopped and was explanted. since the motor current spiked to 30,000, the cause of the failure was electrical-related. at that time, the pump was ran for 279 hours ie almost 12 days which is past the design life cycle of the cp pump per design trace matrix. the returned product was tested for electrical resistances and all motor phases were found to be open. v+/v- was 5600 ohms. the catheter and patient cable to kink protects bonds were checked and found to be sufficient. no kinks in the catheter or kink protector were observed. evidence of a detached monkey fist bond and motor cable twisting within the red handle. evidence of damage to the conductive material in the red and green motor cables and evidence of necking and conductive material damage to the white motor cable. the root cause of the pump stop issue was most likely an motor cable line open electrical due to excessive force applied to the catheter. d.9 date device returned/information was added. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26436. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11363 | 1220648-2025-26437 | 21495345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21495345&pc=OZD | 02/08/2025 | 02/28/2025 |  | 02/28/2025 |  |  | 02/08/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591588 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp pump via the non-abiomed provided sheath. the sheath was the sentrant 14 french and then the 18 french sentrant when the 14 french failed to allow the impella cp to deliver through. the limb became ischemic. the discussion was made for a perfusion catheter but the family made the choice not to escalated care and instead withdrew care. therefore, the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11364 | 1220648-2025-26438 | 21495759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21495759&pc=OZD | 02/09/2025 | 03/31/2025 |  | 02/28/2025 |  |  | 02/09/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523050 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after placement of impella cp, the impella cp had continuous suction. they were unable to resolve with repositioning, fluids, and p level decrease. therefore, the patient was cannulated for extracorporeal membrane oxygenation at bedside and the impella cp was removed. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of positioning issues has been completed. the data was not returned for review. the pump was returned in damaged condition without pump plug. visual inspection found a kink on can about 15 mm from of cage &and can bonding site. no other issues were found on the rest of the pump. the root cause of low flow was most likely kinked cannula. note: changed failure mode (fm) from positioning issue to low flow as clinical description confirmed pump was in proper position and had suction alarms and low flow for the whole case. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11365 | 1220648-2025-26439 | 21496969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21496969&pc=OZD | 02/01/2025 | 02/28/2025 |  | 02/28/2025 |  |  | 02/01/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605362 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported the implant of impella cp device to a patient (unknown race and ethnicity) in cardiogenic shock (cgs) due to an acute myocardial infarction was unsuccessful. the patient with st-segment elevation myocardial infarction had successful deployment of percutaneous coronary intervention to the left anterior descending artery. while completing case for transfer to intensive care unit, the patient arrested and decision to place impella cp for cgs. the initial pump attempt to place with a 14fr x 13cm sheath was unsuccessful. the physician removed and discarded the impella cp pump and the 14fr x 13cm sheath was exchanged for 14fr x 25cm sheath over stiffer wire. the replacement impella cp was placed over abiomed 0.018¿ wire and started on auto. shock alert called with potential plans to transfer. however, within two hours, the patient became hemodynamically unstable when into cardiac arrest and expired in the procedural area. the family stopped care. based on the information at hand, the patient was very sick. the physician for some reason which is unknown (however could have been due to vessel characteristics etc.; details are unknown) was unable to deliver the impella cp device. the patient needed support, and there was a consequential delay with providing impella mechanical circulatory support. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11366 | 1220648-2025-26440 | 21497058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21497058&pc=OZD | 02/08/2025 | 03/23/2025 |  | 02/28/2025 |  |  | 02/08/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721589 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that automated impella controller (aic) was being used with a patient when the aic shut down without warning. the aic was replaced and the patient remained stable. | [5 empty] | The investigation for the console (aic) power issues has been completed. data logs were reviewed finding ¿shutdown requested by user, but pump is connected.¿ it is likely that at this point, the aic screen was blank, leading the customer to assume that the ¿aic shut down without warning,¿ which may have caused the user to press the power button multiple times to turn on the console. subsequently, the console was forcefully shut down, as indicated by the log entry: ¿emergency shutdown imminent ¿ alarm.¿ after the console rebooted, another log entry recorded: ¿unexpected controller shutdown ¿ alarm,¿ due to the force shutdown. this confirms the reported failure mode. console was returned for evaluation. this console was tested after arriving at fs. the reported failure mode could not be reproduced while running the optical test pump. by wiggling the cables between the display and the 4-in-1, no issues were found with the display. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). the root cause of the blank screen (which the user assumed was a console shutdown) was not determined due to the inability to reproduce the issue. corrections to the initial mfr report: h5 should have been no. h8 should have been reuse instead of initial use of device. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11367 | 1220648-2025-26441 | 21497059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21497059&pc=OZD | 02/02/2025 | 03/27/2025 |  | 02/28/2025 |  |  | 02/08/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587856 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed sepsis and passed away during impella cp support. the relationship between the impella and cause of death is unknown. no other information was provided. | [5 empty] | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26441 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11368 | 1220648-2025-26442 | 21497081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21497081&pc=OZD | 02/09/2025 | 04/09/2025 |  | 02/28/2025 |  |  | 02/09/2025 | 02/28/2025 |  |  |  |  | 02/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608338 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient experienced some bleeding / oozing from their access site during impella cp support. it was documented that when the patient became hypotensive, their access site started to bleed extensively. upon returning to a stable blood pressure, the bleeding resolved and both the position and access site were checked by the cardiologists. one unit of packed red blood cells was given to the patient to help treat the condition. support continued uninterrupted. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be patient condition because the patient became very hypertensive causing the bleed and bleeding was resolved after the blood pressure was in range again. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26442 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. e.1 name prefix/title, first name and last name are unknown. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11369 | 1220648-2025-26371 | 21477999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21477999&pc=PYX | 02/05/2025 | 09/19/2025 |  | 02/27/2025 |  |  | 02/02/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025587556 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella rp had placement signal not reliable alarms, while on patient support. several days later, the pa pressure on the flex abruptly increased with matching motor current spike. flows dropped. xray ordered and position stable and appropriate. the patient relatively unchanged with the drop in flow. patient labs will be drawn every hour and if patient tolerates wean, the impella will be removed. mdr report required for the malfunction. | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. low pump flow presented during support. the patient continued on support until the device was successfully weaned. there was no patient harm reported. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. optical signal issue. . log review showed a drop in cvp and signal not reliable (snr) with optical sensor related placement signal not reliable (psnr) alarms and suction alarms in the first day of support. the cvp was below 0 mmhg for much of this time indicating a lack of pressure and therefore suction. although the snr decreased intermittently, the rest of the optical parameters did not indicate sensor issues. no more psnr alarms occurred past the first day of support. the root cause of the optical signal issue was most likely patient condition related since the drop in cvp was occurring alongside suction alarms and no other abnormalities were found that would suggest sensor damage or issues. low pump flow. log review showed a large mc spike with subsequent drop in flows (~4.1l/min to 2.8l/min), step up in placement signal, and rise in purge pressure. the pump was then successfully weaned and explanted within 7 hours. it was noted that the patient was not heparinized due to pericardial drain. also, biomaterial was observed in the outflow cage upon explant. the root cause of the low pump flow was most likely biomaterial since there was a large motor current spike in the data logs prior to the abrupt drop in pump flow. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. no product was returned. log review showed a large motor current spike with subsequent drop in flows (approximately 4.1 liters per minute (l/min) to 2.8 l/min), step up in placement signal, and rise in purge pressure. the pump was then successfully weaned and explanted within seven hours. it was noted that the patient was not heparinized due to pericardial drain. also, biomaterial was observed in the outflow cage upon explant. the root cause of the low pump flow was most likely biomaterial since there was a large motor current spike in the data logs prior to the abrupt drop in pump flow. b.3 revised date of event as it was previously entered incorrectly on the initial manufacture device report number 1220648-2025-26371. g.3 date of awareness should of reflected 5-feb-2025 on the initial manufacture device report number 1220648-2025-26371.
11370 | 1220648-2025-26372 | 21478378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478378&pc=OZD | 02/03/2025 | 04/08/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600297 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implant with an impella cp device for mechanical circulatory support and approximately 13 days after start of the support the pump stopped. the pump was unable to be restarted, which lead to explanted and replacement of the impella cp device. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. the impella cp was returned for evaluation. upon visual inspection a large purge gap indicating bearing wear was observed on the returned pump. no data logs were returned for analysis. clinical details noted that the pump stop occurred on day 14 of support and pump was unable to be restarted after initial pump stop. the root cause of the pump stop was due to bearing wear beyond intended design life of the pump per design trace matrix.
11371 | 1220648-2025-26373 | 21478389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478389&pc=OZD | 02/03/2025 | 03/21/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594323 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and 28 days after start of the support high purge pressure developed. this day there was an acute drop in purge flow. tissue plasminogen activator was attempted in the purge. however, the purge issue remained unresolved. the patient was cannulated for venoarterial extracorporeal membrane oxygenation and the impella 5.5 device was explanted for concern of impending pump stop. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. product was not returned for investigation. data logs were reviewed, though only real time (rt) logs were available. the purge pressure starts to show instability and the ps shows an increased number of spikes. the purge pressure begins to trend up and purge flows trend down gradually. finally, there is a sharp 10 minute rise in the purge pressure, resulting in an alarm. through none of this trend up in the purge pressure are the motor current or speed ever affected. the purge pressure remains high for the remainder of the case as was reported in clinical details. based on the purge pressure trending up for 4 hours without any other abnormalities being noted in the data logs, the root cause of the high purge pressure was most likely biomaterial build up. added h6 impact code 4644.
11372 | 1220648-2025-26374 | 21478460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478460&pc=OZD | 02/03/2025 | 03/14/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608902 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility noted a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and the following day during repositioning of the implanted impella cp pump, the sterile sleeve at the distal end of the catheter tore. tegaderm was placed to cover the tear and support continued uninterrupted. there was no patient harm. | [5 empty] | The investigation product damage (sterile sleeve damage) has been completed. the returned pump was inspected and the sterile sleeve separation was confirmed. there were no other physical abnormalities of the pump. the clinical information mentioned that this occurred during pump repositioning. the root cause of the product damage (sterile sleeve damage) was most likely use-related. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11373 | 1220648-2025-26375 | 21478505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478505&pc=OZD | 02/01/2025 | 02/27/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601917 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a/an (unknown gender, race and ethnicity) patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and after the impella cp was placed, the patient experienced bleeding. the patient was taken to the operating room and left ventricle perforation was found, requiring intervention. the patient subsequently expired. the physician subsequently commented, the perforation had nothing to do with the impella. the bleeding must have occurred before and it did come from the right heart. no further details were made available after the physician's comment and considering this, there is still not enough evidence to exclude the impella device as an associated factor. | [5 empty] | The device was not returned for evaluation. if new information or the device is received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
11374 | 1220648-2025-26376 | 21478548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478548&pc=OZD | 01/28/2025 | 02/27/2025 |  | 02/27/2025 |  |  | 01/28/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582347 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Atrial Fibrillation | Tachycardia | No Clinical Signs, Symptoms or Conditions | [10 empty] | Life Threatening | Required Intervention | Death |  |  |  | PROTEK DUO AND VV ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced delivery difficulties with the guidewire. the guidewire bent during delivery of the impella 5.5 pump. a second guidewire was obtained, and it bent as well. it was noted that the patient was obese with a very thick chest that may have compromised the anatomy/angles being traversed. upon use of a third guidewire, the implant was successful, and the patient was able to receive impella mcs. | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced delivery difficulties with the guidewire. the patient would also experience arrhythmic events and subsequently expire. the guidewire bent during delivery of the impella 5.5 pump. a second guidewire was obtained, and it bent as well. it was noted that the patient was obese with a very thick chest that may have compromised the anatomy/angles being traversed. upon use of a third guidewire, the implant was successful, and the patient was able to receive impella mcs. after placement of the impella 5.5 device, the patient was implanted with a competitor's right-sided device (protek duo), which was advanced to the pulmonary artery to treat right ventricular failure. following the patient experienced atrial fibrillation with rapid ventricular response and was cardioverted twice. the patient was sent to the unit for continued care. it was noted that at some point the patient was placed on veno-venous extracorporeal membrane oxygenation (vv ecmo) as well. while on support approximately five days later, the patient went into ventricular tachycardia and required shocks. following the patient would experience intermittent bouts of ventricular tachycardia even right before his demise. care was eventually withdrawn and the patient expired. the patient had multiple devices, left side impella, right side (protek duo) and ecmo, and given the arrhythmic events it is unclear which device may have possibly caused or contributed to the event and its outcome. medical safety review of the event noted there is not enough evidence to exclude impella 5.5 pump as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Additional details on the complaint were reviewed by medical safety and updates to the type of reportable event and corresponding fields were required. these updates have been made to the following sections: b1, b2, b5, h1 and h6 (health effect and medical device problem codes) accordingly.
11375 | 1220648-2025-26377 | 21478556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478556&pc=OZD | 02/03/2025 | 03/19/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after insertion, the device alarmed placement signal not reliable. the impella was repositioned as it was too deep. the impella still a little deep but was flowing excellent and the physician felt that the position was good. the patient was noted to be hemodynamically stable after repositioning. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. complaint device not returned for analysis. however, clinical information was sufficient to determine the root cause. the cause of the positioning issue most likely was use related since the impella was looked deep after insertion and placement signal not reliable alarmed.
11376 | 1220648-2025-26378 | 21478568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478568&pc=OZD | 02/03/2025 | 03/12/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585410 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced limb ischemia after impella cp was inserted. external fem-fem bypass was performed. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation." | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. e.1 revised to add us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26378. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26378 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11377 | 1220648-2025-26379 | 21478577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478577&pc=OZD | 01/31/2025 | 02/27/2025 |  | 02/27/2025 |  |  | 01/31/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603986 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support (mcs) and the following day escalated impella 5.5 device. the patient was intubated and sedated on impella 5.5 support. the pump had one suction, and it was noted the patient might have coughed right before impella in aorta alarms per nursing staff. upon checking echocardiogram images, it was noted the impella was migrated to aorta. this occurred one day after start of the impella 5.5 support. performance level was dropped to p-2 and the surgeon was called to repositioned. the surgeon was unable to cross back to aortic valve and required to explant. the surgeon requested to check if catheter lock was appropriate because felt a little loose than expected give past use experience. it is unknown if the impella 5.5. was replaced or if mcs therapy was subsequently required. however, outcome at explant noted the patient survived and was in stable condition. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11378 | 1220648-2025-26380 | 21478605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478605&pc=OZD | 01/31/2025 | 04/08/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517827 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device and following the implant, a drop in flow rate occurred. fluoroscopic imaging discovered a kink in the cannula proximal to the motor. attempts to adjust the device were unsuccessful and the pump was ultimately replaced with another impella cp pump. additional comments noted according to the physician, there was no resistance felt when inserting the device, and placement was not difficult. there was no patient harm as a result of the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. | The investigation of low pump flow has been completed. visual inspection found a kink on can about 15 mm from of cage & can bonding site. no other issues were found on the rest of the pump. pump ran without issue under bench test. the root cause of the low flow was most likely kinked cannula. the root cause of the product damage (cannula kink) issue was a kink due to patient anatomy (aortic arch) issue. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26380 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11379 | 1220648-2025-26381 | 21478633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478633&pc=OZD | 02/03/2025 | 04/04/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585370 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. during the procedure for protective pci, the distal end of the 14frx25cm peel-away sheath kinked. the physician was able to manipulate the equipment around the kink and finish the procedure with no harm to the patient. the impella cp device was removed at the end of the case. perclose x 2 and manual pressure held to achieve hemostasis, and the patient was transferred to the unit for postop monitoring. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the introducer damage has been completed. no product was returned and logs are not applicable. the root cause of the introducer damage was most likely a sheath kink due to the patient's tortuous vessels. | D.9 date device returned/information was added. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26381. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26381 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 medical device problem code 2682 was inadvertently omitted on the follow up 1 manufacturer device report number 1220648-2025-26381-1.
11380 | 1220648-2025-26382 | 21478635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478635&pc=OZD | 02/03/2025 | 03/14/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569279 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the following day after start of support, here were high purge pressure/purge system blocked alarms. purge cassette tubing was changed without success. the patient moved or coughed, and the purge flow decreased, and the purge pressure increased. tissue plasma activator was added to the purge. all waveforms and numbers on impella were within normal limits and have normalized out. the patient continued on impella mcs until removal the following day. the impella was weaned down, and the device was successfully explanted. thereafter, the patient was taken back to the unit and was noted to be doing well. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm. if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump was not returned for investigation. the data log was only provided for the first few hours of the case run, and no abnormalities were reported during that time. according to the clinical notes, high purge pressure was reported on the second day of pump use. the cause of the high purge pressure issue was not determined since the product was not returned and sufficient clinical information was not provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26382 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revise manufacturing site name and address as they were inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26382. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26382 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
11381 | 1220648-2025-26383 | 21478641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478641&pc=OZD | 02/03/2025 | 03/14/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5,5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578562 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during the implant, there was resistance around aortic arch. device removed from patient and found few pieces of intima from vessels at inlet. the doctor decided to abort the implant. | [5 empty] | The investigation of thrombus and delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy and was aborted. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26383 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect - clinical code 4440 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26383. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11382 | 1220648-2025-26384 | 21478666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478666&pc=OZD | 02/03/2025 | 03/14/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594093 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 95 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful. impella cp insertion was attempted several times. physician decided to abort impella implant and proceed with the percutaneous coronary intervention. the patient was noted to be in stable condition following the event. | [5 empty] | The investigation of delivery issues has and been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy and was aborted. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a/d6b were inadvertently omitted on the initial manufacturer device report number (b)(4). e4 should have been left blank on the initial manufacturer device report number (b)(4) as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number (b)(4). | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11383 | 1220648-2025-26385 | 21478673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478673&pc=OZD | 02/03/2025 | 10/01/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025577841 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the device exhibited an optical signal issue. | The user facility reported a patient (unknown race) with cardiomyopathy was initially implanted with an impella cp device and subsequently escalated to an impella 5.5 device for continued mechanical circulatory support and bridge to transplant. approximately two months into support on the impella 5.5, there was high purge pressure. tissue plasma activator was added to the purge and support continued. approximately seven days later it was noted that the patient refused heart transplant and further treatment. the patient was removed from heart transplant list and the impella was successfully explanted. the patient survived the explant and was reportedly discharged home. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. according to the given clinical data, the high purge pressure returned to normal after tpa was administered. additionally, psnr was reported on 02/07. on 02/03, the data log shows a gradual rise in purge pressure over 5 hours, followed by a high purge pressure alarm. the high purge pressure gradually returned to normal limits after 10 hours. furthermore, on the 63th day of pump use, the gradual decrease in placement signal and 5s optical signal values verified the sensor drift pattern, which triggered the loss of placement signal with a psnr 1048 alarm. the cause of the placement signal issue was sensor silicone wear beyond the intended design life of 60 days, based on data log review. the cause of the high purge pressure issue was biomaterial buildup, based on data log review and provided clinical details. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the cause of the high purge pressure issue was biomaterial buildup, based on data log review and provided clinical details. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26385 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11384 | 1220648-2025-26386 | 21478687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478687&pc=OZD | 01/31/2025 | 03/12/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589337 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was initially treated with an impella cp and extracorporeal membrane oxygenation (emco) and then escalated to an impella 5.5 device. after approximately 11 days on support, the patient experienced a right-sided stroke which was confirmed on computed tomography scan. the patient had various bouts of bleeding and anticoagulation was held. the plan was to bring the patient back to the operating room later in the day for removal of venoarterial ecmo. however, the family decided to pull support over the weekend, and the patient expired two days later on a sunday. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11385 | 1220648-2025-26387 | 21478703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478703&pc=OZD | 02/03/2025 | 02/27/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576644 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. and following the impella implant, "significant" bleeding was noted coming from a tear in the patient's axillary artery. a surgical repair was completed, and two units of blood were transfused. it was then noted that a small blood leak was found coming from the seam of the red impella plug. damage to the catheter was suspected; however, the decision was made to maintain support while monitoring the pump function. the patient was noted to be doing well. after a total of approximately 23 days on support, the pump was successfully weaned and explanted in the operating room. the patient was on modest dose of dobutamine; however, otherwise the patient was clinically stable upon pump exit. the patient was on minimal respiratory support, and patient had been able to maintain an appetite and increased purposeful engagement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11386 | 1220648-2025-26388 | 21478793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478793&pc=OZD | 02/03/2025 | 03/10/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1368535 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a system self-check failure appeared on an automated impella controller (aic) during preparation for patient support. the aic was swapped. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the data analysis showed that the console logs had a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring communication channel. the root cause of the "system self check failure" was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
11387 | 1220648-2025-26389 | 21478855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478855&pc=OZD | 02/03/2025 | 03/21/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553833 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device was presenting saturated flow and high pure pressure alarms. the impella was explanted and clot was found in the outflow cage. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist, section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation has been completed. the impella 5.5 was returned for evaluation. upon visual inspection, biomaterial was present around the impeller blade and purge gap. the pump was connected to the console and purged for approximately 10 minutes. purge pressure steadily rose and eventually triggered high purge pressure alarms after 15 minutes. the pump was attempted to be run but immediate pump stop occurred. no blood was present in the purge line just proximal from the motor housing. the pump was cut right before the motor housing and high purge pressure resolved. data logs were reviewed and long term log showed a slow gradual buildup of purge pressure over a period of approximately three days. just prior to slow increase in purge pressure, a motor current spike was observed. real time log at time of motor current spike showed a motor current spike with a large speed deviation and momentary drop in pump flow. pump flow after motor current spike becomes saturated. clinical details noted that on day eight of support, the pump flows were not within the normal range for p-4. purge flows were observed to be low with a high purge pressure. motor current was also observed to be slightly high for p-4. it was noted that patient was not on heparin for two days prior to high purge pressure, only sodium bicarbonate in purge solution. the root cause of the high purge pressure was most likely due to biomaterial ingestion that lead to gradual buildup of purge pressure over an extended period of time. as it relates to the saturated flow, the pump was attempted to be run in the lab but an immediate pump stop occurred. data logs were reviewed and long term log showed a gradual increase in motor current with flows becoming more saturated throughout end of support. directly before motor current begins to rise and flows begin to saturate, a motor current spike is present. real time log at time of motor current spike shows the motor current spike occur with a speed deviation and momentary drop in pump flow. flows were clipping and saturated for six minutes prior to motor current spike. clinical details noted that the pump was flowing higher than normal values at p-4 and motor current was observed to be higher than p-level. the root cause of the saturated flow was most likely due to biomaterial. thrombus was removed and investigated under other failure modes listed as the root cause of the "high purge pressure" and "saturated flow" due to biomaterial. b.1 revised since the initial manufacturer device report number 1220648-2025-26389 was submitted to product problem due to investigation results. b.2 selection was removed since the initial manufacturer device report number 1220648-2025-26389 was submitted due to investigation results. b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26389. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26389. h.1 revised since the initial manufacturer device report number 1220648-2025-26389 was submitted to malfunction due to investigation results. h.6 revised health effect - clinical code since the initial manufacturer device report number 1220648-2025-26389 was submitted due to investigation results.
11388 | 1220648-2025-26390 | 21478870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21478870&pc=OZD | 02/03/2025 | 03/06/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  | 01/31/2026 |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453255 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in postcardiotomy cardiogenic shock post coronary artery bypass graft surgery one week prior was implanted with an impella 5.5 device for mechanical circulatory support and experienced a left ventricle perforation and would expire. further details noted the graft for direct impella 5.5 implant was tied too close to the aortic valve at 5 centimeters. relocating the graft was recommended. the physician wanted to make the current location work and was advised this was not recommended and if the decision was to move forward with the current location, the impella device must be clear of the graft. the patient was able to go on full impella support and come off bypass. however, upon closing the sternum and/or trimming graft, a decline in the patient and flows was noted. a left ventricle perforation was discovered. this led to coding and many direct shocks to heart. the patient expired. | [5 empty] | The investigation for the ventricle perforation has been completed. no product was returned. the root cause of the ventricle perforation was not determined as insufficient information on the cause and timing of perforation was provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
11389 | 1220648-2025-26391 | 21479010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479010&pc=OZD | 02/03/2025 | 04/01/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589163 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. after two days of support, the patient experienced a brain bleed. heparin was in the pump purge. subsequently, care was withdrawn and the patient expired. no additional information was provided. | [5 empty] | The investigation of stroke/cerebrovascular accident (cva) has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11390 | 1220648-2025-26392 | 21479049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479049&pc=OZD | 02/03/2025 | 02/27/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546616 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that shortly after impella cp started, impella motor current high alarm with decreased flows occurred. the cp was removed and upon inspection of catheter, a clot was observed within the cannula. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smartassist catheter insertion. section: axillary insertion of the impella cp with smartassist catheter. section: use of impella with transcatheter aortic valves: in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.
11391 | 1220648-2025-26393 | 21479096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479096&pc=OZD | 02/03/2025 | 04/01/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087127 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during impella mechanical circulatory support it was noticed that the automated impella controller (aic) was missing the gray door that covers the impella plug site and replacement of the door was requested. there was no patient harm or indication of interruption in therapy as a result of this event. | [5 empty] | The investigation for the console (aic) damage during therapy has been completed. data logs were not applicable to this event. this console was returned for evaluation and it was confirmed that the gray cover of the front panel optical connector was missing. this confirms the reported failure mode. the envelope (without the test cavity) was noisy, and the (snr) was marginally passing at 2405, unrelated to the reported failure mode. the root cause for missing purge door was most likely use issue. corrections to the initial mfr report: h5-should have been no. h8-should have been reuse. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11392 | 1220648-2025-26394 | 21479106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479106&pc=OZD | 02/03/2025 | 04/03/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553160 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed a hematoma while moving from bed to med flight. manual pressure was applied, femstop, and an antegrade sheath were applied. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations: assess access site for bleeding and hematoma. section: entering cardiac output. use of the repositioning sheath and peel-away introducer: remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding is determined to be patient movement because the hematoma was noted after moving the patient.
11393 | 1220648-2025-26395 | 21479173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479173&pc=OZD | 02/03/2025 | 02/27/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600093 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp was removed due to left ventricle thrombus developing. the patient was noted as stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves: in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.
11394 | 1220648-2025-26396 | 21479292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479292&pc=OZD | 02/04/2025 | 03/06/2025 |  | 02/27/2025 |  |  | 02/05/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563741 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed with an impella 5.5 pump to support for acute myocardial infarction. while on support, the patient was transferred to a different medical center/tertiary center. upon transfer, the team observed a product damage. the sterile sleeve was fully detached. to date the pump has not been removed/replaced. the pump was successfully weaned and patient survived the event. | [5 empty] | The investigation of product damage (sterile sleeve damage) has been completed. no product was returned for analysis. the root cause of the product damage (sterile sleeve) was not determined as insufficient information regarding failure was provided and no product was returned. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11395 | 1220648-2025-26397 | 21479359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479359&pc=OZD | 02/05/2025 | 03/18/2025 |  | 02/27/2025 |  |  | 02/05/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571258 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support for acute myocardial infarction. during support, the patient had a large hematoma at the access site. manual pressure was held. the clinical representative confirmed that heparin was not given to the patient. however, it was unsure if it was due to the hematoma. the patient was eventually weaned and the patient survived the procedure. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11396 | 1220648-2025-26398 | 21479423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479423&pc=OZD | 02/03/2025 | 03/07/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1390016 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during training, the impella automated impella controller had a startup self-check error alarm. there was no patient involvement. | [5 empty] | The investigation for the automated impella controller (aic) self check error issue has been completed. the console logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the console was returned for evaluation and the reported failure mode was reproduced during testing. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. a.1 revised as none was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26398. d.2 revised common device name since the initial manufacturer device report number 1220648-2025-26398 was submitted in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26398 as the product was returned at that time. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11397 | 1220648-2025-26399 | 21479598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479598&pc=OZD | 02/06/2025 | 04/02/2025 |  | 02/27/2025 |  |  | 02/06/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603686 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had limb ischemia. a femoral-femoral bypass was recommended but denied due to the patient's critical condition. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26399.
11398 | 1220648-2025-26400 | 21479628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479628&pc=OZD | 02/03/2025 | 10/01/2025 |  | 02/27/2025 |  |  | 02/03/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597936 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The customer reported that the device exhibited an optical signal issue. | The user facility reported a patient with cardiomyopathy in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and bridge to transplant. approximately 14 days after start of the support, it was noticed the sterile sleeve was recently ripped. concern for bacteremia, the patient was prophylactically started on antibiotics. the patient was stable following the event, and impella mcs continued. patient status update 20 days later noted the patient received the heart transplant. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product damage: clinical details report that the sterile sleeve ripped, but it does not appear clear when/how the rip occurred or if excessive force was used. since neither product nor photos were returned for investigation, the root cause of the product damage could not be determined. optical signal issue: product was not returned for investigation. data logs confirmed the placement signal not reliable (psnr) alarm first triggered on 2/15. at this time, snr dropped to zero, contrast began oscillating between +/-3200, lamp increased, and gain and light dropped. the complication was intermittent for the rest of the case as rt logs showed all optical signal metrics suddenly normalizing for short periods before railing again. upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issues "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The investigation of product damage (sterile sleeve damage) has been completed. since neither product nor photos were returned for investigation, the root cause of the product damage could not be determined. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11399 | 1220648-2025-26401 | 21479649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21479649&pc=OZD | 02/06/2025 | 03/20/2025 |  | 02/27/2025 |  |  | 02/06/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025609817 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 pump was attempting to be delivered but the device was unable to advance through the graft and axillary artery. after multiple failed attempts, the implant was aborted and an impella cp device was placed for patient support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the delivery issue was most likely patient condition related, the impella 5.5 was unable to pass through anastomosis/graft. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11400 | 1220648-2025-26402 | 21480108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21480108&pc=OZD | 02/05/2025 | 03/21/2025 |  | 02/27/2025 |  |  | 02/05/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577754 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after impella cp was placed in a 78-year-old male, staff were unable to doppler pulses on the right lower extremity. a formal duplex study was ordered of right lower extremity (rle) showing no flow down the leg. at this time, the patient had poor prognosis and as a result, the family decided to make patient comfort care and withdraw supportive measures. the patient subsequently expired. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 220648-2025-26402. e.1 fax number reported on the initial manufacturer device report number 220648-2025-26402 was incorrect. | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11401 | 1220648-2025-26403 | 21480626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21480626&pc=PYX | 02/06/2025 | 03/19/2025 |  | 02/27/2025 |  |  | 02/06/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025560860 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 42 year old female patient with rp flex support ongoing for right heart failure. during support, there was hemolysis noted. there was a drop in hemoglobin, hemolyzing labs, and concerns for clot ingestion. the team troubleshot with blood product transfusion and pump explant after 28 hours. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the hemolysis has been completed. data was not return for review. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of hemolysis was most likely due to biomaterial ingestion based on the clinical description and returned product. the instructions for use for the rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller states the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact middle name should have been left blank. it was incorrectly stated as abiomed, inc.
11402 | 1220648-2025-26404 | 21480822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21480822&pc=OZD | 02/06/2025 | 03/24/2025 |  | 02/27/2025 |  |  | 02/06/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604717 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp pump for the patient admitted in cardiogenic shock when after one day of the support the signs of hemolysis prompted the team to remove and replace the pump which resolved the hemolysis. the patient remained stable. | After one day of support the impella cp was exchanged. two days after new impella was implanted the patient's condition deteriorated and the decision was made by the family to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the product and data were not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. b2 was revised from the initial manufacturer device report number 1220648-2025-26404 to reflect the patient outcome of death and date of death was added. b5 revised from the initial manufacturer device report number 1220648-2025-26404 to reflect the patient outcome. h1 revised from the initial manufacturer device report number 1220648-2025-26404 to reflect the patient outcome of death. h6 health effect - impact code revised from the initial manufacturer device report number 1220648-2025-26404 to reflect the patient outcome of death.
11403 | 1220648-2025-26405 | 21480926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21480926&pc=OZD | 02/06/2025 | 03/11/2025 |  | 02/27/2025 |  |  | 02/06/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621754 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient admitted in st-segment elevation myocardial infarction. the pump was placed but there was oozing at the site. the team gave four units of blood after a loss of over one unit. the patient was anticoagulated with heparin, integrilin, and brilinta. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was most likely patient movement since the patient kept moving prior to bleeding most likely leading to the bleed.
11404 | 1220648-2025-26406 | 21481103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21481103&pc=OZD | 02/06/2025 | 03/21/2025 |  | 02/27/2025 |  |  | 02/06/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522070 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump was inadvertently pulled back across the aortic valve during repositioning. the pump was removed, sterilized and then reinserted, but once again, the pump was unable to recross the valve. the decision was ultimately made to explant the pump. there was no patient harm. | [5 empty] | The investigation of positioning issues has been completed. no product was returned. reported initially placing the pump per instructions for use (ifu) but pulled back across aortic valve when attempting to reposition and unable to recross the valve. based on the clinical details, the root cause of the positioning issue was most likely use related since the pump was pulled across the aortic valve when being repositioned. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11405 | 1220648-2025-26407 | 21481417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21481417&pc=OZD | 02/02/2025 | 03/25/2025 |  | 02/27/2025 |  |  | 02/07/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563083 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Renal Failure | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed's long-term outcome and quality indicator (loqi) impella registry that a patient experienced multiple complications during impella cp support. during the course of support, the patient suffered from renal failure, thrombocytopenia, hemolysis, and frequent premature ventricular complexes. each of the conditions were believed to be possibly related to the use of the impella. the patient was escalated to impella 5.5 support. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of renal failure, thrombocytopenia, hemolysis, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11406 | 1220648-2025-26408 | 21483328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21483328&pc=OZD | 01/31/2025 | 03/19/2025 |  | 02/27/2025 |  |  | 01/31/2025 | 02/27/2025 |  |  | 06/30/2025 |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367319 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a bedside echocardiogram showed an implanted impella 5.5 pump was pointing towards the mitral valve. attempts to reposition the pump were unsuccessful and the device was inadvertently pulled back across the valve. attempts to recross the valve were unsuccessful and the pump was subsequently replaced. there was no resulting patient harm. | [5 empty] | The investigation of positioning issues has been completed. the echocardiogram showed pump pointing towards the mitral valve (mv), the pump pulled out across the valve and unable to reposition. no product returned. the cause of the positioning issue was most likely use related as it was pulled across the valve during repositioning and was unable to be resolved leading to pump explant. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26408 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11407 | 1220648-2025-26409 | 21483480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21483480&pc=OZD | 02/05/2025 | 03/13/2025 |  | 02/27/2025 |  |  | 02/05/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607181 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 74 year old male patient that had an impella 5.5 that was not able to be placed due to the patient¿s anatomy and calcified vessels. after several unsuccessful attempts the physician decided to place the impella cp. the patient survived the procedure. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy and calcified vessels and was aborted. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact facility name and reporting contact fax number as they were inadvertently omitted on the initial manufacturer device report number h.10 added some instructions for use that were inadavertently omitted from the initial manufacturer device report number. | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
11408 | 1220648-2025-26410 | 21483493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21483493&pc=OZD | 02/05/2025 | 03/17/2025 |  | 02/27/2025 |  |  | 02/05/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593404 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a 43-year-old male patient had a impella cp for support during high-risk surgery. it was reported that after impella cp started, limb ischemia was noted. external fem-fem bypass was done. the patient was successfully weaned and explanted. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26410 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26410. | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11409 | 1220648-2025-26411 | 21483497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21483497&pc=OZD | 02/06/2025 | 03/18/2025 |  | 02/27/2025 |  |  | 02/06/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601907 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that axillary implant attempted for the impella 5.5 was unsuccessful after numerous attempts. due to difficult anatomical vasculature, catheter shaft was damaged. a new impella 5.5 was used and there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the pump was returned for inspection. upon investigation kink in the catheter was noted. the root cause of delivery issue is determined to be use issue because second device implanted without complication. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26411 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11410 | 1220648-2025-26412 | 21483501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21483501&pc=OZD | 02/06/2025 | 03/18/2025 |  | 02/27/2025 |  |  | 02/06/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600095 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed ischemia during impella cp support. the patient's foot was cold and no pulse was present in the extremity. levo was weaned in response to the ischemia and the condition resolved. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation. | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26412 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11411 | 1220648-2025-26413 | 21483508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21483508&pc=OZD | 02/07/2025 | 03/26/2025 |  | 02/27/2025 |  |  | 02/07/2025 | 02/27/2025 |  |  |  |  | 02/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527806 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, blood was seen leaking from the junction between the blue hub and white port. dermabond was applied to the area to stop the bleed, but a hematoma was noted at the groin site with additional bleeding seen from the access site. manual pressure was applied, but the bleeding persisted. the decision was ultimately made to explant the pump to manage the condition. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the reported blood leaking at the junction of blue butterfly and white port site and pump was removed. evidence of blood found within gap between butterfly and white anchor housing. no other damages or abnormalities. the root cause of the access site bleeding - major was most likely use related since blood was found between the butterfly and white anchor housing indicating the blue butterfly was inserted into the vasculature. d.6b revised explant date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26413. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26413 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.3 if the device was not evaluated should of been reflected on the initial manufacturer device report number 1220648-2025-26413 as device evaluation anticipated, but not yet begun. h.6 code 1884 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26413. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11412 | 1220648-2025-26335 | 21468316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468316&pc=OZD | 02/01/2025 | 03/11/2025 |  | 02/26/2025 |  |  | 02/01/2025 | 02/26/2025 |  |  | 11/30/2025 |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422942 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was concern of new left ventricle thrombus, so bival was started. the patient refused to allow impella time for heart recovery and requested for support to be discontinued. care was withdrawn and subsequently the patient expired. no additional information has been provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2025-26335 in accordance with updated procedures.
11413 | 1220648-2025-26336 | 21468395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468395&pc=OZD | 02/02/2025 | 03/19/2025 |  | 02/26/2025 |  |  | 02/02/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585395 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was initially implanted with an impella cp device for mechanical circulatory support (mcs) for four days before being escalated to an impella 5.5 device for bridge to recovery, left ventricular assist device or transplant. well into the support, approximately 30 days, the pump had high purge pressure. tissue plasma activator was added to the purge and support continued. however, two days later the device was replaced as a result of the event. the patient was noted to be in stable condition following the event, continuing with impella mcs. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: using the automated impella controller with the impella catheter, ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. pump was not returned for investigation. data logs were investigated and found purge pressure high (hpp) followed by purge system blocked alarms were observed. the purge pressure was seen to be rising from previous 38 hours to the point of hpp. saturated flows were also observed at the same p-level. no motor current spikes were observed. prior to the rise in purge pressure, several purge cassette changes exceeding the 2 minute recommendation were observed. as per the clinical information, tissue plasminogen activator (tpa) was started in the purge solution which resolved the high purge pressure alarms after 31 hours. these characteristics are similar to that of biomaterial build up which was resolved when tpa was added to the purge. the root cause of the high purge pressure issue was most likely use issue due to long cassette changes leading up to the gradual rise in purge pressure. added-h6 impact code 4644.
11414 | 1220648-2025-26337 | 21468435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468435&pc=OZD | 02/04/2025 | 03/18/2025 |  | 02/26/2025 |  |  | 02/04/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621611 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon removal of the impella cp pump, the patient experienced persistent oozing/bleeding from their access site after the perclose's failed. after forty minutes of manual pressure, vascular intervention was performed to repair the artery and achieve hemostasis. the patient was stable following the intervention. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11415 | 1220648-2025-26338 | 21468455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468455&pc=OZD | 02/02/2025 | 03/14/2025 |  | 02/26/2025 |  |  | 02/02/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550681 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced oozing and a hematoma at the access site. manual pressure was held and the impella cp was explanted and escalated to an impella 5.5. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a/d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26338. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11416 | 1220648-2025-26339 | 21468520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468520&pc=OZD | 02/04/2025 | 03/17/2025 |  | 02/26/2025 |  |  | 02/04/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563821 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, blood was noted to weeping from the hemashield gold graft. large albumin was given and one unit of packed red blood cells were given due to blood loss. the doctor assessed the subclavian pocket and noted that the graft was weeping more than average. the decision was made to explant the impella device in attempts to prevent a bleeding complication. the procedural outcome was the patient survived the surgery. | [5 empty] | The root cause of the access site bleeding issue was most likely use related as hemashield gold graft and also non abiomed products, plus st 0.018 in guidewire was used and 8 x 25 cm pinnacle sheath were used. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11417 | 1220648-2025-26340 | 21468558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468558&pc=OZD | 02/02/2025 | 04/03/2025 |  | 02/26/2025 |  |  | 02/02/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601180 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device experienced access site bleeding, hemolysis and subsequently expired. the following day after start of support, bleeding at the access site was noted and low-pressure bandage was applied. it does not appear heparin was not used in iv drips around the time of the bleeding event, and it is, however, unknown if the heparin was withheld due to the bleeding. additionally, it was noted the patient developed hemolysis, and the performance level was reduced to p-4. notedly, the pump was running without alarms and the position was checked. however, the patient's condition was listed as very unwell. the following day, characterization of the patient's condition remained the same, and it was noted the patient entered liver and kidney failure. the patient would expire this day; care was withdrawn. the cause of death is unknown, and the impella cp device association to the event outcome (patient's demise) is unclear. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11418 | 1220648-2025-26342 | 21468692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468692&pc=OZD | 02/03/2025 | 03/20/2025 |  | 02/26/2025 |  |  | 02/03/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573825 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhagic Stroke | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hemorrhagic stroke coinciding with hemolysis during impella cp support. plans were made to explant the pump and the patient was bridged to bi-ventricular assist devices. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported stroke and hemolysis has been completed. the product was not returned. the cause of the hemolysis issue could not be determined since the complaint device not returned for analysis and insufficient clinical data was provided. data log was not available to review. as this clinical information was not provided, the true root cause of the stroke could not be determined.
11419 | 1220648-2025-26343 | 21468759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468759&pc=OZD | 02/03/2025 | 03/25/2025 |  | 02/26/2025 |  |  | 02/03/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561844 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp setup, the automated impella controller (aic) presented a "purge disc not recognized" error. the purge cassette was changed without success. the backup console presented with the same error. therefore, implant not possible. | [5 empty] | The investigation of purge cassette failure has been completed. the returned purge cassette was accidentally dropped from return bag by investigator. it is unknown if damage occurred before or after purge cassette was returned. the purge cassette was inserted into the automated impella controller (aic) and the purge cassette defective alarm reproduced on multiple consoles. the logs confirm purge cassette failure alarms. insufficient logs were returned. the root cause of the priming issue was determined to be due to rfid programming issues as issue reproduced on returned product. | The purge cassette was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11420 | 1220648-2025-26344 | 21468864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468864&pc=OZD | 02/03/2025 | 03/24/2025 |  | 02/26/2025 |  |  | 02/03/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589933 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fracture | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that placement signal unreliable alarm occurred. upon inspection of catheter, it was noted to have a fracture at the red plug. catheter was stabilized as to not have any further damage. the fractured 5.5 continued to function, the patient was supported and hemodynamically stable. the physician spoke extensively to the patient about exchanging the impella, but the patient declined the plan to exchange impella. the patient was bridged to a left ventricular assist device placement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (hole in catheter) has been completed. pump was not returned for investigation. without the pump the failure cannot be further investigated. therefore, the root cause of the product damage (hole in catheter) issue was not determined since product was not returned. data logs were investigated. at the time of the placement signal not reliable (psnr) alarm, there snr decreased to 0, contrast decreased, gain and lamp increased and light was stable. the placement signal never came back and the pump was explanted after a few days. clinical information provided also mentions that there was a fracture noted on the catheter near the red handle. these symptoms are similar to that of a damaged fiber line. therefore, the root cause of the optical signal issue was most likely a damaged optical fiber line caused due to excessive force.
11421 | 1220648-2025-26345 | 21468907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468907&pc=OZD | 02/02/2025 | 04/02/2025 |  | 02/26/2025 |  |  | 02/02/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584354 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock and hemodynamic decompensation was implanted with an impella cp device for mechanical circulatory support and experienced high purge pressure. the patient was intubated, vasoactive drugs were administered, and diagnostic catheterization and angioplasty of the left main coronary artery was performed. an intra-aortic balloon pump (iabp) was placed after revascularization and due to poor evolution, it was decided to escalate to an impella cp device which was successfully implanted. however, during transfer from the catheterization lab to the intensive care unit, the purge pressures started to increase, and the purge flow was reduced. at the same time, suction alarms started with reduced peak systolic ventricular curve, raising suspicion of malposition or thrombus. position was checked by echocardiogram. the purge pressure was about 1000 mmhg with low purge flow. the possibility of thrombolytic treatment was considered but given that the patient was unstable, and the thrombotic suspicion was not so strong, it was considered a better and quicker option to replace the impella cp device, which was completed without incident. material suggesting the presence of tissue debris was removed from the explanted impella inlet. the patient was reported to be in critical condition following this event. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the high purge pressure and low pump flow (as evidenced by suction alarms) has been completed. data logs were not returned for investigation. returned product was inspected, and after soaking the pump, there was what appeared to be biomaterial inside the cannula near the impeller. additionally, there appeared to be potential signs of biomaterial below the impeller in the purge gap. however, when the pump and purge cassette were primed and run in a bucket for several hours, high purge pressure did not reproduce and pump flows were within specification. clinical details report that tissue plasminogen activator (tpa) was not run in the purge, but that material was removed from the pump after explant. since data logs were not available and the returned product did not reproduce the issue, the root cause of the high purge pressure could not be determined. clinical details stated that suction alarms occurred. however, when the pump and purge cassette were primed and run in a bucket for several hours, low flows did not reproduce. clinical details report that position was confirmed with echo. additionally, material was removed from the pump after explant. since data logs were not available and the returned product did not reproduce the issue, the root cause of the low pump flow could not be determined. added h6 impact code 4644 added, h6 med dev prob code 1248.
11422 | 1220648-2025-26346 | 21468986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21468986&pc=PYX | 02/02/2025 | 07/09/2025 |  | 02/26/2025 |  |  | 02/02/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529027 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock following coronary artery bypass graft surgery one day prior was implanted with an impella rp flex for mechanical circulatory support. two days after start of the support, a high purge pressure alarm was triggered. the purge cassette was changed, and pressure remained greater than 1000 mmhg. the following day, it was noted that tissue plasminogen activator added to the purge system resolved the high purge pressure. however, this day the pump stopped. the team attempted to reduce the performance level to p-2 for turndown echocardiogram before possible explant. however, when the performance level was reduced and the pump stopped. an attempt was made to restart without success. the automated impella controller was turned off, then back on and performance level increased, which was successful. however, motor current spikes occurred below p-5. the decision was made to explant the rp flex device. the patient remained stable throughout the event and survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of high purge pressure and pump stop has been completed. the pump was returned for investigation. no evidence of physical damage was found on the returned product. no biomaterial was observed on the impeller or inflow cage. the pump passed all electrical testing and was able to run in a bucket of water with saturated pump flow and no purge-related abnormalities. the data log indicates a gradual rise in purge pressure, followed by several motor current spike alarms. according to clinical observations, the purge pressure issue was resolved after the infusion of tissue plasminogen activator (tpa). additionally, on the third day, reports indicated that the pump stopped functioning, accompanied by several motor current spikes. the pump was explanted after this point. the cause of the high purge pressure issue was most likely biomaterial, based on data on data log review and provided clinical details. the cause of the pump stop issue was most likely biomaterial, based on data log review and returned product investigation. d9 device returned to manufacturer date added.
11423 | 1220648-2025-26347 | 21469067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469067&pc=OZD | 01/31/2025 | 03/07/2025 |  | 02/26/2025 |  |  | 02/01/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella cp pump placed for support during cardiogenic shock. the pump was placed after percutaneous coronary intervention. the pump was placed but there was an access site hematoma. the hematoma was treated by manual hold and with four units of packed red blood cells. the patient was eventually successfully weaned and survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11424 | 1220648-2025-26348 | 21469147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469147&pc=OZD | 01/30/2025 | 03/10/2025 |  | 02/26/2025 |  |  | 01/30/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025596387 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and it was observed that there was no purge fluid exiting. the line was de-aired and reconnected. however, the issue remained. the decision was made to explant and replace the impella 5.5 device, which was successfully completed. there was patient harm as a result of the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of priming issue has been completed. insufficient data logs were provided. the logs from when the pump and purge cassette were inserted were not available the data logs showed that purge failing to prime issue was resolved and elevated purge pressures were trending down. the root cause of the priming issue was not determined as no product was returned for analysis and insufficient data logs were provided. e.1 revised us state as it was previously reported incorrectly on the initial manufacturer device report number. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 codes 4627 and 4629 were reported on initial manufacturer device report number and were incorrect.
11425 | 1220648-2025-26349 | 21469163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469163&pc=OZD | 01/31/2025 | 03/18/2025 |  | 02/26/2025 |  |  | 01/31/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594955 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Separation | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Swelling/ Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the repositioning sleeve has separated from repositioning unit, exposing the catheter to a non sterile environment. the sleeve was reattached with tegaderm. | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) as a bridge to transplant and three days after start of the support the patient experienced a hematoma. there was a small amount of swelling to the right shoulder. heparin was held and the patient was monitored for improvement before restarting. it was identified that physical therapy was having the patient use the right arm with a resistance band. following, several days later, a hematoma was again noticed, and heparin was held once again. the patient continues on impella mcs. |  |  |  |  | The investigation of access site bleeding and product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 revised to reflect both adverse event and product problem, based on the additional information received. b5 revised to reflect additional information received. h6 medical device problem code revised based on the additional information received. | The impella device was not received from the customer as the device remains implanted, supporting the patient, and therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11426 | 1220648-2025-26350 | 21469168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469168&pc=OZD | 02/03/2025 | 03/20/2025 |  | 02/26/2025 |  |  | 02/03/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564011 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella device used: ¿(b)(6)2025: hemolytic indices elevated today with pfh 160 (<30) and ldh 351 (293), not on diuretics (hemolysis) (b)(6)2025: hemolytic indices elevated today with pfh 60 (<30) and ldh 310 (233); of note, pt has been receiving lasix 40mg iv bid; remains on p4 (hemolysis) (b)(6)2025: 2d echo with impella in good position (b)(6)2025: pfh 50; hold diuretics (hemolysis) (b)(6)2025: pfh 50; give 500ml ivf (hemolysis) (b)(6)2025 bp low this am, map: 66-76, milrinone 0.375, bival, pfh: 40, ldh: 226, +5515 ml in last 24 hrs., wt. 79.3 kg (b)(6)2025 ldh 192 (b)(6)2025 oht¿. the patient recovered with no sequelae. | The complainant reported the patient was on impella 5.5 support and during support, the pump could not flow above p 4. the pump appeared to be in proximity to septal wall. neither the low pump or positioning issues were resolved. there was no patient harm and the the patient remains on support. |  |  |  |  | B1 revised to reflect both adverse event and product problem based on the additional information received from the clinical site. b2 revised to reflect the additional information received from the clinical site. b5 revised to reflect the additional information received from the clinical site. d6b explant date added h1 type of reportable event revised to reflect serious injury based on the additional information received from the clinical site. h6 revised to reflect the additional information received from the clinical site. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26350. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11427 | 1220648-2025-26351 | 21469223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469223&pc=OZD | 01/31/2025 | 04/02/2025 |  | 02/26/2025 |  |  | 01/31/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611057 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella 5.5 support and during support, the patient had hemolysis. the impella flow was decreased and the pump was repositioned. the patient was also given four units of packed red blood cells. the patient was eventually weaned and survived the procedure. | Additionally, it was reported that the locking mechanism of the 5.5 became more and more difficult to use. more force was needed to reposition, making it difficult to do this as carefully as they wanted to. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | B1 revised to reflect both adverse event and product problem based on the additional information received from the clinical site b5 revised to include the additional information received from the clinical site h6 revised to include the additional information received from the clinical site
11428 | 1220648-2025-26352 | 21469236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469236&pc=OZD | 01/31/2025 | 03/12/2025 |  | 02/26/2025 |  |  | 01/31/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603504 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 44 year old female patient was on impella 5.5 support and during support, the patient had a hematoma at the access site. sandbag was applied and support continued. additionally, the patient had ventricular tachycardia (vt) while on support. this occurred during repositioning. tte showed the pump in proper position. support continued. the patient was eventually weaned and survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of vt/vf has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the vt/vf was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26352 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 mdr reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26352. g1 mdr reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26352. h6 code 1888 was reported incorrectly on manufacturer device report 1220648-2025-26352.
11429 | 1220648-2025-26353 | 21469243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469243&pc=OZD | 02/01/2025 | 03/18/2025 |  | 02/26/2025 |  |  | 02/01/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603904 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and upon start up, there were low pump flows with suction. there were then high flows and high motor current. there was a suspected clot ingestion so the pump was explanted. at explant, there was visible thrombus in outlet, impeller and cannula. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the complaint device not available for analysis. data logs were reviewed and suction occurred. no positioning issue was seen. drop in flow was observed as soon as the pump was started, with possible biomaterial ingestion at starting support. additionally, clinical details confirmed biomaterial observed. the cause of the low pump flow most likely was biomaterial ingestion. as the necessary information was not provided, the root cause of the thrombus was not determined. e.1 added us phone number and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26353. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26353 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name and reporting contact fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26353. h.6 code 4629 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26353 and a new code was added. | The impella device will not be returned from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
11430 | 1220648-2025-26354 | 21469252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469252&pc=PYX | 02/04/2025 | 03/14/2025 |  | 02/26/2025 |  |  | 02/04/2025 | 02/26/2025 |  |  | 07/31/2025 |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024380048 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella rp fail to deliver. the team could not pass the rp fully and so it was replaced by a rp flex. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella unable to pass through the vasculature as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
11431 | 1220648-2025-26355 | 21469388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469388&pc=PYX | 02/01/2025 | 03/20/2025 |  | 02/26/2025 |  |  | 02/01/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025543761 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  |  |  | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced signs of limb ischemia and hemolysis for 78 hours before being explanted. the patient continued with extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ | The investigation of hemolysis and limb ischemia has been completed. clinical states placement signal (ps) trending up and flows trending down. hemolysis also reported. unknown if blood products given. repositioning did not resolve issue. data logs showed as reported, on 2/1 the ps was trending up, flows and motor current (mc) trending down but no mc spikes outside of p level changes. no other abnormalities or relevant alarms. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the failure mode low pump was added and the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26355 as it is unknown if the initial reporter also sent the report to the food and drug administration. h4 device manufacture date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26355.
11432 | 1220648-2025-26356 | 21469405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469405&pc=OZD | 02/02/2025 | 03/10/2025 |  | 02/26/2025 |  |  | 02/02/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487213 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and soon after the patient complained of numbness in the right lower extremity. a runoff was performed and revealed a calcium shelf that was obstructing flow distally. the impella was subsequently explanted to resolve the ischemia. the patient had tolerated the wean without any compromise to hemodynamics. the right femoral artery post closed with perclose x1 with great hemostasis. the patient was transferred to the unit with stable hemodynamics. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation). | The investigation was completed. the device was not returned. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26356 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11433 | 1220648-2025-26357 | 21469415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469415&pc=OZD | 02/02/2025 | 03/19/2025 |  | 02/26/2025 |  |  | 02/02/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569227 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support high purge pressure was noted. the pc was replaced, but the high purge pressure remained. tissue plasminogen activator and pump replacement were discussed. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. based on data logs and returned product, the root cause of the high purge pressure was determined to most likely be biomaterial build up. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26357 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26357.
11434 | 1220648-2025-26358 | 21469420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469420&pc=OZD | 02/02/2025 | 03/19/2025 |  | 02/26/2025 |  |  | 02/02/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596328 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was brought to the catheterization lab and found to be hemodynamically unstable. the patient was intubated and impella cp was implanted for mechanical circulatory support (mcs), and radial access was used for percutaneous coronary intervention to the left anterior descending artery. after removing the peel away sheath and advancing reposition sheath, a hematoma began to form. pressure was applied, and patient transferred to unit on impella mcs. once on the unit, bleeding complications continued. a vascular surgeon was consulted as the bleeding in groin was uncontrolled. consequently, the impella cp device was removed with femoral artery repair. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26358 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11435 | 1220648-2025-26359 | 21469424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469424&pc=OZD | 02/03/2025 | 03/31/2025 |  | 02/26/2025 |  |  | 02/03/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112226 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The automated impella controller (aic) presented ¿purge system has stopped. switch to back up controller¿ alarm. the aic was replaced successfully and the procedure was then completed. no patient was involved. | [5 empty] | The investigation for the console (aic) purge issue has been completed. data logs were reviewed which showed immediately after bootup, the console displayed the message: ¿controller failure (purge pressure sensor defective) ¿ alarm,¿ event #418. after the console was shut down and restarted manually, it again displayed event #418, this confirms the reported failure mode. this console was returned for evaluation and the controller failure alarm could not be reproduced. the purge pressure accuracy did not meet the criteria per the pm procedure (0042-7309). at an applied external pressure of 100 mmhg, the console measured 87 mmhg, which is below the minimum required of 90 mmhg per the criteria. the original purge pressure sensor output voltage at line 7 of connector j3.4 was 0.40v (without the disc inserted), while the output of the known good purge pressure sensor on the test fixture was 0.49v (without the disc inserted). upon replacing the original purge pressure sensor with a known good sensor on (b)(6), the sensor output without the disc inserted was 0.49v. the root cause for the controller failure was most likely defective purge pressure sensor. corrections to the initial mfr report: d2 corrected common device name | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11436 | 1220648-2025-26360 | 21469429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469429&pc=OZD | 02/03/2025 | 03/11/2025 |  | 02/26/2025 |  |  | 02/03/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363394 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) presented a system self-check failure error. this issue occurred during use in a cath lab. no further details were provided. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the data analysis showed that the console logs had a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring communication channel. the root cause of the "system self check failure" was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. d2 common device name revised on manufacturer device report 1220648-2025-26360 in accordance with updated procedures. e1 initial reporter prefix and name revised on manufacturer device report 1220648-2025-26360 to reflect the biomedical engineer name. e1 initial reporter phone number revised as it was not included on the initial manufacturer device report 1220648-2025-26360. e2 revised to no; on manufacturer device report 1220648-2025-26360. e3 revised to biomedical engineer on manufacturer device report 1220648-2025-26360. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26360 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11437 | 1220648-2025-26361 | 21469434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469434&pc=OZD | 01/31/2025 | 03/17/2025 |  | 02/26/2025 |  |  | 01/31/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603434 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and developed bleeding at the impella access site. heparin was held and one unit of red blood cells was given. after much discussion with the family and the patient, a decision was made to withdraw care with family at his side. medical support was turned off and the patient passed peacefully with family at his side. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be anticoagulation management because the heparin was supratherapeutic and heparin was kept on hold for 1 hour to control the bleeding. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
11438 | 1220648-2025-26362 | 21469436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469436&pc=OZD | 02/04/2025 | 04/03/2025 |  | 02/26/2025 |  |  | 02/04/2025 | 02/26/2025 |  |  | 01/31/2026 |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468867 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a 5.5 pump to support a patient admitted in with cardiogenic shock and cardiomyopathy. the pump was supporting for just under 7 days when the pump stopped. the stop had occurred after the perfusion team member had tried to release the twist in the cord and had unplugged it. all attempts to restart the pump after were unsuccessful and due to patient condition, the team made choice to replace the 5.5 pump. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of pump stop has been completed. the failure mode (fm) "pump stop" was changed to "pump not detected," since it was reported that the pump ran fine and it was unplugged to untwist the cable and when it was plugged back in "impella defective" occurred. the data logs show the pump ran fine up until it was disconnected on 02/04. the logs show after the pump was unplugged and then plugged back in, "impella defective" alarms occurred. however, due to the lack of the product retuned, the true root cause of the pump not being detected by the console, cannot be determined. | The device was not received for evaluation. should additional information or the device be received, a supplemental mdr will be filed. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11439 | 1220648-2025-26363 | 21469534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469534&pc=OZD | 01/31/2025 | 03/11/2025 |  | 02/26/2025 |  |  | 01/31/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560166 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed with an impella cp pump for support during non st elevated myocardial infarction. the pump was supporting but the limb accessed by the impella cp lost pulses. the patient was thought to be subtherapeutic and the plan was made for heparin and pump explant. the patient was successfully weaned and survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11440 | 1220648-2025-26364 | 21469642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469642&pc=OZD | 01/31/2025 | 03/19/2025 |  | 02/26/2025 |  |  | 01/31/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601876 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed on impella cp support with acute myocardial infarction. it was reported that after the sheath exchange, the access site continued to bleed. the staff tried side closure and the perclose failed twice. the impella cp was then removed. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11441 | 1220648-2025-26365 | 21469778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469778&pc=OZD | 02/05/2025 | 03/20/2025 |  | 02/26/2025 |  |  | 02/05/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571314 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a purge system open alarm appeared during impella cp support. upon visual inspection of the setup, it was discovered that the purge line broke between the purge filter and purge reservoir. the nurse believed the patient may have moved in their sleep which possibly caused the purge filter/reservoir to crack. the staff was informed that a purge bypass was required to resolve the condition and plans were made to escalate to an impella 5.5 support. there was no resulting patient harm. | [5 empty] | The investigation for the purge leak-pump has been completed. the product was not returned for investigation. data logs were reviewed and the purge system open alarms were confirmed the morning of 5-feb-2025. during this time, the purge pressure dropped to zero for roughly 30 minutes. purge related metrics returned to normal after this period, indicating the purge filter bypass was done. clinical details reported that the damage was done to the sidearm between the reservoir and filter. additionally, it was suspected that the patient caused the damage when rolling over in their sleep. since the product was not returned, the nature of the damage to the sidearm could not be determined. however, based on clinical details and data log review, the root cause of the purge leak was determined to most likely be a damaged sidearm. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
11442 | 1220648-2025-26366 | 21469825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469825&pc=OZD | 02/05/2025 | 03/21/2025 |  | 02/26/2025 |  |  | 02/05/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593410 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 had a purge system blocked alarm while on patient support. the impella was exchanged with a new 5.5. there was no harm to the patient. | [5 empty] | The investigation for the high purge pressure has been completed. pump was not returned for investigation. data logs were returned and it was observed that after 92 hours, there was a gradual rise in purge pressure observed with a small motor current spike. the pump had "pump in aorta" alarms after the motor current spike. no suction alarms were present. the purge pressure was seen to be rising for the next 14 hours leading to high purge pressure alarms and eventually purge system blocked alarms. the flows were saturated around that time with no change in p-level. these characteristics are similar to that of gradual biomaterial buildup. therefore, the root cause of the high purge pressure issue was most likely biomaterial buildup. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11443 | 1220648-2025-26367 | 21469875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21469875&pc=OZD | 02/04/2025 | 03/20/2025 |  | 02/26/2025 |  |  | 02/04/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594170 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. despite dilating the vessel multiple times, the inserted impella 5.5 pump was unable to advance through the axillary artery and into the aorta due to the patient's vessel size. multiple attempts were made, but the decision was ultimately made to abort the implant and place an impella cp pump for mechanical circulatory support. it was noted that the repeat attempts resulted in blood loss prompting packed red blood cells to be administered. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of access site bleeding and delivery issues has been completed. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature despite multiple attempts. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
11444 | 1220648-2025-26368 | 21471656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21471656&pc=OZD | 02/05/2025 | 03/23/2025 |  | 02/26/2025 |  |  | 02/05/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025609736 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Insufficient Information | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the left ventricle and aorta waveforms were separated. the staff were unable to go higher than p4 without suction events and the device was repositioned several times with fluid bolus¿s given. the family decided to put patient on comfort care and the patient expired shortly after. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of low pump flow has been completed. the product and data were not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. d6a and d6b were inadvertently entered incorrectly on the initial manufacturer device report 1220648-2025-26368.
11445 | 1220648-2025-26369 | 21471876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21471876&pc=OZD | 02/06/2025 | 03/18/2025 |  | 02/26/2025 |  |  | 02/06/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573874 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and the patient had a small small porcelain aorta which made pushing the impella cp into left ventricle impossible. the implant was aborted and provided proceed with percutaneous coronary intervention. shortly after the percutaneous coronary intervention, the patient went into a sustained ventricular tachycardia storm, care was withdrawn later and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, patient had bilateral tortuous aneurysmal external illiacs. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11446 | 1220648-2025-26370 | 21474344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21474344&pc=OZD | 02/05/2025 | 04/07/2025 |  | 02/26/2025 |  |  | 02/05/2025 | 02/26/2025 |  |  |  |  | 02/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579954 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support placed for the patient admitted in with in acute myocardial infarction and cardiogenic shock. the access site was bleeding and the team gave one unit of blood, adjusted the angle and a new dressing was placed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the pump was not returned for review (the cassette was returned). the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26370 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11447 | 1220648-2025-26298 | 21457625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21457625&pc=OZD | 01/30/2025 | 03/19/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025602047 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the impella cp was advanced, but due to tortuosity, the sheath kinked. the sheath was replaced with the short sheath and the impella was placed without issue. no patient harm was reported. | [5 empty] | The investigation for the introducer damage has been completed. no product was returned for evaluation. clinical details noted that long impella introducer sheath kinked due to patient tortuosity and was exchanged for a short impella introducer sheath that allowed the pump to be inserted without issue. the root cause of the introducer damage - mechanical was most likely due to a sheath kink as a result of the patient's tortuous anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11448 | 1220648-2025-26299 | 21457656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21457656&pc=OZD | 01/30/2025 | 04/23/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600023 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. the pump failed to be delivered and the procedure was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be most likely patient condition because the insertion was aborted due to patient anatomy.
11449 | 1220648-2025-26300 | 21457702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21457702&pc=OZD | 01/30/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561481 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced ventricular fibrillation during support. the patient required cardioversion to return to normal sinus rhythm. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11450 | 1220648-2025-26301 | 21457796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21457796&pc=OZD | 03/17/2024 | 02/25/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  | 09/30/2025 |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411277 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced ischemia during impella cp support. the pump was subsequently removed due to the right leg ischemia, but the condition persisted. an endarterectomy of the right leg was completed to restore perfusion. the patient recovered with no sequelae. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11451 | 1220648-2025-26302 | 21457880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21457880&pc=OZD | 09/01/2024 | 03/14/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524134 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced hemolysis during impella cp support. the hemolysis was troubleshoot by p level adjustment and position assessment. | [5 empty] | The investigation for the hemolysis has been completed. complaint device not returned for analysis. no data logs available to review. no positioning issue stated in the clinical data. only stated with decreasing the p-level to p-3 hemolysis got better. the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. added e4 should have been no. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11452 | 1220648-2025-26303 | 21457908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21457908&pc=OZD | 01/31/2025 | 10/31/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1387424 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. placement signal not reliable alarm occurred. troubleshooting was completed, but the issue persisted. the staff were advised to replace the console for ongoing support which the doctor will do with perfusionist. the patient outcome was noted to be stable. | [5 empty] | Data analysis (based on rlfs logs): during support, the console displayed ¿placement signal not reliable (psnr) (opm invalid signal, optical pump connected) during the psnr alarm, the contrast was oscillating between -3276.5 and +3276.7. this confirms the reported failure mode. device analysis: not applicable, as the console was not returned for investigation. root cause: the root cause for the psnr alarm was not determined, as the console was not returned for investigation. the root cause of the controller error was due to an aic software issue. this was entered into jama with defect/ bug. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis (based on rlfs logs): ¿ the case associated with the reported complaint was started on (b)(6) 2025, with the pump sn: (b)(6). ¿ during support, on the complaint date of (b)(6) 2025, the console displayed ¿placement signal not reliable (opm invalid signal, optical pump connected) ¿ alarm,¿ event #1036. ¿ during the psnr alarm, the contrast was oscillating between -3276.5 and +3276.7. this confirms the reported failure mode. ¿ the psnr alarm was resolved automatically during support on (b)(6) 2025. ¿ additionally, there was a ¿controller error (opm comm stopped) [optical pump connected] ¿ alarm,¿ event #1043, which was resolved automatically in 18 seconds. device analysis: this console was tested after arriving at fs. the average power at the optical bench connector p4 was 4.8 v, as intended. although the contrast, gain and snr were within specifications, the lamp power was out of spec, measuring only 2.55 w, which is supposed to be greater than 2.7 w according to the pm procedure (b)(4). root cause: ¿ the root cause of the psnr alarm was a lamp malfunction. ¿ the root cause of the controller error and loss of placement signal was due to an aic software issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 updated. h3 updated. h6 updated. d10 concomitant medical products and therapy dates updated.
11453 | 1220648-2025-26305 | 21457986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21457986&pc=OZD | 02/01/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/01/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607038 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and developed hemolysis. the patient presented to the emergency department with chest pain and st-segment elevation myocardial infarction was ruled in. the patient underwent successful percutaneous transluminal coronary angioplasty and stenting to the left anterior descending artery. post intervention, the patient experienced ventricular tachycardia arrest requiring defibrillation and intubation. an impella cp device was placed for hemodynamic support. later in the day, after start of the support, the patient was noted to have hemolysis/hematuria. the patient's labs were hemolyzing and urine was bloody. consequently, the patient was successfully weaned and the impella cp device was explanted the following day in the cardiac catheterization lab. post explant, the patient did well. the patient did not require drips and no further escalation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11454 | 1220648-2025-26304 | 21457987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21457987&pc=OZD | 01/31/2025 | 05/07/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574106 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the team was unable to deliver the 14fr sheath and the team observed peripheral artery disease. the 14fr sheath was removed and replaced with another non-peel away sheath. no patient harm was reported. | It was reported that the patient expired after 43 hours of support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | B1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26304. b2 outcome death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26304. b5 patient outcome was inadvertently omitted from the narrative on the initial manufacturer device report number 1220648-2025-26304. h1 type of event was erroneously selected on the initial manufacturer device report number 1220648-2025-26304. h6 was erroneously selected on the initial manufacturer device report number 1220648-2025-26304. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition since the team was unable to deliver the 14fr sheath and the team observed peripheral artery disease (pad). the team removed the provided 14fr sheath and replaced it with another non-peel away sheath. a2 patient age was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26304. a3 patient gender was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2025-26304. a5 patient race and ethnicity was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26304. d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26304. e1 initial reported first and last name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26304. g1 reporting contact address was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26304. g1 manufacturing site name and address was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2025-26304.
11455 | 1220648-2025-26306 | 21458012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21458012&pc=OZD | 01/31/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575746 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the insertion of an impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful due to the anatomy. the device was unable to cross the aortic valve after multiple attempts, and the implant case was aborted. there was no report of adverse effects to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11456 | 1220648-2025-26307 | 21458089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21458089&pc=OZD | 01/31/2025 | 03/13/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the us complainant had a 7fr companion sheath being placed at the single access but, there was an inability to pass wires through the 7fr companion sheath. there was no noted harm or injury. the vessel was noted to be small. no additional information was provided. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d1, d2 brand name, product code, common device are unknown. d4- catalog number, serial number, lot number, expiration date, udi number are unknown. h4 device mfr date is unknown. | The investigation of introducer damage ¿ mechanical has been completed. the returned introducer and companion sheath were inspected and damage along the companion sheath tip was observed. additional information from clinical team mentions observations of the physician getting caught up and having to use more force to get companion sheath through the peel away sheath. there were no patient vessel issues or conditions. the cause of the 7fr companion sheath tip damage was most likely user related. d9 device returned date was inadvertently not included on the initial manufacturer device report number 1220648-2025-26307. g1 reporting contact facility name was inadvertently not included on the initial manufacturer device report number 1220648-2025-26307.
11457 | 1220648-2025-26308 | 21458170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21458170&pc=OZD | 01/31/2025 | 03/18/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586073 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. when physician advanced reposition sheath he also advanced catheter. this caused the cannula to prolapse and kink the kink led to pump failure due to high motor current. decision made to remove pump and replace with new cp. no patient harm was reported. | [5 empty] | The investigation of positioning issues and saturated flow has been completed. the data logs were unavailable. the product was returned and evaluated. there were two kinks found in the cannula. the pump was run in a bucket of water with no abnormalities and the pump stop was not reproduced. no other abnormalities were found. the root cause of the positioning issue was most likely use-related as the physician inadvertently advanced the pump into the left ventricle (lv) while advancing the repositioning sheath. pump stop. the failure mode "saturated flow" was updated to "pump stop" since pump failure due to high motor current was reported. the root cause of the pump stop was most likely use-related since there was evidence of cannula kinking on the returned product that matched the clinical details. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11458 | 1220648-2025-26309 | 21458387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21458387&pc=OZD | 01/28/2025 | 04/29/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588685 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump was supporting for 20 hours when the patient expired from multiple organ failure due to intraperitoneal bleeding. the bleeding originated at the extracorporeal membrane oxygenation (ecmo) site. the wire used during ecmo insertion damaged the mesenteric artery, resulting in intraperitoneal bleeding. the contribution of the heparin in the impella purge and impella use and any contribution to the death is not known. the bleeding made stabilization of the patient difficult and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11459 | 1220648-2025-26310 | 21459183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21459183&pc=OZD | 01/31/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547841 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient (unknown race and ethnicity) in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and developed limb ischemia. the patient would eventually expire. the patient presented to the emergency department three days prior to the impella implant procedure complaining of abdominal pain, increased thirst for two days along with nausea and vomiting. initially the patient was treated for diabetic ketoacidosis and septic shock. the patient's ekg showed st elevation in the posterior leads and the patient was taken to the cath lab as a code st-segment elevation myocardial infarction at such time. the following day support was discontinued as patient drip requirements and labs were headed in the right direction. there was an attempt to transfer the patient out, but two hospitals turned down the admission at this time. the decision was made, two days later, to put the patient on support again due to worsening labs and increased drip requirements as well as possible transfer for higher level of care. the impella cp device was placed; however, it was noted no intervention was required at this time. the patient was transferred to the unit on impella mcs. the next day the patient was cannulated on ecmo. bilateral pulses were absent per the nurse, and it was unsure if either from the clamp down or compromised perfusion. continuous renal replacement therapy was started. the next day it was noted a reperfusion catheter was placed but the vascular suggested ischemia did not resolve after six hours which required a fasciotomy which was completed at bedside. the following day the impella performance level was turned down to p-3 due to suction events. also, on this day, the patient was noted to have poor flow to the bilateral lower extremities, the team suspected compartment syndrome. the following day the impella was having suction events; performance level was lowered to p-2. it was noted ecmo was "chugging." the patient continued to be supported with all therapies (crrt, ecmo and impella). the vascular surgeon recommended the right lower extremity amputation but decided to wait until patient became stable. however, the patient would expire, it appears, before the amputation after seven days total support. care was withdrawn. the patient was very ill and was not doing well going into the procedure for impella mcs. the patient had a combination of therapies including ecmo. the cause of death is unknown. there is not enough information, however, to exclude the impella cp device as an associated factor to the overall outcome when considering the impella related events (ischemia leading to amputation). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11460 | 1220648-2025-26311 | 21459402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21459402&pc=OZD | 01/29/2025 | 04/23/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510975 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was on impella 5.5 support and developed hemolysis. there is no clear reason for hemolysis as all parameters were in range and the impella 5.5 position was verified. the next day, the impella 5.5 was successfully weaned and removed. hemolysis was not the reason for the explant and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the returned pump was severed and only the cannula was returned. the pump had no abnormalities and could not be tested. no data logs were provided. the root cause of the hemolysis was not determined as no data logs were provided for the case, insufficient clinical information was provided and partial product was returned d9 date device returned to manufacturer added.
11461 | 1220648-2025-26312 | 21459420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21459420&pc=OZD | 06/06/2024 | 03/17/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  | 02/28/2025 |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325809 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiomyopathy | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The clinical study further reviewed the patient information and determined that the ischemic cardiomyopathy was determined to be not related to the impella or impella procedure. | A notification was received from the protect iv study and found the patient had ischemic cardiomyopathy with a definite relationship to the impella device and a probable relationship to the impella procedure. the patient was with new onset symptoms of chest burning sensation discomfort and shortness of breath with activity and exertion. the patient was referred for a cardiac catheterization and possible percutaneous coronary intervention. in-stent restenosis was found to the distal left main, mid left anterior descending and chronic total occlusion to the left circumflex. the patient underwent impella assisted pci, which was complicated by relatively asymptomatic hypotension and acute kidney injury now resolved. the acute kidney injury was noted as a probable relationship to the impella procedure and a probable relationship to the impella device. entresto and spironolactone were held to re-assess renal function. the following day, it was noted the aki was improving. entresto was discontinued in favor of valsartan. spironolactone was reinitiated. and as noted above, the aki had resolved (after three days from on-set) and the patient was discharged. |  |  |  |  | B.5 corrected information was received that cardiomyopathy should of been omitted from the initial manufacturer device report number 1220648-2025-26312. h.6 due to corrected information received, health effect impact code 1764 should of been omitted from the initial manufacturer device report number 1220648-2025-26312. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11462 | 1220648-2025-26313 | 21459553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21459553&pc=OZD | 02/01/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/01/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568544 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Fluid/Blood Leak | [5 empty] | Insufficient Information | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and after the implant low flows were observed. four units of packed red blood cells were given along with multiple other products for volume. the suction and low flow issues resolved and support was continued. additionally, a leak was found coming from the purge side arm. the leak stopped and the purge flow and pressure were in range. support was continued and there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
11463 | 1220648-2025-26314 | 21459654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21459654&pc=OZD | 01/31/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/01/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025591060 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support experienced some bleeding from their access site. femostop was applied to achieve hemostasis and one unit of blood products was transfused. support continued uninterrupted following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11464 | 1220648-2025-26315 | 21459783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21459783&pc=OZD | 01/17/2025 | 09/29/2025 |  | 02/25/2025 |  |  | 02/02/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584946 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | [11 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that there were placement signal not reliable alarms. the pump was working correctly and there were no other concerns. | Additional information received that the patient developed profound neurological changes and a large brain bleed was discovered. care was withdrawn and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported a patient was on impella 5.5 support and during support, the patient had bleeding. the clinical representative could not confirm the location of the bleeding. it occurred after the impella implant. one unit of blood products were given to the patient. the procedural outcome was the patient survived the surgery. |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue, vascular damage - peripheral vessel, and stroke. no product was returned for analysis. the root cause of the vascular damage - peripheral vessel was not determined as no product was returned and limited clinical information related to failure was provided. the data logs confirm placement signal not reliable alarm and show controller error alarm. the root cause of the optical signal issue could not be definitively determined as no product was returned for investigation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-26315. d10 concomitant medical products and therapy dates updated. | Instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b2 patient outcome revised to reflect death and date of death which inadvertently was not included on the initial manufacturer device report 1220648-2025-26315. b5 revised to reflect the patient outcome which inadvertently was not included on the initial manufacturer device report 1220648-2025-26315. h6 revised to reflect additional information and patient outcome which inadvertently was not included on the initial manufacturer device report 1220648-2025-26315. h10 revised to reflect instructions for use which inadvertently was not included on the initial manufacturer device report 1220648-2025-26315. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11465 | 1220648-2025-26316 | 21459858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21459858&pc=OZD | 02/02/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/02/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562926 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp and during support, the patient had ventricular tachycardia. the patient was cardioverted and converted to normal sinus rhythm. the pump was repositioned and the impella cp was at proper depth at the time of the event. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11466 | 1220648-2025-26317 | 21460266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21460266&pc=OZD | 02/01/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/01/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582388 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support experienced rhythm changes overnight resulting in cardioversion. impella subsequently repositioned with echocardiogram. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11467 | 1220648-2025-26318 | 21460610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21460610&pc=OZD | 02/04/2025 | 05/01/2025 |  | 02/25/2025 |  |  | 02/04/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573056 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while a patient was on impella 5.5 support, a red alarm purge system was blocked and there were no kinks in the line. the decision was made to explant the pump and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella 5.5 was returned for evaluation. high purge pressure: upon visual inspection, biomaterial was present in the purge gap. no blood was present in the purge gap. pump was connected to console in the lab with returned purge cassette and ¿high purge pressure¿ alarms were triggered after approximately 10 minutes. pump was cut right below motor housing and high purge pressure was able to immediately resolve. data logs were reviewed and log of full case shows ¿high purge pressure¿ alarms occurring on last day of support. purge pressure rise observed over a period of approximately 8 hours. no motor current spikes observed prior to purge pressure rise. purge pressure remained high for remainder of support with ¿high purge pressure¿ alarms being triggered. clinical details noted that pump weaning had been initiating on day 4 of support when high purge pressure alarms were triggered. purge at time of alarms was d5w with sodium bicarbonate. no kinks present on product noted in the field. the root cause of the high purge pressure was most likely due to biomaterial buildup in the purge gap. d9 date device returned to manufacturer added.
11468 | 1220648-2025-26319 | 21460697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21460697&pc=OZD | 02/02/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/02/2025 | 02/25/2025 |  |  | 12/31/2025 |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456122 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the 63-year-old female patient was on impella cp support and during support, the patient lost pulses in the right lower extremity. no intervention was performed as there was already a plan to explant the pump. it was further reported that care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11469 | 1220648-2025-26320 | 21461040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21461040&pc=OZD | 02/03/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/03/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584615 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella in aorta alarm appeared during impella 5.5 support for a 60 year old man with acute decompensated cardiomyopathy. in response to the noted failure, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11470 | 1220648-2025-26321 | 21462302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21462302&pc=OZD | 01/28/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 01/28/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592302 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and developed thrombus, limb ischemia and would subsequently expire. the patient was receiving dialysis and coded. cardiopulmonary resuscitation (cpr) with multiple shocks was delivered. the patient went to the emergency department and extracorporeal membrane oxygenation (ecmo) was placed. computed tomography (ct) scan revealed a thoracic bleed and three broken ribs from the cpr. the patient was transported to the catheterization lab and catheterization showed chronic total occlusion of the circumflex and "severely" right coronary artery and a small left anterior descending artery was open and possibly clamped down due to pressors. the decision was made to implant an impella cp device which was successfully placed. the patient was noted as very sick and was not a candidate for ventricular assist device or transplant. due to internal bleeding from the cpr, three units of blood and volume was given. the following day after start of impella mcs, thrombus and limb ischemia was noted to the impella leg. an attempt to place an antegrade sheath was noted with the possibility of a thrombectomy as well. the patient has a history of "severe" peripheral vascular disease and multiple interventions on both legs. the following day, the antegrade sheath that was placed clotted off. severe peripheral vascular disease with no distal outflow was noted. continuous renal replacement therapy (crrt) was started with noted issue of circuit clotting as well. the next day the unit update noted the patient had a gastrointestinal bleed. the patient expired the same day. ct scan showed no neural activity and care was withdrawn. the patient was very ill and on a combination of therapies including ecmo. there was a lot going on with the patient now in addition to the thrombus and limb ischemia related to the impella, and with this, there is not enough information to exclude the impella cp device as an associated factor to the outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death."
11471 | 1220648-2025-26322 | 21462839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21462839&pc=OZD | 02/04/2025 | 03/22/2025 |  | 02/25/2025 |  |  | 02/04/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582257 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during impella cp initiation was unable to obtain adequate flows. suction was above .5l while having great position and preload. patient had a difficult anatomy, and the physician felt like something may have caused potential damage/malfunction during insertion. the impella cp was removed and placed with new impella. new impella placed without issue and obtained adequate flows. impella cp saved to be sent back for further investigation and there was no patient harm. | [5 empty] | The investigation for the low pump flow has been completed. data logs confirmed low pump flow when pump started and remained to the rest of the case. visual inspection found a kink on can about 15 mm from of cage and cannula bonding site. no other issues were found on the rest of the pump. pump ran without issue under bench test. the root cause of low flow was most likely kinked cannula. corrections to the initial mfr report: h5 should have been yes. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11472 | 1220648-2025-26323 | 21463102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463102&pc=OZD | 01/30/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595471 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient experienced bleeding from their impella cp access site following implant. femostop was applied to manage the condition and three units of packed red blood cells were given to the patient. support continued uninterrupted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11473 | 1220648-2025-26324 | 21463107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463107&pc=OZD | 01/30/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1739501 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was found to have a battery failure. there was no patient involvement. | [5 empty] | The investigation into the reported issue has been completed. console logs were partially consistent with what was reported in the complaint. the complaint date was listed as (b)(6) 2025 but no logs exist from that date. however, there were battery failure alarms observed from (b)(6) 2025. there a single instance of battery failure (transport connection blown/missing) alarm observed from this date. console was tested after arriving in field service. the reported battery failure issue was unable to be reproduced. no battery failures or charging issues were observed while testing this console. additionally, results of running a batttest on this console showed that there were no issues with the batteries. the reported battery failure issue was determined to be use related. logs show that the battery failure alarm occurred during first boot of this console after preventative maintenance by field service. routinely, field service disengages the battery transport connection switch prior to shipping the console back to the customer after functional check or preventative maintenance procedure. it is the responsibility of the customer to re-engage the transport connection switch. a reminder note is adhered to the screen of the console to remind them of this duty. the complaint description mentioned that the alarm did not resolve after the team was advised to push the battery transport switch on the bottom of the console and charge it overnight but it is likely that the transport switch was still not engaged properly prior to charging it overnight. a clicking noise should be heard when the switch is engaged properly. before sending this console back to the customer, field service was instructed to verify power system functionality and perform a full functional check on the console and optical system.
11474 | 1220648-2025-26325 | 21463111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463111&pc=OZD | 01/30/2025 | 03/07/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599883 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the 0.018 wire was inserted, but the impella cp pump failed to deliver beyond the valve. delivery failed and was aborted. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, unable to pass impella through atrioventricular due to aortic stenosis. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26325 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11475 | 1220648-2025-26326 | 21463117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463117&pc=OZD | 10/20/2024 | 03/18/2025 |  | 02/25/2025 |  |  | 01/30/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546807 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced rental failure during impella 5.5 support. it was noted that the patient required continuous renal replacement therapy after not responding to bumex. the patient was transitioned to hemodialysis to manage the condition. | [5 empty] | The investigation of renal failure has been completed. the impella device was not returned for evaluation. the cause of the renal failure was not determined due to insufficient clinical details. b.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26326. b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26326. e.1 added us phone number and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26326. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26326 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.2 added report source study as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26326. h.5 revised selection as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26326. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11476 | 1220648-2025-26327 | 21463121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463121&pc=OZD | 01/31/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080191 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump had signal loss and the team tried to troubleshoot. the console was replaced, but the team had to explant and replace with an intra-aortic balloon pump. | [5 empty] | The investigation into the reported issue has been completed. console logs from the reported complaint date are consistent with the complaint. this console was tested after arriving in field service. however, the reported optical sensor system errors could not be reproduced. the console functioned normally while testing. the reported optical sensor system error issue was the result of a pump detection issue due to a rare occurring crystal failure. the impellatronic does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a mitigation in place. before sending this console back to the customer, field service was instructed to replace the impellatronic board and perform a full functional check on the console and optical system.
11477 | 1220648-2025-26328 | 21463128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463128&pc=OZD | 01/31/2025 | 03/18/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502397 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis during support. repositioning was performed to troubleshoot the condition, but the issue persisted. in response to the hemolysis, the decision was ultimately made to explant the pump. weaning was successful. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the returned complaint cp device visually inspected. no biomaterial observed. no broken blade or sharp edge observed. hemolysis test conducted successfully with mih=5.72 which passes the hemolysis test. data log reviewed: no motor current (mc) spike outside of the p-level changes occurred. no placement signal trend observed. no suction or positioning issue alarms observed. clinical stated: pump contacted mitral apparatus, but repositioning didn't solved the reported issue. the cause of reported hemolysis issue cannot be determined due to lack of clinical details. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26328 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11478 | 1220648-2025-26329 | 21463141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463141&pc=OZD | 01/31/2025 | 03/12/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601889 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was on support for 22 hours and the device was explanted due to limb ischemia. a new impella cp was placed on the opposite side, the patient outcome was noted to be stable. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. e.1 revised phone number and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26329. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26329 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 compont code 925 was reported incorrectly on initial manufacturer deivce report number 1220648-2025-26329 and a new code was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
11479 | 1220648-2025-26330 | 21463142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463142&pc=OZD | 01/31/2025 | 03/17/2025 |  | 02/25/2025 |  |  | 01/31/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603556 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient experienced non-access site bleeding. heparin was removed and the impella cp was explanted. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. as per the clinical information provided, the patient was bleeding from a non-impella site. heparin was reversed and pump was removed. the cause of the vascular damager peripheral vessel issue was not determined since product was not returned and lack of clinical details, and the relation to impella is unknown. d4 catalog number was inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26330. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26330 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11480 | 1220648-2025-26331 | 21463149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463149&pc=OZD | 01/31/2025 | 07/17/2025 |  | 02/25/2025 |  |  | 02/01/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | GM |  | 2025599056 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support with plan for bridge to left ventricular assist device. after 10 days on the support the pump stopped. it was further noted that the motor current increased and the pump then stopped with unsuccessful repeated attempts to restart. consequently, the impella cp device was explanted with a decision for escalation in therapy to an impella 5.5 device which was noted to be successfully completed. the patient was noted to have survived the explant to escalation; however, the patient's status condition was unknown. | [5 empty] | The investigation of pump stop has been completed. the pump was not returned for investigation. the data log indicates a pump stop accompanied by a motor current spike, resulting in alarm 13. although the pump was able to restart and operated for another 26 hours, it experienced saturation in pump flow and slightly elevated purge pressure on 10th day. eventually, the pump could not restart and was explanted. additionally, a clinical image was not provided. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26331 as it is unknown if the initial reporter also sent the report to the food and drug administration. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11481 | 1220648-2025-26332 | 21463150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463150&pc=OZD | 02/02/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/02/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572067 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support was taken to the cath lab for distal limb ischemia. retrograde perfusion line was placed in right femoral artery. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
11482 | 1220648-2025-26333 | 21463156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463156&pc=OZD | 02/04/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/04/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508822 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that following impella 5.5 implant, a purge system blocked alarm appeared. tissue plasminogen activator was added to the purge in response to the alarm and support continued uninterrupted. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11483 | 1220648-2025-26334 | 21463160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21463160&pc=OZD | 02/04/2025 | 02/25/2025 |  | 02/25/2025 |  |  | 02/04/2025 | 02/25/2025 |  |  |  |  | 02/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849119 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant had a impella 5.5 pump support ongoing when there was an automated impella controller (aic) damage. the team unplugged the console to transport patient to computed tomography imaging and found wire/cord damage. the aic was replaced and there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11484 | 1220648-2025-26270 | 21444774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21444774&pc=OZD | 01/31/2025 | 03/19/2025 |  | 02/24/2025 |  |  | 01/31/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532057 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and the following day after start of the support, the patient developed hemolysis. as a result of the hemolysis coinciding with a co deficiency, the patient also experienced an acute kidney injury. dialysis was initiated and the patient was escalated to impella 5.5 support. the patient was reported to be stable and nimble following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and the renal failure has been completed. no pump was returned for evaluation. no data logs were returned for the case. the root cause of hemolysis was not determined as no product or logs were returned for analysis and insufficient clinical details were provided. the cause of the issue was not determined due to insufficient clinical details.
11485 | 1220648-2025-26271 | 21444804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21444804&pc=OZD | 01/30/2025 | 04/01/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1748660 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while attempting to setup an automated impella controller (aic) for patient impella mechanical support, the silver cog within the purge would not spin when approached by the purge cassette. the purge cassette also failed to lock in place, forcing the staff to manually manipulate the device to properly align the purge cassette. support was successfully started with the aic following manual troubleshooting. there was no report or indication of patient harm or delay in therapy. | [5 empty] | The investigation for the console (aic) purge issue has been completed. data logs were reviewed finding that upon the initial insertion of the purge cassette, the log recorded a continuous purge communication error. after the cassette was removed and reinserted twice, there was no purge communication error. then, the impella cp was inserted, and support was provided for 1 hour and 35 minutes. the console was shut down and manually restarted. through the ¿case start¿ wizard, the same cassette was reinserted multiple times, but a continuous purge communication error was recorded. this aligns with the reported failure mode. the console was returned for evaluation. all the cables between the 4-in-1 and the pud pca were connected appropriately. upon inserting the test cassette multiple times, the purge motor spindle spins each time. unable to reproduce the reported failure mode. a minor dextrose residue was found on the purge motor gasket. the power to the pud pca was 23.8v under ac power and 16v under battery power, which are the typical working voltages that match the known good test fixtures. upon initial console boot-up, the console displayed a battery failure alarm, and the batttest confirmed the cell imbalance with cell 2 of battery 1. unrelated to the reported failure mode. upon replacing the original batteries with known good batteries, the console was able to charge the known good batteries, confirming that there was no issue with the pbm. the root cause for the purge motor spindle not rotating was not determined due to the inability to reproduce the reported failure mode. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11486 | 1220648-2025-26272 | 21444859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21444859&pc=OZD | 01/30/2025 | 03/11/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531628 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) and developed an access site hematoma and bleed. the impella cp was implanted and the pci to the saphenous vein graft to obtuse marginal was completed. the patient was being prepared for transfer to the unit; performance level of p-8 with a flow of 3.4 lieters per minute. distal reperfusion sheath was in place. as the team was preparing the patient for transfer, a large hematomas were noted at the impella right femoral artery access site and at right internal jugular swan site. the fellow held pressure and hematoma was resolved. however, bleeding occurred around the impella sheath. manual pressure was held for 30 minutes and the fellow deployed a purse string suture. the site was still bleeding. the team was unable to obtain hemostasis. heparin was discontinued. the physician elected to remove the impella and implant an intra-aortic balloon pump in the opposite groin. the patient was reported to be in critical condition following the event. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the root cause of the access site bleeding was most likely patient condition since the patient had hematomas at multiple sites.
11487 | 1220648-2025-26273 | 21444935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21444935&pc=OZD | 01/31/2025 | 02/27/2025 |  | 02/24/2025 |  |  | 01/31/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562875 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp for approximately three days and then escalated to an impella 5.5 device for continued mechanical circulatory support. while on the impella 5.5 device high purge pressure was encountered resulting in device replacement. the impella was successfully implanted. the wire was removed, and support was initiated a performance level p-2. the vascular surgeon was called to explant the impella cp device. the impella 5.5 was gradually increased to performance level p-5; position and depth were check on transesophageal echocardiogram. following, the automated impella controller alarmed for purge pressure high. purge flows were 1.8-1.9 ml/hr. with pressures 1050-1100 mmhg, the purge cassette was changed; flow stopped and started twice all without improvement. the decision was made to explant the impella 5.5 device and replace it with a new one. the outcome at explant/replacement noted the patient survived the event. however, the status of the patient's condition was indicated as unknown. | [5 empty] | The investigation of high purge pressure has been completed. the returned product was inspected and there was no biomaterial observed in the purge gap or on the impeller. there were traces of possible biomaterial remains on the outlet cage. the returned pump was run in the lab and high purge pressure did not reproduce. the data logs confirm high purge pressure alarms. the pump was initiated and there was a little motor current spike with a speed deviation shortly after. the purge flow also decreased. there was a gradual rise in purge pressure until alarms were triggered. the root cause of the high purge pressure was most likely due to biomaterial ingestion. d6a revised on manufacturer device report. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11488 | 1220648-2025-26274 | 21444956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21444956&pc=OZD | 01/31/2025 | 03/14/2025 |  | 02/24/2025 |  |  | 01/31/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578655 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and the following day, the physician decided to remove the impella cp device due to worsening hematoma and drop in hemoglobin from 9.5 g/dl to 8.5 g/dl over four hours. an echocardiogram was obtained, and the physician noted the patient would not benefit further from impella mcs in comparison to risks with bleeding. echocardiogram showed improved function to ejection fraction of 50% and total cardiac output of 7.5 liters per minute. the patient underwent a computed tomography angiogram (cta) to rule out a retroperitoneal bleed and other complications. the cta was clear. however, the physician still elected to explant the impella cp device. it was noted there were no issues with explant, and hemostasis was achieved with two perclose and manual pressure. one unit of packed red blood cells was transfused, and the patient was reported to be in stable condition following the event. | [5 empty] | The investigation of access site bleeding has been completed. the root cause of the access site bleeding cannot be determined due to insufficient clinical details were provided. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11489 | 1220648-2025-26275 | 21445125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445125&pc=OZD | 02/01/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 02/01/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491072 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and during of scheduled explant of the impella cp, thrombus was noted. the impella cp was explanted in the catheterization lab. radial access was used to place balloon at the arteriotomy. the sheath was removed and balloon inflated. perclose x2 closed with success. completion angiogram showed successful closure but revealed a thrombus. consequently, the patient was taken to the operating room for a thrombectomy. the patient was reported as stable following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
11490 | 1220648-2025-26276 | 21445282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445282&pc=OZD | 01/29/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/29/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577977 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient suffered bilateral loss of doppler pulses following impella cp implant. it was noted that the patient already had peripheral artery disease and was ischemic while on intra-aortic balloon pump support prior to impella implant. following implant, the ischemia persisted until the family made the decision to withdraw care. the patient passed away shortly after. it is unknow if support contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11491 | 1220648-2025-26277 | 21445397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445397&pc=OZD | 01/27/2025 | 03/12/2025 |  | 02/24/2025 |  |  | 01/27/2025 | 02/24/2025 |  |  | 02/28/2025 |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313375 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received regarding a study patient that experienced ventricular tachycardia, refractory to lidocaine, and successfully treated with direct current cardioversion to sinus tachycardia, albeit with frequent ectopy. the patient prognosis was poor. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26277 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source; study was missing from manufacturer device report 1220648-2025-26277.
11492 | 1220648-2025-26278 | 21445411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445411&pc=OZD | 01/28/2025 | 03/07/2025 |  | 02/24/2025 |  |  | 01/28/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559441 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient's right leg didn't have doppler or palpable pulses. the impella was removed. positive pulses and flow post intervention was noted by vascular surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26278 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11493 | 1220648-2025-26279 | 21445433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445433&pc=OZD | 01/28/2025 | 03/12/2025 |  | 02/24/2025 |  |  | 01/28/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536024 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient lost pulses to the right lower extremity. the pump was then explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. e.1 revised us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26279. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26279 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11494 | 1220648-2025-26280 | 21445439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445439&pc=OZD | 01/29/2025 | 03/14/2025 |  | 02/24/2025 |  |  | 01/29/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601372 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the physician was repositioning impella 5.5 as under echo impella was visualized as deep. the impella was pulled back and while rotating it there was artifact on the echo that appeared as railroad tracks. when the artifact did not move, it was noted that it may not be impella. probe moved and showed impella pulled back across the aortic valve. tried to re-advance but unable to cross. patient stayed hemodynamically stable with no changes in blood pressure. the physician made decision to explant 5.5 at bedside. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical detail, the cause of the positioning issue most likely was use related since impella pulled back across the aortic valve after impella was visualized as deep under echo, tried to re-advance but unable to cross. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26280 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11495 | 1220648-2025-26281 | 21445440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445440&pc=OZD | 01/29/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601387 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation | [11 empty] | Required Intervention | Disability | Life Threatening | Death | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.5 device as elective placement for mechanical circulatory support. the patient was undergoing a left ventricular thrombectomy and an aortic valve replacement (avr) surgery. after the surgery, the impella was directly inserted via the aorta. the left ventricular perforation occurred while the catheter was inserted into the left ventricle, and its position was adjusted. after repair of the adverse event, there was no flow was obtained even after impella 5.5 was started. repair was attempted again; however, bleeding could not be controlled. the impella 5.5 was explanted in the operating room. the patient expired as a result of the adverse event. the physician further commented the following: the reason the impella was inserted using a direct aorta was the chest was opened after avr, and the blood vessel diameter was not large enough to insert 5.5. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11496 | 1220648-2025-26282 | 21445454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445454&pc=OZD | 12/04/2024 | 03/14/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568583 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Ischemia | Pain | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured a gastrointestinal (gi) bleed and atrial fibrillation with a "possible" relationship to the impella device used. the gi bleed developed during support and was treated with three units of packed red blood cells. it was reported that the patient/event has not recovered/not resolved. approximately one month after explant of the impella device, the patient experienced atrial fibrillation. the patient was documented to have recovered from atrial fibrillation event with no sequelae. | [5 empty] | The root cause of the thrombosis and ischemia was unable to be determined due to insufficient clinical details. the cause of the pain was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vascular damage - peripheral vessel and atrial fibrillation has been completed. the pump was not returned for investigation. according to the clinical notes, a gastrointestinal (gi) bleed was noted on the fifth day of pump use. additionally, atrial fibrillation was reported one month after the impella explant. the cause of the vascular damage issue was not determined since the product was not returned and sufficient clinical information was not required. the relation between gi bleed and impella was unknown. the cause of the atrial fibrillation was unable to be determined due to insufficient clinical details. the relationship between the impella device and the atrial fibrillation is unrelated, as the issue occurred one month after the explant. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26282 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26282.
11497 | 1220648-2025-26283 | 21445460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445460&pc=OZD | 11/30/2024 | 03/05/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529902 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient (unknown ethnicity) who experienced monomorphic ventricular tachycardia with a "possible" relationship to the impella device used. impella was implanted and single access protected percutaneous coronary intervention was completed with "great" success and stability. the device was weaned and explanted. three days after the impella device was explanted, the patient experienced two separate sustained runs of monomorphic ventricular tachycardia. the first lasting ~2 minutes, and the second lasting 6 minutes. the patient was asymptomatic and spontaneously returned to sinus tachycardia after a rapid response team was called. electrophysiologist was consulted; intravenous potassium and magnesium replacement was ordered as well as intravenous amiodarone drip and increased dose of metoprolol. the patient was scheduled for biventricular implantable cardioverter-defibrillator placement, which occurred two days later before discharge from index hospitalization. the patient recovered with sequelae. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26283 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11498 | 1220648-2025-26284 | 21445466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445466&pc=OZD | 01/31/2025 | 03/13/2025 |  | 02/24/2025 |  |  | 01/31/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218788 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) displayed a persistent controller error alarm. the internal cables were reseated; however, the controller error alarm remained unresolved. there was no noted patient involvement at the time of the reported failure. further, it was noted the complaint was not related to a clinical procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console was powered up on with a defective optical sensor lamp as bad lamp recovery had prompted a 5s measuring light at 0.49 volt which is equivalent to no light from the optical bench. this was reproduced three more times in the field before the console was returned for investigation. the failure mode was reproduced during testing. replacing the optical bench with a known-good alleviated the failure mode. the original bench was tested individually and a break in the fiber was observed which explains why the optical bench was reading 0.49 volt (no light) despite confirming the lamp was on. the cause of the controller issue was due to a broken fiber on the optical bench causing no light to be outputted. a.1 and a.3 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26284. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26284 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11499 | 1220648-2025-26285 | 21445470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21445470&pc=OZD | 02/01/2025 | 03/17/2025 |  | 02/24/2025 |  |  | 02/01/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577626 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced oozing at the impella access site post implant same day. the site was reinforced with new sutures and angle match dressing reapplied. the following day the patient remained critical and noted to be unstable requiring maximum pressors for hemodynamic support. comfort care measures were discussed with family and the patient was made a do not resuscitate. explant of the impella cp device was planned for the following day, waiting for family to withdraw care. the next day, the impella cp was weaned and explanted per the plan. the patient had episodes of atrial fibrillation requiring minimal levophed support. it was noted levophed would be eventually weaned once patient returns to normal sinus rhythm. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related, new reinforced sutures and angle match dressing reapplied. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26285 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11500 | 1220648-2025-26286 | 21446665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21446665&pc=OZD | 01/30/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582334 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had an ooze at the access site. heparin was withheld and the doctor may have placed an additional suture. furthermore, the patient had an magnetic resonance imaging and found, "tiny acute versus subacute cortical cvas bilateral parietal lobes and cerebral hemispheres". support was continued and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11501 | 1220648-2025-26288 | 21446866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21446866&pc=OZD | 01/29/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/29/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589819 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that multiple complications were encountered during impella cp support. upon placement of a new purge cassette, a low purge pressure message appeared. the purge cassette was replaced and the issue resolved. it was additionally noted that the patient developed a hematoma around their access site following implant. blood products were infused at the time of the condition. support continued uninterrupted and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11502 | 1220648-2025-26287 | 21447097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21447097&pc=PYX | 01/29/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/29/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025575829 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella rp flex support ongoing when there was purge pressure/purge flow alarms that prompted the team to add tissue plasminogen activator to the purge. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
11503 | 1220648-2025-26289 | 21447405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21447405&pc=OZD | 01/29/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591656 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed hemolysis during impella cp support. the condition was diagnosed via plasm free hemoglobin levels and the pump was explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11504 | 1220648-2025-26290 | 21447580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21447580&pc=OZD | 01/30/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1088166 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that there was an automated impella controller error and the alarm was resolved without intervention. the console showed absent placement signal prior to alarm. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11505 | 1220648-2025-26291 | 21449090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21449090&pc=OZD | 01/25/2025 | 04/23/2025 |  | 02/24/2025 |  |  | 01/25/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1221065 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a/an (unknown age) patient with post-cardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs). prior to impella mcs it was noted the patient was on venoarterial extracorporeal membrane oxygenation (va ecmo), inotropes/vasopressors and respiratory (bilevel positive airway pressure [bipap] or continuous positive airway pressure [cpap]) support. approximately three days after start of the impella mcs a call from perfusionist was received due to an observed short breakdown from automated impella controller (aic) with independent restart and alarm "console failure¿ and noted ¿now pump works well.¿ it was recommended that the aic be taken out of service, which appears to have occurred after the implant experience (exact date is unclear, unknown). there were no adverse effects or harm to the patient because of the aic console failure based on the reported information from the reporter. the patient continued on the impella 5.5 mcs. the patient would expire three days later, however, as care was withdrawn. additional information was sought regarding details around the patient¿s demise and the following, although limited, was noted: regarding the aic, after its independent reboot, the pump motor resumed at the previous flow level, and it was stated this did not have anything to do with the patient¿s demise. it was noted that therapy was stopped (care was withdrawn) because of ¿other disease/comorbidities, which were not heart related,¿ and ¿there was no way for a good outcome.¿ these other issues were noted as multi-organ failure, massive cerebral hemorrhage and intestinal ischemia. it is unknown, however, exactly when these adverse events occurred. care was withdrawn for adverse experiences not caused by the aic and given the information, at least one of the adverse events is listed as a potential outcome with use of the impella pump (cerebral vascular accident cva/cerebral hemorrhage). with no alternative cause for the cva event identified as the likely cause, there is not enough information to exclude the impella 5.5 as an associated factor in the patient¿s demise. medical safety review of the event noted for the aic, there was hemodynamic support maintained and regarding the impella 5.5 device, however, there is not enough data to assess the cva if impella pump related; it cannot be excluded from the outcome (demise) based on the limited, incomplete details. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console logs showed the qnx process io-graphics crashed which caused the aic software to reboot. the console rebooted in response followed by an unexpected shutdown alarm confirming what was reported clinically. therapy was resumed before the pump was unplugged and the console was shutdown. device analysis: the failure mode was not reproduced during investigation on engineering lab bench. the io-graphics was force shutdown using a telnet command and the console responded the same as in the field. root cause: the cause of the aic issue was the 3rd party hardware/software that runs io-graphics. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26291 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The automated impella controller was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. this is one of two manufacturing device reports associated with this event. refer to the following medical device report for automated impella controller (b)(6) with date of event 25-jan-2025 and patient identifier ending in 6593 (this report) medical device report for impella 5.5 (b)(6) with date of event (b)(6) 2025 and patient identifier ending in (b)(6).
11506 | 1220648-2025-26292 | 21449798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21449798&pc=OZD | 01/25/2025 | 04/23/2025 |  | 02/24/2025 |  |  | 01/25/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601495 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Life Threatening | Death |  |  |  |  | ECMO AND BIPAP/CPAP | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a/an (unknown age) patient with post-cardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs). prior to impella mcs it was noted the patient was on venoarterial extracorporeal membrane oxygenation (va ecmo), inotropes/vasopressors and respiratory (bilevel positive airway pressure [bipap] or continuous positive airway pressure [cpap]) support. approximately three days after start of the impella mcs a call from perfusionist was received due to an observed short breakdown from automated impella controller (aic) with independent restart and alarm "console failure¿ and noted ¿now pump works well.¿ it was recommended that the aic be taken out of service, which appears to have occurred after the implant experience (exact date is unclear, unknown). there were no adverse effects or harm to the patient because of the aic console failure based on the reported information from the reporter. the patient continued on the impella 5.5 mcs. the patient would expire three days later, however, as care was withdrawn. additional information was sought regarding details around the patient¿s demise and the following, although limited, was noted: regarding the aic, after its independent reboot, the pump motor resumed at the previous flow level, and it was stated this did not have anything to do with the patient¿s demise. it was noted that therapy was stopped (care was withdrawn) because of ¿other disease/comorbidities, which were not heart related,¿ and ¿there was no way for a good outcome.¿ these other issues were noted as multi-organ failure, massive cerebral hemorrhage and intestinal ischemia. it is unknown, however, exactly when these adverse events occurred. care was withdrawn for adverse experiences not caused by the aic and given the information, at least one of the adverse events is listed as a potential outcome with use of the impella pump (cerebral vascular accident cva/cerebral hemorrhage). with no alternative cause for the cva event identified as the likely cause, there is not enough information to exclude the impella 5.5 as an associated factor in the patient¿s demise. medical safety review of the event noted for the aic, there was hemodynamic support maintained and regarding the impella 5.5 device, however, there is not enough data to assess the cva if impella pump related; it cannot be excluded from the outcome (demise) based on the limited, incomplete details. | [5 empty] | The date of onset for the cerebral hemorrhage is unknown at time of this mdr writing. provisional dates for both the aware and best estimate event dates are being aligned, conservatively captured as the same dates for the automated impella controller. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two manufacturing device reports associated with this event. refer to the following medical device report for automated impella controller (b)(6) with date of event 25-jan-2025 and patient identifier ending in (b)(6). medical device report for impella 5.5 (b)(6) with date of event 25-jan-2025 and patient identifier ending in (b)(6). (this report) instructions for use for the related event are as follows: 5.5 with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26292 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11507 | 1220648-2025-26293 | 21451737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21451737&pc=OZD | 01/27/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/27/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595796 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Defective Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and during the implant procedure difficulty with the device luer was experienced. post implant heparin was discontinued given the patient's bleeding from liver laceration and purge system block alarm was experienced. the patient presented earlier in the day following cardiac arrest and motor vehicle accident involving a head on collision. the patient was diagnosed with st-segment elevation myocardial infarction, coded for an extended time in the field with achieved return of spontaneous circulation. the patient was intubated prior to transfer to the emergency department and subsequently found to have significant liver laceration on computed tomography scan. the patient decompensated throughout the day requiring multiple pressors. later in the day, evening, the patient was sent to the cardiovascular lab sent for urgent impella cp insertion which was completed without patient complication but difficulty with the luer. upon insertion of impella cp, physician noted that clear luer on the dilator did not separate after turning counterclockwise. the physician elected to cut the clear luer off with scissors during the procedure in order to remove the dilator from the peel-away and continue the operation. the patient was sent to the unit in "good" condition intubated, sedated. later that evening heparin was discontinued due to the liver laceration. the trauma surgeon, cardiovascular surgeon and intensivist educated the team on appropriate anticoagulation management. the patient continued on support. two days later, purge system block alarm was triggered, and the team was advised to place tissue plasminogen activator in the purge or replace the impella cp pump. the care team decided to leave in the pump for the high-risk percutaneous coronary intervention planned for the following day. education, however, on steps to take if pump was to fail was provided with the care team. the following day, the device was successfully explanted, and the patient was noted to have survived the explant. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11508 | 1220648-2025-26294 | 21453271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21453271&pc=OZD | 01/30/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116070 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | It was reported that a automated impella controller (aic) had a battery failure error message and would not charge past 50%. they tried turning the buttons on and off on the bottom of the aic and tried restarting and leaving plugged in, however the error message could not be resolved. this complaint was not related to a clinical procedure therefore there was no patient involvement. the aic was sent back to field service for further investigation. | [5 empty] | The investigation into the automated impella controller has been completed. the device was returned for evaluation and data logs were made available. console logs were partially consistent with what was reported in the complaint. the complaint date was listed as 2025.01.30 but no logs exist from this date. however, multiple instances of battery failure (transport connection blown/missing) [event set #360] and battery level low (50%) [event set #23] alarms were observed from the date of (b)(6) 2025. the console was tested after arriving in field service. the battery failure issue was able to be reproduced. results of running battery utility revealed that cell 2 in battery 2 had a voltage that was significantly less than the rest of the cells in the battery. the root cause of this issue was an internal cell imbalance observed in battery 2. b3 date of event; the customer reported occurrence date of 1/30/2025 but data logs shows occurrence date of (b)(6) 2025.
11509 | 1220648-2025-26296 | 21453313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21453313&pc=OZD | 01/31/2025 | 02/24/2025 |  | 02/24/2025 |  |  | 01/31/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597635 |  | 1000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). approximately five days after start of the support, the pump became malpositioned. the physicians attempted to reposition back across the valve but was unsuccessful. the patient remained hemodynamically stable on inotropes. the decision was made to explant the device in the operating room which was successfully completed with no adverse effects thereafter. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11510 | 1220648-2025-26295 | 21453425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21453425&pc=OZD | 01/30/2025 | 09/25/2025 |  | 02/24/2025 |  |  | 01/30/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586229 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | Detachment of Device or Device Component | Device Sensing Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that there was a placement signal not available alarm. | The complainant reported that multiple complications were encountered during impella 5.5 support. four days into support, the patient was returning from the operating room, when their pump was inadvertently pulled into the aorta. the patient returned to the operating room where the pump was successfully repositioned into place. however, during this process, the sterile sleeve was damaged and a tegaderm dressing was utilized to repair the component. the following day, a kink was discovered on the shaft next to the blue catheter lock. it was noted that excessive force had been applied during repositioning the day prior. the staff opted to monitor the motor current and positioning via other means and support continued uninterrupted. there was no resulting patient harm. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue: data logs were reviewed and the placement signal not reliable (psnr) alarm was confirmed. returned product came back with a cut in the catheter, so none of the standard optical signal issue test procedures could be done. upon review of data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-26295. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The investigation of product damage (cannula kink), product damage (sterile sleeve damage), and positioning issues has been completed. clinical details report that the pump was accidentally pulled back into the aorta while the patient was being moved for a computed tomography scan. the issue was resolved by bringing the patient into the operating room for repositioning. data log review showed that there were six triggers of position in aorta alarms. once the last resolved, there were no further triggers. the returned pump had a cut in the catheter just proximal to the catheter lock. there were no visual abnormalities relevant to the positioning issues reported. based on clinical details provided and data log review, the root cause of the positioning issue was determined to most likely be a use issue. data logs are not relevant to the complication. the returned product had a cut in the catheter just proximal to the catheter lock and the sterile sleeve had been removed from it, so it was unable to be inspected. clinical details report that while repositioning in the pump in the operating room, excessive force was applied, and the sterile sleeve was damaged. it was managed by using tegaderm. based on clinical details provided, the root cause of the sterile sleeve damage was determined to most likely be a use issue. data logs are not relevant to the complication. the returned product was investigated and the catheter was cut just proximal to the catheter lock. there were no visible catheter kinks as product was returned. the blue button of the catheter lock was able to be pressed and released without issue, and the catheter lock could advance back and forth smoothly over the catheter. the portion of the catheter that had been returned underneath the catheter lock was cleaned of blood and inspected, but there were no visual abnormalities/kinks. this suggests the reported catheter kink was likely on the part of the catheter that wasn't returned. clinical details report that excessive force was applied during the repositioning procedure. based on clinical details provided, the root cause of the product damage was determined to most likely be a catheter kink due to excessive force. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11511 | 1220648-2025-26297 | 21455223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21455223&pc=OZD | 01/31/2025 | 03/18/2025 |  | 02/24/2025 |  |  | 01/31/2025 | 02/24/2025 |  |  |  |  | 02/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572969 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that the automated impella controller (aic) was alarming placement signal low. the physician attempted to move the cp back, but the cp got pulled out. a new impella cp was placed. no patient harm was reported. | [5 empty] | The investigation of positioning issues has been completed. clinical reported the physician pulled the pump out when attempting to reposition the pump after getting placement signal low alarms. pump was returned with the cannula bent in the packaging which caused a kink, unrelated to the failure mode. the inflow cage was also deformed. the root cause of the positioning issues was most likely use related since the physician pulled out the pump. d9 date device returned to manufacture was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26297. h3 selection was inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26297, since the device was returned at the time of the initial report. h6 type of investigation code 4114 was inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26297. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11512 | 1220648-2025-26241 | 21432021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21432021&pc=PYX | 01/27/2025 | 03/12/2025 |  | 02/21/2025 |  |  | 01/27/2025 | 02/21/2025 |  |  | 01/31/2026 |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517138 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient had bleeding at the access site. manual pressure was held and 3 units of red blood cells were given to the patient. support continued. additionally, the patient experienced ventricular tachycardia (vt) during support. amiodarone drip was initiated. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist instructions for use & clinical reference manual, section: using the automated impella controller with the impella rp flex smartassist. system catheter: use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of access site bleeding and ventricular tachycardia (vt) has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary information was not provided, the root cause of the vt was not determined. d.4 revised serial number as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26241. e.1 revised fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26241. h.4 revised device manufacture date as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26241. h.6 code 1721 was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26241 and a new code was added.
11513 | 1220648-2025-26242 | 21432339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21432339&pc=OZD | 01/28/2025 | 03/03/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551920 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had oozing at the access site. heparin was withheld and a femstop was applied. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of the access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11514 | 1220648-2025-26243 | 21432629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21432629&pc=OZD | 01/28/2025 | 03/07/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025483523 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered limb ischemia while on impella 5.5 support. the patient's right arm was pulseless and discolored as there was no flow noted from elbow down. vascular surgery was done to perform embolectomy prior to pump removal. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11515 | 1220648-2025-26244 | 21432738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21432738&pc=OZD | 01/28/2025 | 03/18/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597220 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Crack | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp cannula kinked, and low flows were present throughout the case. the impella 14fr peel away sheath cracked at the orange ring/ white butterfly junction. significant bleeding was noted from crack in sheath. repositioning sheath was placed, and the physician sealed the crack with dermabond. the patient is stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation has been completed. the pump and a short 14 french x 13 centimeter introducer were returned for investigation. data logs were not provided for the investigation. according to the clinical details, the cannula was found to be kinked, which resulted in low pump flow throughout the case. additionally, the doctor reported bleeding from the orange introducer hub and applied dermabond to control the bleeding. no physical damage was observed on the returned pump, except for the kinked cannula. the pump was tested according to specifications in a bucket of water, and it functioned correctly. the returned introducer had dermabond applied on the orange introducer hub. orange part of introducer has exhibits some surface damage where the glue was applied. the introducer was further tested for leaks and met the pressure specification of 180 mmhg. the introducer bleeding failure mode was removed and addressed under the access site bleeding failure mode as the root cause. additionally, the reported product was changed from the pump to the introducer for the access site bleeding failure mode. the root cause of the access site bleeding was most likely due to valve damage on the introducer hub. based on the damage evident on the returned product, the damage could be attributed to interaction with another device. the cause of the low pump flow issue was a kinked cannula, based on returned product investigation and given clinical details.
11516 | 1220648-2025-26245 | 21432793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21432793&pc=OZD | 01/20/2025 | 03/27/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570125 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a study patient was implanted with an impella device for mechanical circulatory support. reintubation was performed as part of the surgical treatment of the hemothorax and respiratory insufficiency. the patient remained on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact address was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26245. g1 manufacturing site name and address was erroneously entered on the initial manufacturer device report number 1220648-2025-26245.
11517 | 1220648-2025-26246 | 21433176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21433176&pc=OZD | 01/29/2025 | 03/17/2025 |  | 02/21/2025 |  |  | 01/29/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583403 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that the patient coded with recurrent ventricular tachycardia (vt). during this time, the impella came out of the left ventricle. attempts to recross the aortic valve were unsuccessful and resulted in a kinked cannula. the pump was explanted and exchanged. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the positioning issue most likely was use related since during the code the impella was accidentally pulled out of the left ventricle (lv), unsuccessful attempted to recross the atrioventricular (av) buckled and kinked the cannula. the cause of the vf/vt/af/at was unable to be determined due to insufficient clinical details.
11518 | 1220648-2025-26247 | 21433426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21433426&pc=OZD | 01/30/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/30/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604832 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and during the support the patient experienced a number of complications: access site bleeding, high purge pressure, and a pump stop ultimately going on comfort care. the patient was noted to have multi-vessel disease and "severe" peripheral vascular disease and was placed on an intra-aortic balloon pump (iabp). the patient was a surgical turndown and planned for an impella cp device via the axillary graft for protected high-risk percutaneous coronary intervention (pci) once stable. the impella cp was successfully implanted; however, the next day the patient was moving and felt a ¿pop¿ followed by bleeding at impella (axillary) site. patient was also having some bleeding from the previous iabp (femoral) site. overnight patient received three units of blood. labs were not drawn at this time to confirm low hemoglobin; however, due to visible volume lost, the decision was made to transfuse. patient was stable with plans to go for the high-risk pci. the following day the patient's condition was unchanged/stable; the patient continued to deteriorate. right heart failure was getting worse with no plans for right-sided support. three days later the patient underwent pci. the impella wean was slowed over the next 24-48 hours. the patient still had minor bleeding from the axillary site of note. the patient was continued on the impella mcs and now seven days later high purge pressure was encountered which was resolved with a cassette change. however, shortly thereafter, the pump stopped and was restarted at performance level p-5. the decision was made to wean and explant the impella cp device and place the patient on comfort care, which was successfully executed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11519 | 1220648-2025-26248 | 21433503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21433503&pc=OZD | 12/26/2024 | 03/05/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1459449 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the automated impella controller (aic) had a blackout during insertion. after connecting the pump to the aic and priming was completed, a sudden blackout occurred while inserting it into the patient. it was rebooted and then things went back to normal. there was no patient harm. | [5 empty] | The investigation for the console (aic) power on issue has been completed. the console logs from the complaint date were consistent with what was reported in the complaint description. the clinical staff reported that there was a blackout during insertion of the pump but likely meant to report that there was a loss of placement signal. there was an instance of controller error (opm comm crc invalid) alarm observed in the logs from the reported complaint date. in addition to the controller error, there was also a placement signal not reliable alarm observed in the logs two days prior to the complaint date but this alarm was not persistent. the console was tested. the placement signal issue was unable to be reproduced. the console was tested with a pump and purge cassette kit but functioned normally without any issues. the reported loss of placement signal issue was due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the automated impella controller software operated as designed. the cause of the observed placement signal not reliable issue from two days prior to the complaint date could not be determined due to the inability to reproduce. d.9 revised as date returned was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26248. e.1 revised facility name as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26248. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26248 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11520 | 1220648-2025-26249 | 21433509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21433509&pc=OZD | 01/28/2025 | 03/06/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578503 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp placement was aborted due to the patient vascular anatomy. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11521 | 1220648-2025-26250 | 21433517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21433517&pc=OZD | 01/30/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/30/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599920 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced hemolysis and saturated pump flow. the impella was successfully implanted and the following day, it was noted the impella's depth was too deep. the patient had lab work and hematuria indicating hemolysis. the physician repositioned the pump to 3.8 cm and 30-45 minutes later, pump pulled back into aorta and was unable to be repositioned with echocardiogram at the bed side. the patient was brought back to the catheterization lab where the reposition was successful under fluoroscopy. however, despite proper positioning motor current was apulsatile and flows were reading 2l/min despite a designated performance level. the flows were indicating flow saturation. decision decided to remove the impella cp and replace it with a new device. the patient was reported to be critical following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. if the device or new information is received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11522 | 1220648-2025-26251 | 21433995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21433995&pc=OZD | 01/27/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/27/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578460 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Difficult to Remove | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that there was a ventricle perforation. after insertion of the impella cp, the patient's pressure as well as the pump's flows decreased. additionally, large amounts of blood was drained. several angiograms and a thoractomy were done which showed no perforation. however, after the patient's chest was opened and placed on bypass the doctor identified a hole in the left ventricle for which then decision to explant the impella cp was made. the patient expired before the impella could be removed. during impella cp explant, there were additional removal issues encountered as the pump got stuck and was unable to be removed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ orrectly in the left ventricle after cpr with echocardiography guidance.
11523 | 1220648-2025-26252 | 21434083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21434083&pc=OZD | 01/27/2025 | 04/23/2025 |  | 02/21/2025 |  |  | 01/27/2025 | 02/21/2025 |  |  | 01/31/2026 |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464939 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when removing the peel away, the impella cp went out of the aortic arch. reinsertion was unsuccessful so another impella cp was placed and the patient remained stable. | [5 empty] | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump pulled out of the left ventricle (lv) upon peeling away sheath. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11524 | 1220648-2025-26253 | 21434121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21434121&pc=OZD | 11/21/2024 | 02/21/2025 |  | 02/21/2025 |  |  | 01/30/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534750 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient (unknown race) that experienced renal failure with a "possible" relationship to the impella cp device used. the patient was admitted for pacemaker removal and replacement due to infection and post pacemaker removal the patient decompensated. ejection fraction went down from 20 to 10%; the patient was intubated and started on multiple drips and scheduled for the impella implant for hemodynamics support, which was successfully completed. additionally, on admission, the came in with acute kidney injury and chronic kidney disease which worsened during the impella mechanical circulatory support (mcs) noted approximately five days after start of the mcs. continuous renal replacement therapy was started to manage the condition. as per the registry, the patient/event has not recovered/not resolved. however, the patient was noted to have survived the explant of the impella cp device at the end of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. if the device or additional information is received, a supplemental mdr will be filed.
11525 | 1220648-2025-26254 | 21434158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21434158&pc=OZD | 01/27/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/27/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585525 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and post implant, same day, the patient experienced limb ischemia. femorofemoral bypass was completed to address the limb ischemia. the patient was reported to be in stable condition following the event. the following day, the patient was planned for transfer to an outside hospital for possible escalation to an impella 5.5 device. it is unknown if the plans were executed as noted. however, it was noted the impella cp was weaned and explanted this day, and the outcome at explant reflected the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11526 | 1220648-2025-26255 | 21434321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21434321&pc=OZD | 01/27/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/27/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591558 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was unable to flow due to debris picked up during implant. the impella cp was explanted and the procedure outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
11527 | 1220648-2025-26256 | 21434388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21434388&pc=OZD | 01/30/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/30/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603531 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and experienced ventricular tachycardia requiring cardioversion and a stroke. the patient would later expire. the impella was placed, and the patient was sent to the unit on support. three days later, the patient experienced a ventricular tachycardia run that was resolved with a shock. an echocardiogram was ordered, and the device was repositioned, pulled back one centimeter. approximately six days later the patient developed "severe" left side facial droop. the patient had experienced a stroke noted on a computed tomography scan. the patient was moved to comfort care, and four days later the patient expired. the family withdrew care; the pump was turned down to performance level p-2, and the patient passed away. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11528 | 1220648-2025-26257 | 21434513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21434513&pc=OZD | 01/27/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/27/2025 | 02/21/2025 |  |  | 08/31/2025 |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385082 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient became unresponsive during support and was found to have multiple severe subarachnoid hemorrhages of unknown etiology. impella support was discontinued in response to the condition and the patient passed away. it is unknown if support contributed to the noted condition and death. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11529 | 1220648-2025-26258 | 21434723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21434723&pc=OZD | 01/30/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/30/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539859 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and following the device was inadvertently pulled back across the aortic valve. as a result of the positioning issue, the pump was fully removed, and the patient had to be re-wired for reinsertion. there was no patient harm, and the pump was placed successfully for support which was maintained on performance level p-6 throughout the pci procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11530 | 1220648-2024-06959 | 21434990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21434990&pc=OZD | 09/23/2022 | 02/21/2025 |  | 02/21/2025 |  |  | 09/23/2022 | 02/21/2025 |  |  | 06/30/2024 |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238390 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cpwas placed in a 74-year-old female patient for mechanical circulatory support. oozing at the impella access was noted. dressing was taken down and the reposition sheath was re-sutured with forward tension, angel matched and was redressed. no new oozing was noted. heparin was withheld after bleeding was again noted from the access sites. it was further reported that care was withdrawn and the patient subsequently expired at explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as a part of the retrospective review of historical records.
11531 | 1220648-2025-26259 | 21435217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21435217&pc=OZD | 01/27/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/27/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534513 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced positioning issues. the positioning issue occurred the evening of the implant day into the next morning. the patient had gone to the operating room for a coronary artery bypass graft (cabg) surgery. placement was confirmed on transesophageal echocardiogram before leaving the operating room. later that night into the early hours of the next morning, the pump migrated into the aorta. the intensivist and the interventional cardiologist discussed plans. the patient was hemodynamically stable, and the plan was to remove the pump the next morning anyway. there were no perclose in place. therefore, a vascular consult was called for removal and repair. the physician pulled the impella back into the descending aorta and turned it to support level p-2 for the rest of the night until the patient could go back to the operating room the next morning for the pump removal. it was stated there were no problems, adverse events, nor was there any harm to the patient. the pump was successfully removed in the operating room with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11532 | 1220648-2025-26260 | 21435322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21435322&pc=OZD | 01/28/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601377 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a hematoma with bleeding at the access site. two units of blood products were given to the patient. additionally, the staff was unable to find pulses in the patient's right foot. arterial ultrasound was completed and showed positive blood flow at dorsalis pedis. however, there was no blood flow at the posterior tibial. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11533 | 1220648-2025-26261 | 21435664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21435664&pc=OZD | 01/08/2025 | 05/13/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1724721 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during impella 5.5 mechanical circulatory support to a patient with an acute myocardial infarction/cardiogenic shock, placement signal was loss overnight. audio alarms were silenced, and the patient was asleep and on support level p-4 at such time. it was noted that the flow appeared more variable at such point. abiomed's investigation engineer has now requested the return of the automated impella controller (aic) to further investigate if this issue could have been a result of the aic. there was no report or indication of patient harm regarding this event. | [5 empty] | The investigation of automated impella controller (aic) - loss of opm data has been completed. the console logs from the reported day of event are not consistent with complaint description: there were no ¿prevent retrograde flow¿ alarms on (b)(6) 2024. nevertheless, throughout the case ((b)(6) 2024 to (b)(6) 2025) there were several instances of ¿loss of opm data¿ events and alarm. placement signal issues were not reproduced during testing; however optical signals pattern associated with ¿oscillating contrast¿ and resulting was consistent with momentary loss of light due to optical bench malfunction (either lamp or electronics). real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) were represented as -16384 values due to the crc error. the cause of the aic issue was a software issue. d9 date device returned to manufacturer added. | The automated impella controller device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11534 | 1220648-2025-26263 | 21436420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21436420&pc=OZD | 01/28/2025 | 04/23/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025598914 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had a hematoma. two units of blood products were given. additionally, high purge pressure was noted. tissue plasma activator was added to the purge and support was continued. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of access site bleeding and high purge pressure has been completed. the pump and data log were not returned for investigation. on the 4th day, a hematoma was noted, and the patient was treated with 2 units of blood. also, on the 111th day, high purge pressure was reported and resolved after tissue plasma activator (tpa) was administered. the cause of the high purge pressure issue was most likely biomaterial buildup, based on the given clinical details that tpa administration was able to help resolve the issue. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number.
11535 | 1220648-2025-26262 | 21436537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21436537&pc=OZD | 01/28/2025 | 03/18/2025 |  | 02/21/2025 |  |  | 01/28/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577870 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the provided flows were lower than expected. on fluoroscopic viewing, a kink was discovered in the cannula. the pump was subsequently replaced to resolve the failure. there was no patient harm. | [5 empty] | The investigation of low pump flow has been completed. the returned complaint device visually inspected. cannula kink near outlet observed. no biomaterial, no broken blade observed. the cause of the low pump flow issue was cannula kink due to improper technique. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11536 | 1220648-2025-26264 | 21436755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21436755&pc=OZD | 01/29/2025 | 09/25/2025 |  | 02/21/2025 |  |  | 01/29/2025 | 02/21/2025 |  |  | 01/31/2026 |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458129 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had dark red urine. the pump position was appropriate but the ventricle was small contributing to the hemolysis. the patient was then escalated to an impella 5.5 due to unrelated reasons. the procedural outcome was the patient survived the surgery. | The complainant reported the patient was on impella cp support and during support, there were placement signal issues. staff monitored. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: after device implantation, patient transport to second hospital. second hospital stated placement signal (ps) unreliable alarms occurred with loaner console, despite appropriate positioning confirmed. pump not returned. data log review showed inconsistent 5s parameters from case start. imc logs identified calibration and pump offset issue with cause unknown. confirmed ps unreliable alarms. no known optical failure patterns identified. the cause of the optical signal issue was not determined since product was not returned and lack of clinical details. failure mode placement signal issue was changed to optical signal issue due to cp pump use. hemolysis: after device implantation, patient transport to second hospital. despite ps unreliable alarms, rn stated pump position was appropriate but patients ventricle was small, likely contributing to the hemolysis. pump not returned. data log reviewed showed no motor current spikes outside of p-level changes. no suction alarms. no other relevant abnormalities. the cause of the hemolysis was not determined as no product was returned and lack of clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
11537 | 1220648-2025-26265 | 21437064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21437064&pc=OZD | 01/29/2025 | 03/14/2025 |  | 02/21/2025 |  |  | 01/29/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1152333 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella support when the automated impella controller (aic) had ¿controller error¿ and ¿placement signal not reliable¿ alarms. pulsatile motor current and appropriate flows were maintained. the doctor stated they gently ¿jiggled¿ the console on the stand and noted that the waveforms returned and both alarm conditions registered as resolved. staff monitored and support continued. there was no patient harm. | [5 empty] | The investigation for the console (aic) loss of opm data and controller error-yellow alarm has been completed. data logs were reviewed which found the console displayed the following error: ¿controller error (opm comm crc invalid) [optical pump connected] ¿ alarm,¿ event #1045, which was resolved within 6 minutes. also, there was a psnr alarm for a few seconds, which resolved automatically. the psnr alarm was likely caused by event #1045. after 3 hours and 40 minutes, event #1045 reoccurred and was resolved in 7 minutes. after 7 hours, there was a final occurrence of event #1045, after which support continued for an additional 6 days. this confirms the reported failure mode. during each event #1045 occurrence, the rt logs confirmed that all signals received from the optical bench (snr, contrast, lamp, gain) were recorded as -16384 values due to the crc error. the console was returned for evaluation. the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11538 | 1220648-2025-26266 | 21437503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21437503&pc=OZD | 12/19/2024 | 10/31/2025 |  | 02/21/2025 |  |  | 01/31/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1631552 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and as a bridge to transplant. orthotopic heart transplant or left ventricular assist device was to be decided. seven days after start of the support, inconsistent changes in waveforms not relating to monitor aortic pressure were noted. the diastolic flows significantly decreased intermittently. the physician followed up with biomedical. the automated impella controller (aic) displayed the impella position unknown warning when issue occurred. it was recommended the physician complete an echocardiogram when issue occurs to see where the pump is positioned. the impella 5.5 device remains implanted supporting the patient until the bridge to transplant. abiomed's investigation engineer has now requested the return of the aic to further investigate if this issue could have been a result of the aic. there is speculation of a defect on the charge-coupled device array. | During investigation of 5.5, the engineer noted as the pressure shifts, the signal-to-noise ratio lowers drastically, which could indicate an optical bench with localized defect on charge-coupled device array on the automated impella controller (aic) and requested that it be returned for further evaluation. |  |  |  |  | The automated impella controller device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show 300+ impella position unknown warnings (#35 due to suction and #33 due to low pulpability. there¿s also one isolated occurrence of momentary loss of optical data (event 1045), which promptly self resolved. ltlog and rtlog data show periods of erratic placement signal, however, without any decrease in snr. investigation 24-39923-1 addresses any possible relation to the pump or device use. it is undetermined whether low optical power contributed to placement signal changes. prior to the investigation, console ic9457 arrived for pm procedure, where its optical bench was replaced due to optical power out-of-spec. after successfully passing all pm testing, console was sent back to the customer. the cause of the issue was not determined since issue was not reproduced. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal description. h6 updated. d10 concomitant medical products and therapy dates updated.
11539 | 1220648-2025-26267 | 21437790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21437790&pc=OZD | 01/29/2025 | 02/21/2025 |  | 02/21/2025 |  |  | 01/29/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597208 |  | 25-41050-1 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support. while on support, the patient's leg did not have distal pulses. the patient later had pedal dopplerable pulses in left leg where impella is. however, the patient still did not have pulses in right leg where impella is not, per nurse. the impella cp was replaced with the impella 5.5 and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
11540 | 1220648-2025-26268 | 21439210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21439210&pc=OZD | 01/29/2025 | 03/17/2025 |  | 02/21/2025 |  |  | 01/29/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607851 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during insertion of an introducer for impella cp implant, the device perforated the femoral artery. manual pressure was held and vascular surgery was consulted to manage the condition. the right femoral artery was ballooned to treat the perforation. the patient was stable following the intervention and the implant was aborted. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. during insertion of the introducer sheath and dilator into the femoral artery wirelessly, the artery was perforated. the artery was reported to be very calcified and resistance was met during insertion. no product was returned. the pump was never inserted. the cause of the vascular damage - peripheral vessel was most likely use related due to the clinical details provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26268 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿.
11541 | 1220648-2025-26269 | 21439504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21439504&pc=OZD | 01/29/2025 | 04/14/2025 |  | 02/21/2025 |  |  | 01/29/2025 | 02/21/2025 |  |  |  |  | 02/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1676151 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller had a controller failure alarm. this event occurred during commercial use and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) ¿ controller failure (red alarm) has been completed. according to the data and device analysis the root cause for controller error was defective optical bench. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26269 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 component code 925 was erroneously entered on the initial manufacturer device report number 1220648-2025-26269. h8 usage of device selection was erroneously entered on the initial manufacturer device report number 1220648-2025-26269. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11542 | 1220648-2025-26211 | 21420015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21420015&pc=OZD | 01/26/2025 | 03/05/2025 |  | 02/20/2025 |  |  | 01/26/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599703 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had a distal perfusion sheath that clotted off so there were no distal pulses, vascular came and fixed it. there are positive distal pulses after vascular intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
11543 | 1220648-2025-26214 | 21420300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21420300&pc=OZD | 01/25/2025 | 03/28/2025 |  | 02/20/2025 |  |  | 01/28/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575275 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and approximately 15 days after start of support the pump stopped. further information provided noted there was no sign that the pump was about to stop; it suddenly stopped. an attempt was made to restart the pump according to the manual; however, the pump did not restart. consequently, the impella cp device was explanted. the device was inspected and no blood clots and other adhesion with found. outcome at end of unit support noted an additional impella device was needed; however, implant of another device is unknown. there was no patient harm as a result of the pump stop. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of pump stop has been completed. no product/logs were returned. the pump stop occurred on the fifteenth day of support. the root cause of the pump stop could not be determined since no product/logs were returned. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e.1 added address line 2 as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26214. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26214 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26214. h.10 added instructions for use as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26214.
11544 | 1220648-2025-26215 | 21420301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21420301&pc=OZD | 04/24/2024 | 09/08/2025 |  | 02/20/2025 |  |  | 01/28/2025 | 02/20/2025 |  |  | 12/31/2024 |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024305879 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hemothorax | Pleural Effusion | Ventricular Fibrillation | Cardiac Tamponade | Thrombosis/Thrombus | Intraoperative Cardiac Injury |  |  |  |  | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported via the impact eu study a patient experienced a mediastinal hemorrhage with a definite relationship to impella device. the adverse event narrative explained the patient was going to be discharged the day of the event (which was four days after explant of the impella 5.5 device; the patient recovered to native heart function). the patient was unresponsive after the second walk to the room on the unit. the patient's pulse dropped, and code blue was activated. cardiopulmonary was administered and return of spontaneous circulation achieved. there was a significant growing left pleural effusion. left chest tube was placed with return of 1500ml of blood. the patient received 4 units of packed red blood cells, 1 fresh frozen plasma and 1 platelet. arterial blood gas revealed acidosis, 4 amps. bicarbonate was given. the patient was taken to the operating room for mediastinal washout. small bleeding from the side of the right coronary button was found and stitch was placed. the patient was noted to have survived the event. | Additionally, the patient experienced ventricular fibrillation during code on (b)(6) 2024, leading to cardiac arrest. around 12:45, patient was walking, felt light-headed, had pre-syncope, became hypotensive, and code blue was called due to suspected cardiac arrest. 12:53, second code blue called, patient unresponsive, pea noted on monitor. cpr started. ns 1000 ml free flow started, levo drip started, patient intubated, central and arterial lines placed. echo and tee done. 13:02; rosc achieved, cpr stopped. amiodarone started at 1 mg and discontinued coreg. the patient had a large left hemothorax noted. the patient underwent emergency mediastinal re-exploration. tee was used to diagnose the presence of tamponade, evacuation of clots, and changes in ventricular or valvular function during the surgery. during the procedure, the patient had a repair of the aortic hole with a pericardial patch and chest tube for hemothorax. per the principal investigator (pi), cardiac arrest with the ventricular fibrillation possible related to the impella procedure and the impella device. the left hemothorax(bleeding) definite related to the impella device and the impella procedure. |  |  |  |  | The investigation of vf/vt/af/at and access site bleeding has been completed. the impella device was not returned for evaluation. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. access site bleeding: the cause of the access site bleeding - major was not determined since product was not returned and insufficient clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vascular damage - great vessel has been completed. patient had 5.5 implanted and explanted without issue, with successful weaned from support. four days after explant, patient experienced mediastinal hemorrhage. during or intervention for mediastinal washout, with small bleeding from the side of the right coronary button. no product returned. data logs returned but not applicable to this failure mode. the cause of the vascular damage - great vessel was not determined since product was not returned and lack of clinical details. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26215 is it is unknown if the initial reporter also sent the report to the food and drug administration.
11545 | 1220648-2025-26216 | 21420330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21420330&pc=OZD | 01/28/2025 | 03/17/2025 |  | 02/20/2025 |  |  | 01/28/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599093 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient presenting with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced bleeding, a hematoma and low pump flow. the impella was placed without complications. extracorporeal membrane oxygenation (ecmo) was placed as well. percutaneous coronary intervention was performed to the distal left main and proximal left anterior descending and right coronary arteries. there was bleeding from the first 14fr 23cm sheath used; therefore, the physician switched to a short sheath which at the time of discussion was hemostatic but due to the previous bleeding the physician did not want to peel it away. the limitations of leaving the peel away sheath in was explained, and the decision made to peel away with close attention to technique. the peel-away sheath was removed, and the repositioning sheath advanced with no bleeding at the groin access site. the patient was sent to the unit on impella mcs. later overnight the patient received multiple blood products for suction and chugging on the ecmo circuit. three units of packed red blood cells and 2l albumin were given. further noted, the patient had a large, firm neck hematoma, and bruising around the abdomen that were suspicious for bleeding sources. the patient continued on supported noted to be in stable condition. the impella cp device supported the patient for a total of three days until successfully escalated to the impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation for the access site bleed and low pump flow has been completed. data logs confirmed the failure mode and clinical narrative. logs showed after pump started, low flow occurred hat triggered auto suction response & suction alarms. this event remained for 6.5 hours then was resolved. pump then ran without issue to the rest of the case. pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. the root cause of the low flow was most likely a patient condition as no issue was found on the pump and low flow was resolved after fluid compensation. the root cause of access site bleeding was most likely patient movement related as clinical team reported impella movement during transport causing minor hematoma to site. introducer was not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. added d9 device return date.
11546 | 1220648-2025-26217 | 21420455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21420455&pc=OZD | 01/28/2025 | 03/14/2025 |  | 02/20/2025 |  |  | 01/28/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595517 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention (pci) was unsuccessful. while attempting to pass the impella across bioprosthetic aortic valve using a buddywire, the physician elected to remove device because ¿it feels wrong and won't torque around the aortic arch. something is not right with this device; i can feel it.¿ three attempts were made to cross aortic valve without success. a new impella cp device was obtained and successfully implanted for the protected pci procedure. there was no patient harm as result of this event. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to aortic valve stenosis and aortic arch calcification. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11547 | 1220648-2025-26218 | 21420546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21420546&pc=OZD | 01/28/2025 | 03/17/2025 |  | 02/20/2025 |  |  | 01/28/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597951 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device and then successfully escalated to an impella 5.5 device for continued mechanical circulatory support (mcs). the patient was sent back to the unit vented and sedated. update post impella 5.5 implant noted insertion site with 2 jackson-pratt drains were in place and oozing was noted; no hematoma noted. heparin was withheld. the patient was on levophed, vasopressin, amiodarone, lidocaine, and milrinone, and a plan to restart continuous renal replacement therapy overnight and rest the heart. creatinine was 3.7 and urine output was a dark tea color less than 30ml/h. position was checked, and performance level was p-9 with flows at 5.3l/min. the groin site from impella cp removal was check and noted to be appropriate and without bleeding/hematoma. peripheral pulses all intact. the following day the patient aroused and followed commands, there was minimal drainage and no hematoma. heparin was restarted. the patient continued and remains on impella mcs. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. the data logs showed that there were multiple instances of suction positioning alarms triggered during the case. there was no motor current spike or trend in the placement signal during support. due to lack of product returned and limited clinical details, the root cause of the hemolysis could not be determined. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. d.6b explant date was added. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26218. h.10 revised instructions for use as they were reported incorrectly from the initial manufacturer device report number 1220648-2025-26218.
11548 | 1220648-2025-26219 | 21420715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21420715&pc=OZD | 01/29/2025 | 03/11/2025 |  | 02/20/2025 |  |  | 01/29/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1208378 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported at the start of a case the staff started to prime the impella cp pump. when the impella plugged was inserted into the into the aic, priming stopped. the pump was never fully primed. the aic was exchanged was the team was able to successfully prime the impella cp pump and support the patient. there was no harm to the patient as a result of this event. abiomed clinical consultant further reported later that afternoon before return of the console, demo equipment was used with the aic and when the impella plug was plugged into the aic, it alarmed "controller error." the unit was returned for evaluation and testing. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of priming issue and automated impella controller (aic) - controller error (yellow alarm) has been completed. according to that data and device analysis the root cause for the controller error (defective optical sensor lamp) was most likely due to the defective optical bench or the lamp assembly the root cause of priming issue was not related to any problems with the aic. no issues were found within the aic. h6 was revised to reflect the failure mode priming issue as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-26219.
11549 | 1220648-2025-26191 | 21420752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21420752&pc=OZD | 01/29/2025 | 03/17/2025 |  | 02/20/2025 |  |  | 01/29/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603535 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient with acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support but encountered placement signal issue, low flow and/or pump stop during three attempts with three impella cp devices. the patient was admitted and found to have received three shocks from the automated implantable cardioverter device for ventricular tachycardia (vt) and atrial flutter (af). the patient was taken to the electrophysiology lab for vt/af ablation and during this procedure, the patient coded received one shock with return of spontaneous circulation achieved. the decision was made to place and impella cp device. the impella cp (pump 1 sn(b)(6)) was inserted and initial flows of 2.6l were noted with no placement signal available. the elected to replace the impella cp device (pump 1) in order to have a placement signal. new impella cp device (pump 2 sn(b)(6)) was primed on different automated impella controller and was placed with no issues on insertion. pump 2 initial flows of 3.6l were noted and were being secured when placement signal suddenly disappeared, and flows started to decrease. the placement signal not reliable alarm triggered and then flows stopped. therefore, the physician decided yet again to replace the impella cp device (pump 2). another impella cp device (pump 3 sn(b)(6)) was primed on a different automated impella controller and on startup there was immediate suction noted left ventricular placement signal was abnormal and unable to obtain adequate flows. all troubleshooting performed, and the team was advised of possible corevalve transcatheter aortic valve replacement interaction. the decision made to remove the impella cp (pump 3) for a final time, and the physician elected to then place an intra-aortic balloon pump. additional information noted the following confirmations: adequate anticoagulation status was verified, transesophageal echocardiogram was performed, with no left ventricle thrombus, all sheaths were excessively flushed, and there was no clot noted upon removal of all pumps. all three devices are expected to be returned for analysis. the was no harm and the patient was reported to be stable following the event. | [5 empty] | The investigation of optical signal issue has been completed. signal not reliable (snr) was nearly zero and contrast was below specifications, but the pump passed the laser continuity test. the sensor was removed and evaluated. the sensor face was damaged and appeared to be peeled up. data logs show the sensor operating without issue until just after the pump was started. multiple motor current spikes were seen upon startup. the root cause of the optical signal issue was most likely a damaged sensor resulting from interaction with the transcatheter aortic valve replacement (tavr) mentioned in clinical details. ¿¿low pump flow was added as an failure mode (fm) since the clinical details state that the initial pump flows were 2.6 l/min and the pump was started at p-9 based on the data logs. during inspection of the returned pump, the impeller blades were found to be damaged with one blade bent in the corner and the other blade missing a large piece. the pump was run in a bucket of water and maintained low flows (~2.7 l/min at p-9) for 30 minutes. data logs show pump flow to be a stable ~2.7 l/min upon startup at p-9. the pump was run for 10 minutes before being explanted. the root cause of the low pump flow was most likely a damaged impeller resulting from interaction with the tavr mentioned in clinical details. d10 revised to reflect transcatheter aortic valve replacement (tavr) interfered with the impella device on manufacturer device report number e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of three device manufacturing reports associated with this event. refer to the following: medical device report for impella cp pump 1 sn(b)(6) with date of event 29-jan-2025 and patient identifier ending in (b)(6) (this report). medical device report for impella cp pump 2 sn(b)(6) with date of event 29-jan-2025 and patient identifier ending in (b)(6). medical device report for impella cp pump 3 sn(b)(6) with date of event 29-jan-2025 and patient identifier ending in (b)(6).
11550 | 1220648-2025-26220 | 21421269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21421269&pc=OZD | 01/24/2025 | 03/23/2025 |  | 02/20/2025 |  |  | 01/24/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 79 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Information was obtained from a literature, titled experience with impella supported off-pump coronary artery bypass grafting for acute coronary syndrome. it was reported that a nstemi patient with 3-vessel disease and heart failure with ejection fraction at 40%, so an impella cp was inserted. after impella was removed a surgical site infection was observed at the impella insertion site. treatment of the infection is unknown and the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the full literature citation and english translated article were unable to be obtain at the time of writing this manufacture report. all literature information provided has been noted in this report. a.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11551 | 1220648-2025-26221 | 21421680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21421680&pc=OZD | 01/27/2025 | 03/19/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573070 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump could not be advanced through the patient's axillary artery due to the narrow passage. the implant was aborted and an impella cp was subsequently placed for patient support. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through a narrowing in the course of the axillary artery. g.1 added reporting contact address line 1, reporting contact us city, reporting contact state, reporting contact us zip code and reporting contact country as they were indavertently omitted from the initial manufacturer device report number 1220648-2025-26221. revised manufacturing site name and address as they were previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26221. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11552 | 1220648-2025-26222 | 21421908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21421908&pc=OZD | 01/27/2025 | 03/14/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135089 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported by biomed that the console with purge system was not priming. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of priming issue has been completed. during device analysis the console performed as expected and had no priming issues similar to the console's performance in the field. during the data log review was unable to confirm the reported issue. root cause: there was no issue found during the case. d4 unique identifier number was omitted on the initial manufacturer device report number 1220648-2025-26222. e2 was inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26222. e3 was inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26222. h3 was inadvertently selected as yes, when at the time of the initial manufacturer device report number 1220648-2025-26222, the product had been received. h8 usage of device was inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26222.
11553 | 1220648-2025-26223 | 21422170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21422170&pc=OZD | 01/27/2025 | 03/17/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593301 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) and encountered low pump flow (suction alarms) requiring blood transfusion one day after start of support and thereafter experienced device malpositioning. during the basic angiogram, the patient quickly decompensated with hypotension and shortness of breath. dopamine was started via intravenous piggyback. an impella cp was decided and initiated for mcs. one day after start of support, low pump flow was encountered. the patient received a unit of blood which improved pressures and resolved some suction alarms. impella mcs continued. the following day, the nurse reported the patient accidentally completely pulled the impella out of the left ventricle with the pump in the descending aorta confirmed on echocardiogram. it was noted that the patient was extubated earlier this day and became increasingly confused overtime which led to the patient moving around a lot and then accidentally pulling out the pump from its position. the physician noted as the patient was hemodynamically stable with an ejection fraction now of 45%, the device would not be repositioned. it would remain in the aorta on surgical mode with removal later in the day morning. the device was successfully weaned and explanted. following the explant, the patient was noted to have survived the procedure and was in stable condition. | [5 empty] | The investigation for the low pump flow and the positioning issue has been completed. no product returned. reported patient accidentally pulling the pump out of the left ventricle into the aorta. the root cause of the positioning issue was most likely patient movement since the patient pulled the pump. clinical reported suction alarms that were resolved by receiving a unit of blood. data logs show suction alarms throughout the case and slightly lower flows. p-level was being adjusted throughout, staying around p-4 to p-6. there were no motor current spikes outside of p-level changes. the root cause of the low pump flow was most likely patient condition related since receiving volume reduced suction alarms. corrections to the initial mfr report: h5 should have been yes. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11554 | 1220648-2025-26224 | 21422266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21422266&pc=OZD | 01/24/2025 | 03/21/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554010 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support the peel away sheath was replaced with the repositioning sheath and bleeding occurred. the bleeding stopped initially, but the site began bleeding again the following day. the impella cp was explanted and the patient remained stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
11555 | 1220648-2025-26225 | 21422515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21422515&pc=OZD | 01/28/2025 | 03/07/2025 |  | 02/20/2025 |  |  | 01/28/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599862 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) as a bridge to transplant. five days after start of support, an ultrasound of the right upper extremity confirmed deep vein thrombosis to the right basilic vein and right internal jugular. systemic heparin was increased. two days following a hematoma of the access site was noted. treatment to resolve is unknown. reportedly, the patient continued moving the arm against medical advice. impella mcs continue for another 21 days until the impella 5.5 device was weaned down and explanted, and the patient was successfully bridged to heart transplant. the patient was noted to be hemodynamically stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported thrombus and hematoma has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus and hematoma could not be determined as the device was not returned nor sufficient clinical details.
11556 | 1220648-2025-26226 | 21422587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21422587&pc=OZD | 01/27/2025 | 03/14/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520564 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient's labs hemolyzed during impella cp support. the patient was given blood products to treat the condition, but was later reported to have expired. it is unknown if the reported condition contributed to the patient's passing. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. h.1 revised type of reportable event as it was previously selected incorrectly on the initial manufacturer device report number 1220648-2025-26226. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11557 | 1220648-2025-26227 | 21422882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21422882&pc=OZD | 01/29/2025 | 03/07/2025 |  | 02/20/2025 |  |  | 01/29/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578977 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced limb ischemia. the patient presented with chest pain and was a surgical turndown. while intervening on one of the vessels, the patient went into ventricular fibrillation arrest. cardiopulmonary resuscitation was initiated and return of spontaneous circulation was achieved. a temporary pacer wire was implanted and requiring levophed, epinephrine and dobutamine. a decision was made to implant the impella cp device which was successfully completed. after the insertion of the impella cp, the physician decided to leave in the peel-away sheath. contrast was injected to view and no flow down the leg was seen. consequently, a distal perfusion catheter was placed with a successful outcome. the patient was sent to the unit on impella mcs in stable condition. the device was successfully weaned and explanted the following day. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
11558 | 1220648-2025-26228 | 21423175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21423175&pc=OZD | 01/29/2025 | 03/20/2025 |  | 02/20/2025 |  |  | 01/29/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1838437 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) had a controller failure (red alarm) while connected to an impella 5.5 device providing mechanical circulatory support to a patient with acute myocardial infarction/cardiogenic shock. it appears the aic was not exchanged as the report indicated the outcome of the event was successful with this device. reportedly, there was no harm to the patient as a result of this event. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the aic controller failure -red alarm has been completed. data logs were reviewed and after 20 hours since the pump started, the log recorded a continuous pud communication issue error. due to the persistent pud communication issue there was pud reset. after a couple of pud resets, the console displayed ¿controller failure (pud comm loss) ¿ alarm,¿ event #1020. this confirms the reported failure mode. this console was returned for evaluation. upon initial bootup, the console displayed a ¿system self-check failed¿ screen. the fw check via telnet showed that the pud hardware revision was n.a. it was confirmed that the pud firmware (fw) was corrupted by comparing it with the production pud firmware. the root cause of the controller error was the corrupted pud firmware (fw). the cause of the corrupted pud fw was not determined.
11559 | 1220648-2025-26229 | 21423229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21423229&pc=OZD | 01/29/2025 | 03/07/2025 |  | 02/20/2025 |  |  | 01/29/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127440 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility biomedical engineer reported an automated impella controller (aic) displayed a system self check error at start up. the event was noted to have occurred in use in the catheterization lab but not related to a clinical procedure. there was no patient involvement. additionally, there was no report or indication of delay in therapy support as a result of this event. | [5 empty] | The investigation for the console (aic) self check issue has been completed. data logs were reviewed. during initial boot up, the console restarted automatically due to a pmd communication issue. due to a persistent pmd communication issue, the console rebooted several times and displayed ¿system self check failed.¿ upon manual shutdown and reboot, the console rebooted again due to the pmd communication issue and displayed ¿system self check failed¿. this confirms the reported failure mode. console was returned for evaluation and during testing, reproduced the reported failure mode. during the initial boot-up, the console restarted automatically due to the pmd communication issue, resulting in the ¿system self check failed¿ screen. the fwcheck via telnet confirmed the problem was due to the impellatronic pca board, as the hw revision was n.a. for the sau_motor_1. it was identified that the firmware was corrupted upon comparing the extracted firmware with a known good. the root cause for the ¿system self check failure¿ was corrupted firmware on the impellatronic pca board. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11560 | 1220648-2025-26230 | 21425788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21425788&pc=OZD | 01/27/2025 | 03/07/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594181 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the implant was unsuccessful due to tortuosity and calcification. decision was made to abort 5.5 placement and instead place an impella cp. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue was determined to be patient condition because of tortuosity and calcification. the decision was made by the medical doctor to abort the impella 5.5 and place and impella cp and the impella cp was placed successfully. b.7 added information was that inadvertently omitted from the initial manufacturer device report number 1220648-2025-26230.
11561 | 1220648-2025-26231 | 21426238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426238&pc=OZD | 01/27/2025 | 03/17/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590638 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, the pump locked in place but positioned into aorta and was removed and replaced. additionally, oozing was noted to be present at the access site; however, the angle was not being maintained due to fem stop. the doctor wanted fem stop removed. the nurse was educated on angle matching and forward tension sutures when removed to assess if oozing present was due to access versus angle. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and positioning issues has been completed. product was not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review.
11562 | 1220648-2025-26232 | 21426370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426370&pc=OZD | 01/24/2025 | 03/14/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2025593782 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the device kinked while it was in use. as a result, the decision was made to remove the impella cp and exchange it for a different cp pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (cannula kink) has been completed. the impella device was not returned for evaluation. product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition.change failure mode (fm) to delivery issue as the cannula kink is the result of delivery failure. e4 should have been left blank on the initial manufacturer device report number (b)(4) as it is unknown if the initial reporter also sent the report to the food and drug administration.
11563 | 1220648-2025-26233 | 21426379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426379&pc=OZD | 01/27/2025 | 03/25/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589361 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp placement was not possible due to kinking in the catheter. it was noted that resistance had been met in the innominate artery which resulted in the observed kinking. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the implant of the pump was aborted due to the patient's anatomy. e.1 revised address line 1 and city as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26233. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26233 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 added a medical device problem code as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26233.
11564 | 1220648-2025-26234 | 21426380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426380&pc=OZD | 01/27/2025 | 03/17/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583036 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient had hemolysis and acute renal failure. the patient suffered shock liver and acute renal failure with increasing pressor requirements and persistent lactic acidosis. hemolysis was also noted. device was repositioned as it was initially deep. however, it was unclear if hemolysis was resolving or not given anuria, but lactate dehydrogenase was trending up. the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26234 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11565 | 1220648-2025-26235 | 21426393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426393&pc=OZD | 01/27/2025 | 03/07/2025 |  | 02/20/2025 |  |  | 01/27/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539652 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the doctor was unable to advance an impella 5.5 pump during delivery due to a calcified axillary artery. after multiple failed attempts, the implant was aborted. there was no patient harm reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related. multiple attempts to advance impella through but aborted insertion due to patient anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26235 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 code 4611 was reported incorrectly on manufacturer device report 1220648-2025-26235. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11566 | 1220648-2025-26236 | 21426400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426400&pc=PYX | 01/28/2025 | 03/17/2025 |  | 02/20/2025 |  |  | 01/28/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025585352 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella rp flex and that they encountered low purge flows. as a result, tissue plasminogen activator was started to resolve the purge flow issue. the patient was stable. | [5 empty] | The investigation of high purge pressure (hpp) has been completed. the impella device was not returned for evaluation based on the clinical details the cause of the issue was most likely use related since a 4 minute bag change exceeding the 2 minute change recommendation was observed in the data logs prior to the hpp event which was resolved with tissue plasminogen activator (tpa). e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26236 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11567 | 1220648-2025-26237 | 21426405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426405&pc=OZD | 01/28/2025 | 03/18/2025 |  | 02/20/2025 |  |  | 01/28/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1881868 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) has a loose port were the catheter plug goes in. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - damage before therapy has been completed. there was no evidence in logs for this complaint. while running a test pump, though there was no difficulties with the electrical connections, a placement signal not reliable alarm was received and aic logs showed a 1036 error. real time logs confirmed raw optical sensor dropped to -3276.8 at this time. this is indicative that there was an air gap and the pump was noted to be a bit loose in the blue receptacle. the root cause of the reported issue was a defective optical connector that was loose resulting in optical placement signal issues during testing. e4 should have been left blank on the initial manufacturer device report number (b)(4) as it is unknown if the initial reporter also sent the report to the food and drug administration.
11568 | 1220648-2025-26239 | 21426528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426528&pc=OZD | 01/29/2025 | 03/14/2025 |  | 02/20/2025 |  |  | 01/29/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1629362 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported during preparation for an impella implant, a controller error appeared on the automated impella controller (aic). the aic was replaced in response to the failure. there was no report or indication of patient harm or delay in therapy. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) controller error-yellow alarm has been completed. console logs from the reported day of event are consistent with complaint and show an controller error alarm (opm comm crc error ¿ event set # 1045) and controller error (optical bench ambient pressure out of range) [optical pump connected] - alarm issued triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the console was returned for evaluation. voltage from the pbm at connector (j3) was 4.95v. the voltage on the optical bench lamp connector (p4) was 4.87v, and the lamp power was 3.66uw. visible light was detected from the optical pump connector in the console, and "5s" parameter testing confirmed that the optical system was within specification. the controller error was most likely due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data) and controller error (optical bench ambient pressure out of range) [optical pump connected] - alarm issued" was most likely to be a firmware issue.added h6 impact code 4629
11569 | 1220648-2025-26240 | 21426538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21426538&pc=OZD | 01/30/2025 | 03/13/2025 |  | 02/20/2025 |  |  | 01/30/2025 | 02/20/2025 |  |  |  |  | 02/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167416 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported automated impella controller failed to recognize an installed purge cassette. there was no report of patient involvement at the time of the event. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. according the device and data analysis the root cause of this issue could not be determined due to the inability to reproduce. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26240 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
11570 | 1220648-2025-26188 | 21409675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409675&pc=OZD | 01/25/2025 | 03/11/2025 |  | 02/19/2025 |  |  | 01/25/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571362 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a patient admitted in with cardiogenic shock. the pump was supporting when after two days there were signs of hemolysis. the labs hemolyzed, urine color had changed, and pfhg levels were high. the pump was explanted the following day and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11571 | 1220648-2025-26169 | 21409812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409812&pc=OZD | 01/25/2025 | 03/27/2025 |  | 02/19/2025 |  |  | 01/25/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570744 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient admitted in acute myocardial infarction/cardiogenic shock. the pump was successfully implanted. while removing the peel away sheath, the pump dislocated and had to be repositioned again. there was no harm to the patient. | [5 empty] | The investigation of positioning issues has been completed. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related because the pump was dislocated while user removing the peel away sheath. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26169 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11572 | 1220648-2025-26189 | 21409830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409830&pc=OZD | 01/26/2025 | 03/07/2025 |  | 02/19/2025 |  |  | 01/26/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589215 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient was brought to the operating room for hematoma evacuation at the impella site. the patient remained stable. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11573 | 1220648-2025-26190 | 21409862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409862&pc=OZD | 01/24/2025 | 04/04/2025 |  | 02/19/2025 |  |  | 01/24/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592154 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex and impella cp for mechanical circulatory support (mcs). initially, the impella cp was implanted and during the percutaneous coronary intervention, the patient began to have frequent suction alarms. right heart catheterization was performed and a pulmonary artery pulsatility index was calculated of 0.5. the decision was then made to place an impella rp flex for bipella support. within approximately three to four hours after returning to unit, the patient began to lose pulsatility on the impella cp. mean arterial pressures were sustained in the 40¿s and 50¿s despite current support and vasopressor medications. the impella cp was repositioned at bedside which appeared to be stuck in the papillary muscle. flows were diminished at 2l on performance level p-8. impella rp flex flows reduced to avoid pulmonary edema. decision was made to transport patient to catheterization lab to cannulate for venoarterial extracorporeal membrane oxygenation (va ecmo). after arriving, the patient went into cardiac arrest. cardiopulmonary resuscitation was initiated. the impella cp was removed and left femoral artery cannulated for ecmo. the left femoral vein was also cannulated. the impella rp flex was being explanted at time of ecmo cannulation and initiation. at time of impella rp flex explant, air was induced into ecmo venous line, which air locked and stopped ecmo circuit. no pulse was assessed. decision was made to cease cardiopulmonary and resuscitation efforts and the patient expired. | The complainant reported that the impella cp was repositioned at bedside which appeared to be stuck in the papillary muscle. flows were diminished at 2l on p8. medical safety review noted bipella support. impella cp repositioned appeared to be stuck in the papillary muscle. flows were diminished at 2l on p8 and is product malfunction. the demise outcome is associated with the impella rp flex device in related manufacturer device report number 1220648-2025-26163. |  |  |  |  | The investigation of low pump flow has been completed. the product and data were not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26190 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two device manufacturing reports associated with the patient implant experience. refer to initial medical device report for the impella rp flex serial number (b)(6) with date of event 25-jan-2025 and patient identifier ending in (b)(6). | The complainant reported that the impella cp was repositioned at bedside which appeared to be stuck in the papillary muscle. flows were diminished at 2l on p8. medical safety review noted bipella support. impella cp repositioned appeared to be stuck in the papillary muscle. flows were diminished at 2l on p8 and is product malfunction. the demise outcome is associated with the impella rp flex device in related manufacturer device report number 1220648-2025-26163.
11574 | 1220648-2025-26192 | 21409866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409866&pc=OZD | 01/27/2025 | 03/10/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604718 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was being implanted with an impella 5.5 device and the delivery was unsuccessful due to a patient's anatomy. the implant was subsequently aborted with no noted harm to the patient. no further details on subsequent treatment were provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a/d6b revised to reflect the implant and explant dates that were not included on the initial manufacturer device report number 1220648-2025-26192. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26192 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 revised to reflect the instructions for use that were not included on the initial manufacturer device report number 1220648-2025-26192.
11575 | 1220648-2025-26193 | 21409877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409877&pc=OZD | 01/27/2025 | 03/03/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522358 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and within hours the patient developed pulsatile bleeding from the left groin site. consequently, the impella 5.5 was explanted, and the patient was noted to have survived the explant reportedly in stable condition. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11576 | 1220648-2025-26194 | 21409880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409880&pc=OZD | 01/27/2025 | 03/12/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552313 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) and was implanted with an impella cp device for mechanical circulatory support. the catheterization revealed multi-vessel disease: heavily calcified lesions to proximal circumflex, proximal left anterior descending and distal left main arteries. the impella was successfully inserted. however, de-coupled waveforms on the automated impella controller (aic) were displayed for which repositioning was attempted. during repositioning, impella in aorta alarms appeared, and team was unable to advance the pump back across the valve. additionally, the team noted a kink in the cannula above the inlet cage. the team was unable to remove kink and unable to remove the impella device from the sheath. a snare was inserted through femoral sheath contralaterally, and it was snared around pigtail portion of impella cp, which was then removed from peel-away sheath, while straightening kink with tension contralaterally. another impella cp was successfully placed for the high-risk pci. there were no adverse effects to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. the pump was not returned for investigation. the data log showed low pump flow after the case started, accompanied by an active suction alarm. the pump was not able get proper flow on p2. according to the clinical notes, the device was found in the aorta during reposition the device. the pump was unable to advance back across the valve due to a kink in the cannula above the inlet cage. the impella was removed, and a second device was inserted without issue. the reported cannula kink failure mode was replaced with low pump flow failure mode since the data log provided evidence that low pump flow occurred during the cannula kink. the cause of the low pump flow issue was most likely kinked cannula, based on data log review and provided clinical observation. the cause of the positioning issue was most likely use related to improper technique, based on given clinical details. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26194 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 a new medical device problem code was added due to investigation results. due to investigation results, code 2889 is not necessary and was on the initial manufacturer device report number 1220648-2025-26194.
11577 | 1220648-2025-26195 | 21409897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409897&pc=OZD | 01/27/2025 | 03/22/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576974 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support (mc) experienced a pump stop and thrombus of the left ventricle. at the beginning of said month the patient was with st-segment elevated myocardial infarction (stemi). drug eluting stents x 2 were placed in the left anterior descending artery (lad) and one des was placed in the obtuse marginal (om) artery. ejection fraction was 20%. the patient was sent home on lifevest but return the next morning due to chest pain. stemi alert called, patient was taken to the catheterization lab and found previous stents occluded and lad dissection. penumbra was done, des x3 to the lad. the patient went into ventricular tachycardia during ballooning of the om, left ventricular gram was completed and end-diastolic pressure was 45 and ejection fraction went down to 10%. the decision made to place the impella cp device at this time which was successfully completed. the patient was sent to the unit on impella mcs. seven days later the patient was being escalated to an impella 5.5 placement. however, in the morning of this day, a hard pump stop occurred. after multiple attempts to restart, a stat echocardiogram was started and the physician decided to pull the impella cp pump back over the valve and go emergently to operating room for escalation to an impella 5.5. left ventricle thrombus was noted on transesophageal, and the decision was made to place the patient on cardiopulmonary bypass, complete an aortotomy and remove the thrombus. thrombus successfully removed and impella 5.5 was then placed via axillary approach. the patient was successfully weaned from cpb on impella 5.5 mcs. the patient was sent to the unit in stable condition. | [5 empty] | The investigation of pump stop and thrombus has been completed. no product was returned for evaluation. data logs prior to the product malfunction were not returned for review. data logs were reviewed and lt log at time of pump stop shows a motor current spike and an alarm impella stopped ¿ motor current high¿ triggered at time if pump stop. prior to pump stop, purge pressure observed to be dropping slightly with momentary increases in purge flow. motor current and pump flow also observed to be increasing momentarily at the same time as purge flow increases. real time (rt) log at time of pump stop shows one minute before pump stop alarm, a motor current spike and step up in motor current with a speed deviation. pump flow also becomes fully saturated after motor current spike. purge pressure increases slightly with a decrease in purge flow right after motor current spike. pump was also attempted to be restarted multiple times with pump unable to be restarted. clinical details noted that patient had left ventricle (lv) thrombus present at time of pump stop. the root cause of the pump stop was most likely due to sudden biomaterial ingestion. the root cause of "thrombus" was removed and investigated under "pump stop" as the pump stop was caused by biomaterial ingestion. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26195 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
11578 | 1220648-2025-26196 | 21409906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409906&pc=OZD | 01/26/2025 | 04/24/2025 |  | 02/19/2025 |  |  | 01/26/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589261 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | Positioning Problem |  |  |  |  | Unspecified Infection | Tachycardia | Thrombocytopenia | No Clinical Signs, Symptoms or Conditions | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was purge system blocked alarms and pressure was increasing. the staff added tissue plasma activator to the purge and the patient remains on support. | Additionally, the patient was heparin-induced thrombocytopenia (hit) positive during support. furthermore, the patient also had ventricular tachycardia (vt) during support. also, there were suction alarms throughout support. additionally, the pump was found deep. the pump was eventually repositioned. furthermore, the patient had an infection during support. the source of the infection is unknown. the patient expired on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. as per the data log review, the root cause of the high purge pressure issue was most likely biomaterial based on data log review and the clinical details provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26196 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | B1 revised to reflect both adverse event and product problem based on the additional information received from the clinical site. b2 revised to reflect death based on the additional information received from the clinical site. b5 revised to include the additional information received from the clinical site. d6b explant date added. h1 type of report revised to reflect death based on the additional information received from the clinical site. h6 revised to include the additional information received from the clinical site. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11579 | 1220648-2025-26197 | 21409909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21409909&pc=OZD | 01/26/2025 | 03/05/2025 |  | 02/19/2025 |  |  | 01/26/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574293 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had a right frontoparietal intraparenchymal hemorrhage. the patient exhibited left sided weakness and heparin was held. per the neurosurgeon no surgical intervention would be done. the patient was able to protect airway and did not require intubation. support continued and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26197 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 was revised to include instructions for use which were inadvertently not included on the initial manufacturer device report number 1220648-2025-26197. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11580 | 1220648-2025-26198 | 21410008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21410008&pc=OZD | 01/27/2025 | 03/10/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515198 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported initial insertion attempts of an impella 5.5 and impella cp device to a patient for protected intervention was unsuccessful. however, a final attempt with an impella cp was made and successful. the heart team evaluation led to decision for an impella 5.5 for protected high-risk percutaneous coronary intervention (pci) due to severe peripheral artery disease. the patient refused redo coronary artery bypass grafting surgery, left ventricular assist device and extracorporeal membrane oxygenation (ecmo). the angiogram showed adequate caliber for right axillary impella 5.5 access; however, there is heavy calcification at innominate origin. cut down on right axillary was without issue. activated clotting time of 298 seconds. gelweave 10mm x 30 cm beveled and anastomosed without an issue. the left ventricle was accessed with some difficulty, and the impella 5.5 was prepped, inserted, however, the impella 5.5 was unable to pass the innominate ostial calcification into the ascending aorta. the decision was made to convert to an impella cp implant via the axillary graft. the wire access in the left ventricle was maintained and the impella cp was prepped, inserted with activated clotting time of 350 seconds. the impella cp inlet and most of the cannula did cross into the ascending aorta; however, the motor was unable to pass. a buddywire technique was attempted without success. the decision was made to use shock wave on calcium. the motor was still unable to cross the calcification. the team, therefore, decided to remove the impella cp and wires and start with new wires and new impella cp device, which was successfully implanted. the patient was sent to the unit on impella cp mechanical circulatory support with plans to complete the protected pci the following day. there was no report or indication of any adverse effects to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue was most likely patient condition since the implant was aborted due to patient had severe pad and known calcification of ostial innominate artery. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this is one of two device manufacturing reports associated with this event. refer to initial medical device report for impella cp serial number: (b)(6) with date of event 27-jan-2025 and identifier ending in (b)(6).
11581 | 1220648-2025-26199 | 21410241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21410241&pc=OZD | 01/27/2025 | 03/10/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported initial insertion attempts of an impella 5.5 and impella cp device to a patient for protected intervention was unsuccessful. however, a final attempt with an impella cp was made and successful. the heart team evaluation led to decision for an impella 5.5 for protected high-risk percutaneous coronary intervention (pci) due to severe peripheral artery disease. the patient refused redo coronary artery bypass grafting surgery, left ventricular assist device and extracorporeal membrane oxygenation (ecmo). the angiogram showed adequate caliber for right axillary impella 5.5 access; however, there is heavy calcification at innominate origin. cut down on right axillary was without issue. activated clotting time of 298 seconds. gelweave 10mm x 30 cm beveled and anastomosed without an issue. the left ventricle was accessed with some difficulty, and the impella 5.5 was prepped, inserted, however, the impella 5.5 was unable to pass the innominate ostial calcification into the ascending aorta. the decision was made to convert to an impella cp implant via the axillary graft. the wire access in the left ventricle was maintained and the impella cp was prepped, inserted with activated clotting time of 350 seconds. the impella cp inlet and most of the cannula did cross into the ascending aorta; however, the motor was unable to pass. a buddywire technique was attempted without success. the decision was made to use shock wave on calcium. the motor was still unable to cross the calcification. the team, therefore, decided to remove the impella cp and wires and start with new wires and new impella cp device, which was successfully implanted. the patient was sent to the unit on impella cp mechanical circulatory support with plans to complete the protected pci the following day. there was no report or indication of any adverse effects to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this is one of two device manufacturing reports associated with this event. refer to initial medical device report for impella 5.5 serial number (b)(6) with date of event (b)(6) 2025 and identifier ending in (b)(6). | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the motor was unable to cross calcification in the vessel. g1 manufacturing site name and address was revised as incorrect information was inadvertently reported on the initial manufacturer device report number.
11582 | 1220648-2025-26200 | 21410510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21410510&pc=OZD | 01/20/2025 | 03/13/2025 |  | 02/19/2025 |  |  | 01/20/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582370 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella 5.5 implant, the patient experienced persistent bleeding from the axillary pocket. multiple blood products were given and the pocket was packed to manage the bleed. the site was stapled shut and kcentra was given overnight in response to the bleed. support continued with the implanted pump following the intervention. | [5 empty] | The investigation of access site bleeding has been completed. team attempting to close bleeding axillary pocket by giving blood and packing pocket. stated to not be successfully managed. patient has coagulopathy disorder. the root cause of the access site bleeding issue was most likely patient condition related as the patient had coagulopathy. d6b added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11583 | 1220648-2025-26201 | 21410943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21410943&pc=OZD | 01/09/2025 | 02/27/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567468 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | Unspecified Hepatic or Biliary Problem | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured recurring non-sustained ventricular tachycardia with a "possible" relationship to the impella device used. the patient developed recurring runs of self-limiting non-sustained ventricular tachycardia. potassium and magnesium replaced per hospital protocol. the event occurred two days after start of support. it was reported that the patient/event had not recovered/not resolved. additionally, the registry noted that approximately 16 days later, that patient endured hepatic failure with a "possible " relationship to the impella device used. as well, the patient/event had not recovered/not resolved. the impella 5.5 device was implanted for mechanical circulatory support as a bridge to transplant. two days after the hepatic failure event, the patient received a heart transplant. in the operating room, the impella was weaned to support level p-2 right before transplant and before bypass. the patient hemodynamically tolerated the wean and therefore, the impella 5.5 device was explanted prior to sternal opening. the patient was noted to have survived the explant. | Additional information via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured event with a "possible" relationship to the impella device used: ¿the patient has a history of chronic kidney disease; however, the patient developed worsening acute kidney injury, eventually decompensating enough to initiate continuous renal replacement therapy via hemodialysis catheter on (b)(6) 2025 which was removed (b)(6) 2025. admission creatinine 1.25 mg/dl, peak 4.75 mg/dl (b)(6) 2025. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Instructions for use renal failure. impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b5 revised to reflect additional information not included on the initial manufacturer device report 1220648-2025-26201. h6 revised to reflect the additional information received not included on the initial manufacturer device report 1220648-2025-26201. h10 revised to reflect instructions for use not included on the initial manufacturer device report 1220648-2025-26201.
11584 | 1220648-2025-26202 | 21410961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21410961&pc=OZD | 01/27/2025 | 10/02/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569292 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and may have encountered positioning issue. the patient was admitted with fluid volume overload and over the past several weeks had worsening heart failure and cardiogenic shock requiring intra-aortic balloon pump (iabp) placement. the decision was made to upgrade to an impella 5.5 device, which was placed in normal fashion with no issues through a 10 millimeter gelweave graft. fluoroscopy was used for insertion of the diagnostic catheter and guidance of the impella over a 0.018 guidewire. once the guidewire was removed, the impella was started at performance level p-2. the impella appeared deep and, therefore, was pulled back slightly measuring 4.2 centimeters. once insertion steps were completed, the iabp was removed and the impella was slowly ramped up to p-7. the implant was complete and the patient was sent to the unit on support. approximately 24 days post start of support, the automated impella controller displayed the following indication "impella position unknown" despite having great pulsatility. the physician thought maybe the optical sensor was damaged. the patient continued and currently remains on impella 5.5 mcs. there was no patient harm as a result of the event. | [5 empty] | The investigation of positioning issues has been completed. the pump was not returned for investigation. according to the clinical report, the pump position was found to be deep following the initial placement. additionally, several "impella position unknown" alarms were reported during the case, despite the pump being in good position. the purge case replacement resolved the purge casette failure alarm. the cause of the positioning issue was most likely use related since pump was found deep initial placement. | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of writing this manufacturer device report (mdr), and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. according to the data log, the signal-to-noise ratio (snr) and contrast dropped below the normal ranges of 1500 and 15, with no corresponding changes to the p-level from the beginning of the case. no abnormalities were observed in the other 5s optical signal parameters. the cause of the optical signal problem remains not determined, as the pump was not returned for examination and it is unclear whether the pump potentially contributed to the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-26202 d10 concomitant medical products and therapy dates updated.
11585 | 1220648-2025-26203 | 21411223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21411223&pc=OZD | 01/27/2025 | 04/02/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582535 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced pump malpositioning, low pump flow and subsequently expired. the next day after start of support, the staff was proning the patient when the impella cp device was pulled back into the aorta due to the tuohy-borst being not completely closed. an impella in aorta alarm was triggered with impella flows 2.8/2.6 (2.7) liters per minute suddenly dropping to flows of 0.1/0.0 (0.0) liters per minute. there were still no signs of pulsatile motor current and the flows were still 0.1/0.0 (0.0) liters per minute. an attempt was made to reposition the pump. however, it would not move. several times lost flows were pointed out, and the pump could have been stuck/caught somewhere and/or have a kink. the position nor flows could be restored. the left ventricle was distended and contractility was very bad. a discussion on how to proceed was made and steps taken thereafter are unknown. however, the patient was reported to have passed away the next day. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues and low pump flow has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the low pump flow issue was most likely improper positioning of the device. the root cause of the positioning issue was touhy-borst related to improper technique d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26203.
11586 | 1220648-2025-26204 | 21411400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21411400&pc=OZD | 01/25/2025 | 03/18/2025 |  | 02/19/2025 |  |  | 01/25/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025596371 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was placed and overnight the patient had ectopy. on echocardiogram the impella was pulled back and ectopy resolved. it was further reported that on day 11 of impella support, the impella 5.5 had high purge pressure. the cassette was changed with no recovery of the flow. tissue plasminogen activator was ordered to be run through the purge line. the next day, the impella 5.5 was removed as planned and there was no harm to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the clinical details report tissue plasminogen activator (tpa) was used in the purge, but it's unclear how quickly this was done after the onset of the issue. additionally, the patient was not on systemic heparin. based on the gradual nature of purge pressure rise and clinical details provided, the root cause of the high purge pressure was most likely due to biomaterial build up in the purge gap. as the necessary information was not provided, the root cause of the vt/vf was not determined."
11587 | 1220648-2025-26205 | 21411637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21411637&pc=OZD | 01/14/2025 | 03/13/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  | 12/31/2025 |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445234 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Cardiogenic Shock | Gastrointestinal Hemorrhage | [9 empty] | Death | Required Intervention | Life Threatening |  |  |  | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE QXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing an ablation was implanted with an impella cp device for mechanical circulatory support (mcs). impella was implanted using best practice. and the ablation was performed. the decision made to keep the impella in place. the peel-away removed, repositioning sheath advanced, angle matched, and impella secured. the patient was sent to the unit on impella mcs. plan to continue fluid resuscitation, monitor heart function and hemodynamics. later in the date it was reported the patient had worsening hemodynamics. the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va ecmo), and blood was transfused with fluids. approximately 13 later, the patient had a gastrointestinal bleed and blood products: 1 unit of packed red blood cells, platelets 2 units and fresh frozen plasma 1 unit were given. heparin was used in the impella purge system, and it is unknown if the use of the heparin impacted or exacerbated the outcome of the bleed. the patient would expire three days later. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured progression of cardiogenic shock with a "definite" relationship to the impella device used: ¿despite maximal efforts, family decided to proceed with withdrawal of care.¿. the patient outcome was fatal. additionally, it was reported that the patient endured hemorrhagic shock with a "definite" relationship to the impella device used: ¿post impella implant, patient developed a large hematoma in the left groin. lactic began increasing, required va ecmo cannulation. multiple blood products administered.¿. the patient is not recovered and issue was not resolved. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B.5 updated with additional information received from the clinical site. h.6 updated to reflect the additional information received from the clinical site. h.10 updated to add instructions for use for the h.6 codes that were added. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.2 revised as life threatening was incorrectly selected on initial manufacturer device report number 1220648-2025-26205. b.3 revised event date due to information in b.5 since the initial manufacturer device report number 1220648-2025-26205 was submitted. d.3 e-mail address should have been left blank on the initial manufacturer device report number 1220648-2025-26205. d.4 revised lot number and unique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26205. g.1 revised manufacturing site name and address as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26205. g.2 added study as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26205. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26205.
11588 | 1220648-2025-26206 | 21411653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21411653&pc=OZD | 01/26/2025 | 03/19/2025 |  | 02/19/2025 |  |  | 01/26/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025575630 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 had high purge pressure, purge system blocked alarm while on patient support. anticoagulation had not been in range for the last few hours. motor current was in range. the physician was advised to perform tissue plasminogen activator-lysis. it was further reported that the patient expired while on impella support. there is not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. the product was not returned. the purge pressure was stable for 23 hours and then there were some fluctuations in the purge pressure and the purge flow started to decrease. there were no motor current spikes or long bag/cassette changes. the fluctuations lasted for three hours and then there was a rapid rise over 38 minutes, which triggered high purge pressure/blocked alarms. it remained elevated for 12 hours and then it starts to slowly decrease and is stable throughout the rest of the case. this is associated with biomaterial buildup in the purge gap. the root cause of the high purge pressure is most likely due to biomaterial buildup based on the data log analysis.
11589 | 1220648-2025-25747 | 21411678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21411678&pc=OZD | 01/05/2025 | 03/12/2025 |  | 02/19/2025 |  |  | 01/06/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570304 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and approximately 16 hours after start of support, it was reported the pump's position was completely in the aorta and repositioning was impossible. patient critical was taken back to the catheterization lab, and the pump was explanted and replaced with new impella cp device. the patient was returned to the intensive care unit on impella mcs. however, within a couple of hours on the second impella cp device, it was noted to be dislocated; this pump slipped into the aorta as well. according to the physician, the tuohy borst valve was fixed. consultation was made with the chief physician, and the decision was made to explant and place another new (now third implant) impella cp device as it was noted the patient was in ¿very¿ critical condition. the third replacement pump was successfully placed. however, the following day within some hours the patient expired. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports related to the patient implant experience. refer to manufacturing report number 1220648-2025-25744. | The investigation of positioning issues has been completed. product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25747 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11590 | 1220648-2025-26207 | 21411723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21411723&pc=OZD | 01/26/2025 | 03/07/2025 |  | 02/19/2025 |  |  | 01/26/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597047 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had a drop in platelets. heparin-induced thrombocytopenia panel was sent. the patient was transfused with platelets. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined.
11591 | 1220648-2025-26208 | 21412084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412084&pc=OZD | 01/27/2025 | 03/10/2025 |  | 02/19/2025 |  |  | 01/26/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582258 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | Thrombosis/Thrombus | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during and following impella cp support. following planned explant of the device the following day, the patient went into atrial fibrillation and a stroke alert was called. in addition, a thrombus was discovered in the left ventricle. the patient's overall condition following the report was considered to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus, atrial fibrillation, and stroke has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the thrombus, atrial fibrillation, and stroke was not determined. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.3 revised event date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26208. g.3 date of awareness was reflected incorrectly on the initial manufacturer device report number 1220648-2025-26208 and should have reflected (b)(6)2025. h.10 some instructions for use wer inadvertently omitted from the initial manufacturer device report number 1220648-2025-26208 and were added.
11592 | 1220648-2025-26209 | 21412579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412579&pc=OZD | 01/28/2025 | 03/12/2025 |  | 02/19/2025 |  |  | 01/28/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568618 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5p device for mechanical circulatory support. while on support, approximately 42 days later, purge pressure/purge system blocked alarm triggered. no kinks were observed. tissue plasminogen activator was added to the purge to address the issue. the device remains in use supporting the patient. there was no report or indication of adverse effects to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump was not returned for investigation. the data log recorded several air in purge alarms and purge system open alarms during the de-airing process, which the doctor took three minutes to complete. this was followed by a purge system block alarm. several de-airing and bag/cassette change procedures were performed to resolve the issue. the purge pressure problem was resolved after 23 hours, showing a gradual decreasing trend. the clinical report confirms the tissue plasminogen activator (tpa) administration. longer de-airing and an open purge cassette line allowed blood to ingress into the motor housing, blocking the purge way. the cause of the high purge pressure issue was most likely related to biomaterial, specifically due to a longer priming procedure that allowed the blood to clot inside the motor housing. d6b explant date added.
11593 | MW5166432 | 21412927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412927&pc=QOC | 01/30/2025 | 02/12/2025 |  | 02/19/2025 |  |  |  | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONSOLE |  |  | QOC |  |  | 0042-0000-US |  |  | ABIOMED, INC. | [8 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | NA |  | Male |  |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | During implantation of 5.5 impella to axillary artery, the screen of the impella console went blank.
11594 | 1220648-2024-22743 | 21412960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412960&pc=OZD | 06/01/2022 | 06/26/2025 |  | 02/19/2025 |  |  | 06/01/2022 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383042 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was placed mid procedure due to multiple episodes of sustained ventricular tachycardia. defibrillation was required and therapy was delivered over 30 times in the procedural area, including several rounds delivered from patients implantable cardioverter-defibrillator. upon plugging in the automated impella controller (aic), it was reported that the aic failed to recognize the pump. the instructions for use (ifu) was followed and a backup controller was used. additional information noted that the care was withdrawn and the patient expired. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. automated impella controller (aic) malfunctioned in which the aic failed to recognize the pump. the instructions for use (ifu) was followed and a backup controller was used. additional information noted that the care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Additional information was provided in b.5 and h.10. additional review of this complaint event was conducted via abiomed's medical safety office. while abiomed is unable to dissociate the patient's outcome to the impella device, our safety office believes the product issue is unlikely to be related to the patient outcome. the death was likely due to patients clinical condition (pre-procedural cardiogenic shock and sustained vt that required cardioversion prior to initiation of impella therapy). the delay in support was transient and therapy was initiated on the back-up console. therapy continued for approximately 8.5 hours until the patient expired. d.2 revised common device name in accordance with updated procedures since initial manufacturer device report number 1220648-2024-22743 was submitted. g.1 revised reporting contact email in accordance with updated procedures since initial manufacturer device report number 1220648-2024-22743 was submitted. h.6 code 4629 was added as it was inadverently omitted from the initial manufacturer device report number 1220648-2024-22743. h.10 corrected information was provided. | The investigation for the reported pump not detected has been completed. the product was returned and the pump not detected issue was confirmed. the console was thoroughly tested on an engineering lab bench. the epm crystal issue was easily reproduced but was effectively mitigated by pmd reset. the pump detection issues that occurred in the field were most likely due to the epm crystal issue. this report is being filed as part of a retrospective review of historical records.
11595 | 1220648-2024-22072 | 21412961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412961&pc=OZD | 04/27/2022 | 11/05/2024 |  | 02/19/2025 |  |  | 04/27/2022 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232827 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant reported that the automated impella controller (aic) observe a yellow aic alarm. the aic was not used during patient use. the team followed the instructions for use (ifu) and the team replaced the aic with a backup controller. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, (b)(6) was fully investigated and tested on an engineering bench. the reported alarm was reproduced upon boot up without pump/ purge connected. impellatronic board was temporarily replaced with a known good one, hence the alarm was resolved after power cycle. therefore, the root cause of aic alarm ¿impella defective¿ even without pump connected was due to a defective impellatronic pca. the impellatronic board was replaced, the blue receptacle was cleaned, and a functional check was performed on the console and the optical system; before returning the console to the customer. this report is being filed as part of a retrospective review of historical records.
11596 | 1220648-2024-22724 | 21412963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412963&pc=OZD | 01/06/2022 | 11/08/2024 |  | 02/19/2025 |  |  | 01/06/2022 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245459 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure, battery failure (red alarm). there was no impact as this did not occur during patient support. | [5 empty] | E1, e3, e4: the name of the initial reporter and occupation is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure and battery communication error alarms was due to a defective battery (rapid decline in cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
11597 | 1220648-2024-22733 | 21412964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412964&pc=OZD | 09/26/2022 | 11/08/2024 |  | 02/19/2025 |  |  | 09/26/2022 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023230233 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a purge disc not reading alarm. the staff tried to troubleshoot; however, it was unsuccessful. the device was replaced with a new aic. it was reported that the procedure was completed successfully with no harm to patient related to this malfunction. the procedural outcome is patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the patient's date of death is unknown. this device was returned for evaluation. the consoles purge assembly was inspected where it was observed that the purge disc retainer had a large fracture in it but was still intact on the purge assembly.
11598 | 1220648-2024-22736 | 21412965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412965&pc=OZD | 09/18/2022 | 11/08/2024 |  | 02/19/2025 |  |  | 09/18/2022 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1158015 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the console was stuck on priming during a cassette change. the device was replaced with another aic and the procedure was successful. no report of patient harm or injury related to this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown. the device was returned for evaluation. console was inspected and tested on an engineering lab bench. glucose buildup was observed in the purge insert tray but did not impact the ability to run a known-good (kg) pump and purge (no issues observed). the root cause of the purge fluid not detected error was most likely due to improper cassette placement (use issue). this report is being filed as part of a retrospective review of historical records.
11599 | 1220648-2024-22747 | 21412966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412966&pc=OZD | 05/08/2022 | 11/08/2024 |  | 02/19/2025 |  |  | 05/08/2022 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087350 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced controller failure (red alarm). issue was resolved by switching to backup aic. additional information noted that the care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) has been completed. the product was returned and the aic - controller failure (red alarm) was confirmed. the cause of the aic issue was most likely due to a defective purge assembly, however the overall root cause could not be determined because of the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
11600 | 1220648-2024-22749 | 21412968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412968&pc=OZD | 12/08/2021 | 11/08/2024 |  | 02/19/2025 |  |  | 12/08/2021 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, US |  |  | ABIOMED, INC. |  |  | US |  | 1117649 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an automated impella controller (aic) for mechanical circulatory support. while on support, it was reported there was a battery failure. the device was replaced with another aic. no report of patient harm or injury. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of battery-related alarms was a communication glitch between charger (pbm) and batteries, most likely due to defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
11601 | 1220648-2024-22744 | 21412969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412969&pc=OZD | 09/28/2022 | 11/08/2024 |  | 02/19/2025 |  |  | 09/28/2022 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 2023234685 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a controller error alarm. the device was replaced with another aic, the procedure was successful with no reported harm to patient regarding this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the patient death is unknown. the cause of the issue was not determined since product was not returned.
11602 | 1220648-2024-22737 | 21412976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412976&pc=OZD | 03/31/2023 | 11/08/2024 |  | 02/19/2025 |  |  | 03/31/2023 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1166653 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | No Pressure | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the aic experienced purge disc/flag issues with no resolution specified. the patient ultimately expired as care was withdrawn, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the cause of the issue was broken/ missing purge disk retainer. this report is being filed as part of a retrospective review of historical records.
11603 | 1220648-2024-15573 | 21412986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412986&pc=OZD | 05/08/2021 | 08/17/2024 |  | 02/19/2025 |  |  | 05/08/2021 | 02/19/2025 |  |  | 03/31/2023 |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022059125 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, there were positioning issues and aorta alarms. action taken to resolve the issue are unknown. the impella cp was explanted the same day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
11604 | 1220648-2024-17720 | 21412987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412987&pc=OZD | 08/16/2024 | 03/28/2025 |  | 02/19/2025 |  |  | 08/16/2024 | 02/19/2025 |  |  | 09/30/2024 |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276807 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had ongoing hemolysis. the pump was removed due to hemolysis. the doctor wanted to know the cause of the hemolysis. the patient survived the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for hemolysis has been completed. the returned complaint device was visually inspected. biomaterial was observed around the impeller. no sharp edge or broken blade was found. data log reviewed was completed. many suction alarms occurred followed by impella position in aorta and ventricle alarms. despite multiple repositions and noting pump was in a good position, the hemolysis issue did not resolved. the cause of the hemolysis issue was biomaterial ingestion. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2024-17720. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-17720 was submitted in accordance with updated procedures.
11605 | 1220648-2024-13198 | 21412991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412991&pc=OZD | 07/10/2024 | 11/18/2025 |  | 02/19/2025 |  |  | 07/11/2024 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127918 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the automated impella controller (aic) would not turn off. upon restarting the aic an error message showing system error. the aic was replaced, and the issue resolved. there was no reported patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. d9 device return date added. updated h3 to device evaluation anticipated. h6 type of investigation code added. | The investigation has been completed. data analysis: on (b)(6) 2024, ic2240 had pmd communication issues resulting in a pmd reset. the following day, the console had cpld communication issues that prevented the console from shutting down. before the console was sent back for investigation, the console was booted up into system self check as captured in (b)(4). device analysis: this complaint was investigated alongside (b)(4) as both described symptoms of impellatronic communication failures. the failure mode was not reproduced despite power cycling. root cause: the root cause of the system self check failure after rebooting was a cpld communication issue. the cause of the cpld communication issue could not be determined due to the inability to reproduce it.
11606 | 1220648-2024-15460 | 21412993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21412993&pc=OZD | 09/03/2021 | 05/12/2025 |  | 02/19/2025 |  |  | 09/03/2021 | 02/19/2025 |  |  | 06/30/2023 |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022106504 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a small hematoma noted around impella insertion site. pressure was held and hematoma pushed out by nurse and hemostasis team. the patient received 2 units of red blood cells. the patient developed another hematoma after being turned and bathed. the hematoma was pressed out and bleeding minimal. the patient is awake and is continually moving around the bed which causes bleeding to increase throughout the night. the impella was weaned and removed, however during removal cts attempted to doppler right femoral artery, without success. a call was placed to vascular surgeon and thrombectomy was performed, artery was repaired with patch. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
11607 | 1220648-2024-13196 | 21413017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21413017&pc=OZD | 07/10/2024 | 03/28/2025 |  | 02/19/2025 |  |  | 07/10/2024 | 02/19/2025 |  |  | 02/28/2026 |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481106 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. it was further reported that fluid leakage occurred from the purge pressure transmitter of the purge set. it has been less than five days since it was replaced. additionally, the pump stopped suddenly. the impella was exchanged as a result of the pump stop, and it was noted that the patient was stable after the pump was exchanged. | The complainant reported a patient of unknown age and gender with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that blood was leaking from the catheter plug of the impella 5.5. the blood leakage has subsided, and the 5.5 is still in use. this event is being monitored by the staff. there was no harm to the patient. |  |  |  |  | B.5 new information was received. d.6b explant date was added. d.9 device returned date/information was added. h.6 a new medical device problem code was added due to new information received. h.10 instructions for use were added for the new medical device problem code. the investigation has been completed. impella 5.5 and purge cassette were returned for evaluation. electrical testing was performed on the pump and two motor cable lines were found to be above specification. the pump red handle was opened and blood was present in the red handle. a hole in the catheter was present at the 20 centimeter (cm) mark. the returned purge cassette was connected to console and the pump was purged for approximately 20 minutes. in the first few minutes of purging, purge pressure was low with high flow but de-air of purge cassette was performed again and all connections were secured and purge pressure was able to normalize. no leaks were observed from the purge cassette while running in lab. no data logs were returned for analysis. clinical details noted that blood leakage was coming from catheter plug and leaking was observed from purge cassette. the reported failure mode of "purge leak - pump" was a result of the blood leakage coming from the red handle and therefore there was no issue with purge leaking from the pump itself. no problem was found with the purge leak - pump. the root cause of the pump stop was due to an electrical short caused by blood ingress into the red handle from a hole in the catheter. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ a.3 revised as selection was inadvertently omitted on the initial manufacturer device report number 1220648-2024-13196. e.1 revised address line 2 as it was inadvertently omitted on the initial manufacturer device report number 1220648-2024-13196. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-13196 was submitted in accordance with updated procedures. g.2 selection was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2024-13196. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
11608 | 1220648-2025-26210 | 21414206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21414206&pc=OZD | 11/20/2024 | 05/09/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548993 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. after replacing the purge cassette, the purge pressure fluctuated wildly, and when checking the trend, the purge flow rate was also unstable. the purge cassette was replaced. after that, the purge flow rate stabilized. additionally, water droplets were found on the inside of the exterior of the purge side arm. crystallized glucose was also found on the outside of the connection between the purge side arm and the purge line. although there was no change in the purge pressure, the service life had been exceeded and, considering the risk of the pump stopping, the impella 5.5 was replaced. there was no patient harm. | The patient expired while on support. per the medical safety officer review there was no hemodynamic consequence because of the issue. |  |  |  |  | The pump and purge cassettes were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of priming issue and purge leak has been completed. the pump was returned and there was biomaterial all in the cannula. no damages were visualized to the sidearm. the pump was soaked in warm water with tergazyme to attempt to remove the biomaterial but most remained. the pump was purged and a leak wasn't able to be reproduced. due to the lack of data logs and the issue not able to be reproduced, the root cause of the purge leak pump and priming issue could not be determined. the patient expired while on support. per the medical safety officer review there was no hemodynamic consequence because of the issue.
11609 | 1220648-2025-26212 | 21415202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21415202&pc=OZD | 01/24/2025 | 03/17/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  |  |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599654 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump inserted to support a patient admitted in for surgery. the plan was to escalate from intra-aortic balloon pump to the 5.5 pump in coronary artery bypass grafting. the team moved/repositioned the pump as they believed it was too deep, they pulled back and lost position and had to place new introducer kit to deliver to the left ventricle. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of positioning issues has been completed. during the implantation, the echocardiogram indicated device was too deep causing the physician to pull the device across the valve back into the aorta. the physician stated the device was actually too shallow. device was returned and inspected. no data logs were returned. visual inspection of the returned product showed no abnormalities. the cause of the positioning issue was most likely use related due to clinical details stating the device was pulled back through the aortic valve. b3 date of event revised from the initial manufacturer device report number 1220648-2025-26212, as it was inadvertently entered incorrect. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26212 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact fax number was inadvertently omitted on initial manufacturer device report number 1220648-2025-26212.
11610 | 1220648-2025-26213 | 21415209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21415209&pc=OZD | 01/27/2025 | 03/24/2025 |  | 02/19/2025 |  |  | 01/27/2025 | 02/19/2025 |  |  | 11/30/2025 |  | 02/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430602 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had high purge pressure alarm. there was no kink noted. there was also a position alarm, but echocardiogram showed impella in correct position in the left ventricle. the cp was exchanged and the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of high purge pressure and optical signal issue has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26213 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11611 | 1220648-2025-26166 | 21393121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21393121&pc=PYX | 01/25/2025 | 03/11/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025578258 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella rp flex to support a patient admitted in with right heart failure and preoperative. the rp was explanted after 8 days due to low flows issue and questions of either clot ingestion or sensor damage. no other information was provided and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of low pump flow has been completed. the reported motor current spike and 0 flow. biomaterial was found around impeller upon investigation. data logs showed motor current spike followed by low flows and shift in placement signal. the cause of the low pump flow was most likely biomaterial since there was a motor current spike and biomaterial on the impeller. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11612 | 1220648-2025-26159 | 21394203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21394203&pc=OZD | 01/25/2025 | 03/03/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589258 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown ethnicity) was implanted with an impella 5.5 device for mechanical circulatory support and as a bridge to transplant. twenty-five days into the support, signs of hemolysis were noted. the patient's urine was red, and lactate dehydrogenase (ldh) was greater than 1000. hemoglobin and hematocrit and platelets levels were stable. an echocardiogram was completed and showed the inlet too close to the mitral valve and consequently the pump was repositioned. the urine cleared within 12 hours. the ldh continued to rise, however, and was noted to be greater than 3000 two days later. hemoglobin dropped from 8g/dl to 7g/dl. platelets level and activate partial thromboplastin time were stable and consistent. the patient was escalated to a status 1 awaiting heart transplant. the following day, the device was explanted, and the patient survived the explant. whether or not the patient received the heart transplant or the impella was replaced is unknown. | [5 empty] | The investigation of hemolysis has been completed. no product returned. hemolysis occurred on day 25 of support. echo showed inflow cage close to mitral valve. repositioning resolved the issue. data logs show stable motor current. brief impella in aorta alarms on day 10 of support. the cause of the hemolysis was most likely improper positioning of the device. b1 revised to reflect adverse event only on manufacturer device report 1220648-2025-26159. h6 code 2616 was reported incorrectly on manufacturer device report 1220648-2025-26159. new code was added to medical device problem code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11613 | 1220648-2025-26168 | 21394352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21394352&pc=OZD | 01/27/2025 | 02/28/2025 |  | 02/18/2025 |  |  | 01/27/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546197 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 as escalation of care from an impella cp for continued mechanical circulatory support. eleven days after start of the impella 5.5 support, the patient developed a hematoma and suffered from disseminated intravascular coagulation. one unit of packed red blood cells was infused and support continued with the implanted pump until successful weaned and explanted without incident seven days later. the patient survived the explant and was noted to be in stable condition. | [5 empty] | The investigation of access site bleeding and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of thrombus could not be determined. based on the clinical details the root cause of the access site bleeding was most likely patient condition since patient had coagulation disorder. d6a/d6b revised to reflect the correct implant and explant dates on manufacturer device report 1220648-2025-26168. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11614 | 1220648-2025-26162 | 21394484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21394484&pc=OZD | 05/12/2023 | 03/25/2025 |  | 02/18/2025 |  |  | 01/24/2025 | 02/18/2025 |  |  | 01/31/2025 |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317526 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A publication was received titled "a case in which a patient with cardiogenic shock due to extensive transmural infarction with a two-vessel lesion was saved by the combination of impella and ino" which noted the patient experienced renal failure and hemolysis post implant of an impella cp device for mechanical circulatory support. in summary, the following was noted: the patient presented to the hospital with the main complaint of chest pain. an electrocardiogram showed st elevation in i, avl, and v2-6 as well as complete atrioventricular block, and cardiac ultrasound showed decreased wall motion of the anterior wall. as the patient had acute myocardial infarction with circulatory dynamics collapse, a circulatory assist pump catheter, impella cp, was promptly inserted, and percutaneous coronary stent placement was performed to treat the in-stent complete occlusion of the left anterior descending coronary artery and 99% stenosis of the right coronary artery. after returning to the intensive care unit, with impella cp total support, the patient developed circulatory failure, ci 1.6 l/min/m2, and lac of 3 mmol/l. hemolysis, acute renal failure, and right heart failure occurred, leading to cardiogenic shock, so on the second day of hospitalization, the device was upgraded to impella 5.5 and ino combination therapy was initiated. thereafter, with partial support, the patient achieved a ci of 2.2 l/min/m2 and became stable, and on the 10th hospital day, the impella was weaned and a ci of 2.4 l/min/m2 was achieved by the patient's own heart. in this case, the impella cp inserted in an emergency was unable to maintain sufficient circulation and right heart failure occurred. therefore, it was decided to upgrade to an impella 5.5 without missing the opportunity, and by using ino in combination. extracorporeal membrane oxygenation (ecmo) insertion was avoided and successfully saved the patient's life by breaking free from cardiogenic shock. | [5 empty] | The investigation of hemolysis and renal failure has been completed. no product or data logs were returned for evaluation. it was noted that on day two of patient hospitalization, the patient experienced hemolysis and acute renal failure and the patient was escalated to impella 5.5 support. the root cause of the hemolysis could not be definitively determined as limited clinical details were provided. the root cause of the renal failure could not be definitively determined due to insufficient clinical details. e.1 revised address line 2 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26162. | The date the event occurred is unknown, and therefore, the implant date was entered as the provisional date. additionally, a copy of the publication/literature is inaccessible at the time of writing this manufacturer device report. however, once accessible, it will be provided as an attachment on a supplemental/follow-up manufacturer device report. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11615 | 1220648-2025-26157 | 21394823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21394823&pc=OZD | 01/24/2025 | 03/11/2025 |  | 02/18/2025 |  |  | 01/24/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600742 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support and while on support the patient developed red urine and an increased in lactate dehydrogenase(ldh) post impella insertion. the p-level was reduced but ldh levels kept trending upwards. the patient was escalated to impella 5.5 and the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11616 | 1220648-2025-26160 | 21395055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21395055&pc=OZD | 01/25/2025 | 03/07/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555934 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race) in cardiogenic shock from an acute myocardial infarction underwent percutaneous coronary intervention was then implanted with an impella cp device for mechanical circulatory support and the following day was escalated to an impella 5.5. the removal of the impella cp and implantation of the impella 5.5 was successfully completed. five days after start of support on the impella 5.5 device the impella insertion site was swollen, and the jackson-pratt drain was filling with blood. the team highly suspected the patient was bleeding around the impella insertion site. the patient was reopened for exploratory surgery to find the location of the bleed. it was indicated the device was explanted this day as well. findings and subsequent treatment after the exploratory surgery are unknown. however, it was noted, the patient survived the explant of the device and was reported to be in stable condition following the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 health effect - impact code was revised to include device explant which was not included on the initial manufacturer device report number 1220648-2025-26160. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11617 | 1220648-2025-26172 | 21396615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21396615&pc=OZD | 01/27/2025 | 03/12/2025 |  | 02/18/2025 |  |  | 01/27/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586015 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and was on support for 3 days before escalation to an impella 5.5. approximately 18 days after the start of impella 5.5 support, the nurse was attempting to reposition the impella 5.5 when the sterile sleeve broke and disconnected near the blue button. there was no harm to the patient. the device remained in use until successfully weaned and explanted four days later. the patient successfully recovered and was transported back to the unit for continued evaluation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sterile sleeve damage) has been completed. based on the clinical details the cause of the sterile sleeve damage issue was most likely use related to improper technique during the repositioning.
11618 | 1220648-2025-26161 | 21397005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21397005&pc=OZD | 01/26/2025 | 03/10/2025 |  | 02/18/2025 |  |  | 01/26/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603552 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Swelling/ Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and one day of impella cp support, the patient developed some swelling in their left thigh and around their access site. one unit of packed red blood cells was given to the patient to treat the condition. the day after, oozing at the impella access site was observed and, however, noted the site was "under control." following, now a couple of days later, the patient started to bleed more from the impella site. four units of red blood cells and one unit of platelets were given. the dressing was changed and impella access angle of entry was maintained with gauze under the site; manual pressure applied. the patient was reported to be in critical condition following the event. unit support eventually ended and had an outcome of the impella cp device successfully weaned and explanted after a total of six days on support with the patient having survived the explant procedure. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11619 | 1220648-2025-26158 | 21397274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21397274&pc=OZD | 01/25/2025 | 04/01/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573247 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that three days into impella cp support, the patient began to experience persistent bleeding from their access site. the repositioning sheath was adjusted and femostop was applied at the site, but the bleed continued. in addition, it was documented that the pump was unable to provide adequate support due to suction being encountered at levels above p5. as a result of the patient's worsening condition and the restricted support level, discussions were initiated on the discontinuation of support. the patient expired at the end of support. it is unknown if the complications directly contributed to the patient's passing. | [5 empty] | The investigation of access site bleeding and low pump flow has been completed. the pump was not returned for investigation. as per the clinical information provided, the patient was bleeding from multiple site and the at the moment activated clotting time's were not available. therefore, the root cause of the access site bleeding issue was most likely patient condition related. data logs were not returned as well. without the pump or the data logs the failure could not be further investigated. therefore, the root cause of the low flow issue was not determined since the product was not returned and data logs were not returned as well. b.2 required intervention was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26158. g.1 added address line 1, city, state, zip code and country were added as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26158. h.6 a health effect - impact code was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26158. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11620 | 1220648-2025-26171 | 21398393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21398393&pc=OZD | 01/26/2025 | 03/17/2025 |  | 02/18/2025 |  |  | 01/26/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576640 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) while waiting to bridge to a left ventricular assist device (lvad). approximately 13 days after start of the support, oozing from the impella access site was noted and heparin was discontinued. additionally, on this day an optical malfunctioned occurred. there was a false placement signal low alarm. the placement signal appeared to have drifted. the impella 5.5 device was exchanged with a new impella catheter using the existing graft at the right axillary and the patient's mcs continued while waiting for the bridge to a lvad. | [5 empty] | The investigation of access site bleeding and optical signal issue has been completed. the product was not returned. log review showed "placement signal (ps) low" alarms only from (b)(6) 2025, not on the reported day 1/26/2025. there was a downward trend in ps starting around (b)(6) 2025. on (b)(6) 2025, the ps dropped ~50 mmhg over 30 minutes before drifting below 0 mmhg within the next 5 hours. the ps ultimately went out of range. signal not reliable (snr) and contrast trended downward with the ps. the rest of the 5s parameters were unaffected. a placement signal not reliable (psnr) alarm triggered when the ps went out of range. the optical signal issue was reported 6 days into the case, which is within the intended design life of 60 days per design traceability matrix. the root cause of the optical signal issue was most likely sensor silicone wear since the ps downward trend matches the known data log pattern for sensor wear. the root cause of the access site bleeding was not determined since insufficient clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11621 | 1220648-2025-26170 | 21398742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21398742&pc=OZD | 01/26/2025 | 03/07/2025 |  | 02/18/2025 |  |  | 01/26/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593280 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race) in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and six days following the implant, low pump flow was encountered. the device was explanted with an unsuccessful attempt to place and impella 5.5 however another impella cp was implanted. three days following, the replacement impella cp device would have suction alarms, patient experience ventricular fibrillation and expires. the patient presented to the hospital with chest pain and vomiting and found to have st-segment elevation myocardial infarction and cardiogenic shock with ejection fraction 10-15%. the patient had a stent placed five days prior. the patient decompensated in the emergency department and was taken back to the catheterization lab for an impella cp device placement, which was successfully completed. the patient was transferred to the unit on support. approximately six days after start of the impella mcs, continuous suction alarm was reported. there were low flows at high p levels, but no known hemolysis. the patient was transferred to another hospital for a higher level of care and advanced therapies evaluation. the decision was made to escalate to an impella 5.5 device. patient brought to operating room, prepped/draped for left axillary impella 5.5 device implant. left axillary cut down, graft sewn directly onto artery and tunnel inferior. left ventricle was then accessed via wire/catheter technique under fluoroscopy/transesophageal echocardiogram (tee). the .018 guidewire was loaded and placed into lv and catheter removed. the impella 5.5 was then backloaded onto wire; however, there were failed attempts to cross the anastomosis. impella 5.5 then removed and an impella cp was opened and placed. placement confirmed with tee, measuring at 3.7 cm. impella cp ramped to p-8 with flows 2.3 l/min. chimney graft/repo sheath sutured to the left chest. sterile dressing applied to left axillary site per their protocol. the patient was transferred to the bed for transport back to the unit, but upon leaving the room, the patient went into ventricular fibrillation. the patient was shocked two times and returned to normal sinus rhythm. tee probe was re-inserted and placement confirmed and the patient was taken to the unit. same day, later that night, on the unit, the automated impella controller displayed/sounded continuous suction alarm. replacing pa catheter and plan to obtain an echo as soon as possible to confirm placement. the next day, axillary impella cp continued on p-8 and continuous suction alarms continue. audible alarms turned off. flows of 2.5. urine output is clear/yellow. follow up the next day noted impella mcs continued with patient's prognosis as guarded. ecpella (extracorporeal membrane oxygenation and impella mcs) was considered but the patient was not a candidate for ecmo. the plan was to wean the patient and assess the patient's response. however, the patient would expire; care was withdrawn. medical safety review of the event determined there is not enough evidence to exclude the replacement impella cp as an associated factor in the patient's outcome. | [5 empty] | The investigation for the ventricular fibrillation (vf) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of three devices related to the patient's implant experience. refer to manufacturing report number 1220648-2025-26083 for the impella 5.5 and refer to initial medical device report for the initially implanted impella cp serial number (b)(6) with date of event on 23-jan-2025 and patient identifier ending in (b)(6).
11622 | 1220648-2025-26173 | 21399046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21399046&pc=OZD | 01/20/2025 | 03/18/2025 |  | 02/18/2025 |  |  | 01/21/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025497937 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | Hemodynamic instability | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient implanted with an impella 5.5 device for mechanical circulatory support (mcs) experienced a hemorrhagic stroke and would subsequently expire. the patient was admitted in for an emergency percutaneous coronary intervention to the right coronary artery approximately 15 days prior to the impella 5.5 device implant procedure. the next day following the pci, the patient went into ventricular fibrillation and underwent pci to the left anterior descending artery. intraoperative blood clots arose and left main artery infarction occurred; extracorporeal membrane oxygenation (ecmo) and intra-aortic balloon pump (iabp) were placed. sometime thereafter, the patient was taken off ecmo and iabp and crashed in a few hours. the patient was placed on ecmo and iabp again. eventually, the patient was placed on the impella 5.5 device for mcs. nine days later after start of the impella 5.5 mcs, however, the patient would experience a hemorrhagic stroke. the team was unable to maintain the patient's hemodynamics, and the patient would expire this day. actions taken were use of sodium bicarbonate in the purge, and blood coagulation management. it was further noted activated clotting time was not excessively prolonged, and the hemorrhage was an unpredictable accident. the relationship between the impella device and stroke is noted to be unknown, and the patient expired from multi-organ failure due to the cerebral hemorrhage. there is not enough evidence to exclude impella as an associated factor. per the report, the relationship between the device and hemorrhagic stroke is unknown and therefore the demised outcome cannot be dissociated. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 the selection life-threatening was inadvertently selected on the initial manufacturer device report number 1220648-2025-26173 h3 yes was inadvertently selected for evaluation summary on the initial manufacturer device report number 1220648-2025-26173, the investigation results were not included in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings, cautions & precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
11623 | 1220648-2025-26174 | 21399416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21399416&pc=OZD | 01/27/2025 | 03/10/2025 |  | 02/18/2025 |  |  | 01/27/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592063 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported insertion of an impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful due to the patient's anatomy. the impella cp was loaded onto a 0.18 guidewire and was unable to get past the tip of the sheath due to heavy calcium. the 13 centimeter (cm) sheath was exchanged for a 14 x 25cm sheath and the pump was still unable to pass. multiple sheath exchanges with stiffer wires were completed and all were unsuccessful. the impella cp implant was aborted and bailout to extracorporeal membrane oxygenation (ecmo) if needed. there was no report of adverse effects to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the delivery issue was most likely patient condition related. /they were unable to advance the impella due to subclavian stenosis.
11624 | 1220648-2025-26164 | 21400151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21400151&pc=OZD | 01/24/2025 | 03/12/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025602021 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported impella 5.5 pump support was ongoing for a patient admitted in with cardiogenic shock. the pump had signal loss and the team tried to troubleshoot. the console was replaced, but the team had to explant the impella 5.5 and replace with an intra-aortic balloon pump. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue has been completed. the cause of the optical signal issue was not determined since product was not returned.
11625 | 1220648-2025-26165 | 21400504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21400504&pc=OZD | 01/26/2025 | 03/12/2025 |  | 02/18/2025 |  |  | 01/26/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574211 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock with plans for coronary artery bypass grafting was implanted with an impella 5.5 device as escalation of care from an impella cp device for long-term mechanical circulatory support. approximately 12 days after the start of the impella 5.5 device support, the patient was noted having multiple delirious episodes and grabbed onto the impella device and pulled it 8.5 centimeters back into the aorta. the patient was emergently taken for surgical removal of the impella 5.5. the graft was surgically cut above the anastomosis of the graft and artery and sutured closed. the patient¿s chest was then closed with sutures. the patient survived the explant of the impella device and was reported to be in stable condition following the event. | [5 empty] | The investigation of positioning issues has been completed. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely patient movement related. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11626 | 1220648-2025-26177 | 21401166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21401166&pc=OZD | 01/25/2025 | 03/19/2025 |  | 02/18/2025 |  |  | 01/29/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575294 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | Other |  |  |  |  | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device via the left femoral artery for mechanical circulatory support (mcs). two days after start of support it was noted circulation could not be maintained, and the patient expired. therapy started prior to impella mcs was noted as venoarterial extracorporeal membrane oxygenation (va ecmo). problems with the impella device operation was noted as "none." no further information is available at this time. as the impella device was present, providing mechanical circulatory support to the patient, when the event occurred or otherwise associated with the event and no alternative cause for the event is identified as a likely cause of such event, the demised outcome cannot be dissociated from the impella device. per the reporter, an autopsy is being completed, and the cause is currently being investigated. upon receipt of additional information, a supplemental report will be submitted accordingly. | [5 empty] | The investigation for the unspecified death has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of death could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11627 | 1220648-2025-26180 | 21401303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21401303&pc=OZD | 01/29/2025 | 03/18/2025 |  | 02/18/2025 |  |  | 01/29/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603537 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient with acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support but encountered placement signal issue, low flow and/or pump stop during three attempts with three impella cp devices. the patient was admitted and found to have received three shocks from the automated implantable cardioverter device for ventricular tachycardia (vt) and atrial flutter (af). the patient was taken to the electrophysiology lab for vt/af ablation and during this procedure, the patient coded received one shock with return of spontaneous circulation achieved. the decision was made to place and impella cp device. the impella cp (pump 1 (b)(6)) was inserted and initial flows of 2.6l were noted with no placement signal available. the elected to replace the impella cp device (pump 1) in order to have a placement signal. new impella cp device (pump 2 (b)(6)) was primed on different automated impella controller and was placed with no issues on insertion. pump 2 initial flows of 3.6l were noted and was being secured when placement signal suddenly disappeared, and flows started to decrease. the placement signal not reliable alarm triggered and then flows stopped. therefore, physician decided yet again to replace the impella cp device (pump 2). another impella cp device (pump 3 (b)(6)) was primed on a different automated impella controller and on startup there was immediate suction noted left ventricular placement signal was abnormal and unable to obtain adequate flows. all troubleshooting performed, and the team was advised of possible corevalve transcatheter aortic valve replacement interaction. the decision made to remove the impella cp (pump 3) for a final time, and the physician elected to then place an intra-aortic balloon pump. additional information noted the following confirmations: adequate anticoagulation status was verified, transesophageal echocardiogram was performed, with no left ventricle thrombus, all sheaths were excessively flushed, and there was no clot noted upon removal of all pumps. all three devices are expected to be returned for analysis. the was no harm and the patient was reported to be stable following the event. | [5 empty] | The investigation for the low pump flow and the optical signal issue has been completed. the product was returned and evaluated. both impeller blades were broken off. no other abnormalities were found. low flow was reproduced in the lab. field data logs show a large motor current (mc) spike shortly after the pump is started at p-9. an immediate drop in pump flow followed the mc spike and flows remained below spec the rest of the case. the root cause of the low pump flow was most likely a damaged impeller resulting from interaction with the tavr mentioned in clinical details. the returned product's 5s parameters did not indicate any hard failures of the sensor with snr at ~2700 and contrast at ~14.8%. the sensor passed the laser continuity test and there were no kinks found in the catheter. the pump was run and there were no abnormalities with the optical parameters. the sensor was removed and there was no visible damage to the sensor face. data logs showed a few instances where the lv waveform drops below zero. one of these instances occurs around the time of the motor current spike associated with the drop in pump flow. comparison was made to another broken impeller lades complaint but the data logs did not resemble the same optical signal issue. the root cause of the optical signal issue was not determined since the returned product did not show significant abnormalities. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of three device manufacturing reports associated with this event. refer to the following: medical device report for impella cp pump 1 (b)(6) with date of event (b)(6) 2025 and patient identifier ending in (b)(6). medical device report for impella cp pump 2 (b)(6) with date of event (b)(6) 2025 and patient identifier ending in (b)(6). medical device report for impella cp pump 3 (b)(6) with date of event (b)(6) 2025 and patient identifier ending in (b)(6) (this report).
11628 | 1220648-2025-26167 | 21401317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21401317&pc=OZD | 01/27/2025 | 03/04/2025 |  | 02/18/2025 |  |  | 01/27/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398229 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility biomedical team reported a console displayed a controller error during use. there was no report or indication of patient harm. no further information was provided. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of aic - controller error (yellow alarm) has been completed. the boot up before the complaint, the console booted up with light = 0.49 volt which equates to no light output. after the second 5s confirms this measurement, there was an event for a controller error (defective optical sensor lamp). this reproduced on the complaint date as again two 5s measurements with 0.49 volt which led to the controller error described in the complaint. the console was returned for investigation but the failure mode was not reproduced as lamp was outputting light fine with proper current and voltage into the optical bench and the lamp. no amount of console manipulation could recreate the failure mode. the controller error was due to a lamp not outputting light but the cause of the lamp not outputting light could not be determined as the failure mode was not reproduced during testing. e.1 revised facility name as it was inadvertently omitted from the initial manufacturer device report number. revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number.
11629 | 1220648-2025-26184 | 21401617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21401617&pc=OZD | 01/28/2025 | 02/20/2025 |  | 02/18/2025 |  |  | 01/28/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1451015 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon turning on an automated impella controller for a patient going on impella mechanical circulatory support, a system self-check failure appeared. the console was exchanged as a result of the failure and the report noted there was no patient harm. there was no report or indication in delay in therapy. | [5 empty] | The automated impella controller has been received from the customer and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the console and data logs were received for investigation. according to the data and device analysis the root cause of the ¿system self check failed¿ was power battery manager (pbm) corrosion. h10 investigation results were inadvertently omitted on the initial manufacturer device report 1220648-2025-26184.
11630 | 1220648-2025-26163 | 21401644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21401644&pc=PYX | 01/25/2025 | 03/14/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025548081 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Difficult to Remove | Use of Device Problem | [5 empty] | Hemodynamic instability | [12 empty] | Life Threatening | Death |  |  |  |  | VA EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex and impella cp for mechanical circulatory support (mcs). initially, the impella cp was implanted and during the percutaneous coronary intervention, the patient began to have frequent suction alarms. right heart catheterization was performed and a pulmonary artery pulsatility index was calculated of 0.5. the decision was then made to place an impella rp flex for bipella support. within approximately three to four hours after returning to unit, the patient began to lose pulsatility on the impella cp. mean arterial pressures were sustained in the 40¿s and 50¿s despite current support and vasopressor medications. the impella cp was repositioned at bedside which appeared to be stuck in the papillary muscle. flows were diminished at 2l on performance level p-8. impella rp flex flows reduced to avoid pulmonary edema. decision was made to transport patient to catheterization lab to cannulate for venoarterial extracorporeal membrane oxygenation (va ecmo). after arriving, the patient went into cardiac arrest. cardiopulmonary resuscitation was initiated. the impella cp was removed and left femoral artery cannulated for ecmo. the left femoral vein was also cannulated. the impella rp flex was being explanted at time of ecmo cannulation and initiation. at time of impella rp flex explant, air was induced into ecmo venous line, which air locked and stopped ecmo circuit. no pulse was assessed. decision was made to cease cardiopulmonary and resuscitation efforts and the patient expired. | [5 empty] | The investigation for the removal issues has been completed. clinical details state that the patient went into cardiac arrest after continued decline on both rp and cp support. the cp was explanted and ecmo was cannulated on the left femorals and initiated. after explant of the rp flex, air was introduced into the ecmo venous line. this was caught by the air lock but the patient ultimately expired. no product was returned and logs were not available for the case. the cause of the unspecified (allegation) - death as it relates to the removal of the device was not determined since product was not returned and lack of clinical details. the instructions for use for the impella rp smartassist system with automated impella controller stated the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: g1-reporting contact last name corrected. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two device manufacturing reports associated with the patient implant experience. refer to initial medical device report for the impella cp serial number (b)(6) with date of event (b)(6) 2025 and patient identifier ending in (b)(6) .
11631 | 1220648-2025-26181 | 21402709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21402709&pc=OZD | 01/25/2025 | 03/14/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537969 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, impella cp primed and backloaded and delivered over the abiomed 0.018 wire. however, while the doctor was attempting to advance impella he had difficulty advancing impella catheter with kinks noted to 0.018 wire under fluoro. the doctor stated that the ascending aorta appeared tortuous and he felt that contributed to the kink. the abiomed wire was removed and the impella cp removed. after regaining left ventricle (lv) access with a v18 wire the impella cp was backloaded and delivered over the v18 0.018 wire into proper position. impella waveforms and impella flow were within normal limits but the purge flow was <3 and purge pressure >1000. no kinks noted to purge tubing. purge tubing and cassette completely replaced. purge flow and pressures did not improve. the decision was made to replace the impella cp with a new one. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the high purge pressure has been completed. clinical details state that 10 minutes into device use, a purge pressure high alarm appeared. troubleshooting commenced with an inspection for kinking and replacing the purge cassette which did not resolved the alarm. the device was explanted two hours after implanted. the data logs confirmed the high purge pressure alarms after a gradual trend in purge pressure immediately after pump activation with no motor current spikes.the device was returned. visual investigation showed dextrose build up at the purge gap. the alarms seen in the field were not able to be recreated in the lab when the pump was soaked and run at p-9 for 10 minutes. the cause of the high purge pressure was most likely biomaterial buildup since a gradual trend in purge pressure was observed immediately after the pump was activated. no motor current spikes. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number (b)(4).
11632 | 1220648-2025-26186 | 21402764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21402764&pc=OZD | 01/25/2025 | 03/13/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565432 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that they encountered access site bleeding and bleeding from all areas including nose and mouth. as a result, the insertion site dressing was changed, proper insertion angle was maintained, blood products were given, heparin order was revised, and a fem stop was placed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. no product was returned for evaluation. clinical details noted that patient was bleeding from multiple sites and required multiple units of volume. the root cause of the access site bleeding - major was most likely patient condition related as patient was bleeding from multiple sites. the failure mode of "vascular damage - peripheral vessel" was removed and investigated under "access site bleeding - major" as no damage to vasculature was reported. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26186 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11633 | 1220648-2025-26179 | 21402944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21402944&pc=OZD | 01/28/2025 | 03/10/2025 |  | 02/18/2025 |  |  | 01/28/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1765949 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) was received with an inoperable front access door and noted that preventative maintenance on the device had been completed in january 2025. there was no patient involvement with this event. | [5 empty] | The automated impella controller has been received from the customer and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge door won't close has been completed. the console was returned for further investigation. initially, the purge assembly door was sealed, and upon closer inspection of the purge assembly, excessive purge fluid residue was found. this residue indicated a purge fluid leakage during clinical use. the root cause of the purge door inoperable, not opening was due to the glucose ingress.
11634 | 1220648-2025-26175 | 21403474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21403474&pc=OZD | 01/25/2025 | 03/13/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572027 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered positioning issues. there was an 'impella in aorta' alarm and the so the patient was taken back to cath lab to reposition the pump. however, the impella was too far in the aorta to reposition. additionally, the patient became restless and moving in bed and the impella inlet and pigtail were across the aortic valve. as the patient was stable the doctor decided that further support was not needed and the decision was made to remove the impella. the patient remained stable. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue could not be determined as there was insufficient information provided for why the impella was too deep in the aorta. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26175 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11635 | 1220648-2025-26185 | 21403671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21403671&pc=OZD | 01/29/2025 | 03/07/2025 |  | 02/18/2025 |  |  | 01/29/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601955 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support but encountered placement signal issue, low flow and/or pump stop during three attempts with three impella cp devices. the patient was admitted and found to have received three shocks from the automated implantable cardioverter device for ventricular tachycardia (vt) and atrial flutter. the patient was taken to the clectrophysiology lab for vt/atrial flutter ablation and during this procedure, the patient coded receiving one shock with return of spontaneous circulation achieved. the decision was made to place an impella cp device. the impella cp (pump 1 serial number (b)(6)) was inserted and initial flows of 2.6 liters were noted with no placement signal available. they elected to replace the impella cp device (pump 1) in order to have a placement signal. a new impella cp device (pump 2 serial number (b)(6)) was primed on a different automated impella controller and was placed with no issues on insertion. pump 2 initial flows of 3.6 liters were noted and were being secured when placement signal suddenly disappeared, and flows started to decrease. the placement signal not reliable alarm was triggered and then flows stopped. therefore, the physician decided yet again to replace the impella cp device (pump 2). another impella cp device (pump 3 serial number (b)(6)) was primed on a different automated impella controller and on startup there was immediate suction noted that the left ventricular placement signal was abnormal and was unable to obtain adequate flows. all troubleshooting was performed and the team was advised of possible corevalve transcatheter aortic valve replacement interaction. the decision was made to remove the impella cp (pump 3) for a final time, and the physician elected to then place an intra-aortic balloon pump. additional information noted the following confirmations: adequate anticoagulation status was verified, transesophageal echocardiogram was performed, with no left ventricle thrombus, all sheaths were excessively flushed, and there was no clot noted upon removal of all pumps. all three devices are expected to be returned for analysis. the was no harm and the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of three device manufacturing reports associated with this event. the other two reports are for: medical device report for impella cp pump 1 serial number (b)(6) with date of event 29-jan-2025 and patient identifier ending in (b)(6). medical device report for impella cp pump 3 serial number (b)(6) with date of event 29-jan-2025 and patient identifier ending in (b)(6). | The investigation of low pump flow, pump stop, and optical signal issue has been completed. the clinical details mention core valve interaction as a contributing factor for failures reported. the pump stop failure mode was removed because there was no pump stop and instance when pump was turned down was likely a use inputted troubleshoot. low pump flow: no product was returned for analysis. there was a motor current spike with speed deviation three minutes after pump was started. the logs show several more successive motor current spikes before the pump was turned down to p0. after the pump was restarted to p9 there was low pump flow. the root cause of the low pump flow was most likely due to biomaterial optical signal issue: the data logs show several placement signal spikes after the pump was started. the analysis of the 5s parameters from the data logs revealed an identified pattern of a fiber break possibly due to excessive pull force. at the instant placement signal went out of range, there was a drop in signal not reliable (snr) to 0 and a drop in contrast. lamp increased to 100 and gain increased. the clinical details mention pump was being secured when this occurred. the root cause of the optical signal issue was most likely due to a fiber break when securing the device
11636 | 1220648-2025-26176 | 21404836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21404836&pc=OZD | 01/25/2025 | 03/18/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590675 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after the impella cp was placed and while placing a left common femoral vein hemodialysis line, an impella position in aorta alarm occurred. attempts were made to recross the valve several times and all were unsuccessful. the impella cp was removed and the patient remained stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related.
11637 | 1220648-2025-26182 | 21405180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21405180&pc=OZD | 01/25/2025 | 03/17/2025 |  | 02/18/2025 |  |  | 01/25/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597094 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the sterile sleeve of an implanted impella cp pump was damaged. anti-contamination sleeve was torn. the pump was replaced to resolve the damage. there was no patient harm. | [5 empty] | The investigation of product damage (sterile sleeve damage) has been completed. no product was returned for evaluation. clinical details noted that the sterile sleeve was torn. it was not noted if excessive force was applied or if tear occurred when attempting to reposition pump. the root cause of the product damage (sterile sleeve damage) cannot be definitively determined as limited clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11638 | 1220648-2025-26178 | 21405615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21405615&pc=OZD | 01/26/2025 | 03/19/2025 |  | 02/18/2025 |  |  | 01/26/2025 | 02/18/2025 |  |  | 06/30/2025 |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357268 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced device malpositioning and low pump flow. the patient would later expire. the patient was status post coronary artery bypass graft (cabg) surgery x 3 vessels one day prior to the impella cp implant. cabg, there was a left internal mammary artery and obtuse marginal artery lima and om graft failure, poor targets, and ejection fraction was reduced from 55-60% pre-operative to 20-25% post operative. the decision was made to do high-risk percutaneous coronary intervention with the impella cp mcs for recovery. the patient had severe disease in the left main, left anterior descending and left circumflex arteries. the impella cp was successfully implanted, and percutaneous coronary intervention was completed to the three vessels, the patient remained stable throughout the procedure and was transported to the unit on impella mcs in critical condition. the next day venoarterial extracorporeal membrane oxygenation (va ecmo) with impella cp mcs was planned and placed with perfusion sheaths to the lower legs. four days later, it was reported there was a brief suction event in the morning of said day after repositioning the patient. the patient remained on ecmo with flows of 4.5-5.5 liters per minute. there were occasional placement in aorta alarms as the morning continued. motor current remained pulsatile. with simultaneous ramp trial, bedside transthoracic echocardiogram showed deep position with pigtail in capillary muscle although images suboptimal due to body habitus. there was decoupling of waveforms when performance level was increased with ecmo flow reduced. the device was repositioned pulled back in 0.5-1.0 increments to a depth of 4.02 centimeters across the aortic valve and approximately 10 minutes after repositioning, impella position in aorta alarm presented despite no movement from the patient. waveforms and images were reviewed, discussed and it was noted the impella cp pump position appeared to be shallow. the decision was made not to reposition again but rather disable the alarm. the risk of disabling the alarm was discussed with the physician and it was suggested to closely monitor pulmonary artery catheter waveforms, urine output, motor current, and vasopressor requirements. additionally, it was further reported on this same day, the patient started having a gastrointestinal (gi) bleed. platelets were low and heparin-induced thrombocytopenia panel was sent with negative results. all anti-coagulation was stopped. an endoscopy was completed and revealed an upper gi bleed from an active ulcer. cautery was performed by the gastroenterologist and several blood products had been given since the start of the bleed. it was noted the physician did not attribute any of the patient¿s bleeding to the impella cp device, and of note, there was no heparin in the impella purge system. one day later, ecmo was weaned down with impella cp increased with decannulation and escalation to an impella 5.5 planned for the next day. the next day, the patient lost all pulsatility and continued to ooze from nose, mouth, and triple lumen. due to no pulsatility, the decision was made to go back up on ecmo and decrease impella. a discussion was made with the family and the patient continued ecmo and impella mcs with the decision to withdraw care instead of escalating to an impella 5.5. the patient expired the following day. the cause of death is unknown, not provided. there is not enough information to exclude the impella as an associated factor to the overall event outcome although [it should be noted] escalation was rejected, and care was withdrawn. | [5 empty] | The investigation for the low pump flow has been completed. clinical details states that suction alarms occurred when ecmo weaning. no product and data were returned. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11639 | 1220648-2025-26183 | 21405959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21405959&pc=OZD | 01/27/2025 | 03/17/2025 |  | 02/18/2025 |  |  | 01/27/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573226 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had ventricular tachycardia (vt). the doctor believed that when the patient woke up, the impella 5.5 moves and caused vt. support was continued. additionally, the impella 5.5 showed to be very close to the septal wall. however, the doctor did not want to reposition. support was continued. furthermore, upon explanting the pump, when bleeding was back, a large clot was noted. the patient remained stable. | [5 empty] | The investigation of thrombus, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. the failure mode positioning issues removed since there was no complaint reported and support continued indicating that any standard troubleshooting for the positioning issue was sufficient. as the necessary information was not provided, the root cause of the thrombus and vt/vf were not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
11640 | 1220648-2025-26187 | 21406147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21406147&pc=OZD | 01/28/2025 | 03/14/2025 |  | 02/18/2025 |  |  | 01/28/2025 | 02/18/2025 |  |  |  |  | 02/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589373 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following impella 5.5 implant, oozing/bleeding was noted from the ancillary cut down site. the patient was taken to the operating room for a washout and two units of blood were transfused. the jp drain was removed and support continued. four days later, the pump was inadvertently pulled into the aorta and the decision was made to explant the device. the patient was stable following explant. | [5 empty] | The investigation of positioning issues and access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of access site bleeding was most likely use related as physician confirmed placement of jp drain pressing on a vessel causing some oozing. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26187 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11641 | 1220648-2025-26137 | 21383708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21383708&pc=OZD | 01/24/2025 | 04/01/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571385 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and there was a kink in the cannula. pump was brought into the left ventricle, the placement was not correct, pump slipped towards the aorta and then bent so much for the motor housing between intake and outlet window that the motor no longer worked. the implant was aborted and there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation has been completed. the failure mode delivery issues was changed to positioning issues since the pump was able to pass the valve. reported getting the pump into the left ventricle but in the incorrect placement. the pump slipped toward the aorta then kinked. cannula kink was found near the outlet cage. no other damages or abnormalities. the root cause of the positioning issue was not determined due to insufficient clinical details. d.9 date device returned/information was added. a.5 should have been left blank on the initial manufacturer device report number 1220648-2025-26137, as the ethnicity and race are unknown. h.6 medical device problem code was revised since the initial manufacturer device report number 1220648-2025-26137 was submitted due to investigation results.
11642 | 1220648-2025-26138 | 21383801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21383801&pc=PYX | 01/24/2025 | 06/12/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  | 01/31/2026 |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517128 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support. while on support, it was noticed the pa placement signal was increasing and the flow was decreasing on p6. flows were 3l when starting, but down to 2l. it was evident that the pump ingested a clot. heparin was on hold. patient started showing markers of hemolysis including requiring blood transfusions and lactate dehydrogenase increasing from 3000-4000 to upwards of 6000. the pump was replaced with a new rp flex and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of thrombus and hemolysis has been completed. the data was not returned to review. visual inspection found no biomaterial in the pump housing. kinked can was found about 23 mm from delo bonding site of can and if cage. no other issues were found on the rest of the pump. pump was conducted hemolysis test and& passed the test with mih: 6.3. the root cause of hemolysis was unable to be determined since no issue was found on the product and insufficient evidence from data log and clinical review. as the necessary information was not provided, the root cause of the thrombus was not determined | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11643 | 1220648-2025-26139 | 21383976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21383976&pc=OZD | 01/25/2025 | 03/03/2025 |  | 02/14/2025 |  |  | 01/25/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552962 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device as part of bipella mechanical circulatory support (mcs) and experienced access site bleeding. the patient presented with syncope due to ventricular tachycardia/ventricular fibrillation with automated implantable cardioverter-defibrillator (aicd) activation. the patient was started on continuous renal replacement therapy for severe metabolic acidosis and acute renal failure. the patient was taken to the catheterization lab for an impella rp flex and impella cp insertion (bipella mcs). four days later the patient was brought operating room for impella cp escalation to axillary impella 5.5 which was successfully completed. the next day, however, excessive bleeding was noted from the axillary pocket to the jackson-pratt drain. consequently, the patient was taken back to the operating room for a washout and direct anticoagulant application. the patient was reported to be stable following the event and remains on bipella support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26139 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11644 | 1220648-2025-26140 | 21384627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21384627&pc=OZD | 08/02/2024 | 02/14/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103719 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic optical sensor failed, but the motor current and flow remained stable and the support continued. it was noted that the patient was declining and the family withdrew care. the patient expired after 21 days of support. the patient was quite ill. however, the aic did not contribute to the patient's demise. there was no interruption of support. there was note of a controller alarm/optical sensor (unknown if user was aware). however, there is no report or indication of aic exchange. the patient was still getting support. the information indicates the aic supported the patient through the implant experience. | [5 empty] | The investigation for the reported aic - controller error (yellow alarm) issue has been completed. the product was returned. the console logs confirmed that the controller error (defective optical sensor lamp) - alarm was issued. section 23 of the preventative maintenance procedure was performed after cleaning the front optical connector. the console was found to be within specification. the root cause was for the controller error (defective optical sensor lamp) was most probably a defective lamp or/and defective optical bench that is intermittent failure.
11645 | 1220648-2025-26141 | 21384673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21384673&pc=OZD | 01/24/2025 | 03/10/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562302 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and leaking was noted from the sidearm. the impella catheter, just proximal to the purge disc was taped by the registered nurse. the tape was removed and the sidearm was slightly damp with no visible dripping. the team was advice to raise the dextrose level if the impella 5.5 was not going to be replaced. the concentration level was raised and leak stopped. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the root cause of the purge leak - pump cannot be definitively determined as no product was returned for evaluation.
11646 | 1220648-2025-26142 | 21384825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21384825&pc=OZD | 01/24/2025 | 03/12/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586204 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during preparation for impella cp implant, no fluid was noted exiting the pump after priming. a second pump was setup and primed without issue for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of priming issue has been completed. no product was returned for evaluation. data logs were reviewed and real time log of pump priming showed a sudden drop in purge pressure with an increase in purge flow. a ¿purge system open¿ alarm was noted. motor current, motor speed, and placement signal all with values of 0 indicating pump was never inserted into the body or turned on. the logs from the case start showed priming was completed. purge system open alarm was triggered after priming step of case start was completed. clinical details noted that no purge fluid was exiting the impella after priming. the purge cassette was not changed and no leaking was observed from the pump. the root cause of the priming issues could not be definitively determined as no product was returned for evaluation. b.7 added some information that was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12650. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12650 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11647 | 1220648-2025-26143 | 21384828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21384828&pc=OZD | 01/24/2025 | 03/07/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573889 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding from the impella access site. the patient required several dressing changes, and a quick clot was applied to stop the oozing. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26143 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11648 | 1220648-2025-26144 | 21384837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21384837&pc=OZD | 01/24/2025 | 03/10/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  | 11/30/2025 |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428138 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Infusion or Flow Problem | [5 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with presented with elevated troponins and underwent cardiac catheterization which showed multi-vessel disease. the patient was planned for and underwent coronary artery bypass grafting three times (left internal mammary artery to the left anterior descending artery; saphenous vein graft to the obtuse marginal and saphenous vein graft to the right coronary artery) and impella 5.5 placement postoperatively which was successfully completed. the impella was ramped up to support level p-6 with the patient coming off cardiopulmonary bypass. the position was confirmed and the chest was closed. the patient was transported to the unit on impella mechanical circulatory support (mcs). three days after starting support, during a proactive call to the account, pump metrics showed signs of flow saturation (p-9 and 6.0l). the team and the physician were notified of the potential for a pump stop. further on this day, the patient with sustained ventricular tachycardia (140s) was cardioverted one time. two days thereafter, an echocardiogram was completed, and it showed the atrial and mitral valves not opening. the impella inlet appeared to be in the chord of the mitral. the physician decreased the performance level to p-2, broke suction. mitral valve then started opening which was confirmed on echocardiogram. color mosaic was noted at the outlet of the impella 5.5 per echocardiogram. the plan was to administer volume, albumin 500, and get the central venous pressure to 10, and if the volume did not increase the central venous pressure, an epinephrine drip would start. related to the patient's status/condition, it was noted that the patient had just returned from the operating room for repair of arterial cannulation site of extracorporeal membrane oxygenation (ecmo) and received 10 units of packed red blood cells, four units of fresh frozen plasma, platelets, and one unit of cryo. additionally noted the automated impella controller was not displaying left ventricule waveform on p-4. non pulsation motor current. performance level was decreased. the flow of the impella now equaled minimum and maximum flow. echocardiogram showed the pump across valve, but it was noted images were not "great." formal echocardiogram was ordered. to note, the patient was started on another therapy: continuous renal replacement therapy. ten days later, the patient experienced intermittent bouts of atrial fibrillation. following there were suction alarms at performance level p-7 and the need for volume was discussed and ordered. however, the next day the team noted performance level p-7 still could not be maintained without suction. the plan is to continue current therapies. the patient remains on impella mcs. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at, saturated flow, and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11649 | 1220648-2025-26145 | 21384908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21384908&pc=OZD | 01/24/2025 | 03/24/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1187996 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller fell to the floor and housing broke. this issue occurred during patient use and there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - damage during therapy has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the damaged housing was most likely a use issue. e2 and e3 selections were inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2025-26145.
11650 | 1220648-2025-26146 | 21385083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21385083&pc=OZD | 01/25/2025 | 03/12/2025 |  | 02/14/2025 |  |  | 01/25/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582345 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support, and three days into support, the pump stopped. it was further reported wires were frayed/disconnected between the red handle and gray portion of the impella 5.5. the staff was unsure what caused the disconnection. the pump was explanted with a plan to replace. the patient remained hemodynamically stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop and product damage (bond failure) has been completed. the pump was returned for evaluation. upon visual inspection, catheter was separated at the kink protector from the catheter to the red handle. kink protector was found to be torn from red handle. glue joint for catheter to wire connections remained intact around purge line and optical fiber. no data logs were returned for evaluation. clinical details noted that the red handle and grey connector had become disconnected and wires were visible on pump. it was not clear how disconnection occurred but it was noted that patient was sitting up and was agitated. the root cause of the pump stop was most likely due to an electrical open as a result of excessive force on the pump. the failure mode of "product damage (bond issue)" was removed and investigated under "pump stop" as the pump becoming separated at the red handle and kink protector was due to excessive force.
11651 | 1220648-2025-26148 | 21385471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21385471&pc=OZD | 01/20/2025 | 03/05/2025 |  | 02/14/2025 |  |  | 01/20/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566555 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient in cardiogenic shock from an acute myocardial infarction implanted with an impella cp device for mechanical circulatory support experienced limb ischemia and access site bleeding requiring a surgical cutdown and massive blood transfusion. the patient would later expire, but prior the patient was escalated to an impella 5.5 with continued complication. the patient arrived at the hospital via life flight and was intubated. patient with st-segment elevation myocardial infarction to the catheterization lab for intervention coding multiple times in route and on arrival. the impella cp was placed via the right femoral artery. left ventricle access was challenging due to moderate to severe aortic stenosis. after multiple attempts, the physician was finally able to cross the aortic valve into the left ventricle. the impella cp was inserted without an issue. flows 3.7l, mean arterial pressure was 66 on epinephrine and levophed. multiple blood products were infused at this time due to patient hemoglobin of 4.9 at the start of the case. two stents were implanted to the left anterior descending artery. post percutaneous coronary intervention, the 14f x 13 cm peel-away sheath introducer was removed and repositioning sheath placed. digital subtraction angiography was checked to assess for right lower extremity flow; occlusion noted. fem-fem antegrade sheaths were placed and perfusion restored. a large hematoma was noted and profound bleeding around the repositioning sheath occurred. decision made to explant the impella cp in catheterization lab with vascular surgery at lab table for possible dry closure for right common femoral artery (rcfa). the decision was made to take the patient to the operating room for a surgical cut down to rcfa. the patient's hemoglobin was 4.9 on arrival and other bleeding sources were questioned. no further scans were completed prior to arriving in the cath lab. the patient received multiple units of pack red blood cells, fresh frozen plasma and cryoprecipitate were given. it was noted no issues or resistance were noted during insertion nor during the case; patient's hemodynamics improved with the impella cp. however, the patient was in critical condition. the following day, the patient was placed on an impella 5.5. during the procedure the left ventricle access was obtained and the 0.018 guidewire placed. the impella 5.5 was placed in left ventricle and initially was caught up on the mitral valve. it was maneuvered under echo guidance and ultimately was 4.63 cm in the central left ventricle. the device was noted to be working well. additional lines were placed at the end of the case which caused several episodes of ventricular tachycardia which required shock. once the swan was placed, the patient went systolic. a transvenous pacing wire was placed and ultimately rhythm was restored as well as return of spontaneous circulation. impella 5.5 flows increased back to 4.5 lpm. the patient had significant metabolic acidosis leaving the operating room which continued to be corrected. three days later the patient would expire. medical safety review of the event determined there is not enough evidence to exclude impella cp as an associated factor in the patient's outcome. whether the patient recovered or not with escalation is inconsequential because the patient was very sick and experienced "profound" bleeding as a result of the impella cp requiring multiple blood products, and the demised outcome cannot be excluded from impella cp even long term. | [5 empty] | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two device manufacturing reports associated with this event. refer to manufacturing report number 1220648-2025-26084 for the impella 5.5 device report. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation).
11652 | 1220648-2025-26149 | 21385828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21385828&pc=OZD | 01/24/2025 | 03/07/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594137 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that overnight after impella 5.5 was placed some bleeding was noted from the insertion site. pressure dressing was placed, and anticoagulant was stopped. the patient remained stable and on support. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11653 | 1220648-2025-26151 | 21385981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21385981&pc=OZD | 01/24/2025 | 03/04/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591438 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient in cardiogenic shock from an acute myocardial infarction was initially implanted with an impella rp flex device for medical circulatory support (mcs) and unexpectedly the pump flow started decreasing with inaccurate placement signal numbers and then the pump failed. the impella rp flex was emergently replaced with bleeding at the access site. the patient would expire the next day. however, prior the patient was implanted with an impella cp device for bipella support and would experience limb ischemia. the patient came presented to the emergency room with st-segment elevation myocardial infarction and was emergently taken to the catheterization lab. three stents were implanted in the right coronary artery with patient coding. the decision was made to implant an impella rp flex after putting in a swan and seeing signs of right ventricular failure. the patient was transferred to the unit on impella mcs. approximately 10 hours later, the patient needed left ventricular unloading with a left sided impella cp device. the patient was sent back to the unit now on bipella mcs. the plan was to unload and rest the left ventricle, trend down drips as able, and keep up with volume. the patient was to be placed on continuous renal replacement therapy if urine output remained low. the patient was noted to be in very critical condition with no long-term plan, as poly substance use excludes from some long-term therapies. constant discussions were going on with family and physicians to determine the next steps if the patient starts improving after bipella support. it was also noted that the patient had multiple bleeds, some of unknown origin. chest and abdomen were distended in multiple areas showing signs of bleeding from the internal jugular impella rp flex access site and other major bleeding in multiple areas of unknown origin. albumin, and four units of blood were given. eventually, a massive transfusion protocol was started. later in the day, approximately one day after start of impella rp flex support, impella rp flex flows and catheter were running as expected, and then suddenly started dropping rapidly with inaccurate placement signal numbers displaying. the flow dropped from 3.0-3.5, and then dropped down to 2.5 and kept dropping down over a few minutes eventually down to 0.0. there were no alarms, and motor current was showing the motor was running. the performance level was still p-9, and was brought down. the performance level was dropped to p-0, then increased back up to p-4 as an attempt to correct the issue but this was unsuccessful. there was no effect. after the pump stopped working and failed, the patient was brought down to the catheterization lab to replace with another impella rp flex, attempting to use the same site first, and if not successful then move to a new site in the right femoral vein. the physician commented that when informed of the situation, the doctor quickly said to replace the device. further, the physician mentioned the intensive care unit (icu) did not start systemic heparin which he was unaware of. however, the icu said due to the lab values, it was felt additional heparin was not needed for the patient at such time. when the patient came down for replacement of the failed impella rp flex catheter, the distal right extremity was noted to be slightly mottled but had a pulse when checked previously while on the unit. at the end of the impella rp flex replacement, the leg looked worse and more mottled. an angiogram was performed showing blockage of flow passed the impella repo sheath insertion site. an external bypass for distal perfusion was discussed. however, before getting access for it, the plan was changed to remove the current impella cp and place a new one on the left side. after removal on the right, and placement of a new impella cp on the left, both extremities had adequate distal flow with pulses confirmed with doppler and the color of the right lower extremity improved. during the swap of the impella rp flex, a large hematoma developed from swapping the impella rp flex catheters at the same site requiring two units of blood. the patient now is back on bipella support with a newly placed impella cp. however, flows lower than expected despite plenty of volume going in and full impella rp flex flow at performance level p-9. discussion was made regarding possible major bleeds and chest/abdomen very distended in multiple areas. the family was informed of critical status of the patient and urgency of end-of-life discussions if situation did not dramatically improve very quickly. the patient would expire this same day. medical safety review of the event determined there is not enough evidence to exclude the impella rp flex that failed as an associated factor, leading to the demised outcome. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this is one of three device manufacturing reports associated with this event and this report is for the impella cp. refer to the following related two initial medical device reports: impella rp flex serial number (b)(6) with date of event 23-jan-2025 and patient identifier ending in (b)(6). impella rp flex serial number (b)(6) with date of event 24-jan-2025 and patient identifier ending in (b)(6).
11654 | 1220648-2025-26150 | 21385990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21385990&pc=PYX | 01/24/2025 | 03/20/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025507415 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient in cardiogenic shock from an acute myocardial infarction was initially implanted with an impella rp flex device for medical circulatory support (mcs) and unexpectedly the pump flow started decreasing with inaccurate placement signal numbers and then pump failed. the rp flex was emergently replaced with bleeding at the access site. the patient would expire the next day. however, prior the patient was implanted with an impella cp device for bipella support and would experience limb ischemia. the patient came presented to the emergency room as a st-segment elevation myocardial infarction and was emergently taken to the catheterization lab. three stents were implanted in the right coronary artery with patient coding. the decision was made to implant an impella p flex after putting in a swan and seeing signs of right ventricular failure. the patient was transferred to the unit on impella mcs. approximately 10 hours later the patient needed left ventricular unloading with a left sided impella cp device. the patient sent back to the unit now on bipella mcs. the plan was to unload and rest left ventricle, trend down drips as able, and keep up with volume. patient to placed on continuous renal replacement therapy if urine output remained low. the patient was noted to be in very critical condition with no long-term plan, as poly substance use excludes from some long-term therapies. constant discussions were going on with family and physicians to determine next steps if patient starts improving after bipella support. it was also noted the patient had multiple bleeds, some of unknown origin. chest and abdomen distended in multiple areas, showing signs of bleeding from internal jugular rp flex access site and other major bleeding in multiple areas of unknown origin. albumin, and several units of blood were given (4 units of packed red blood cells). eventually a massive transfusion protocol was started. later in the day now approximately one day after start of impella rp flex support, rp flex flows and catheter were running as expected, then suddenly started dropping rapidly with inaccurate placement signal numbers displaying. the flow dropped from 3.0-3.5, then dropped down to 2.5 and kept dropping down over a few minutes eventually down to 0.0. there were no alarms, and motor current was showing the motor was running. the performance level was still p-9, and was brought down. the performance level was dropped to p-0, then increased back up to p-4 as an attempt to correct the issue but this was unsuccessful; there was no effect. after the pump stopped working, failed, patient was brought down to the cath lab to replace with another impella rp flex, attempting to use the same site first, and if not successful, move to a new site in the right femoral vein. the physician commented: when informed of the situation, quickly said to replace the device. further, the physician mentioned the intensive care unit did not start systemic heparin which he was unaware of; however, the icu said due to the lab values, it was felt additional heparin was not needed for the patient at such time. when patient came down for replacement of the failed rp flex catheter, the distal right extremity was noted to be slightly mottled but had a pulse when checked previously while on the unit. at the end of the rp flex replacement, the leg looked worse and more mottled. an angiogram was performed showing blockage of flow passed the impella repo sheath insertion site. an external bypass for distal perfusion was discussed; however, before getting access for it, the plan was changed to remove the current impella cp and place a new one on the left side. after removal on the right, and placement of a new impella cp on the left, both extremities had adequate distal flow with pulses confirmed with doppler and the color of the right lower extremity improved. during the swap of the rp flex a large hematoma developed from swapping the rp flex catheters at the same site requiring two units of blood. the patient now back on bipella support with a newly placed impella cp. however, flows lower than expected despite plenty of volume going in and full rp flex flow at performance level p-9. discussion was made regarding possible major bleeds and chest/abdomen very distended in multiple areas. family was informed of critical status of patient and urgency of end-of-life discussions if situation did not dramatically improve very quickly. the patient would expire this same day. medical safety review of the event determined three is not enough evidence to exclude the rp flex that failed as an associated factor, leading to the demised outcome. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was most likely patient condition since patient had bleeding from multiple sites. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this is one of three device manufacturing reports associated with this event and this report is for the impella rp flex replacement with access bleeding. refer to the following related initial medical device reports: impella rp flex serial number (b)(6) with date of event 23-jan-2025 and patient identifier ending in (b)(6). impella cp serial number (b)(6) with date of event 24-jan-2025 and patient identifier ending in (b)(6). impella rp flex serial number (b)(6) with date of event 24-jan-2025 and patient identifier ending in (b)(6).
11655 | 1220648-2025-26147 | 21388083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21388083&pc=OZD | 01/24/2025 | 03/10/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 10903701 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had several different types of alarms and it was not able to be powered down. a different aic was used and there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. on (b)(6) 2025 the pump was switched from one autommated impella controller to another, immediately unplugged, then plugged back in. after the pump was plugged back in, the logs show two commands "processsamplingdatafromopticalbench: sentstopacquisitioncmd ack" then "clearallopticaldatasharedmemory: 0x4000". this results in "handleopticalerrorconditionswithvalidsample: calculator placement signal 130" which is indicative of the failure mode loss of optical data due to a software defect. all of the failure modes listed in the complaint including inability to shutdown were due to the console not recognizing pump disconnect. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the root cause of the controller error was due to an automated impella controller software issue. b.7 atrial fibrificiency was reported incorrectly on the initial manufacturer device report number. atrial fibrillation was added.
11656 | 1220648-2025-26153 | 21388571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21388571&pc=OZD | 01/15/2025 | 02/26/2025 |  | 02/14/2025 |  |  | 01/15/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570545 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support, and 19 days after start of the support, the physician encounter repositioning difficulty. reportedly, when repositioning the pump, advancing forward and pulling back, the physician stated there was "something" that almost felt like it was catching and impeding repositioning. the device remained implanted supporting the patient until approximately 22 days later when it was successfully weaned and explanted. the patient was reported to be stable following the event and explant of the device. | [5 empty] | The investigation for the positioning issue has been completed. data logs are not relevant to the complication. returned product was evaluated and there were several abnormalities noted around the repositioning unit. first of which, the blue butterfly was not aligned parallel with the cath anchor button, as it would be out of box. additionally, after cutting back the sterile sleeve, a kink could be seen in the catheter about 0.5 cm proximal to the cath anchor. the blue button was able to press and release without problem. when attempting to advance and pull back the catheter through the hub, it required some force, but was able to move. the cath anchor was slid down toward the motor housing, and the section of the catheter that was returned inside the hub was cleaned of dried blood. at this point, it was able to advance back and forth easily over that section. however, when trying to advance the kink in the catheter that came back 0.5 cm from the hub housing forward through it, excessive force was required. there were no failures of the product to meet design requirements based on the fact that the blue button functioned properly and the catheter was able to slide through the hub normally, especially once the catheter was cleaned. based on returned product, the root cause of the positioning issues were determined to be a cath anchor use issue. what caused the initial resistance is unclear. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11657 | 1220648-2025-26152 | 21388856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21388856&pc=OZD | 01/23/2025 | 03/11/2025 |  | 02/14/2025 |  |  | 01/23/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597607 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and six days following the implant, low pump flow was encountered. the device was explanted with an unsuccessful attempt to place an impella 5.5. however, another impella cp was implanted. the patient presented to the hospital with chest pain and vomiting and was found to have st-segment elevation myocardial infarction and cardiogenic shock with ejection fraction of 10-15%. the patient had a stent placed five days prior. the patient decompensated in the emergency department and was taken back to the catheterization lab for an impella cp device placement, which was successfully completed. the patient was transferred to the unit on support. approximately six days after start of the impella mcs, a continuous suction alarm was reported. there were low flows at high p levels, but no known hemolysis. the patient was transferred to another hospital for a higher level of care and advanced therapies evaluation. the decision was made to escalate to an impella 5.5 device. the patient was brought to the operating room, prepped/draped for left axillary impella 5.5 device implant. the left axillary was cut down and graft was sewn directly onto the artery and tunnel inferior. left ventricle was then accessed via wire/catheter technique under fluoroscopy/transesophageal echocardiogram (tee). the .018 guidewire was loaded and placed into the left ventricle and the catheter was removed. the impella 5.5 was then backloaded onto the wire. however, there were failed attempts to cross the anastomosis. the impella 5.5 was then removed and an impella cp was opened and placed. the placement was confirmed with tee, measuring at 3.7 centimeters. the impella cp was ramped to p-8 with flows of 2.3 liters per minute. chimney graft/repo sheath was sutured to the left chest. sterile dressing was applied to the left axillary site per their protocol. the patient was transferred to the bed for transport back to the unit, but upon leaving the room, the patient went into ventricular fibrillation. the patient was shocked two times and returned to normal sinus rhythm. tee probe was re-inserted and placement was confirmed and the patient was taken to the unit. same day later that night on the unit, the automated impella controller displayed/sounded continuous suction alarm. replacing pa catheter and plan to obtain an echocardiogram as soon as possible to confirm placement. the next day, axillary impella cp continued on p-8 and continuous suction alarms continue. audible alarms were turned off. flows of 2.5 were noted. urine output was clear/yellow. follow up the next day noted impella mcs continued with patient¿s prognosis as guarded. ecpella (extracorporeal membrane oxygenation and impella mcs) was considered but the patient was not a candidate for ecmo. the plan was to wean the patient and assess the patient¿s response. however, the patient would expire and care was withdrawn. medical safety review of the event determined there is not enough evidence to exclude the replacement impella cp as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of three devices related to the patient's implant experience. refer to manufacturing report number 1220648-2025-26083 for the impella 5.5. initial medical device report for the replacement impella cp serial number (b)(6) with date of event 26-jan-2025 and patient identifier ending in (b)(6). | The investigation of low pump flow has been completed. product was not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
11658 | 1220648-2025-26154 | 21389618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21389618&pc=OZD | 01/24/2025 | 03/12/2025 |  | 02/14/2025 |  |  | 01/24/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487692 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had access site bleeding. the patient lost an estimated two liter of blood therefore four units of packed red blood cells were given to the patient. the bleeding was controlled and the patient remained stable. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related since non- abiomed product was used. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26154 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11659 | 1220648-2025-26155 | 21389630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21389630&pc=OZD | 01/25/2025 | 03/06/2025 |  | 02/14/2025 |  |  | 01/25/2025 | 02/14/2025 |  |  |  |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599729 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support a patient in for surgery. the pump was supporting for three days when the pump stopped and could not be restarted despite all troubleshooting measures. the medical doctor at bedside discontinued impella, surgical closure and stapled shut. the family decided to withdraw care as the patient status continued to decline and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the pump stop has been completed. the pump was inspected and biomaterial was found in the outlet cage and around the impeller. the biomaterial did not clear upon soaking. the data logs from the case showed the impella stopped - motor current high alarms. there was a motor current spike with speed deviation and placement signal shift characteristic of biomaterial ingestion. the root cause of the pump stop was determined to be due to biomaterial ingestion. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26155. h.6 added code to health effect - impact code as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26155. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11660 | 1220648-2025-26156 | 21389844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21389844&pc=PYX | 01/23/2025 | 03/25/2025 |  | 02/14/2025 |  |  | 01/23/2025 | 02/14/2025 |  |  | 02/28/2026 |  | 02/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025478682 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Pumping Stopped | Failure to Sense |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a patient in cardiogenic shock from an acute myocardial infarction was initially implanted with an impella rp flex device for medical circulatory support (mcs) and unexpectedly the pump flow started decreasing with inaccurate placement signal numbers and then pump failed. the rp flex was emergently replaced with bleeding at the access site. the patient would expire the next day. however, prior the patient was implanted with an impella cp device for bipella support and would experience limb ischemia. the patient came presented to the emergency room as a st-segment elevation myocardial infarction and was emergently taken to the catheterization lab. three stents were implanted in the right coronary artery with patient coding. the decision was made to implant an impella p flex after putting in a swan and seeing signs of right ventricular failure. the patient was transferred to the unit on impella mcs. approximately 10 hours later the patient needed left ventricular unloading with a left sided impella cp device. the patient sent back to the unit now on bipella mcs. the plan was to unload and rest left ventricle, trend down drips as able, and keep up with volume. patient to placed on continuous renal replacement therapy if urine output remained low. the patient was noted to be in very critical condition with no long-term plan, as poly substance use excludes from some long-term therapies. constant discussions were going on with family and physicians to determine next steps if patient starts improving after bipella support. it was also noted the patient had multiple bleeds, some of unknown origin. chest and abdomen distended in multiple areas, showing signs of bleeding from internal jugular rp flex access site and other major bleeding in multiple areas of unknown origin. albumin, and several units of blood were given (4 units of packed red blood cells). eventually a massive transfusion protocol was started. later in the day now approximately one day after start of impella rp flex support, rp flex flows and catheter were running as expected, then suddenly started dropping rapidly with inaccurate placement signal numbers displaying. the flow dropped from 3.0-3.5, then dropped down to 2.5 and kept dropping down over a few minutes eventually down to 0.0. there were no alarms, and motor current was showing the motor was running. the performance level was still p-9, and was brought down. the performance level was dropped to p-0, then increased back up to p-4 as an attempt to correct the issue but this was unsuccessful; there was no effect. after the pump stopped working, failed, patient was brought down to the cath lab to replace with another impella rp flex, attempting to use the same site first, and if not successful, move to a new site in the right femoral vein. the physician commented: when informed of the situation, quickly said to replace the device. further, the physician mentioned the intensive care unit did not start systemic heparin which he was unaware of; however, the icu said due to the lab values, it was felt additional heparin was not needed for the patient at such time. when patient came down for replacement of the failed rp flex catheter, the distal right extremity was noted to be slightly mottled but had a pulse when checked previously while on the unit. at the end of the rp flex replacement, the leg looked worse and more mottled. an angiogram was performed showing blockage of flow passed the impella repo sheath insertion site. an external bypass for distal perfusion was discussed; however, before getting access for it, the plan was changed to remove the current impella cp and place a new one on the left side. after removal on the right, and placement of a new impella cp on the left, both extremities had adequate distal flow with pulses confirmed with doppler and the color of the right lower extremity improved. during the swap of the rp flex a large hematoma developed from swapping the rp flex catheters at the same site requiring two units of blood. the patient now back on bipella support with a newly placed impella cp. however, flows lower than expected despite plenty of volume going in and full rp flex flow at performance level p-9. discussion was made regarding possible major bleeds and chest/abdomen very distended in multiple areas. family was informed of critical status of patient and urgency of end-of-life discussions if situation did not dramatically improve very quickly. the patient would expire this same day. medical safety review of the event determined three is not enough evidence to exclude the rp flex that failed as an associated factor, leading to the demised outcome. | [5 empty] | The investigation of low pump flow, placement signal issue, and vascular damage - peripheral vessel has been completed. the low pump flow failure mode was removed as reason for reported low pump flow is encapsulated under the placement signal issue. the returned product was inspected and there was no biomaterial in the inlet cage and on the sensor. the data logs were reviewed and dp sensor drift was observed with a rapid drift up of placement signal simultaneous with a rapid drift down of pump flows. placement signal and cvp were both noisy with cvp negative for most of the support duration before trending up during ps and flow drift. the pump was tested in the flow loop and the sensor drift reproduced with #1051 placement signal not reliable - dp sensor alarm. the root cause of the placement signal issue was determined to be due to sensor drift as issue reproduced on returned product. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information related to failure was provided b1 was revised from the initial manufacturer device report number 1220648-2025-26156 to reflect both adverse event and product problem e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26156 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 medical device problem code 1503 was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2025-26156. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. this is one of three device manufacturing reports associated with this event and this report is for the initially implanted impella rp flex device that failed. refer to the following related initial medical device reports: impella rp flex serial number (b)(6) with date of event 23-jan-2025 and patient identifier ending in (b)(6). impella cp serial number (b)(6) with date of event 24-jan-2025 and patient identifier ending in (b)(6). impella rp flex serial number (b)(6) with date of event 24-jan-2025 and patient identifier ending in (b)(6).
11661 | 1220648-2025-26123 | 21373618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21373618&pc=OZD | 01/22/2025 | 03/12/2025 |  | 02/13/2025 |  |  | 01/22/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529402 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Failure to Advance | Detachment of Device or Device Component | [5 empty] | Foreign Body Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support and encountered delivery difficulty. the impella access site was the femoral access despite knowledge of "severe" peripheral artery disease. the long 25 centimeter (cm) sheath failed to deliver due to the iliac artery kinking. the physician instead placed a 13 cm sheath. however, the impella cp was unable to advance through and the decision was made to remove the impella cp. during the removal, the inlet tore, broke off and became stuck in the vessel. consequently, a snare was used to retrieve the separated piece. the physician commented there was no hard pulling and felt immediately the device tore. additionally, the physician commented the pump break occurred when the device was pulled back at the entrance to the stent. the patient was reported to be stable at the end of the procedure. | [5 empty] | The investigation of product damage (detached inflow/outflow ¿ peripheral vessel and introducer damage ¿ mechanical has been completed. the returned pump was inspected and product damage was confirmed. the cannula showed detachment of the inlet. clinical details mentioned calcified vessel with severe peripheral arterial occlusion. details mention the pump got stuck during explant and inlet tore off. the root cause of the product damage (detached inflow/outflow - peripheral vessel) was most likely use related as evidenced by cannula stretching observed on returned product. the root cause of the introducer damage - mechanical was not determined as no product or images of sheath were returned for analysis a5 ethnicity was inadvertently omitted on the manufacturer device report number 1220648-2025-26123. | The pump, a sheath and two snare kits were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11662 | 1220648-2025-26124 | 21373772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21373772&pc=OZD | 01/24/2025 | 03/12/2025 |  | 02/13/2025 |  |  | 01/24/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552568 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and subsequently escalated to an impella 5.5 device. the pump was implanted; however, issues were encountered with the 23fr sheath being delivered and kinking. the team had to place a new graft lock to achieve hemostasis. the pump supported the patient for a total of 6 days and was then weaned and explanted. the patient was reported to be in stable condition following the event and explant of the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation has been completed. the product was not returned for investigation. according to the provided clinical details, the bleeding was resolved after applying an extra graft lock past the introducer kink. the patient was reported to have a tortuous vessel condition. the introducer damage failure mode was removed since it was addressed as the root cause of the access site bleeding against the axillary insertion kit. the cause of the access site beleeing issue was most likely introducer sheath kink due to patient tortuous vessel condition, based on provided clinical detail. b.1 revised to only reflect adverse event since the initial manufacturer device report number 1220648-2025-26124 was submitted due to investigation results. b.7 revised as information was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26124. h.6 revised component code as previous code was reported incorrectly on the initial manufacturer device report number 1220648-2025-26124. revised medical device problem code since the initial manufacturer device report number 1220648-2025-26124 was submitted due to investigation results. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26124 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
11663 | 1220648-2025-26125 | 21373899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21373899&pc=OZD | 01/25/2025 | 03/25/2025 |  | 02/13/2025 |  |  | 01/25/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569940 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support and eight days after start of the support, the pump stopped. reportedly, there was high motor current with saturated flows before the pump stopped. as a result, the pump was explanted. therapy/treatment provided to after explant of the device is unknown. however, the patient was reported to be in stable condition after the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop and saturated flow has been completed. 2the pump was returned. data logs were not recovered. visual inspection showed extensive biomaterial build up and both inflow and outflow cages. the pump was unable to be run with the biomaterial on the impella. the pump was soaked in warm water and tergazyme and biomaterial was removed. the pump was run at both p-4 and p-9 without issue. impella defective alarm not reproduced. pump stop could not be reproduced. the cause of the pump stop was most likely biomaterial as biomaterial was found on returned product and pump stop reproduced; removing the biomaterial enabled the pump to run. the cause of the saturated flow was most likely biomaterial due to biomaterial found on returned product; saturated flows reproduced in the lab. d9 date device returned to manufacturer added. g1 manufacturing site name and address was erroneously on the initial manufacturer device report number 1220648-2025-26125.
11664 | 1220648-2025-26126 | 21375001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21375001&pc=OZD | 12/03/2024 | 03/06/2025 |  | 02/13/2025 |  |  | 01/23/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571055 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding with a "definite" relationship to the impella cp device used: ¿on return from cath lab on night of (b)(6) 2024 for ecmo + impella placement, pt had extensive bleeding around impella site, evaluated by interventional fellow who tried but did not improve. required transfusion of 3 units prbcs, 1 of platelets, 2 of cryo, 1 of ffp. heparin held, continued on asa but plavix held for bleed. were able to eventually control bleeding site and readily has been able to be weaned off pressors. off pressors as of (b)(6) 2024, with groin bleeding improved and stable hemoglobin. hemoglobin 11.3 prior to impella placement, at 8.8 morning of (b)(6) 2024.¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be anticoagulation management because heparin was held and bleeding was controlled at the site. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11665 | 1220648-2025-26127 | 21375366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21375366&pc=OZD | 10/01/2024 | 03/03/2025 |  | 02/13/2025 |  |  | 01/23/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541756 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that multiple complications were encountered during impella cp support. it was documented that the patient endured ventricular arrhythmia with a probable relationship to the impella device. the patient was bolused with 50mg of amiodarone and cardioversion was performed with noted improvement of their heartrate. it was additionally reported that the patient experienced bleeding with a probable relationship to the impella device. the patient was given 2 units of packed red blood cells treat the dropping hemoglobin levels. due to the patient's condition, the family decided to move the patient to comfort care where they expired. it is unknown if the conditions contributed to the patient's passing. | [5 empty] | The investigation for the vascular damage and arrhythmia has been completed. the impella was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the arrhythmia was not determined b.2 added required intervention as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26127. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26127 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11666 | 1220648-2025-26128 | 21375443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21375443&pc=OZD | 01/23/2025 | 03/12/2025 |  | 02/13/2025 |  |  | 01/23/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578582 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, there was a slight hematoma at axillary site. pressure dressing applied and heparin was on hold. a possible gastrointestinal bleed was noted. the patient remains on support. | [5 empty] | The investigation for the access site bleed and vascular damage has been completed. as per the clinical information provided, the doctors resumed the systematic anticoagulation very soon which led to access site bleeding. once the anticoagulation was stopped, the bleeding was stopped as well. therefore, the root cause of the access site bleeding issue was most likely improper anticoagulation management due to patients high act's. as per the clinical information provided mentions that the patient had a possible gastrointestinal bleeding but the information about the location and the reason of the bleeding is unknown. therefore, the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and unknown relation to impella with gi bleed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11667 | 1220648-2025-26129 | 21375947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21375947&pc=OZD | 01/23/2025 | 02/21/2025 |  | 02/13/2025 |  |  | 01/23/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594901 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella 5.5 was inserted a large hematoma was noted at the axillary pocket. the patient was returned to the catheterization lab, and the pocket was reopened to find the graft anastomosis was compromised on the axillary artery. bleeding was controlled by surgical repair of the axillary artery and the impella 5.5 was explanted at that time. transfusion was required and the patient remained stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
11668 | 1220648-2025-26130 | 21376059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21376059&pc=OZD | 01/14/2025 | 03/12/2025 |  | 02/13/2025 |  |  | 01/23/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583029 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient developed hemolysis during impella cp support. it was indicated that there was a definite relationship between the condition and the impella device used. the patient's plasma free hemoglobin levels were elevated at the time of the diagnosis. the patient was transfused with one unit of packed red blood cells to treat the condition. due to the patient's overall condition, the patient's family made the decision to make the patient do not resuscitate and the pump was removed. the patient subsequently passed away. it can not be determined if the hemolysis contributed to the patient's passing. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11669 | 1220648-2025-26131 | 21381392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21381392&pc=OZD | 01/23/2025 | 03/13/2025 |  | 02/13/2025 |  |  | 01/23/2025 | 02/13/2025 |  |  | 12/31/2025 |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443386 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the pump was repositioned and it was accidentally pulled into the aorta. the pump was explanted and replaced. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the product was not returned for review. as per clinical information, the pump was accidently pulled out into aorta. based on clinical description, the root cause of positioning issue was most likely use related due to improper technique. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26131 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11670 | 1220648-2025-26132 | 21381394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21381394&pc=OZD | 01/19/2025 | 03/25/2025 |  | 02/13/2025 |  |  | 01/24/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584214 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | Cardiac Perforation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Complainant reported the patient was on impella cp smart assist support and during support, the patient's left ventricle ruptured. the cause of the rupture is unknown and the relationship between the rupture and the impella is unknown. the ventricle rupture caused the patient to expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. as per the clinical information provided, after two days of pump insertion, there was a ventricle perforation observed ie. left ventricle rupture, but the relation to impella is unknown. therefore, the root cause of the ventricle perforation issue was not determined since the product was not returned and insufficient clinical information was provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26132 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2152 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26132. new code was added to health effect - clinical code.
11671 | 1220648-2025-26133 | 21381399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21381399&pc=OZD | 01/24/2025 | 03/14/2025 |  | 02/13/2025 |  |  | 01/24/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1284914 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller had a system self-check error. this issue occurred during support. no further details were provided. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs from the reported complaint date were consistent with the complaint. during the investigation the reported system self-check issue was not able to be reproduced. the voltage between the power battery manager (pbm) and the impellatronic was measured to see if there was an issue with the power supplied to the impellatronic but no issues were found. the cable between the impellatronic and 4-in-1 was also inspected but no issues were found. the root cause of the system self check issue could not be determined due to the inability to reproduce but the suspected components involved with the communication issues are the impellatronic and 4-in-1. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26133 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11672 | 1220648-2025-26134 | 21381401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21381401&pc=OZD | 01/17/2025 | 03/12/2025 |  | 02/13/2025 |  |  | 01/24/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559984 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Contamination | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was contaminated by a staff member and the cp was discarded. a decision was made not to place an additional impella. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be use issue because the impella was contaminated before the insertion by the staff member. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
11673 | 1220648-2025-26135 | 21381403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21381403&pc=OZD | 01/24/2025 | 03/07/2025 |  | 02/13/2025 |  |  | 01/24/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580929 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, a computerized tomography was performed showing an acute cerebellum infarct. it was reported that the patient went into atrial fibrillation requiring amio bolus and an a iv drip. prior to event, patient had been sub therapeutic on heparin according to partial thromboplastin time. additionally on support, it was noted the patient a small hemorrhagic right parietal lobe infarct following the known ischemic right cerebral event. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the stroke and atrial fibrillation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke and atrial fibrillation could not be determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ a.1 revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26135. d.1 revised brand name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26135. d.4 revised model number, catalog number, lot number, serial number and unique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26135. e.1 revised first and last name as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26135. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26135 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26135. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2025-26135 and were added.
11674 | 1220648-2025-26136 | 21381407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21381407&pc=OZD | 01/25/2025 | 03/12/2025 |  | 02/13/2025 |  |  | 01/25/2025 | 02/13/2025 |  |  |  |  | 02/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563709 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. however, following the implant, it was noted that support was inadequate above performance level p-6. the patient was taken into the intensive care unit where the position was adjusted, and both the volume and heart status were evaluated. during troubleshooting, the pump unexpectedly stopped and was able to be turned back on. however, flows would not increase beyond .3-.5. during this time, the patient coded multiple times. the patient's family ultimately made the decision to withdraw care, and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vf/vt/af/at, low pump flow, and temporary pump stop has been completed. the pump was not returned for investigation. the data log shows low pump flow with an active suction alarm while running on a higher p-level setting. the pump flow issue was resolved after turning to p2. a pump stop was reported with a motor current high alarm, but it was able to restart. however, the pump was unable to achieve proper flow on p2 after the restart. additionally, the patient was coded multiple times while on support. the cause of the temporary pump issue was most likely biomaterial, based on data log review. the cause of the low pump issue was most likely biomaterial, based on data log review. as the necessary information was not provided, the root cause of the vt/vf was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26136 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11675 | 1220648-2025-26100 | 21363344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21363344&pc=OZD | 01/22/2025 | 03/11/2025 |  | 02/12/2025 |  |  | 01/22/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595466 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis following impella cp implant. the catheter was repositioned however the condition persisted. in response to the hemolysis, an attempt was made to escalate to impella 5.5 support, but the patient's anatomy prevented successful delivery of the device. the pump was subsequently left in despite the noted condition and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11676 | 1220648-2025-26101 | 21363430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21363430&pc=OZD | 01/23/2025 | 03/11/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576448 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and on day nine, while in the operating room, the impella cp device was unintentionally pulled from left ventricle during extracorporeal membrane oxygenation (ecmo) decannulation. the heart team decided to go ahead and pull impella cp device. once removed, it was noted the patient initially looked fine but started to decompensate. cardiovascular thoracic surgery and heart failure discussed the possibility of placing an impella 5.5, but patient having hematologic issues. the team ultimately decided to send the patient back to the unit, get the hematologic issues better controlled, and then resume discussions around next steps. therapies after explant were inotropes/vasopressors and respiratory support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related due to improper technique.
11677 | 1220648-2025-26102 | 21363489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21363489&pc=OZD | 01/22/2025 | 02/21/2025 |  | 02/12/2025 |  |  | 01/22/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578333 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a thrombus was discovered in the left ventricle during impella cp support. it is unknown if intervention was required to manage the condition. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
11678 | 1220648-2025-26103 | 21363571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21363571&pc=OZD | 01/22/2025 | 05/22/2025 |  | 02/12/2025 |  |  | 01/22/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084348 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during initial priming of impella cp, placement signal sensor alarm triggered. the case was started and initially the pump was on auto with no placement signal on the automated impella controller (aic). a back up aic was obtained and was successfully switched, connected to the impella cp device. without further issues, the impella cp remained connected to the console until patient escalation to an impella 5.5 device. there was no patient harm and no report or indication of interruption in support. | [5 empty] | The investigation of automated impella controller (aic)- loss of opm data has been completed. bsed on the clinical details and data and device analysis, the root cause of the optical sensor not supported, was traced back to an incorrect hardware revision for sau_motor_1. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | Investigation revised to reflect that while testing the automated impella controller (aic), it was noted that the rotary encoder (push button) was non-operational, due to the rpb malfunction which was caused by loss of vcc_pwr (vcc) due to a blown fuse f13 on the power battery manager (pbm) (defective pbm). the root cause for the rotary encoder (push button) being non-operational was due to defective pbm h6 medical device problem code 4063 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26103. h10 incomplete investigation results 1220648-2025-26103 on the initial manufacturer device report number.
11679 | 1220648-2025-26104 | 21363621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21363621&pc=OZD | 01/24/2025 | 03/04/2025 |  | 02/12/2025 |  |  | 01/24/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1620355 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a system self-check failure occurred when the automated impella controller (aic) was turned on. the aic was immediately removed and replaced. there was no patient harm and no report or indication of delay in support to a patient. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs from the reported day of event are consistent with complaint and show system self-check failures due to persistent communication failure. the console was tested and issue was reproduced. inspection of console¿s electronics revealed corrosion on the power batter manager (pbm) near supercaps affecting serial communication with pbm1. the cause of the aic issue was contamination. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11680 | 1220648-2025-26105 | 21363639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21363639&pc=OZD | 01/24/2025 | 03/13/2025 |  | 02/12/2025 |  |  | 01/24/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591428 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed high purge pressure the following day. the patient was brought to the catheterization lab for balloon aortic valvuloplasty and impella cp placement post intervention which was successfully completed. a transcatheter aortic valve replacement was to be planned for a later date. the patient was transferred to the unit on support. the following day, the console displayed purge flow blocked and purge pressure high alarms. no kinks in tubing or catheter were noted. tissue plasminogen activator was ordered for the purge. it was discussed the patient was stable enough on low impella support and drips to explant, given overall plan which was to try to extubate. the physician was also concern for hemolysis with rising lactate dehydrogenase and felt explanting the device would help with the patient's overall neurological status. the patient was taken to operating room and a surgical cutdown was performed. the impella cp device was weaned to performance level p-2, then to p-0 while pulling across the valve and was explanted. groin was hemostatic prior to leaving operating room. the patient survived the explant and returned to unit for further hemodynamic monitoring. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. the data logs were reviewed and the high purge pressure event was confirmed. on the 24th, purge pressure rose over a 5-minute period until the alarm triggered. it resolved after four hours. leading up to the event, there was no clear motor current spikes. however, pump flows were trending down and out of specification for p-5 and there were some slight disturbances in the motor current, motor speed, and placement signal. the returned product was investigated and there were no abnormalities or signs of biomaterial. the pump was connected to a console, primed, and run in a bucket for ten minutes, and high purge pressure was not reproduced. this aligned with the clinical details provided that tissue plasminogen activator (tpa) was used and resolved the complication. based on tpa resolving the high purge pressure in the field and it not reproducing in the lab, the root cause of the high purge pressure was most likely biomaterial.
11681 | 1220648-2025-26106 | 21363747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21363747&pc=OZD | 01/24/2025 | 03/19/2025 |  | 02/12/2025 |  |  | 01/24/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583011 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. one day after start of support, a suction alarm appeared and flows suddenly decreased. performance levels were decreased, and volume was given. however, the issue persisted and a sudden increase in motor current occurred. clot entrainment was suspected, though not confirmed, and the pump was subsequently replaced. the patient was reported to be stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
11682 | 1220648-2025-26107 | 21363842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21363842&pc=OZD | 01/23/2025 | 03/21/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525915 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was supporting for over 8 days when the pump stopped. the stopped pump was removed and replaced by an impella 5.5 pump. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of pump stop has been completed. product was not returned for investigation, and data logs were not available. clinical details report a high motor current pump stop on the 9th day of impella cp support. since neither product nor logs were returned for investigation, the root cause of the pump stop could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11683 | 1220648-2025-26108 | 21364114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21364114&pc=OZD | 01/24/2025 | 03/12/2025 |  | 02/12/2025 |  |  | 01/24/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605335 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp pump device for mechanical circulatory support. the impella cp was care escalation from an intra-aortic balloon pump. the cp pump was placed; however, the 14fr sheath had a crack, not a the tip, which caused oozing from the access site. the pump supported for 25 hours and was then explanted. reportedly, there was no patient harm. | [5 empty] | The investigation of introducer damage ¿ mechanical has been completed. no information was provided about how the introducer was damaged or patient's vasculature. product was not returned as well. without the product and necessary clinical information, the failure cannot be investigated. therefore, the root cause of the introducer damage issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11684 | 1220648-2025-26109 | 21364181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21364181&pc=OZD | 01/23/2025 | 03/06/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563708 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in acute myocardial infarction and cardiogenic shock when in the intensive care unit there was an observation made of a hematoma. the hematoma was treated by premature explant. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11685 | 1220648-2025-26110 | 21364196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21364196&pc=OZD | 01/24/2025 | 03/14/2025 |  | 02/12/2025 |  |  | 01/24/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337492 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported when the automated impella controller (aic) was started, an unexpected shutdown message was displayed. consequently, the aic was not used. there was no report or indication of patient harm or delay in support. | [5 empty] | The investigation automated impella controller (aic) - shutdown or restart issue has been completed. on 24-jan-2025 while booting up the console, there was failed ossi (optical) process triggering a partial reset of the software (sw) watchdog. the console rebooted but again had failed ossi triggering another partial reset. again the ossi process triggered the sw watchdog but since this was the third watchdog trigger, a full reset was initiated (by design). the console then booted up as expected and triggered an unexpected controller shutdown alarm descriptive of the previous watchdog reboots. the console was returned for investigation and was power cycled twenty times. the first failed boot up during testing was from a qnx process failure that leads to all processes failing and the console automatically rebooting. the second failure was a variety of different communications failing triggering a reboot. the 4-in-1 was replaced with a known good and the console was able to power cycle 20 times as expected. the console was unexpectedly rebooting due to a defective 4-in-1. d.9 date device returned /information was added. d.5 revised operator of device as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26110. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26110 as it is unknown if the initial reporter also sent the report to the food and drug administration. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11686 | 1220648-2025-26111 | 21364205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21364205&pc=OZD | 01/23/2025 | 03/04/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1380525 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller displayed a battery failure alarm. during troubleshooting, the nurse pressed the battery button firmly and the alarm resolved. there was no noted patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the automated impella controller (aic) battery failure-red alarm has been completed. console logs were unavailable since the console was not returned. the logs did not contain any relevant data to the complaint since the issue was resolved in the field and was determined to be use related. the console was not returned. the reported battery failure alarm was resolved in the field and was determined to be use related. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26111 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11687 | 1220648-2025-26112 | 21364225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21364225&pc=OZD | 01/23/2025 | 03/13/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025595630 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the there was an impella stopped alarm that self resolved after going over a bump in the elevator. staff stated that the impella never stopped and the automated impella controller never turned off. there was no patient harm and the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of temporary pump stop has been completed. the product was not returned for investigation. data logs were reviewed and a temporary pump stop did occur, but on (b)(6) 2025. suddenly, without any abnormalities leading up to it, an alarm triggered with a motor current spike to 700 ma. the pump restarted two seconds after the alarm and purge pressure was unaffected. these facts suggest the issue most likely was not an electrical failure nor biomaterial ingestion. clinical details stated that the temporary pump stop occurred during patient movement when rolling over a bump in the elevator. data logs also showed that this same sequence of events occurred on (b)(6) 2025 as well, though it was not noticed/reported. based on clinical details provided and data log review, the root cause of the pump stop was most likely a use issue during patient movement. d.6b explant date was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26112 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11688 | 1220648-2025-26113 | 21364301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21364301&pc=OZD | 01/24/2025 | 03/07/2025 |  | 02/12/2025 |  |  | 01/24/2025 | 02/12/2025 |  |  | 01/31/2026 |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453279 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella 5.5 device to a patient with cardiomyopathy was unsuccessful due to the cardiac anatomy. the device was unable to pass the aortic valve and the case was aborted. there was no report of adverse effects to the patient. the patient was being escalated to an impella 5.5 from an impella cp device. during implant of the impella, through the right axillary approach, once the 0.035 guidewire and catheter were inserted into the ascending aorta, the surgeon stopped and discussed the "severe" angles of the patient¿s anatomy. the surgeon removed the guidewire and shot contrast to better image the anatomy. carotid approach was discussed; however, the surgeon agreed that the angle was too acute for the device and refused sternotomy. therefore, the impella 5.5 device implant was aborted. the patient was transferred back to the unit in stable condition with the impella cp in place and working well. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number. a new code was added.
11689 | 1220648-2025-26114 | 21364508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21364508&pc=OZD | 12/27/2024 | 02/19/2025 |  | 02/12/2025 |  |  | 01/13/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515222 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Detachment of Device or Device Component | Device Sensing Problem |  |  |  |  | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impellla 5.5 was placed on a patient needing/waiting a heart transplant. an impella defective placement signal occurred for a portion of support. the impella was also flipped towards the mitral valve and repositioning was unsuccessful. additionally, the sterile sleeve was no longer intact, so tape was used to tape the sleeve to the white catheter. it was further reported that the patient had a cerebrovascular accident and computed tomography scans twice a day for three days were done. due to these issues, the physician exchanged the impella 5.5 for a new one and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the reported issues has been completed. neither data logs nor the pump were returned for investigation. clinical details reported that the placement signal values on the console were very different from the readings from the blood pressure cuff, but no information was provided that would suggest a cause for the issue. since the product was not returned for investigation, the root cause of the optical signal issue could not be determined. clinical details reported that during an echocardiogram the angle of the pump had appeared to shifted toward the mitral valve. it also mentioned that a previous attempt to reposition the pump was unsuccessful, and no further information was provided regarding what may have caused the position change or how it was troubleshot. since the product was not returned and limited clinical details were provided, the root cause of the positioning issues could not be determined. data logs were not relevant to this complication and the product was not returned for investigation. clinical details reported that the sterile sleeve had separated from its distal hub and was taped down to the catheter. however, it is not reported when or how the sterile sleeve became disconnected, and whether that involved tearing/excessive force. since the product was not returned and insufficient clinical details were provided, the root cause of the product damage could not be determined. as this clinical information was not provided, the true root cause of the cerebrovascular accident could not be determined.
11690 | 1220648-2025-26115 | 21365150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21365150&pc=OZD | 01/23/2025 | 03/13/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1656412 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a controller error during preparation prior to the case and patient use. the aic was turned off and on and it did not reproduce the error. however, the aic was replaced due to potential patient harm. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. after the console was booted up, the optical signals were abnormal; specifically, the light was 0.45 volt and the signal-to noise ratio was 1.6, which resulted in the ¿controller error (defective optical sensor lamp) ¿ alarm¿. this confirmed the reported failure mode. the console was turned on again, and there was no controller error. the console was tested. the reported failure mode was unable to be reproduced. voltage from the power battery manager to the optical bench was 4.9 volt, and the average voltage to the lamp was 4.8 volt. no issue was found with the analog output from the optical bench. the front panel optical connector was found contaminated and was unable to remove the debris upon cleaning (most likely unrelated to the reported failure mode). the root cause of the intermittent controller error was most likely due to a defective lamp. a.3 revised as gender was reported incorrectly on the initial manufacturer device report number 1220648-2025-26115. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26115 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.5 revised labeled for single use as it was reported incorrectly on the initial manufacturer device report number 1220648-2025-26115. h.8 revised usage of device as it was reported incorrectly on the initial manufacturer device report number 1220648-2025-26115. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11691 | 1220648-2025-26116 | 21366724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21366724&pc=OZD | 01/14/2025 | 06/12/2025 |  | 02/12/2025 |  |  | 01/14/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594363 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | Malposition of Device | Infusion or Flow Problem |  |  |  |  | Atrial Fibrillation | Atrial Flutter | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and while on support, the patient experienced atrial fibrillation/flutter and pump saturation with pump failure. the patient was brought to the catheterization lab for a right and left heart catheterizations. during the heart catheterization, the patient decompensated. a lesion was noted in the left anterior descending (lad) artery. the impella cp was implanted and then percutaneous coronary intervention (pci) to the proximal lad and distal left main arteries was completed. after the pci, the impella remained, and the patient was sent to the unit on support in stable condition. the patient was placed on continuous renal replacement therapy. additionally, the nurse reported the patient was cardioverted from atrial fibrillation to normal sinus rhythm. pump flow saturation was indicated as well as a position in ventricle alarm triggered. the nurse reduced the performance level to p-2 and called for an echocardiogram. the next day, the patient was having atrial flutter. the decision was made not to cardiovert as the patient's blood pressure was low. the following day, however, the patient was still with atrial flutter and was shocked two times. eight days later, the nurse inquired about the left ventricle waveform above placement signal. the nurse was advised that the pump was not in a flow saturation situation. flows were equalized, and flows were higher at 4.0 at performance level p-7. the nurse was advised to inform the care team of impending pump failure and have a backup plan available for either pump removal or exchange. the pump eventually failed, and the patient was planned for explant of the device which was completed. treatment plan thereafter is unknown. the patient's outcome after explant was indicated as unknown as well. however, there was no report or indication of non-survival. | [5 empty] | The investigation of saturated flow, pump stop, and vf/vt/af/at has been completed. pump was not returned for investigation. as per data logs, there was a 119 alarm which is an electrical open alarm. there is no clinical information provided about the reason of the pump stop. therefore, the root cause of the pump stop issue was an open electrical line but the reason is unknown. data logs were investigated and there was no motor current spike observed prior to the saturated flow. on p8 the flows observed very over 3.4l/min which is the upper limit for the p-level. there were no placement signal trends observed or any purge spikes observed. the data logs are inconclusive and without the pump the failure cannot be further investigated. therefore, the root cause of the saturated flow issue was not determined since product was not returned and data logs were inconclusive. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11692 | 1220648-2025-26117 | 21367033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21367033&pc=OZD | 01/23/2025 | 03/06/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590620 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a diagnosed st-segment elevation myocardial infarction patient was intubated in the emergency department prior to transporting to the catheterization lab for impella cp placement. the patient coded after gaining access and cardiopulmonary resuscitation was initiated. the return of spontaneous circulation was gained briefly. then the patient went into ventricular tachycardia and then coded for 45 minutes with sustained pulseless electrical activity arrest. the impella cp was opened and primed but never inserted. the patient expired and the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported delivery issue and ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of delivery issue could not be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt was not determined. b.2 revised as a selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26117. h.1 revised as selection was previously incorrect on the initial manufacturer device report number 1220648-2025-26117. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11693 | 1220648-2025-26118 | 21367229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21367229&pc=OZD | 01/23/2025 | 03/24/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603326 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the seal on an impella cp pump set was broken upon arrival. there was no visible external damage on the outer box at the time of delivery. the site requested the pump be replaced as a result of the broken seal. there was no patient involvement. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of packaging issues- sterility has been completed. the cause of the packaging issues - sterility was not determined since product was not returned and lack of clinical details.
11694 | 1220648-2025-26119 | 21368578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21368578&pc=PYX | 01/22/2025 | 03/14/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  | 02/28/2026 |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478658 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced multiple complications during and following impella rp support. it was documented that post op, the patient suffered bleeding complications for which they were given one unit each of platelets, cryo, and fresh frozen plasma. the patient was taken off of heparin, but began hemolyzing and the bleeding continued. the patient's chest was opened to resolve the bleed, during which the pump was swapped for centrimag with protek. an magnetic resonance imaging showed that the patient suffered a major stroke with midline shift coinciding with additional diagnosis of right arm ischemia, active pulmonary bleeding, and renal failure. in addition, a significant clot was removed from the bleeding site at the time of pump explant. the patient survived the intervention. | [5 empty] | The investigation of stroke/cerebrovascular accident (cva), thrombus, limb ischemia, hemolysis, renal failure, vascular damage - great vessel has and been completed. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. as the necessary information was not provided, the root cause of the thrombus and limb ischemia were not determined. the product was not returned for investigation. data logs were returned and at 14 hours there was a motor current spike along with suction alarms present. pump flow was observed to have a dip corresponding the motor current spike. the characteristics are similar to that of a sudden clot ingestion. clinical information also mentioned that there was a clot removed. therefore, the root cause of the hemolysis issue and renal issue was most likely biomaterial. as per the clinical information provided, the patient had pulmonary bleed suggesting damaged vessel but the location of the bleed and the reason of the bleeding is unknown. without the information the failure could not be further investigated. therefore, the root cause of the vascular damage issue was not determined since the product was not returned and insufficient clinical information was provided. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller: section: contraindications (united states): ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26119 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 revised instruction for use as they were reported incorrectly on the initial manufacturer device report number 1220648-2025-26119. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter: use of the repositioning sheath and the 23 fr peel-away introducer: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: contraindications (united states): ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
11695 | 1220648-2025-26120 | 21368602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21368602&pc=OZD | 01/23/2025 | 03/26/2025 |  | 02/12/2025 |  |  | 01/23/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577094 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient was being supported by impella cp. when the impella was checked the 16fr sheath base was loose, so impella was about to come out. it is unknown whether impella had come out of the aortic valve however the cp was exchanged for a new one. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of positioning issues has been completed. the pump was discarded. as per clinical information the impella was checked and a small amount of the 16 french sheath had come out from the base. the sheath was not properly fixed and came loose leading to the impella out of position. the root cause of the positioning issue was most likely use related since the sheath was not properly fixed and loose during support. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26120 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26120. h.6 code 4114 was entered incorrectly on the initial manufacturer device report number 1220648-2025-26120. a new code was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11696 | 1220648-2025-26121 | 21368959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21368959&pc=OZD | 01/17/2025 | 03/03/2025 |  | 02/12/2025 |  |  | 01/17/2025 | 02/12/2025 |  |  |  |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592109 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and post implant, same day, oozing was noted when the patient lifted the impella leg. systemic heparin was not started as the physician was to start systemic heparin drip when activated clotting time is in range of 160-180 as it was 260. impella site was monitored for continued oozing. five days later, the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va ecmo) overnight. the impella cp was at this time being used for left ventricle unloading. in the afternoon, the patient was transferred to an outside hospital for a left ventricular assist device (lvad) workup. patient escalation to impella 5.5 was noted. later, that evening, the patient was taken to the operating room for the impella 5.5 escalation. the physician utilized a double barrel technique to exchange the impella cp for the impella 5.5 with successful flow swap. the impella cp was explanted without complications, re-access port utilized for perclosure. however, post perclosure on right femoral artery for impella cp explant, doppler pulses were extremely monophasic and difficult to find. an order was made for an arterial ultrasound study for right lower extremity. the patient was transferred to the unit on support. the study showed flow. the following day, it was noted the patient was not being worked up for a durable lvad or transplant, as the team stated the patient would only qualify if renal recovery was present and decided to attempt weaning devices first to monitor recovery. the following day an attempt at ecmo weaning/decannulation while continuing to ride 5.5 support was made. patient on continuous renal replacement therapy. the next day the patient was taken to the operating room for decannulation of ecmo; ecmo was removed and impella 5.5 up to support level p-8 supporting the patient without incident. no advanced therapies were offered. the patient went into ventricular fibrillation arrest the following day, coded for approximately 20 minutes before the time of demise was called. although the patient was very ill no advanced therapies available, there is not enough evidence to exclude impella cp device as a possible associated factor. | [5 empty] | The investigation for the limb ischemia and access site bleed has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause for the access site bleeding. the root cause of the access site bleeding issue was most likely patient movement. some oozing was noted after the patient lifted leg. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26121. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26121 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4614 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26121. a new code was added. h.10 part of the instructions for use wer inadvertently omitted from the initial manufacturer device report number 1220648-2025-26121 and were added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11697 | 1220648-2025-26122 | 21369844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21369844&pc=OZD | 01/18/2025 | 03/07/2025 |  | 02/12/2025 |  |  | 01/18/2025 | 02/12/2025 |  |  | 10/31/2025 |  | 02/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423210 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and during support hemolysis was seen with poor positioning. echocardiogram was performed and the impella cp measured 3.5cm with pigtail directed toward to papillary muscle with no plans to reposition. it was noted post implant overnight, the patient's labs hemolyzed and the patient experienced hematuria and increased ectopy. continuous renal replacement therapy was started with the goal to goal to pull 150-200ml/hr. the following morning the patient had sustained runs of ventricular tachycardia with hypotension. lido bolus and drip and vaso drip started. an echocardiogram was completed, and the physician attempted to free the pigtail from the papillary muscle without success. inflow appeared to be unobstructed. the plan noted was to continue on impella mcs with hopes for mitral valve surgery. two days later the patient was transported to the operating room for a planned escalation to impella 5.5 for more support and removal impella cp in which both were successfully completed. the patient was reported to be in stable condition following the event and after escalating to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation has been completed. data logs showed the pump ran without issue during support. there were positioning alarms triggered in the end of the case and resolved. the product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the cause of the arrythmia was unable to be determined due to insufficient clinical details. b.2 revised as an incorrect selection was on the initial manufacturer device report number 1220648-2025-26122. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26122 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11698 | 1220648-2025-26080 | 21354136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354136&pc=OZD | 01/17/2025 | 03/07/2025 |  | 02/11/2025 |  |  | 01/17/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561167 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to the hospital and began decompensating going into pulseless electrical activity. the patient was started on medication (levophed, epinephrine, etc.) after receiving an unknown number of cardiopulmonary resuscitation rounds. the patient was placed on venoarterial extracorporeal membrane oxygenation (va ecmo) during this time in the emergency room. flowing on ecmo was difficult. the patient was found to have bilateral atrial thrombosis. at this time, interventional cardiology decided to urgently bring the patient to catheterization lab for advancement of venous catheter to the superior vena cava and impella cp to vent because the aortic valve was failing to open. an impella cp placement was successful. the patient was sent to the unit on ecmo and impella cp. impella cp performance level was on p-2 with suction. the physician was made aware and per orders the performance level remained at p-2 despite suction. audible alarm was disabled. the following day continuing on ecmo and impella cp, pink colored urine was noted with impella performance level now at p-1 still with suction alarms. an echocardiogram was completed to assess the position of the impella. it is unknown, however, if the impella pump was repositioned. the impella cp had been on p-1 due to continuous suction above the p-1 performance leve. lactate dehydrogenase rose above 2000. the physicians elected to remove the impella cp device and place the patient on an intra-aortic balloon pump to assist while the patient remained on ecmo. the patient was noted to have survived the explant and reportedly was in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and low pump flow has been completed. post impella removal ldh continued to rise over 3,500, which indicates the hemolysis issue most likely not relation to the impella. patient had other underlying conditions (autoimmune hepatitis. patient was found to be influenza a positive and was discharged home on tamiflu). impella had been on p1 due to continuous suction above that p- level. the cause of the hemolysis issue most likely was patient condition related since the ldh continued to rise after impella removal. the cause of the low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. data log not available for review.
11699 | 1220648-2025-26081 | 21354235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354235&pc=OZD | 01/22/2025 | 03/17/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  | 05/31/2025 |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355620 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. once the pump was placed it was turned on and while trimming and suturing the graft to the hub, the physician noticed blood coming through the hub into the sterile sleeve. the physician attempted to readjust the plastic hub/blue button and noticed that the connection was loose. bleeding continued to occur through with drops steadily forming in the sleeve, decision made to replace pump, which was successfully completed. the patient was sent to the unit on impella 5.5 support noted to be in stable condition following the event. | [5 empty] | The investigation of product damage (sterile sleeve damage) has been completed. the product was returned and evaluated. upon initial visual inspection, there was dried blood noted inside of the sterile sleeve with blood residue also seen in connection between butterfly and catheter anchor. the returned pump's butterfly was leak tested up to ~630 mmhg (~84 kpa) and no leaks were observed from the butterfly. the root cause of the access site bleeding was most likely use-related as the blue butterfly acts as a hemostatic component when using best practices in suturing the graft to butterfly and no mechanical abnormalities were found with the returned pump. the failure mode (fm) product damage (sterile sleeve damage) was updated to fm access site bleeding - major since bleeding was reported and no sterile sleeve damage was reported. b1 revised to reflect adverse event according to the investigation results on manufacturer device report number 1220648-2025-26081 h1 revised to reflect adverse event according to the investigation results on manufacturer device report number 1220648-2025-26081 h6 health effect - clinical code revised to reflect adverse event according to the investigation results on manufacturer device report number 1220648-2025-26081. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
11700 | 1220648-2025-26082 | 21354243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354243&pc=OZD | 01/22/2025 | 03/24/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588687 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, a purge pressure low alarm appeared. the purge cassette and automated impella controller were replaced, but the issue persisted. the purge fluid concentration was subsequently increased to 50% glucose and the purge pressure began to normalize. support was continued with the implanted pump with no noted harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the product was returned and evaluated. there were no abnormalities found upon visual inspection. upon running the pump with a lab purge cassette, the low purge pressure and high purge flow reproduced with pp at 200 mmhg and purge flow at 30 ml/hr. there were no leaks observed. the data logs also show low purge pressure from the beginning with purge flow at ~30 ml/hr. the glucose concentration was increased and after that, the purge pressure rose to 350 mmhg. the purge flow slightly decreased multiple times after the glucose increase but remained near 30 ml/hour the entire time. no other abnormalities were seen with the data logs. although the pump passed all checks in manufacturing and was built to specification, it is likely that the tolerance stack up of the specific components of this build lead to less resistance and higher purge flow. the root cause of the low purge pressure was most likely a design limitation of the pump since the purge pressure increased after increasing the purge fluid's glucose concentration. d9 date device returned to manufacturer added.
11701 | 1220648-2025-26083 | 21354296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354296&pc=OZD | 01/23/2025 | 03/06/2025 |  | 02/11/2025 |  |  | 01/23/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580908 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of an impella 5.5 device to a patient with cardiomyopathy in cardiogenic shock was unsuccessful as the device could not be fully deliver beyond the anastomosis. the decision made to instead implant an impella cp device. the patient was noted to be in stable condition following the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the delivery issue was most likely patient condition related, impella 5.5 unable to cross anastomosis. h5 revised to reflect yes on manufacturer device report. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11702 | 1220648-2025-26084 | 21354396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354396&pc=OZD | 01/21/2025 | 02/14/2025 |  | 02/11/2025 |  |  | 01/21/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582857 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Positioning Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 insertion, the impella 5.5 was initially caught up on the mitral valve. it was maneuvered under echocardiogram guidance and ultimately was 4.63 centimeters in the central left ventricle. the impella was working well. additional lines were placed at the end of the case which caused several episodes of ventricular tachycardia which required shock. a transvenous pacing wire was placed and ultimately rhythm was restored as well as return of spontaneous circulation. it was further reported that decision was made by the family to withdrew care and the patient expired. the relationship between the impella and cause of death is unknown. | The user facility reported that a patient in cardiogenic shock from an acute myocardial infarction implanted with an impella cp device for mechanical circulatory support experienced limb ischemia and access site bleeding requiring a surgical cutdown and massive blood transfusion. the patient would later expire, but prior the patient was escalated to an impella 5.5 with continued complication. the patient arrived at the hospital via life flight and was intubated. patient with st-segment elevation myocardial infarction to the catheterization lab for intervention coding multiple times in route and on arrival. the impella cp was placed via the right femoral artery. left ventricle access was challenging due to moderate to severe aortic stenosis. after multiple attempts, the physician was finally able to cross the aortic valve into the left ventricle. the impella cp was inserted without an issue. flows 3.7l, mean arterial pressure was 66 on epinephrine and levophed. multiple blood products were infused at this time due to patient hemoglobin of 4.9 at the start of the case. two stents were implanted to the left anterior descending artery. post percutaneous coronary intervention, the 14f x13cm peel-away sheath introducer was removed and repositioning sheath placed. digital subtraction angiography was checked to assess for right lower extremity flow; occlusion noted. fem-fem antegrade sheaths were placed and perfusion restored. a large hematoma was noted and profound bleeding around the repositioning sheath occurred. decision made to explant the impella cp in catheterization lab with vascular surgery at lab table for possible dry closure for right common femoral artery (rcfa). the decision was made to take the patient to the operating room for a surgical cut down to rcfa. the patient's hemoglobin was 4.9 on arrival and other bleeding sources were questioned. no further scans were completed prior to arriving in the cath lab. the patient received multiple units of pack red blood cells, fresh frozen plasma and cryoprecipitate were given. it was noted no issues or resistance were noted during insertion nor during the case; patient's hemodynamics improved with the impella cp. however, the patient was in critical condition. the following day, the patient was placed on an impella 5.5. during the procedure the left ventricle access was obtained and the 0.018 guidewire placed. the impella 5.5 was placed in left ventricle and initially was caught up on the mitral valve. it was maneuvered under echo guidance and ultimately was 4.63 cm in the central left ventricle. the device was noted to be working well. additional lines were placed at the end of the case which caused several episodes of ventricular tachycardia which required shock. once the swan was placed, the patient went systolic. a transvenous pacing wire was placed and ultimately rhythm was restored as well as return of spontaneous circulation. impella 5.5 flows increased back to 4.5 lpm. the patient had significant metabolic acidosis leaving the operating room which continued to be corrected. three days later the patient would expire. medical safety review of the event determined there is not enough evidence to exclude impella cp as an associated factor in the patient's outcome. whether the patient recovered or not with escalation is inconsequential because the patient was very sick and experienced "profound" bleeding as a result of the impella cp requiring multiple blood products and the demised outcome cannot be excluded from impella cp even long term. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | This is one of two device manufacturing reports associated with this event. refer to initial medical device report for impella cp serial number (b)(6) with date of event 20-jan-2025 and patient identifier ending in (b)(6). respective sections were updated accordingly based on and to include complete details of the reported event sections updated include: b1 with b2, b5, h6 (patient and device codes) h1 (corrected type of reportable event to serious injury) and h10 (to reflected the related manufacturing device report).
11703 | 1220648-2025-26085 | 21354431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354431&pc=OZD | 01/23/2025 | 03/11/2025 |  | 02/11/2025 |  |  | 01/23/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025595660 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient found to have new heart failure dysfunction, atrial fibrillation with multivessel disease underwent coronary artery bypass surgery and had an impella electively placed postoperatively for mechanical circulatory support. approximately 23 hours after start of support, an acute clot entrainment developed. the motor current subsequently spiked and then flattened. the flow dropped to less than 1l on support level p-6 and plans were subsequently made to explant the device. the patient was noted to be hemodynamically stable enough for the impella 5.5 removal and following the explant was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. the data logs confirmed low pump flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. logs also showed pump was in aorta shortly but resolved. after two hours since start of case upward rise/trend in motor current (mc) and small intermittent mc spikes were observed with drop in flows. the mc trend and drop in flows indicate likely ingestion trend. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. ¿note: remove failure mode thrombus as it was a root cause of low flow.
11704 | 1220648-2025-26086 | 21354609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354609&pc=OZD | 01/22/2025 | 03/06/2025 |  | 02/11/2025 |  |  | 01/23/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562872 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy implanted with an impella cp device for mechanical circulatory support (mcs) developed a hematoma around the access site. a dressing change was applied, and the repositioning sheath was adjusted. however, the patient continued to bleed. manual pressure was then held and a femostop was placed to achieve hemostasis. the patient received six units of packed red blood cells and two units of fresh frozen plasma due to the blood loss. the patient was later brought back to the operating room for pump removal, and a new impella cp device was implanted via the opposite leg. the patient was noted to have survived the event and sent back to the unit on impella mcs in stable condition. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11705 | 1220648-2025-26087 | 21354713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354713&pc=OZD | 01/15/2025 | 03/03/2025 |  | 02/11/2025 |  |  | 01/23/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584966 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that developed a hematoma at right axillary insertion site with a "definite" relationship to the impella device used. the patient undergoing high-risk percutaneous coronary intervention was implanted with an the impella 5.5 device for mechanical circulatory support (mcs), and the following day a hematoma was noted, treated by albumin and device explant. the patient was on impella mcs for 72 hours. per the registry, the patient recovered with no sequelae. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. g.2 added study selection as it was inadvertently omitted from the initial manufacturer device report number: 1220648-2025-26087. e.1 added middle name as it was inadvertently omitted from the initial manufacturer device report number: 1220648-2025-26087. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number: 1220648-2025-26087 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11706 | 1220648-2025-26088 | 21354864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354864&pc=OZD | 01/23/2025 | 03/24/2025 |  | 02/11/2025 |  |  | 01/23/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575406 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported at the beginning of the impella cp implant procedure to support a patient undergoing high-risk percutaneous coronary intervention, the guidewire tip was noted to be bent after removal from its sterile packaging. a guidewire from another impella set was used to successfully complete impella cp pump implant procedure. percutaneous coronary intervention was completed and the impella cp device was weaned and explanted. there was no patient harm as a result of this event. | [5 empty] | The impella guidewire has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of packaging issues- sterility has been completed. the guidewire was returned and evaluated. the packaging was also returned with the seal broken and the guidewire inside. the guidewire had been slightly pulled out of the tubing and a kink was visible at the tip. no other abnormalities were found. the root cause of the packaging issue was not determined since the guidewire packaging had been opened in the field and insufficient clinical details were provided. the failure mode (fm) "packaging issues - sterility" was updated to "packaging issues" since it was unrelated to the sterility of the product. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26088 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11707 | 1220648-2025-26089 | 21354865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354865&pc=OZD | 01/23/2025 | 03/18/2025 |  | 02/11/2025 |  |  | 01/23/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601285 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the next day after start of support, he patient impella leg was cold with no perfusion. two units of blood were given to the patient, and the decision was made to explant the pump the following day. the patient's outcome at explant was noted as survived; however, the condition was critical with a poor prognosis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26089 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10: instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26089.
11708 | 1220648-2025-26090 | 21354887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21354887&pc=OZD | 01/23/2025 | 03/06/2025 |  | 02/11/2025 |  |  | 01/23/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601929 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported during an attempt to implant an implella cp device for support during off-pump coronary artery bypass grafting (cabg), the pump was unable to advance due to heavy calcification in the axillary and innominate as well as a porcelain aorta. the chest was subsequently opened for a direct approach; however, again the patient's anatomy prevented successful placement. the implant was subsequently aborted. the patient was placed on cardiopulmonary bypass for the cabg x3. it was noted the impella insertion was attempted at the end of the cabg procedure and was also unsuccessful. the outcome of the patient after the case wrap up is unknown. however, there was no report or indication of any adverse effects to the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. the impella was unable to pass due to patient anatomy, heavy calcification in axillary, innominate, porcelain aorta. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d.6a/d.6b revised as dates were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26090. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26090 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.2 revised as user facility was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26090. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26090 and a new code was added. h.10 instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26090 and were added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11709 | 1220648-2025-26091 | 21355083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21355083&pc=OZD | 01/22/2025 | 03/12/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562911 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient underwent coronary artery bypass graft with impella cp support. impella cp was in surgical mode during the procedure. decoupling of waveforms was noted coming off of the surgical mode. repositioning was attempted under transesophageal echocardiography guidance with no success. the impella cp was kinked and was unable to be repositioned. therefore, the impella cp was exchanged for a new one. there was no patient harm noted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the complaint device was visually inspected. three cannula kinks near outlet, inlet and in the bending area observed. biomaterial around impella and purge gap was observed. the cause of the positioning issue was use related since pump pulled out while attempted to reposition occurred cannula kinks and was unable to reposition.
11710 | 1220648-2025-26092 | 21355274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21355274&pc=OZD | 01/21/2025 | 03/20/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596327 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was unable to be inserted as there was a technical problem during the preparation. the impella cp procedure was aborted and the impella cp was discarded. there was no patient harm reported. | [5 empty] | The investigation of priming issue has been completed. the impella device was not returned for evaluation. based on limited clinical details and no conclusive issues/alarms in data logs, the root cause of the priming issue could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
11711 | 1220648-2025-26093 | 21357263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21357263&pc=OZD | 01/22/2025 | 03/11/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583468 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis during impella cp support. four units of red blood cells were given to the patient to treat the condition. support continued with the implanted pump following the intervention. it was later documented that the patient expired. it is unknown if the condition contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the data logs showed pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved quickly. the product was not returned for review. the root cause of hemolysis was most likely patient condition related.
11712 | 1220648-2025-26094 | 21357548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21357548&pc=OZD | 01/22/2025 | 03/04/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591474 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Device Damaged Prior to Use | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 0.018 wire damaged in a impella cp package. it was partially fractured and unable to be used. an alternate wire was used and there was no patient harm. | [5 empty] | The failure mode product damage (guide wire damage - great vessel) was changed to packaging issues since this was claimed to be damaged in the packaging. the product was returned and the guidewire was found to be kinked at the distal tip, it was still in the casing. the root cause of the packaging issue cannot be determined. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11713 | 1220648-2025-26095 | 21357886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21357886&pc=OZD | 01/22/2025 | 02/11/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576348 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a patient admitted in acute myocardial infarction and cardiogenic shock. the pump was poorly positioned in the left ventricle. the pump was persistently in the mitral and not in the cavity of the left ventricle appropriately. the site of access also developed a hematoma. the team gave blood products and fresh frozen plasma. the pump was then prematurely explanted and escalated to the 5.5 pump after just 17 hours of support. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11714 | 1220648-2025-26096 | 21358608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21358608&pc=OZD | 01/22/2025 | 02/24/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585157 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp pump to support a patient admitted in with a acute myocardial infarction and cardiogenic shock. the pump was placed and allowed for percutaneous coronary intervention. the pump was supporting when after 3 days the pump was explanted to allow for care escalation to the 5.5 pump. after the explant there was an observation made of limb ischemia. due to the lack of flow down the limb, which was treated by angioplasty/thrombectomy. the patient is now supported by the 5.5 pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
11715 | 1220648-2025-26098 | 21359563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21359563&pc=OZD | 01/17/2025 | 03/24/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025598960 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support a patient admitted in for surgery. the pump was placed and enrolled in the impact protocol. the impact form documented thrombocytopenia. the allegation was the ae was "possible" related to the impella procedure and the patient outcome was noted as recovered/resolved. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation of thrombocytopenia has been completed. the impella device was not returned for evaluation. thrombocytopenia reported on last day of support. the cause of the thrombocytopenia was not determined due to insufficient clinical details. d.6b explant date was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26098. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26098 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11716 | 1220648-2025-26099 | 21359664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21359664&pc=OZD | 01/22/2025 | 03/06/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582133 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that delivery of an impella cp pump for mechanical circulatory support was unsuccessful. the implant was aborted after resistance was met due to the patient's anatomy. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature and insertion was aborted. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26099 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11717 | 1220648-2025-26097 | 21360264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21360264&pc=OZD | 01/22/2025 | 03/06/2025 |  | 02/11/2025 |  |  | 01/22/2025 | 02/11/2025 |  |  |  |  | 02/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578487 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the impella site started to ooze. two units of packed red blood cells were transfused. it was decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11718 | 1220648-2025-26054 | 21343045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343045&pc=OZD | 01/20/2025 | 03/06/2025 |  | 02/10/2025 |  |  | 01/20/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585436 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 43 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding. the impella was successfully placed via the left femoral artery with flows and mean arterial pressure within normal limits. the right femoral artery with a six french sheath was used for the percutaneous coronary intervention site in which balloon angioplasty to the mid circumflex followed by three drug-eluting stents and balloon angioplasty to the obtuse marginal one followed by three drug-eluting stents were completed. thesix french sheath was removed and the site was closed using mynx. the impella access site, left femoral artery, and the 14 french peel-away introducer sheath were exchanged for the repositioning sheath without incident. slacked was removed and the site was sutured into place. there was no hematoma or bleeding complications. the patient was transferred to the unit on support. the following day, however, there was bleeding from impella access site as a result of patient movement. four units of packed red blood cells products were given to the patient. a femstop and manual pressure were applied. it was noted that the bleeding resolved the same day the event occurred, and the patient was reported to be in stable condition. following, an echocardiogram was pending for wean and explant of the impella cp, which was successfully completed the next day. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was determined to be patient movement since the bleeding was caused due to patient movement.
11719 | 1220648-2025-26055 | 21343155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343155&pc=OZD | 01/21/2025 | 02/20/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574163 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon performing coronary interventions into the left anterior descending artery (lad), a small dissection of the lad was noted. the physician, therefore, requested the patient remain on impella support for a few hours. subsequently (number of minutes/hours following are unknown), bleeding and oozing were noted from the access site. despite cinching perclose devices, the oozing continued, and a hematoma developed. the physician opted to wean and explant impella in the catheterization lab. the patient remained stable and tolerated removal. hemostasis was achieved with closure devices, and the hematoma resolved with manual pressure. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
11720 | 1220648-2025-26056 | 21343218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343218&pc=PYX | 01/21/2025 | 02/28/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  | 03/31/2025 |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024353698 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of an impella rp device to a patient for mechanical circulatory support was unsuccessful. it was noted that the patient had multiple abdominal surgeries and had scar tissue and adhesions in the abdomen. impella rp was attempted to be placed but the physician struggled tracking the impella rp through difficult anatomy. the impella rp implant was aborted. this said impella rp device was the last one at hand. a request planned for an impella rp flex to be placed once received. there was no report or indication of harm to the patient. the patient was reported to be in stable condition on three inotropes/pressors with a mean pressure of 70. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella was unable to pass through the vasculature. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11721 | 1220648-2025-26057 | 21343262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343262&pc=OZD | 01/21/2025 | 03/10/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561484 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp and escalated to an impella 5.5 device for mechanical circulatory support as a bridge to left ventricular assist device (lvad). the patient was admitted into the hospital post percutaneous coronary intervention and with mitral regurgitation and low ejection fraction of 10-15%. the impella 5.5 pump supported the patient while awaiting lvad. on day 18 the pump was being explanted for the bridge to lvad. post explant of impella 5.5 device and during durable lvad placement, it was noted the patient had wide open aortic insufficiency. patient cross-clamped and noted to have impella sized hole in aortic leaflet. the physician commented the following: during an ascending aortic replacement (aar), the physician stated that before valve excised that a hole appeared the exact size of impella. the physician further stated that it would have had to been accidently pierced by wire and device loaded over it due to appearance of hole. no further details were noted regarding the event. however, the patient did survive the event and was reported to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (unites states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter: ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subangular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation of heart valve damage has been completed. as per the clinical information provided, the doctor suspects that the hole noted on the aortic valve was likely the size of the impella pump which could possibly have accidently pierced the valve. also, it was mentioned that there was no issues observed with insertion or other practices. the images showed the heart valve was pierced. therefore, the root cause of the heart valve damage issue was not determined since the product was not returned and due to insufficient clinical evidence. b.2 revised as an incorrect selection was made on the initial manufacturer device report number 1220648-2025-26057.
11722 | 1220648-2025-26058 | 21343264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343264&pc=OZD | 01/21/2025 | 02/27/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578951 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and after two days on support, it was reported that the placement signal disappeared with placement signal not reliable alarms. the automated impella controller was reviewed, and pulsatile motor current was noted. an echocardiogram confirmed the position. the patient already was on the operating room schedule for an impella 5.5 escalation and had the impella cp device explanted and replaced with the impella 5.5. this day, the patient was also cannulated for venoarterial extracorporeal membrane oxygenation. the impella 5.5 device would eventually be successfully weaned after 72 hours with ejection fraction improvement noted from 10% upon admission to 63% prior to explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue has been completed. no product was returned for analysis. the data logs confirm placement signal not reliable alarms and show placement signal out of range at 3000mmhg. concurrently the signal not reliable (snr) dropped to zero, gain and light decreased, lamp increased and contrast increased and became more pulsatile. the observed pattern has been characterized as a fiber break. the root cause of the optical signal issue was most likely due to a fiber break, the exact location and cause of the break could not be determined without product return and clinical details
11723 | 1220648-2025-26059 | 21343344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343344&pc=OZD | 01/22/2025 | 03/04/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1727786 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) displayed a purge disc not detected alarm during impella mechanical circulatory support to a patient. the aic was replaced. there was no report or indication of interruption in therapy and/or harm to the patient. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the automated impella controller (aic) purge disc/flag issues has been completed. data logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was returned for evaluation. a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number.
11724 | 1220648-2025-26060 | 21343348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343348&pc=OZD | 01/07/2025 | 03/24/2025 |  | 02/10/2025 |  |  | 01/20/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583517 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient was on venoarterial extracorporeal membrane oxygenation (va ecmo) therapy less than one hour prior to initial impella mcs. after insertion of the pump, purge fluid leak alarm sounded with no visible leak. in response to the issue, the purge cassette, purge housing was replaced, but the issue remained unresolved. consequently, approximately four hours after start of initial mcs, the impella cp pump was explanted and replaced with a new impella cp device. the impella timing was bailout with percutaneous coronary angioplasty to the left main artery. one day after the impella cp support, the patient was escalated to an impella 5.5 device due to the need of long-term support. the patient's hemodynamics did not improve, and the patient expired of multi-organ failure six days later. there is not enough evidence to exclude impella cp as an associated factor in the outcome given initial support with the purge system failure (purge alarms trigger with no visible leak resulting in the need for additional impella intervention procedure). however, of note, the patient received uninterrupted mcs for seven days. hemodynamics did not improve (medical history and any comorbidities are unknown). per the report the patient expired from multi-organ failure due to an unknown cause. | [5 empty] | The impella pump with purge cassette has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of purge leak ¿ pump has been completed. the product was returned and evaluated. there were no abnormalities found upon visual inspection. upon running the pump with a test purge cassette using dextrose five percent (5%) in water, the low purge pressure and high purge flow reproduced with purge pressure at approximately 200 mmhg and purge flow at approximately 30 milliliter per hour. there were no leaks observed. the data logs also showed low purge pressure from the beginning with purge flow at approximately 30 milliliters per hour. the glucose concentration was 5% in field. no other abnormalities were seen with the data logs. although the pump passed all checks in manufacturing and was built to specification, it is likely that the tolerance stack up of the specific components of this build lead to less resistance and higher purge flow. the root cause of the low purge pressure was most likely a design limitation of the pump. d.10 therapy dates should have been left blank as information was erroneously entered on initial manufacturer device report number 1220648-2025-26060. e.1 revised address line 2 as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26060. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26060 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11725 | 1220648-2025-26061 | 21343365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343365&pc=OZD | 01/22/2025 | 03/10/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563952 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and escalated to an impella 5.5 device. a delivery difficulty was encountered, however. right axillary cutdown and 8 millimeter x 30 centimeter hemashield platinum graft beveled appropriately and anastomosed to vessel. however, the impella 5.5 device was unable to advance over 0.018" guidewire past the subclavian bend to innominate. multiple attempts were made. the physician tried arm manipulation to abduction, mineral oil for lubricant and several different wires and techniques to no avail. the physician noted there was "kinking of catheter near motor" when trying to advance and was uncomfortable using said impella 5.5 device for reinsertion. therefore, another impella 5.5 device was obtained and successfully implanted to support the patient. there was no report or indication of adverse effects to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. based on clinical description, the root cause of delivery issue was most likely patient condition due to diseased vessel condition.
11726 | 1220648-2025-26062 | 21343368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343368&pc=OZD | 01/22/2025 | 03/07/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589801 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aneurysm | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the decision was made to escalate to an impella 5.5 device. however, impella 5.5 device delivery via right axillary insertion was aborted after discovery of an aneurysm in the subclavian artery. upon discovery of a right brachiocephalic artery and junction dissection, the pump was removed. it was noted that despite being inserted, the pump was never turned on. the patient was transported back to the unit on impella cp support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance | The investigation has been completed. as per the clinical information provided, amidst the delivery of the pump, it was found that the patient had separation in the layers of the arterial wall and a right brachiocephalic artery dissection. the patient had coronary angioplasty done along with stent placement. the delivery was successful but the reason for the dissection is unknown. the root cause of the vascular damage - great vessel issue was not determined since the product was not returned and clinical information provided was not sufficient. as per the clinical information provided, the patient was discovered to have aneurysm in the subclavian artery. the patient had coronary angioplasty. therefore, as per the clinical information provided, the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related to diseased vessels. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26062 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4614 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26062.
11727 | 1220648-2025-26063 | 21343372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343372&pc=OZD | 01/21/2025 | 03/12/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | GM |  | 1152776 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support (mcs). there were no issues during the implant of the impella cp device. however, right after starting the pump there was an automated impella controller "controller failure" alarm (with the noted advice to change the controller) and the placement signal invisible. the pump was noted to be running without any issues. the aic was exchanged; however, said alarm remained. impella mcs continued and there were no report or indication of adverse effects to the patient. eventually, the patient was successfully weaned, and the device explanted with the patient noted to be stable. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller failure has been completed. the console logs (b)(6) from the reported day of the event are consistent with the complaint. the logs show instances of the following error "controller error (defective optical sensor lamp) - alarm issued event" was observed. in response to the controller error, the console was first exchanged for aic (b)(6). after encountering the same error, aic (b)(6) was used. subsequently, aic (b)(6) was used and ran for next 23 hours. console (b)(6) was not returned to field service for further investigation. the root cause for the controller error alarms, was not determined as the aic was not returned. b.1 revised brand name as it was previously reported incorrectly on the initial manufacturer device report 1220648-2025-26063. d.6a/d.6b should have been left blank on the initial manufacturer device report number 1220648-2025-26063. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26063, as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2199 was previously reported incorrectly on the initial manufacturer device report 1220648-2025-26063.
11728 | 1220648-2025-26064 | 21343374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343374&pc=OZD | 01/21/2025 | 02/20/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525452 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support then escalated to an impella 5.5 device. approximately five days after start of the impella 5.5 support, bleeding was noted at the access site. the patient was taken to the operating room to reopen the pocket, and no bleeding was noted in the pocket. following, a computed tomography scan showed a hematoma around the access site. one unit of packed red blood cells were administered. support continued. the patient was eventually weaned and device explanted two days later with a survival outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26064 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26064 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
11729 | 1220648-2025-26065 | 21343473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343473&pc=PYX | 12/20/2024 | 09/24/2025 |  | 02/10/2025 |  |  | 12/20/2024 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560920 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old male patient on an impella rp flex needing mechanical circulatory support for high risk surgery. it was reported that the implanted impella rp flex triggered a placement signal not reliable alarm during patient support. the nurse was walked through how to disable the alarm and support with the implanted pump continued. roughly four days later, flows dropped overnight following swan removal and the motor current fluctuated. the pump was subsequently explanted as a result of a clot ingestion. the patient survived the procedure. there was no additional information provided. | The complainant also reported a placement signal not reliable alarm on their impella rp flex. placement signal waveforms were intermittent with constant alarms, and the motor current was unchanged. |  |  |  |  | The investigation for the low pump flow and the thrombus has been completed. the pump was not returned for investigation. the data log shows the fiber torque trend on the 5s optical signal parameter on the 13th day of use, followed by the placement signal not reliable (psnr) alarm, which is related to an optical signal issue. no symptoms were reported related to the dp sensor issue or motor current trend during the optical signal issue. several motor current spikes were reported. after four days of optical signal issues, low pump flow was reported, along with a hike in the placement signal during the motor current spikes. pump was explanted. clinical details also suggest the biomaterial buildup during the case. the exact location of the damage was unknown without the product being returned. the cause of the low pump flow issue was most likely biomaterial, based on data log review. the instructions for use for the impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller states the following: section: potential adverse events (united states), ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was most likely damaged optical fiber line, based on data log review. the exact location of the damage was unknown without the product being returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5, h6, and h10.
11730 | 1220648-2025-26066 | 21343484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343484&pc=OZD | 01/22/2025 | 03/04/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592975 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and experienced a thrombotic embolism resulting in limb ischemia. the physician obtained right femoral vein and right femoral artery access. vessel noted to be good size, limited calcifications. the physician then attempted to perclose site but was unsuccessful; therefore, the 6fr sheath left in the artery. the physician assessed the left femoral artery about the same, did somewhat of a lower stick, still in good area, and attempted perclose x 1 which failed. the physician reangled and was able to successfully place the perclose. the impella cp was then prepped and inserted in standard fashion without complication. the physician attempted to wire right coronary artery; the physician utilized multiple wires and was successful. an attempt to ballon the initial rca lesion was made with an attempt to wire distally enough for further intervention. however, after multiple attempts the wire was unable to pass. activating clotting time was approximately 200. the impella then was weaned and explanted. attempt to secure perclose was made; however, some failure was noted. reportedly, a clot distally traveled to lower extremities and the patient required thrombectomy with penumbra which was completed with a successful outcome. distal perfusion was intact post procedure. the patient was transferred to the unit, in stable condition, for recovery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26066. h.10 some instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26066 and were added.
11731 | 1220648-2025-26067 | 21343636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343636&pc=OZD | 01/22/2025 | 03/11/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601551 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhagic Stroke | [12 empty] | Disability | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device was mechanical circulatory support and the next day an intracerebral hemorrhage was noted on computed tomography scan of the head. flows were appropriate per performance level. the discussion with the family was planned and would include possible removal of the impella cp. however, the following day the impella cp was removed after an impella in the ventricle alarm triggered. the patient was reported to be hemodynamically stable post explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6: code 4617 was inadvertently omitted on manufacturer device report 1220648-2025-26067. | The investigation of stroke and positioning issues has been completed. the impella device was not returned for evaluation. therefore, the root cause of the positioning issue and stroke was not determined since the product was not returned and insufficient clinical information was provided. b.2 revised selections as some selections were reported incorrectly on the initial manufacturer device report number 1220648-2025-26067.
11732 | 1220648-2025-26068 | 21343721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343721&pc=OZD | 01/21/2025 | 03/14/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1152774 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and just after starting the pump there was a controller failure alarm and the placement signal was not visible. the pump was running without any issues. they changed the controller, but the alarm remained. the staff switched back to the original automate impella controller. this complaint was created to document the replacement controller error. the pump was not explanted and support was continued. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The cause of the issue was not determined since product was not returned. the cause of the issue was not determined since product was not returned. h5 revised from yes to no on manufacturer device report 1220648-2025-26068. h8 revised from initial use of device to reuse on manufacturer device report 1220648-2025-26068.
11733 | 1220648-2025-26069 | 21343834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343834&pc=OZD | 01/22/2025 | 02/10/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1756886 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported an automated impella controller sounded an alarm for battery failure. the device was returned and received for evaluation. it is unknown if the failure occurred during patient support. however, there is no report or indication of interruption in therapy or patient harm as a result of this event. | [5 empty] | The investigation was completed. console logs were partially consistent with what was reported in the complaint. the complaint date was listed as 2025.01.22 but the reported battery failure actually occurred on (b)(6) 2024. there was a single instance of battery failure (transport connection blown/missing) [alarm observed in the console logs from (b)(6) 2024. the console was tested. the reported battery failure issue was able to be reproduced. results of running the battiest utility on telnet revealed that cell 3 in battery 1 had a voltage that was significantly less than the rest of the cells in the battery. after troubleshooting by replacing battery 1 with a known good battery, the alarm resolved and the console charged as normal. the reported battery failure was determined to be caused by the cell imbalance observed in battery 1.
11734 | 1220648-2025-26070 | 21343853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21343853&pc=OZD | 01/20/2025 | 02/21/2025 |  | 02/10/2025 |  |  | 01/20/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578963 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient developed a hematoma at the access site. pressure was held and a hematoma evacuation was performed. the impella cp was removed and the impella 5.5 was placed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11735 | 1220648-2025-26071 | 21344162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21344162&pc=OZD | 01/21/2025 | 03/11/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594954 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. following introducer placement, the staff was unable to advance the impella pump through the 14fr peel away sheath due to the deep vessel and angle of the sheath. as a result of the complication, the decision was made to abort the implant. there was no noted patient harm. | [5 empty] | The investigation has been completed. the pump was returned and evaluated. there was minor kinking found on the cannula, but no other abnormalities were found with the pump. the introducer was not returned. the failure mode "delivery issues" was removed since the pump could not pass through the introducer. a second failure mode "difficulty inserting/removing pump through introducer" was added against the pump since the customer reported the issue against the pump. the minor kinking on the returned pump cannula most likely occurred due to the difficulty inserting through the introducer. the voice of the customer alleges the difficulty passing through the introducer against the pump. however, there was no problems found with the pump since the introducer kink prevented the passing of the pump. the root cause of the difficulty inserting the pump through the introducer was most likely an introducer sheath kink since they noted a sheath kink in the field along with the patient's deep vessels and a steep sheath angle. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h.6 revised medical device problem codes due to investigation result. the codes on the initial manufacturer device report number 1220648-2025-26071 are incorrect. h.10 added instruction for use due to investigation results. the instructions for use on the initial manufacturer device report number 1220648-2025-26071 are incorrect. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11736 | 1220648-2025-26072 | 21345535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21345535&pc=OZD | 01/21/2025 | 03/18/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025602085 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 being placed to support a patient admitted in with cardiomyopathy and cardiogenic shock. the impella 5.5 was unable to be delivered due to anatomy. the team removed and instead placed an impella cp. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy and was aborted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11737 | 1220648-2025-26073 | 21345654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21345654&pc=OZD | 01/12/2025 | 03/13/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528375 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Embolism/Embolus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, subarachnoid hemorrhage was observed in an area without aneurysm. impella was removed because the bleeding was thought to be caused by impella since it was in an area without aneurysm. after removing impella, the 16fr sheath was left in place to keep the patient at rest for the subarachnoid hemorrhage, but a lower limb embolism occurred. apparently the embolism was caused by the 16fr sheath used to insert impella. to treat the lower limb embolism the sheath was removed and blood flow was maintained through collateral circulation. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of stroke/cva and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was inadvertently reported incorrectly on the initial manufacturer device report number 1220648-2025-26073. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
11738 | 1220648-2025-26074 | 21349305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21349305&pc=OZD | 01/15/2025 | 03/06/2025 |  | 02/10/2025 |  |  | 01/15/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592009 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp explant, the patient experienced excessive bleeding from their access site. it is unknown if intervention was required to manage the bleed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11739 | 1220648-2025-26075 | 21349609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21349609&pc=OZD | 01/22/2025 | 03/17/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1106950 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following the case start process, an automated impella controller (aic) repeatedly displayed a purge failure alarm. the aic was swapped and pump implant was successful. there was no patient harm. | [5 empty] | The investigation for the console (aic) purge issue has been completed. data logs were reviewed finding the console then displayed 'purge disc not detected ¿ alarm,' event # 402, likely because the disc was not inserted, which caused the purge pressure transducer click signal to go undetected for more than 10 seconds, resulting in event # 402." after observing the alarm, the user likely inserted the disc and initiated the 'case start' wizard. at step 2, the disc and the cassette were removed, most likely to reinitiate the 'case start' wizard without the cassette and disc being inserted in the first place. the ¿case start¿ wizard was initiated again, the same cassette (s/n (b)(6) was inserted, and the disc insertion was detected successfully. after the pump (s/n: (b)(6) was connected, the console displayed ¿purge flow increased ¿ alarm,¿ event # 415, because the purge flow had increased by more than 2.5 ml/hr. event # 415 is an advisory alarm (just a notification; no action is required per the design). all the above purge-related alarms were most likely misreported as ¿repeated - purge failure.¿ the pump and the cassette were then moved to the second console (ic10795) on january 22, 2025, and support was continued for two days. on december 16, 2024 (before the complaint date), the user operated the console with the demo pump. during the initial run, the log recorded a couple of heartbeat checks and a "watchdog: partial reset," which resulted in 4-in-1 reset. after the 4-in-1 reset, there was a core dump of ossiopsens.core, indicating an optical bench communication issue. again, there were heartbeat checks and a "watchdog: partial reset," resulting in 4-in-1 reset. after the reset, due to the persistent issue, the log recorded a "watchdog: full reset," and the console shut down unexpectedly and rebooted. following the reboot, there was a core dump log (ossiopsens.core), and the console then displayed ¿unexpected controller shutdown ¿ alarm,¿ event # 603, which is unrelated to the reported failure mode. this console was returned for evaluation. the console detects the cassette and disc insertion as intended. the purge pressure accuracy was within specification. additionally, there was no physical damage found on the retainer or the flag. there was no console hardware issue related to the reported failure mode. the unexpected controller shutdown could not be reproduced either (unreported and unrelated to the reported failure mode). the front panel optical connector was found to be contaminated, and the debris could not be removed upon cleaning (this is unrelated to the reported failure mode). the root cause of the repeated purge failure was most likely user issue. the root cause of the intermittent unexpected controller shutdown could not be determined due to the inability to reproduce the issue. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11740 | 1220648-2025-26076 | 21350974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21350974&pc=OZD | 01/16/2025 | 03/07/2025 |  | 02/10/2025 |  |  | 01/16/2025 | 02/10/2025 |  |  | 01/31/2025 |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024296632 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy in cardiogenic shock implanted with an impella 5.5 device for mechanical circulatory support (mcs) experienced episodes of torsades de pointes with defibrillation and device repositioning, and purge pressure issue. the patient would subsequently expire. the patient was admitted and newly diagnosed with heart failure with reduced ejection fraction of 10% and was noted likely in cardiogenic shock. inpatient diuresis was started, and the patient was transferred to out for advanced therapies evaluation with plans were in progress for an impella 5.5 placement and workup for durable left ventricular assist device versus heart transplant. the patient was sedated and intubated on arrival, and the impella 5.5 device was placed via right axillary artery without complication. performance level started at p-2 and gradually increased to p-8, achieving flows of 4.8 l/min. hemodynamics were stable throughout procedure. the patient was extubated at end of case and transferred to unit on mcs with multiple drips infusing (epinephrine 4 mcg/min, cardene 5 mg/hr. milrinone 0.25 mcg/kg/min). approximately 14 days later, the patient was awake, alert and sitting in the chair when episode of sustained torsades was experienced. the patient was defibrillated a total of seven times this day and the impella 5.5 pump was repositioned. the patient continued on support, was placed on venoaterial extracorporeal membrane oxygenation (va ecmo) and was being worked up for heart transplant last noted as a status 2. the following day, high purge pressure was encountered; purge flow went missing and after a minute returned at 3.4ml/hr. tissue plasminogen activator was added, performance level went back up to p-2 and then ramped back up needed support level. following, 8 days later, the patient would expire; care was withdrawn. additional details surrounding the patient's demise was not provided. impella related events are unable to be excluded as a possible contributing factor in the patient's outcome. it is of note, however, the patient had severe heart failure (ejection fraction 10%), subsequently required additional support from ecmo, pending a transplant and considering could likely be the primary for the demise outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of arrythmia and high purge pressure has been completed. the impella device was not returned for evaluation. product was not returned for investigation. as per the clinical report, the patient had an episode of sustained torsades, and the impella was repositioned. the next day, the purge flow dropped to 0, and tissue plasminogen activator was given per protocol. additionally, the purge cassette was replaced after reported low purge volume alarm and an air in purge alarm. the data log showed a gradual rise in purge pressure, which was resolved with a gradual decrease trend over six hours. also, the air in the purge system alarm was observed after the completed bag change procedure. the doctor initiated the de-air procedure, but it did not resolve the air in the purge cassette alarm, and the purge cassette had to be replaced. the cause of the high purge pressure issue was most likely biomaterial, based on data log review and the provided clinical details. as the necessary information was not provided, the root cause of the arrythmia was not determined. b.2 revised as an incorrect selection was reported on the initial manufacturer device report number 1220648-2025-26076. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26076 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11741 | 1220648-2025-26077 | 21351137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21351137&pc=PYX | 01/21/2025 | 03/03/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590398 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support. while on support, patient scheduled for replacement of rp flex after possible ingestion of clot for the right ventricular failure flex placement. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number: 1220648-2025-26077 as it is unknown if the initial reporter also sent the report to the food and drug administration. g4 pma/510(k) number revised as it was inadvertently reported incorrectly on initial manufacturer device report number: 1220648-2025-26077. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11742 | 1220648-2025-26078 | 21351139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21351139&pc=OZD | 01/24/2025 | 03/13/2025 |  | 02/10/2025 |  |  | 01/21/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579016 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced multiple complications during impella cp support. high purge pressure and purge pressure blocked alarms appeared on the automated impella controller. the purge cassette was replaced in an attempt to troubleshoot the issue but the failure persisted. tissue plasminogen activator was then added to the purge and the pressure stabilized. there was no resulting patient harm. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. data logs were reviewed and the high purge pressure event was confirmed. there was an initial step up in the purge pressure from 575 to 841 that aligned with a spike in the placement signal and deviations in the motor current and speed. shortly after, there was a motor current spike to 1098 ma and a speed deviation while at constant p-level. this coincided with the start of a four minute steep rise in the purge pressure that resulted in a high purge pressure alarm. while the purge pressure remained high, there continued to be several more abnormal motor current spikes. these are all signatures of biomaterial ingestion. the purge pressure resolved after seven hours, which aligns with the clinical details provided that the tissue plasminogen activator procedure was initiated. the product was not returned for investigation. based on data log review and clinical details, the root cause of the high purge pressure was most likely biomaterial ingestion. b.3 revised date of event as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-26078. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26078 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.3 date of awareness should of reflected 24-jan-2025 on the initial manufacturer device report number 1220648-2025-26078. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11743 | 1220648-2025-26079 | 21351192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21351192&pc=OZD | 01/22/2025 | 03/11/2025 |  | 02/10/2025 |  |  | 01/22/2025 | 02/10/2025 |  |  |  |  | 02/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565424 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while attempting to reposition the impella since it was too deep, the pump was pulled more than expected. staff attempted to reposition into the ventricle, the cannula kept bending causing suction alarm. staff decided to explant and replace. there was no patient harm. | [5 empty] | The investigation of positioning issues has been completed. the returned pump was visually inspected. there were two cannula kinks near outlet and near bending area observed. there was no biomaterial or broken blades observed. no other abnormality found. pump connected to aic to reproduce low pump flow issue since clinical details stated suctions occurred. pump run for more than 10 minutes in water. no low pump flow reproduced. pump flow was within specified limit. insufficient data log provided. the cause of the positioning issues most likely was use related since the pump was pulled back more than expected after deep insertion, cannula kinks occurred during attempted to reposition. failure mode: low pump flow issue added to the complaint since clinical details stated suctions occurred. the cause of the low pump flow issue was most likely cannula kinks due to improper repositioning technique. d4 udi number revised as it was reported incorrectly on initial manufacturer device report number. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11744 | 1220648-2025-26046 | 21332053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21332053&pc=OZD | 01/20/2025 | 03/22/2025 |  | 02/07/2025 |  |  | 01/20/2025 | 02/07/2025 |  |  |  |  | 02/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568279 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the implant, the sheath was placed successfully, yet after removing the dilator, it kinked. the sheath was replaced with a medtronic sentrant 18 french by 28 centimeter sheath. it was noted that the patient had very tortuous vessels that led to the sheath kinking. there was no report of patient harm. | [5 empty] | The investigation for the introducer damage has been completed. 14fr x 25cm impella introducer sheath and dilator were returned for evaluation. upon visual inspection, no damage to the introducer was observed. dilator tip did have slight curvature. clinical details noted that patient had very tortuous anatomy at insertion site. the root cause of the introducer damage - mechanical was most likely patient condition related as the patient had severely tortuous anatomy at insertion site. added:d9 device return date. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11745 | 1220648-2025-26047 | 21332064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21332064&pc=OZD | 01/21/2025 | 03/03/2025 |  | 02/07/2025 |  |  | 01/21/2025 | 02/07/2025 |  |  |  |  | 02/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1391515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) was giving persistent low purge pressure alarms during setup for patient support. the aic was exchanged as a result of the noted issue. there was no report of patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the controller had a purge pressure low alarm that did not resolve as case start was completing. reviewing the real time logs, the purge pressure was around 200 mmhg correctly triggering the low purge pressure alarm and flow was maxed out at 30 milliliiters. the pump and purge cassette were then moved over to console serial number (b)(6) and the alarms went away. however, after only eight hours of running the purge, the user performed a bag change. this was notable because the average purge flow should be around 10 milliliiters per hour so a 500 milliliiter bag should last 50 hours, not eight. the real time logs on console serial number (b)(6) were as expected. the controller was returned but the failure mode was not reproduced. the console was able to operate a purge/pump setup with the correct purge flows and pressures. additionally, the purge pressure was checked but confirmed to be accurate. the low purge pressure alarm was most likely caused by a leak while using the controller that was fixed when the pump was moved over to another controller. this accounts for the low purge pressure, high purge flow, and early bag change. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11746 | 1220648-2025-26048 | 21332333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21332333&pc=OZD | 01/20/2025 | 05/01/2025 |  | 02/07/2025 |  |  | 01/20/2025 | 02/07/2025 |  |  |  |  | 02/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548423 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | Fluid/Blood Leak | Increase in Pressure |  |  |  |  | Arrhythmia | Stroke/CVA | Septic Shock | Thrombocytopenia | No Clinical Signs, Symptoms or Conditions | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received from the clinical site that the patient was on impella 5.5 support for over two months. the patient was not a candidate for advanced therapies. while on impella support the patient experienced arrhythmias, thrombocytopenia, and septic shock was suspected requiring levo to be increased. the patient had a stroke that occurred after starting on an anticoagulant. the impella also encountered some suctions and p-level was lower. the impella had high purge pressure alarms that were resolved with lysis therapy. the patient expired while on impella support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The user facility reported a patient with an acute myocardial infarction was initially implanted with an impella cp and escalated to an impella 5.5 device. three days after start of support a small leak was noted at distal end of the purge side arm where it connects to the patient side of the driveline. purge flow and purge pressure are within normal limits. the leak is being monitored. the patient continues on mechanical circulatory support with the device and is noted to be in stable condition. there was no report or indication of patient harm as a result of this event. |  |  |  |  | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. log review confirmed that the purge flow and purge pressure remained within normal limits. the root cause of the pump purge leak was most likely a damaged sidearm but the exact location of the leak is unknown without the returned product. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ a.5 revised race as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26048. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26048. h.10 added instructions for use as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26048. | B1 revised to reflect both adverse event and product problem based on the additional information received from the clinical site. b2 revised to reflect death and required intervention based on the additional information received from the clinical site. b5 revised to reflect the additional information received from the clinical site. d6b explant date added h1 type of event revised to death based on the additional information received from the clinical site. h6 health effect - clinical code, health effect - impact code, and medical device problem code revised to reflect the additional information received from the clinical site. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11747 | 1220648-2025-26049 | 21332488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21332488&pc=OZD | 01/21/2025 | 03/06/2025 |  | 02/07/2025 |  |  | 01/21/2025 | 02/07/2025 |  |  |  |  | 02/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025590019 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race) with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and four days after start of support the impella 5.5 axillary site / jp drain site experienced bleeding. anticoagulants were temporarily held, and pressure was applied to achieve hemostasis. support remained ongoing during the event and until the patient underwent transplant surgery approximately five days later. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management , anticoagulants was held to achieve hemostasis as per the clinical description.
11748 | 1220648-2025-26050 | 21332567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21332567&pc=OZD | 01/17/2025 | 02/28/2025 |  | 02/07/2025 |  |  | 01/17/2025 | 02/07/2025 |  |  |  |  | 02/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589816 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Disability | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced limb ischemia and would subsequently expire. the patient was flu positive and with elevated troponins. initially diagnosed as non-st-segment elevation myocardial infarction and admitted to the telemetry unit. the patient converted to sinus rhythm spontaneously, and three days later st-segment elevation myocardial infarction was identified and confirmed. the patient underwent a left heart catheterization. upon engaging coronaries (left main, left anterior descending [lad], left circumflex [lcx] and right coronary arteries), the patient became hypotensive, levo started, and decision was made to place impella cp. after the impella cp was implanted, the patient's blood pressure "significantly" improved, and the physician proceeded with the percutaneous coronary intervention. cardiothoracic surgery was consulted for lad and lcx targets. while in catheterization lab, the physician initially preferred to leave in the peel-away sheath. the sheath shot showed no flow distally. peel-away sheath was exchanged for repositioning sheath. a sheath shot was repeated through re-access port and showed flow. however, the flow was sluggish. the patient was immediately taken to the operating room. coronary artery bypass graft surgery (cabg) x2 was performed. the patient was on multiple vasoactive drugs. the physician inspected the right leg, saw mottling and felt cool temperature. abiomed support recommended transitioning to impella 5.5 since chest was still open and there was option for direct insertion. however, the physician opted to perform an external fem-fem bypass. an 8 fr. retrograde on the left femoral supplying blood 6 fr antegrade via right using male to male connector. initially right foot had a palpable pulse. as the physician completed the cabg, the patient continued to be hypotensive which resulted in increased use of vasoactive medications. the right ventricle function was noted to be impaired. an impella rp was unavailable. dobutamine and milrinone were added for support. the impella cp position was confirmed via transesophageal echocardiogram, and the patient in critical condition was transferred to the unit. approximately six days later the right foot was noted without pulse. fem-fem bypass in place. the next day positive popliteal pulse on right extremity, no dorsalis pedis pulse was noted. five days following an impella rp flex was inserted. the patient now on bipella support. the next day continuous renal replacement therapy was noted. additionally, no flow in the right lower extremity was noted and the physician was unable to flow above p-5 without suction. it was stated that the patient continued to progressively decline with no progress resulting in multiple organ dysfunction (mod) syndrome. family ultimately decided to withdraw all life-sustaining measures. all therapies were stopped, and the patient expired. the patient never fully recovered from the cardiac event. the patient was in a poor state resulting to mod and expired. the patient had significant limb ischemia prior to the insertion of the rp flex device. limb ischemia was treated with fem-fem bypass so the implant was already quite severe. the impella related event for the limb ischemia is the impella cp. and the impella cp related events/periods of limb ischemia with surgical procedure fem-fem bypass cannot be excluded as possibly contributing to the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.1 revised to reflect adverse event only as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26050. b.2 revised selection as incorrect selections were entered on the initial manufacturer device report number 1220648-2025-26050. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26050. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26050 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 health effect - impact code for death was added as it was inadvertentlyomitted from the initial manufacturer device report number 1220648-2025-26050. code 1248 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26050. a new code was added to medical device problem code.
11749 | 1220648-2025-26051 | 21332592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21332592&pc=OZD | 01/20/2025 | 02/12/2025 |  | 02/07/2025 |  |  | 01/20/2025 | 02/07/2025 |  |  |  |  | 02/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132984 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Literature | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a purge disk not detected alarm appeared during preparation of an automated impella controller (aic) for patient support. despite replacing the purge cassette and troubleshooting the unit, the issue persisted. the aic was subsequently replaced to resolve the failure. there was no report or indication of interruption in support and noted harm. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) purge disc/flag issues has been completed. data logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) twice. the console was examined returned for evaluation and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. added h6 impact code 4629.
11750 | 1220648-2025-26052 | 21335265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21335265&pc=OZD | 01/09/2025 | 05/09/2025 |  | 02/07/2025 |  |  | 01/09/2025 | 02/07/2025 |  |  |  |  | 02/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock implanted with an impella cp for mechanical circulatory support encountered positioning issues that resulted in device replacement. the patient was treated with a drug-eluting stent to the left anterior descending artery and impella placement prior to being transferred from another local hospital for the possible surgery or high-risk percutaneous coronary intervention (pci) to other coronary diseased vessel(s). while on the unit, the patient experienced a ventricular tachycardia storm. therefore, the decision was made to bring the patient to the catheterization lab to have venoarterial extracorporeal membrane oxygenation (ecmo) placed and reposition the impella cp pump. the patient was cannulated for va ecmo. during this time, the patient started to have respiratory distress and was intubated. the physician attempted to reposition the impella pump. however, the pump was inadvertently pulled back across the aortic valve. when attempts to recross the valve were unsuccessful, the decision was made to remove the pump, rewire the left ventricle, and implant a new impella cp device. the patient was transferred to the unit on ecmo and impella therapies indicated to be in stable condition. | Upon follow-up from abiomed¿s long-term outcome and quality indicator (loqi) impella registry, it was additionally documented that the patient had endured a bilateral femoral artery access site bleed with a probable relationship to the impella device. it was noted that the patient had significant bleeding at their cannulation sites requiring redressing, several epinephrine/lidocaine, and several units of packed red blood cells. anticoagulation was put on hold to manage the condition. the patient was reported to have recovered from the condition. |  |  |  |  | The investigation for the positioning issue and ventricular tachycardia (vt) has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. the root cause of the vt could not be determined without additional information. corrections to the initial mfr report: d6b explant date was incorrectly noted and should have been the implant date. no explant has been provided at this time. | Device-specific information (other than impella cp) and including the manufacturer device date are unavailable. additionally, the implant date is unclear with the given information. therefore, the respective fields have been left blank accordingly. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B.5 additional information was received from the clinical site. g.2 added report source; study due to new information that was received from the clinical site. h.6 added codes to reflect additional failure mode according to the additional information received from the clinical site. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.4 added catalog number, serial number, lot number, expiration date and unique identifier (udi) # as they were inadvertently omitted on previous manufacturer device report number 1220648-2025-26052. h.4 added device manufacture date as it was inadvertently omitted on previous manufacturer device report number 1220648-2025-26052. h.6 code 4315 was inadvertently omitted on previous manufacturer device report number 1220648-2025-26052 and was added. | The root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the investigation results for positioning issue and ventricular tachycardia were reported in manufacturer device report number 1220648-2025-26052-1.
11751 | 1220648-2025-26053 | 21336159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21336159&pc=OZD | 01/09/2025 | 02/26/2025 |  | 02/07/2025 |  |  | 01/09/2025 | 02/07/2025 |  |  |  |  | 02/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587405 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy in cardiogenic shock implanted with an impella 5.5 for mechanical circulatory support (mcs) went into cardiac arrest and was able to be converted back to normal sinus rhythm. the impella was placed via the right axillary artery without complications, the patient was sent to the unit on support and was noted as a status 1 for heart transplant. three days after start of the mcs, the patient went into cardiac arrest and was shocked three times back to normal sinus rhythm. echocardiogram was completed with repositioning of the device. the patient continued on support until explant of the device 11 days later. the patient's outcome at explant is survived. | [5 empty] | The investigation of cardiac arrest has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the cardiac arrest was not determined. b.2 revised as an incorrect selection was entered on the initial manufacturer device report number 1220648-2025-26053. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26053 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11752 | 1220648-2025-26021 | 21320835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21320835&pc=OZD | 01/18/2025 | 02/12/2025 |  | 02/06/2025 |  |  | 01/18/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1854146 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported while supporting a patient implanted with an impella 5.5 device, the automated impella controller (aic) displayed an air in purge system alarm with audible indicator that randomly started. the team attempted to repair. however, the air in purge system alarm continued. consequently, the aic was exchanged. there was no patient harm as a result of the event. additionally, there was no report or indication of interruption in support. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the console logs from the complaint date are consistent with the reported complaint. the console was tested after arriving in field service. however, the air in purge system alarm could not be reproduced. this console was tested with a known good pump and purge cassette kit but functioned normally without any issue. the root cause of this issue was likely use related. there was likely a kink in the purge line or the purge fluid bag needed to be changed.
11753 | 1220648-2025-26022 | 21320924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21320924&pc=PYX | 01/18/2025 | 03/07/2025 |  | 02/06/2025 |  |  | 01/18/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ENDOVASCULAR AORTIC/ANEURYSM (EVAR) STENT | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an unsuccessful attempt to implant an impella rp device to a patient for mechanical circulatory support. the physician was unable to pass the 23 french sheath and dilator through the right and left femoral veins. the physician noted the patient had a history of endovascular aortic/aneurysm (evar) repair, and the sheath appeared to be getting caught up on the evar stents causing the peel away sheaths to prematurely split at the tip. the decision was made to abort the impella rp implant and place the patient on extracorporeal membrane oxygenation. the patient was noted to be stable following the event. | [5 empty] | A.1, a.3, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the delivery issue and the introducer damage has been completed. clinical details state that the 23fr sheath was unable to be inserted on patient with evar stents. decision made to abort rpsa and put patient on ecmo. no product was identified by serial number despite reaching out to rep. no data logs exist as no product was inserted. the clinical details were reviewed and the cause of the delivery issue was most likely patient condition related to the pre-existing evar stent placements. evar stents prematurely split peel away sheath. patient in stable condition. the clinical details were reviewed and the cause of the introducer damage - mechanical was most likely an interaction with evar stents.
11754 | 1220648-2025-26023 | 21321126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321126&pc=OZD | 01/18/2025 | 03/21/2025 |  | 02/06/2025 |  |  | 01/19/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108704 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. per medical safety officer review there is not enough evidence to exclude the automated impella controller (aic) as an associated factor in the patient's outcome. | The user facility reported while supporting a patient implanted with an impella 5.5 device, the automated impella controller (aic) alarm noise was very low and making a different sound than normal. a silent hum was observed rather than the number beep the aic makes for white informational alarm. the stated plan is to return the aic for repair after the patient completes the impella 5.5 mechanical circulatory support. there was no harm to the patient as a result of this event with not report or indication of uninterrupted support. |  |  |  |  | The investigation of automated impella controller (aic) - buzzer/ alarm inaudible has been completed. the console logs from the reported day of event did not contain any information relevant to the reported issue. the console was received for investigation where it was extensively tested. we were unable to reproduce any audible anomaly, including one that would match specific problem description. the cause of the issue was not determined since issue could not be reproduced. b.5 additional information was received, and abiomed's medical safety officer review was completed. b.1 revised from the initial manufacturer device report number 1220648-2025-26023 to reflect both adverse event and product problem. b.2 revised from the initial manufacturer device report number 1220648-2025-26023 to reflect death and date of death. e.1 fax number was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26023. h.1 revised from the initial manufacturer device report number 1220648-2025-26023 to reflect death. h.6 revised from the initial manufacturer device report number 1220648-2025-26023 to reflect death. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11755 | 1220648-2025-26024 | 21321230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321230&pc=OZD | 01/20/2025 | 03/21/2025 |  | 02/06/2025 |  |  | 01/20/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571840 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had bleeding from the access site after removal of the peel away and insertion of the repositioning sheath. blood products were given and a comprehensive procedure was carried out. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed has been completed. the product was not returned for review. bleeding was reported at the leg access site after inserting the peel-away. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review.
11756 | 1220648-2025-26025 | 21321437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321437&pc=OZD | 01/18/2025 | 03/17/2025 |  | 02/06/2025 |  |  | 01/18/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561267 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and approximately 28 days after start of this support, the pump stopped. consequently, the impella cp device was successfully explanted and replaced with a new one. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation of pump stop has been completed. no product or data logs were returned for evaluation. clinical details noted a pump stop occurred on day 29 of support. no additional details on the pump stop were provided. the root cause of the pump stop cannot be definitively determined as limited clinical details were provided and no logs or product were returned for analysis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
11757 | 1220648-2025-26026 | 21321583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321583&pc=OZD | 01/18/2025 | 03/04/2025 |  | 02/06/2025 |  |  | 01/18/2025 | 02/06/2025 |  |  | 01/31/2025 |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314717 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient with acute myocardial infarction was unsuccessful. the device was unable to be inserted, and the case was aborted. there was no report of adverse effects to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella was unable to pass through the vasculature. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11758 | 1220648-2025-26027 | 21321717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321717&pc=OZD | 01/19/2025 | 03/07/2025 |  | 02/06/2025 |  |  | 01/19/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601305 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient with an acute myocardial infarction was unsuccessful due to the cardiac anatomy and thrombus. the impella was inserted in left groin, the iliac was calcified and the impella was removed. after removal, a clot on the cannula and motor was observed. the impella cp device was exchanged for a new one and was successfully implanted. there were no adverse effects to the patient. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and thrombus has been completed. the pump was inserted in left groin but the iliac was calcified so the pump was removed; clot in cannula and motor was observed. returned complaint device visually inspected. no cannula kink observed. small biomaterial residue around impeller observed. no other abnormality found. the cause of the delivery issue most likely was patient condition related due to calcification vessels. as t he necessary information was not provided, the root cause of the thrombus was not determined.
11759 | 1220648-2025-26028 | 21321758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321758&pc=OZD | 01/08/2025 | 02/26/2025 |  | 02/06/2025 |  |  | 01/08/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568648 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Increase in Pressure | [5 empty] | Insufficient Information | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was a drop in purge flow and purge pressure increased to the 800's. tissue plasma activator was added to the purge and support continued. additionally, there were low pump flows during support. the patient received one unit of packed red blood cells. support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of low pump flow and high purge pressure has been completed. the impella device was not returned for evaluation. low pump flow: in the data logs it can be seen that there are many suction events seen throughout the case occurring more frequently at higher p-levels. the motor current was very flat and there was a trend in the placement signal. it was stated that the patient was on continuous renal replacement therapy (crrt) and having volume issues. the root cause of the low pump flow is most likely due to the patients condition. high purge pressure: the data logs show that on the 8th around 12:30pm there was a motor current spike followed by a 3 minute rise in purge pressure to around 1000mmhg. pump flows became saturated and the purge flows decreased to about 0ml/h. no high purge pressure high alarms were triggered as it did not reach the 1050mmhg threshold. there were purge volume low alarms triggered. the purge pressure remained elevated for an hour but was trending down insinuating biomaterial hit the impeller and then was in the purge gap. tissue plasma activator (tpa) protocol was started immediately at 1:25pm as the patient had been off systemic anticoagulation. the elevated purge pressure resolved and remained stable for the duration of the case. the root cause of the high purge pressure is most likely due to biomaterial. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26028 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11760 | 1220648-2025-26029 | 21321797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321797&pc=OZD | 01/08/2025 | 02/26/2025 |  | 02/06/2025 |  |  | 01/08/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583093 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | Ischemia Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after impella 5.5 was placed the patient had atrial fibrillation in the operation room and was cardioverted. at the end of the case, transesophageal echocardiogram showed that the 5.5 was flipped. the team attempted to repositioned, but the physician decided not to reposition due to the patient ventricular septal defect repair. the team is aware of the positioning issue. additionally, the patient had a computed tomography scan for new facial droopiness and found to have transient ischemic attack. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of positioning issues, stroke, and atrial fibrillation (af) has been completed. data logs showed suction alarms the first few days of the case which resolved. brief impella position unknown alarms at various times throughout support were noted. motor current is stable. the cause of the positioning issue was not determined due to lack of clinical information and no product returned. as this clinical information was not provided, the true root cause of the stroke could not be determined. as the necessary information was not provided, the root cause of the af was not determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26029 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11761 | 1220648-2025-26030 | 21321804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321804&pc=OZD | 01/17/2025 | 02/14/2025 |  | 02/06/2025 |  |  | 01/17/2025 | 02/06/2025 |  |  | 07/31/2025 |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440456 |  | 1000080 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | Malposition of Device |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) experienced thrombus, positioning issue with pump stop. the patient presented to the emergency department for evaluation and an acute st-segment elevation myocardial infarction was diagnosed. "severe" multi-vessel disease found including chronic total occlusion of the right coronary and left anterior descending arteries. the initial decision was made to place an intra-aortic balloon pump (iabp) and plan for coronary artery bypass graft surgery. the patient would decompensate within an hour, was intubated and placed on venoarterial extracorporeal membrane oxygenation (va emco). pressors started for blood pressure. echocardiogram showed significantly reduced left ventricular function with an ejection fraction of 10% and "very poor" right ventricular function. an impella cp device was implanted for mcs and multiple complications were encountered involving the impella cp pump. during support, it was discovered that the pigtail had wrapped around the patient's mitral valve. upon repositioning the device, a large spike in motor current occurred, and the pump stopped. the pump was removed, and the decision was made to not place a new pump due to a high clot burden in the left ventricular outflow tract and aortic arch. the patient survived the event with next steps determining if the patient is a transplant candidate. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation has been completed. the failure mode "pump stop" was changed to "low pump flow" because the pump stop that occurred was initiated by the user turning p-level down to p0. the returned pump was inspected and thrombus was found sitting on the impeller. there were no other physical abnormalities and the pump passed electrical testing. the data logs showed a motor current spike with speed deviation and a drop in pump flow characteristic of biomaterial ingestion. the pump continued running with low flows until the pump was turned down to p0 by the user. the pump was restarted and turned to p3 four seconds later and flows continued within good range for most of the remainder of the case. the root cause of the low pump flow was determined to be due to biomaterial ingestion. the clinical details mention that repositioning was attempted as the pump was directed toward the mitral valve. while repositioning, the pump was pulled back to the descending aorta. the root cause of the positioning issues was not determined as limited clinical information was provided related to how the pump was mispositioned. as the necessary information was not provided, the root cause of the thrombus was not determined. h.6 a new medical device problem code was added due to investigation results. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.9 revised is the device available for evaluation as the incorrect selection was marked on the initial manufacturer device report number 1220648-2025-26030. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26030. code 1503 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26030 due to investigation results. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2025-26030 as some instructions were omitted and some have changed due to investigation results.
11762 | 1220648-2025-26031 | 21321809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321809&pc=OZD | 01/20/2025 | 03/12/2025 |  | 02/06/2025 |  |  | 01/21/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590512 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and shortly after the implant, the pump started to leak from the side arm, this caused low purge pressure. the pump was explanted and there was no reported patient harm. | [5 empty] | The investigation of purge leak ¿ pump has been completed. no product or data logs were returned for evaluation. clinical details noted that a leak was observed from purge sidearm and low purge pressure was observed in the field. the root cause of the purge leak - pump is most likely due to a damaged sidearm but the exact location of the leak could not be determined as no product was returned for evaluation. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26031 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
11763 | 1220648-2025-26032 | 21321812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321812&pc=OZD | 01/17/2025 | 02/28/2025 |  | 02/06/2025 |  |  | 01/17/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597119 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with am impella 5.5 device for mechanical circulatory support (mcs), and approximately seven days after start of support the patient developed a hematoma. the hematoma was in the chest near impella incision. computed tomography angiography of chest and neck showed a large hematoma in pectoral muscle and active arterial bleed. the patient was taken to the operating room for an emergent surgical cutdown. the following day no hematoma was noted, site was clean, and hemoglobin and hematocrit were stable. the patient was reported to be stable following the event. impella mcs continued for an additional two days at which time the patient underwent left ventricular assist device placement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26032 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11764 | 1220648-2025-26033 | 21321857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321857&pc=OZD | 01/18/2025 | 02/24/2025 |  | 02/06/2025 |  |  | 01/18/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561379 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs), and the following start of support, the patient developed limb ischemia and expired. the patient was in a hemodialysis run and started to feel week; systolic blood pressure was in the 40's. heparin and levo drips were started, and the patient was transferred to the hospital. a left heart catheterization was performed for concern of an acute myocardial infarction. "extensive" left anterior descending artery disease was noted and vessel size too small to correct. additionally, there was diffuse disease in right coronary artery. the left ventricular end-diastolic pressure was 25 and the patient's hypotension was worsening. the decision made to place impella cp via left femoral artery which was successfully completed. upon arrival to the unit, left pedal pulse was not palpable. the nurse attempted to find the pedal pulse using a doppler with no success. the physician was informed, it was noted the patient had severe vascular disease, and, however, due to the patient's degree of cardiogenic shock, the patient needed the temporary mcs. reportedly the patient would expire the next day. no further details surrounding the death was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.2 revised selections as disability was reported incorrectly on the initial manufacturer device report number 1220648-2025-26033. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26033 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4614 was reported incorrectly on the initial manufacturer device report number 1220648-2025-26033.
11765 | 1220648-2025-26034 | 21321860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321860&pc=OZD | 01/18/2025 | 03/03/2025 |  | 02/06/2025 |  |  | 01/18/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595538 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient, in cardiogenic shock from an acute myocardial infarction, implanted with an impella cp device for mechanical circulatory support developed limb ischemia and would eventually expire. the patient was transported to the hospital via ambulance and coded during transport. cardiopulmonary resuscitation was performed with return of spontaneous circulation and the patient was intubated. the initial angiogram showed small diameter of right femoral artery and there was uncertainty if the impella sheath would fit. due to patient's decline in status the decision was made to quickly place an impella cp device. percutaneous coronary intervention (pci) was performed to the left circumflex artery. there was another lesion in the left anterior descending artery left untreated as the physician was unable to pass guidewire due to severe calcification. the plan was to cool the patient and bring the patient back for treatment at a later date. at end of case the peel-away sheath was exchanged for the repositioning sheath. the team was unable to doppler any pulses in right foot. lower leg was cold, and bottom of right foot was mottled. decision was made to place an external fem-fem bypass. a 6f antegrade access in right superficial femoral artery below impella sheath and 6f retrograde access in left femoral artery. sheaths were connected and flow was noted. the patient was transferred to the unit for monitoring, on support, but would expire two days later. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11766 | 1220648-2025-26035 | 21321878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321878&pc=OZD | 01/19/2025 | 03/04/2025 |  | 02/06/2025 |  |  | 01/19/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591476 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction was implanted with an impella cp for mechanical circulatory support. after transfer to the unit oozing of the access site was noticed and successfully managed with a femstop after tightening perclose devices with manual pressure failed. the patient continued on support for a total of three days. the patient was successfully weaned, and the device was explanted without incident. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation). | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
11767 | 1220648-2025-26036 | 21321887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321887&pc=OZD | 01/20/2025 | 03/06/2025 |  | 02/06/2025 |  |  | 01/20/2025 | 02/06/2025 |  |  | 07/31/2025 |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371577 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the physician attempted to deliver the impella 5.5 completely into the axillary artery and into proper position with fluoroscopic guidance but was unsuccessful due to patient anatomy. the inlet of the impella 5.5 appeared to be able to navigate through the arteriotomy and into the axillary artery. attempts to advance the 5.5 at this point resulted in kinking the catheter and 0.018 wire. after many attempts, the 0.018 wire was safely removed from the heart, the 5.5 was explanted from the graft, and the graft was stapled, and pocket was closed. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was determined to be patient condition because the insertion was unsuccessful due to patient anatomy. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26036 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
11768 | 1220648-2025-26037 | 21321894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21321894&pc=OZD | 01/20/2025 | 02/24/2025 |  | 02/06/2025 |  |  | 01/20/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594251 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and approximately four days after start of support, the patient developed limb ischemia followed by arterial thrombus. the patient had no distal pulse in either of the legs during the impella cp support. it was additionally noted that an arterial thrombus was present bilaterally below the popliteal. it was noted the event was not believed to be associated with the impella; however, no alternative cause is identified as a likely cause for the event. in response to the noted condition, plans were made to explant the pump which was completed two days later. the patient was noted to be stable following the event and explant of the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details.
11769 | 1220648-2025-26038 | 21322124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21322124&pc=OZD | 12/17/2024 | 03/10/2025 |  | 02/06/2025 |  |  | 01/19/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573328 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs). the patient was on impella mcs for approximately 45 days. five days after start of support, the patient developed thrombocytopenia. a notification was received via abiomed¿s long-term outcome and quality indicator impella registry noting the patient that endured thrombocytopenia with a "possible" relationship to the impella device used. platelet count baseline was 219 at admission. five days later, the platelet count began dropping. days following, the lowest platelet count was at 23. however, day 16 of support, the platelet count increased to 118 and continued to rise. the patient recovered from this event with no sequelae. support continued now on day 36 of support the nurse reported intermittent increases in purge pressure from a twist in the impella catheter just prior to the red impella plug. the impella catheter was untwisted and the increases in the pressure resolved. there were no adverse effects to the patient as a result of the impella catheter twisting. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the thrombocytopenia and product damage has been completed. no product returned. data logs show multiple hpp alarms all gradually rising within a period of 5 minutes to 20 minutes and resolving within a few minutes. no motor current spikes observed. clinical details noted twisted/coiled/torqued patient cable and twisted catheter and unkinking resolved the increased purge pressure. the cause of the high purge pressure was most likely a kinked purge line. the exact location of the kink cannot be determined without product return. the cause of the thrombocytopenia was not determined due to insufficient clinical details.
11770 | 1220648-2025-26039 | 21322247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21322247&pc=OZD | 01/21/2025 | 03/05/2025 |  | 02/06/2025 |  |  | 01/21/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594717 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support for a patient coding and in acute myocardial infarction/cardiogenic shock. the pump was placed and on 2nd day of support an access site bleed was noted. the bleed was treated by three units of blood and pressure being held while medical therapy and femostop placed. family decided to withdraw care and patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be anticoagulation management because anticoagulation was supratherapeutic and heparin was discontinued.
11771 | 1220648-2025-26040 | 21322275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21322275&pc=OZD | 01/21/2025 | 02/24/2025 |  | 02/06/2025 |  |  | 01/21/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089162 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) error occurred while supporting a patient implanted with an impella cp device along with venoarterial extracorporeal membrane oxygenation due to an acute myocardial infarction/cardiogenic shock. the aic was successfully exchanged. there was no report or indication of patient harm or interruption of support which is still ongoing. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. the console logs from the reported day of event are consistent with the complaint and showed multiple recurring controller error alarms due to pmd voltage out of spec, that self-resolve immediately but keep repeating. during testing, the issue was reproduced. power was delivered to impellatronic by the power battery manager and was verified to be within specification. replacing the impellatronic board with a known-good one resolved the issue. the cause of the aic issue was a defective impellatronics board.
11772 | 1220648-2025-26041 | 21322421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21322421&pc=OZD | 01/21/2025 | 03/07/2025 |  | 02/06/2025 |  |  | 01/21/2025 | 02/06/2025 |  |  | 05/31/2025 |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352619 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was unable to advance through the innominate artery during attempted delivery. the implant was aborted and an impella cp pump was placed for planned patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the delivery issue has been completed. impella 5.5 pump was unable to advance through the innominate artery during attempted delivery. returned complaint 5.5 pump visually inspected. cannula kinks observed near outlet and middle of the cannula, likely due to the challenging delivery. the cause of the delivery issue was most likely patient condition related due to vessels condition (pad/pvd).
11773 | 1220648-2025-26042 | 21322442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21322442&pc=OZD | 01/19/2025 | 03/03/2025 |  | 02/06/2025 |  |  | 01/19/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573368 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had ventricular tachycardia (vt). this correlate with patient movement of the right arm. the impella 5.5 was implanted in the right axillary. the pump was repositioned after the vt and support was continued. additionally, the staff believed the patient had a left ventricle thrombus. anticoagulation was given to the patient. the patient was having vt again and the decision was made to place intra-aortic balloon pump and remove the impella 5.5. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis. | The investigation of ventricular tachycardia and thrombus has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the ventricular tachycardia and thrombus was not determined.
11774 | 1220648-2025-26043 | 21324037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21324037&pc=OZD | 01/21/2025 | 03/05/2025 |  | 02/06/2025 |  |  | 01/21/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599952 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 placement was aborted due to the patient's anatomy. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to resistance at innominate artery. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
11775 | 1220648-2025-26044 | 21328796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21328796&pc=OZD | 01/12/2025 | 02/28/2025 |  | 02/06/2025 |  |  | 01/12/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574196 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Crack | Pumping Stopped | [5 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock undergoing an impella cp device implant for mechanical circulatory support (mcs) experienced a pump stop and would later expire. the patient was admitted with distal left main, left anterior descending and circumflex arteries disease and under surgical consult was placed on an intra-aortic balloon pump (iabp) and sent to the unit. the patient began to further decompensate and was escalated to an impella cp from an iabp. the right femoral artery was prepped and the iabp was removed. priming steps were being followed when the automated impella controller (aic) skipped the priming step and went straight to the impella start mode/initial placement screen. on aic performance level was p-0 with the impella stopped retrograde flow followed by impella stopped motor current high. the team attempted to re-prime with de-airing process; no fluid was seen coming at around motor area. at this point, the patient unstable as iabp was already removed. therefore, a new impella cp device was opened, and the physician proceeded with implant which was completed without incident. additionally, the 14fr 13cm peel-away sheath had a cracked at the hub and therefore, it was replaced by a new one. after one day on the replacement impella cp support, it was noted lactate came down but did not completely clear. therefore, a plan was made it to escalate to the impella 5.5 to undergo protected percutaneous coronary intervention with no plan for coronary artery bypass graft surgery. the patient was brought to the operating and peripheral veno-arterial (va) cannulated. the physician lost access for left femoral artery extracorporeal membrane oxygenation (ecmo) site. a retroperitoneal bleed occurred which was managed by surgical cutdown and repaired. impella cp was removed and dilated up to use as arterial access for va. next right axillary artery exposed, 10mm graft anastomosis done. axillary kit was used, ventricular access was done with a jr4 catheter. the impella 5.5 device was advanced and slowly ramped coming on ecmo flows. impella was secured to chest. next peripheral va decannulated. massive blood transfusion was required. the patient received 7 units of packed red blood cells, 1 units of platelets, and 1 units of fresh frozen plasma. the patient continued to code on the unit, massive blood products transfusion going; the patient was bleeding internally per the plan of care updates. after approximately three days on impella 5.5 support the patient would expire. there is not enough evidence to exclude impella cp device initially implanted as an associated factor in the patient's outcome. there was an initial delay in getting the patient on appropriate support. once support was met, the patient received impella mcs for approximately 4 days. patient had procedural complications related to non-abiomed device(s) which most likely resulted in the patient's demise, however, (which included internal bleeding, massive blood transfusion requirements). | [5 empty] | The impella pump with the introducer has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the pump stop and introducer damage has been completed. the pump and short introducer were returned for investigation. the returned pump had the red lumen as it was assembled. the pump was primed and operated according to the specifications after removing the red lumen. no other physical damage was noted on the returned pump. the introducer was investigated, and it was found that one side of the orange hub was detached and missing from the returned package. it was not possible to verify the applied adhesive due to the missing component. no other damage was noted to the introducer. logs provided evidence that the case start was successfully completed and the pump was set to p9 when the pump stop-motor current high alarm was reported. the cause of the pump did not start issue was a red lumen removal issue. red lumen should be removed before starting or implanting the pump. the cause of the introducer damage issue was not determined since sufficient product was not returned. b.2 revised date of death as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26044. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26044 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11776 | 1220648-2025-26045 | 21328805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21328805&pc=OZD | 01/21/2025 | 03/18/2025 |  | 02/06/2025 |  |  | 01/21/2025 | 02/06/2025 |  |  |  |  | 02/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564665 |  | OZD |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was had an attempted impella 5.5 implant and it was aborted due to difficulty in inserting the device and damage to the inhalation site. the implant was aborted and a impella cp was placed instead. the patient remained stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the pump was returned for inspection. upon investigation, it was observed that the cannula was kinked and the catheter was found to be kinked near the end of cannula. the root cause of delivery issue is determined to be patient condition because after aborting 5.5 they were able to insert cp successfully. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26045 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11777 | 1220648-2025-25993 | 21312120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312120&pc=OZD | 01/18/2025 | 02/24/2025 |  | 02/05/2025 |  |  | 01/19/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558930 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia Stroke | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support experienced runs of ventricular tachycardia. later the patient had an ischemic stroke while on impella 5.5 support therefore the decision was made to removed the pump. upon removal the physician noted organized thrombus on outside of cannula. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of ventricular tachycardia, thrombus, and stroke has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
11778 | 1220648-2025-25994 | 21312141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312141&pc=OZD | 04/10/2024 | 02/27/2025 |  | 02/05/2025 |  |  | 01/16/2025 | 02/05/2025 |  |  | 12/31/2025 |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456421 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pseudoaneurysm | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced drifting hemoglobin and subsequent cross-sectional imaging revealed evidence of left femoral artery pseudoaneurysm eight days after the initial impella explant with a "possible" relationship to the impella device used. repair indicated in combination with planned transcatheter aortic valve replacement and removal of right axillary impella. the patient was noted to be stable following the event. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
11779 | 1220648-2025-25995 | 21312240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312240&pc=OZD | 01/17/2025 | 03/18/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592678 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella 5.5 device was in a patients aorta while improving the positioning in the ventricle. the impella cannula was completely folded over in the aorta and the pump stopped. the impella 5.5 device was explanted and replaced for preoperative placement to provide the patient mechanical circulatory support. reportedly, there was no patient harm as a result of the event. | [5 empty] | The investigation of pump stop and product damage (cannula kink) has been completed. the pump was returned for investigation; however, the data log was not provided. a kink in the cannula was noted near the motor, but no other physical abnormalities were reported on the returned product. the pump passed electrical testing and operated as expected when tested in a bucket of water. the pump stop did not reproduce even when it was tested with fully folded cannula at the kink. as per the given clinical details, cannula is completely folded over in the aorta, causing the impella device to stop functioning. no other failure modes were confirmed as a result of the cannula kink, during case run. the cause of the pump stop issue was not determined without sufficient clinical information including data log. pump stop did not reproduce during the test. the cause of the product damage-cannula kink issue was most likely use relate | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11780 | 1220648-2025-25996 | 21312299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312299&pc=OZD | 10/25/2024 | 02/18/2025 |  | 02/05/2025 |  |  | 01/16/2025 | 02/05/2025 |  |  | 01/31/2026 |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025451625 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification via abiomed¿s long-term outcome and quality indicator impella registry was received regarding a patient who experienced non-sustained ventricular tachycardia during impella supported bioprosthetic mitral valve replacement; one shock was administered. additionally, per the registry, the patient experienced atrial fibrillation post bioprosthetic mitral valve replacement and left atrial appendage ligation. cardioversion was done prior to discharge, and the patient recovered with no sequelae. the registry indicated the events are possibly related to the impella device. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11781 | 1220648-2025-25997 | 21312354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312354&pc=OZD | 01/17/2025 | 02/21/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563028 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, there was a lack of flow to the patient's left leg. due to the concern of leg ischemia, external femorofemoral bypass was placed. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. h.6 component code 925 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25997 and a new code was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11782 | 21312359 | 21312359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312359&pc=OZD | 01/12/2025 | 01/30/2025 | 01/30/2025 | 02/05/2025 | 01/30/2025 | 02/05/2025 |  | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 513479 |  | 0048-0003 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | NA |  | Prefer Not To Disclose |  |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During prep of the impella device, the catheter was not recognized by the console. another console was attempted with no connection. the catheter was removed and replaced with a new impella catheter with no problems. the rep took the defective catheter for inspection by abiomed. manufacturer response for catheter, catheter (per site reporter). mfg notified by site.
11783 | 1220648-2025-25998 | 21312408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312408&pc=OZD | 01/18/2025 | 03/10/2025 |  | 02/05/2025 |  |  | 01/18/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597856 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. post implant, same day, the sterile sleeve of the implanted impella 5.5 pump was inadvertently cut after the drapes were removed. a tegaderm was placed over the opening using a sterile technique to manage the damage. support with the implanted pump was uninterrupted. there was no report of indication of patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sterile sleeve damage) has been completed. the pump was not returned for investigation. as per the clinical information provided, it is clear that the sterile sleeve was accidently cut by the doctor and tegaderm was placed over it. therefore, the root cause of the product damage - sterile sleeve issue was use related to improper technique.
11784 | 1220648-2025-25999 | 21312542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312542&pc=OZD | 01/17/2025 | 03/19/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548194 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device mechanical circulatory support. within 24 hours after the high-risk pci after impella support, hemolysis was observed through changed in the patient's urine color. the device was explanted in the procedure area and discarded. the patient was noted to have survived no further report or indicated adverse effects. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. d.3 added us zip code as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25999. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25999. h.6 code 4641 was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25999.
11785 | 1220648-2025-26000 | 21312566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312566&pc=OZD | 01/18/2025 | 03/05/2025 |  | 02/05/2025 |  |  | 01/19/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555104 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Advance | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump being placed to support a patient admitted in acute myocardial infarction/cardiogenic shock. the patient arrested, ventricular fibrillation arrested and had cardiopulmonary resuscitation prior to hospital admission. the impella cp sheath was placed and during the delivery, the patient experienced ventricular fibrillation again and was defibrillated twice and went asystolic. the patient expired before pump was delivered. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, the patient expired before pump delivered via the sheath.
11786 | 1220648-2025-26001 | 21312671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312671&pc=OZD | 01/20/2025 | 03/12/2025 |  | 02/05/2025 |  |  | 01/20/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557936 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the automated impella controller (aic) could not be started. the aic was connected to the power supply the whole time, even the small green light was on. however, there was no reaction on the display. the aic was replaced and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. no console logs were observed from the reported complaint date. this was to be expected given the user was not able to startup the console. however, there were logs observed prior to the complaint date that were indicative of a malfunctioning. console was tested after arriving in field service. the reported startup issue could not be reproduced. the console was able to start up normally. in effort to investigate the cause of the kbd errors, the cables from the 4-in-1 to the kbd and power battery manager (pbm) to 4-in-1 were inspected, but no issues were found. the logs that were indicative of a malfunctioning 4-in-1 subsided when testing this console with a known good 4-in-1. conclusion: the imc-kbd error¿s observed in the logs were caused by a malfunctioning 4-in-1. the cause of the start up issues were not determined due to the inability to reproduce. d6a/d6b should have been left blank on manufacturer device report 1220648-2025-26001, due to the this complaint is for a console which is not an implantable device. h6 code 925 was reported incorrectly on manufacturer device report 220648-2025-26001. new code was added to component code | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11787 | 1220648-2025-26002 | 21312804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312804&pc=OZD | 01/16/2025 | 02/27/2025 |  | 02/05/2025 |  |  | 01/16/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553741 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 95 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an attempt to implant an impella cp device pump for mechanical circulatory support in a patient experiencing acute myocardial infarction/cardiogenic shock was aborted due to the patient's anatomy. the physician decided to complete the percutaneous coronary intervention unprotected. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient condition because of the resistance at femoral artery. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system, section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a/d6b revised on manufacturer device report 1220648-2025-26002, the dates were absent from the initial report. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26002 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 revised on manufacturer device report 1220648-2025-26002, the instructions for use were absent from the initial report.
11788 | 1220648-2025-26003 | 21312817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312817&pc=OZD | 01/17/2025 | 03/03/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288947 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that in preparation, the automated impella controller (aic) had system self check failed alarm. no further information was provided. there was no report or indication of patient harm or delay in support. | [5 empty] | The investigation for the automated impella controller (aic) self check issue has been completed. the console logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the aic was returned for evaluation. the reported failure mode was reproduced during testing. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. a.3 revised as selection was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26003. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
11789 | 1220648-2025-26004 | 21312833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312833&pc=OZD | 01/17/2025 | 02/24/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582367 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella 5.5 device to a patient with postcardiotomy cardiogenic shock was unsuccessful due to the cardiac anatomy. there was difficulty inserting the device over the guidewire. the impella 5.5 was removed and kink on the guidewire was observed. the physician attempted to insert a new abiomed guidewire but was still unable to advance. the decision was made to open and use a new impella 5.5 which successfully implanted. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigatino has been completed. no product was returned. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d.6a and d.6b revised implant and explant dates as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-26004. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26004 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4629 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26004 and was added. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted initial manufacturer device report number 1220648-2025-26004 and were added.
11790 | 1220648-2025-26005 | 21312846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312846&pc=PYX | 01/17/2025 | 02/21/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  | 02/28/2026 |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025480363 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella rp flex for mechanical circulatory support. approximately three days after start of support, following swan removal, it was believed that the implanted impella rp flex pump ingested a clot. due to patient anatomy with scoliosis, the decision was made to replace the impella rp flex in the internal jugular. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. b.7 revised since the initial manufacturer device report number 1220648-2025-26005 was submitted as heart failure was with preserved ejection fraction and not reduced. d.3 revised address line 2 to reflect blank as it should of been blank on the initial manufacturer device report number 1220648-2025-26005. revised us city as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26005. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26005 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address section as it was entered incorrectly on the initial manufacturer device report number 1220648-2025-26005.
11791 | 1220648-2025-26006 | 21312851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312851&pc=OZD | 01/17/2025 | 03/10/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592335 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. approximately four days after start of support, systemic heparin was stopped due the patient bleeding from the mouth. an otolaryngologist was consulted. additionally, an echocardiogram at bedside was completed, and a clot at inlet was noted. there were pump pressure alarms with purge flow <3 l/min and pressure >1100. the cardiothoracic surgeon at bedside explanted the impella 5.5 device in sterile prepped field with betadine with surgical stapler and successful wound closure. a clot was confirmed on the pump after removal. the access site was dry and clean, and the patient was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist systm. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation has been completed. data logs were reviewed and the complication was confirmed. however, the abnormalities start the day before. on (b)(6) 2025 there was a bag/cassette change that exceeded two minutes. later that day, there are two motor current spikes that correlate with motor speed deviations and spikes in the placement signal, followed by a motor current trend up without a change in p-level. then on (b)(6) 2025, there was another long bag/cassette change that appeared to be the onset of a trend in the purge pressure. both bag changes showed purge pressure dropping to 0 mmhg. finally, around 8:00 am on (b)(6) 2025, a clear motor current spike to 1200 ma and motor speed deviation were followed by false position in aorta alarms, saturated flows, and a five minute rise in the purge pressure until alarms were triggered. this timeline suggests that the long bag changes were the onset of the issue, and likely increased the likelihood for the biomaterial ingestion event that followed. clinical details such as the patient's systemic heparin being on pause, extracoporeal membrane oxygenation decannulation, and biomaterial on the explanted pump support the data log review. the returned product also had a large amount of biomaterial in the outflow cage and wrapped around the impeller. based on data logs and returned product, the root cause of the saturated flow was determined to be due to biomaterial, which was likely a result of the long bag/cassette change. clinical details reported that the patient's systemic heparin was put on pause due to bleeding from the mouth. no further details were provided. data logs are not relevant to the complication. the returned product was visually inspected and there were no physical damaged or abnormalities noted. due to limited clinical details, the root cause of the vascular damage could not be determined. the above investigation into saturated flows concluded that the cause was biomaterial. therefore, it is representative of the investigation into thrombus as well. d.9 date device returned/information was added. b.1 revised as product problem was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26006.
11792 | 1220648-2025-26007 | 21312904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312904&pc=OZD | 01/17/2025 | 03/18/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025595623 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) post aortic valve replacement was implanted with an impella 5.5 device for mechanical circulatory support. approximately two days after start of the support, a slight hematoma was noted on surgical access site to right axillary. additionally, it was noted the patient had a decrease in hemoglobin along with hematoma above surgical site and to left groin from previous heart catheterization. the patient was administered one unit of red blood cells with a pressure dressing to the surgical site. the patient survived the event, and the pump was explanted after just under seven days of total support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access siter bleeding has been completed. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26007 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11793 | 1220648-2025-26008 | 21312912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312912&pc=PYX | 01/18/2025 | 03/10/2025 |  | 02/05/2025 |  |  | 01/18/2025 | 02/05/2025 |  |  | 12/31/2025 |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025453771 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during attempted delivery of an impella rp flex pump for mechanical circulatory support. after successfully wiring the right ventricle, the impella rp flex was backloaded for implant. however, upon attempting to advance the device, the pump was unable to fully pass through the sheath. inspection of the introducer noted a kink in the mid section of the sheath prompting sheath removal. a second sheath was placed, but the same complication was encountered and the sheath once again kinked. at this point in time, a bleed was observed and it was discovered that the introducer had perforated the subclavian vein. the perforation was patched and several units of packed red blood cells, fresh frozen plasma, and cryo were given to treat the condition. the implant was ultimately aborted and the patient was hemodynamically stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The investigation for the difficulty insertion and vascular damage has been completed. no product was returned and logs are not applicable. three separate introducers were used in three attempts to advance the pump into the patient. all three attempts failed with a noted sheath kink. the patient's arm positions were manipulated as an attempt to resolve the sheath kinks, but it was unsuccessful. the root cause of the difficulty inserting the pump through the introducers was most likely an introducer sheath kink as kinks were observed in the field. the root cause of the vascular damage was most likely use-related since the 23fr sheath had perforated the subclavian vein upon reinsertion.
11794 | 1220648-2025-26009 | 21312914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312914&pc=OZD | 01/19/2025 | 04/11/2025 |  | 02/05/2025 |  |  | 01/19/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the alternating current (ac) power cable ripped off during unknown age patient support. recommendations were made to replace the automated impella controller; however, the staff declined with hopes the pump would be weaned prior to the battery losing power. upon follow up with the site, it was reported another cable was found to use for ac power. there was no patient harm or interruption of support. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - cabling issue has been completed. the alleged broken power cord was confirmed by pictures sent from the clinical team. logs were not informative for this complaint. the console was eventually returned with the broken power cable. one of the screws in the connector was found to be unscrewed. quality engineering was alerted of this discovery. the root cause of the issue was a defective cable with the untightened screw allowing the wires to pull out when yanked with force. d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26009 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-26009.
11795 | 1220648-2025-26010 | 21312925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21312925&pc=PYX | 01/18/2025 | 03/03/2025 |  | 02/05/2025 |  |  | 01/20/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560802 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella rp flex support and while on support lactate dehydrogenase was rising, and the patient was not making any urine. the patient has a history of kidney disease but not on dialysis. the patient was weaned downed to p3 and is now on dialysis. it is unknown if the hemolysis and renal failure is attributed to the impella. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of hemolysis and renal failure has been completed. patient updates revealed labs hemolyzing. the returned pump was inspected and biomaterial was observed in the inflow cage and around the impeller. the pump was hemolysis tested and passed. the data logs from the case were reviewed and there was a motor current spike with speed deviation and placement signal shift also with a drop in pump flow. this is characteristic of biomaterial ingestion. the root cause of the hemolysis was most likely due to biomaterial. the root cause of renal failure was not determined since the product was not returned. d.6b revised as it was not inadvertently omitted on initial manufacturer device report 1220648-2025-26010. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26010 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11796 | 1220648-2025-26011 | 21313066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21313066&pc=OZD | 01/09/2025 | 03/20/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611418 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs). one day after start of support, intestinal bleeding was confirmed, and anticoagulant therapy, including the heparin used in the impella purging solution, was discontinued. the purging solution was changed to sodium bicarbonate. further examination revealed that the patient had colon cancer, which was causing intestinal bleeding, and coil embolization was used to successfully stop the bleeding. there was still concern about the patient continuing to bleed afterwards; therefore, anticoagulant therapy was not performed. heparin was not administered. the bleeding would eventually stop (date was not provided). transfusion products were used. it is unclear how many units were transfused; however, more than two units of both red blood cells and fresh frozen plasma were transfused. it was stated although impella was not the direct cause, it is possible that the bleeding from the colon cancer was promoted by anticoagulant therapy using the impella, which prolonged the activated clotting time act and activated partial thromboplastin time. impella mcs continues, and the patient's condition was noted as serious. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. one day after device insertion, the patient experienced intestinal bleeding due to colon cancer combined with anti-coagulation therapy. no data logs were recovered. the cause of the vascular damage - peripheral vessel was not determined since the product was not returned and lack of clinical details.
11797 | 1220648-2025-26012 | 21313118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21313118&pc=OZD | 01/17/2025 | 02/28/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594332 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella 5.5 device to a patient (unknown race and ethnicity) as elective placement for mechanical circulatory support after planned coronary artery bypass graft surgery was unsuccessful due to the cardiac anatomy. the device was unable to pass through the axillary artery. the decision was made to abort the impella 5.5 and place an impella cp, which was successfully completed. the patient was noted to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to resistance in axillary artery. corrections to the initial mfr report: added d6a/d6b h6 impact code 4629.
11798 | 1220648-2025-26013 | 21313421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21313421&pc=PYX | 01/17/2025 | 03/05/2025 |  | 02/05/2025 |  |  | 01/17/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025599072 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | SWAN GANZ CATHETER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient post cardiothoracic surgery was implanted with an impella rp flex device for mechanical circulatory support. after successful implantation via the right internal jugular, the physician and elected to place a dialysis catheter in the left internal jugular. during access utilizing the micro puncture kit, the .018 micro catheter guidewire was ingested into the impeller of the pump and subsequently stopped the pump. the pump was able to be restarted with flow rate back to performance level p-9. however, upon further examination under fluoroscopy, it was noted that the micro puncture guidewire was entrapped in the impeller. the decision was made to remove and replace the impella rp flex. the pump was removed utilizing the clap technique to re-access the vessel. the micro puncture guidewire remained intact with the impeller and was removed with successful reinsertion of new pump. the patient was noted to be stable following the event. | [5 empty] | The investigation for the pump stop has been completed. no product was returned. log review showed two motor current spikes leading up to a final motor current spike and pump stop. there were multiple attempts to restart before the pump began running again for a short time before removal. the root cause of the pump stop was most likely interaction with another device as a wire became caught in the impeller during pump use. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
11799 | 1220648-2025-26014 | 21313516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21313516&pc=OZD | 01/20/2025 | 03/05/2025 |  | 02/05/2025 |  |  | 01/20/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537981 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while removing the peel away sheath, the impella cp pump was inadvertently pulled back into the aorta. attempts to recross the valve were unsuccessful and the pump was ultimately replaced. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. impella cp was not stabilized well enough and with the removal of the peel away, the pump was pulled into the aorta. attempts to correct were unsuccessful. the pump was then explanted. the cause of the positioning issue most likely was use related since the issue occurred during removal of the peel away sheath. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
11800 | 1220648-2025-26015 | 21313788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21313788&pc=OZD | 01/19/2025 | 02/25/2025 |  | 02/05/2025 |  |  | 01/19/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597759 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and it was aborted due to patient anatomy. the patient began to rapidly decline and decision was made by the doctor to abort the 5.5 and place impella cp. there was no reported harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. according to the clinical details the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy.
11801 | 1220648-2025-26016 | 21313843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21313843&pc=OZD | 01/20/2025 | 02/28/2025 |  | 02/05/2025 |  |  | 01/20/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577233 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient being evaluated and planned for transplant. the patient was on impella cp support for nearly 12 days when the pump stopped. the patient was stable while on extracorporeal membrane oxygenation support; multiple troubleshooting attempts were made. the patient was then brought to cath lab for removal and replacement of a new impella cp. | [5 empty] | The investigation for the pump stop has been completed. no product was returned. there is a gap in the data logs from 7:39 on 1/20 to 10:59am on 1/20. there is no pump stop in the first data log shown and in the second data log, the pump is attempted to be turned on but keeps getting alarm 13. the purge pressure was also over 1000mmhg, and it looks like the pump is kept in for almost 3 hours without it running. due to insufficient data logs and lack of product returned, the root cause of the pump stop cannot be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11802 | 1220648-2025-26017 | 21315579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21315579&pc=OZD | 01/20/2025 | 03/10/2025 |  | 02/05/2025 |  |  | 01/20/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487638 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that positioning issues were encountered during impella 5.5 support. it was documented that the device had to be repositioned multiple times with noted difficulty when trying to torque the tip towards the apex. flows began to decrease and the pump inadvertently came back across the valve. attempts to re-cross the valve were unsuccessful and the decision was made to explant the pump. there was no patient harm. | [5 empty] | The investigation of positioning issues and low pump flow has been completed. the product and data were not returned for review. based on clinical description, the root cause of the positioning issue was most likely use related due to improper technique as the pump was pulled out during support. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. h.6 code 4628 was inadvertently omitted from the initial manufacturer device report number and was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11803 | 1220648-2025-26018 | 21315850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21315850&pc=OZD | 01/21/2025 | 02/12/2025 |  | 02/05/2025 |  |  | 01/21/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1175803 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | PHYSICIAN ASSISTANT | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the automated impella controller (aic) was returned after the annual service. when unpacking the console, the medical technician noticed a rattling noise inside the console, which indicated a defective part in the housing of the aic. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - damage before therapy has been completed. the console logs did not contain any data relevant to this complaint. the reported rattling noise in the console was confirmed. further inspection of the console revealed that there was solidified glucose ingress built up inside the console. the root cause of the aic issue was contamination.
11804 | 1220648-2025-26019 | 21317944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21317944&pc=OZD | 01/21/2025 | 12/11/2025 |  | 02/05/2025 |  |  | 01/21/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1389732 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that an automated impella controller alarmed with a "battery temperature high" alarm during impella support. the aic was replaced with no further issues noted. there was no patient harm. | [5 empty] | D4 serial number was corrected. | H1: changed from a malfunction to a serious injury. h6: added code. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) battery overheating issue has been completed. console logs from the complaint date revealed the ¿battery temperature high¿ alarm (#49, due to battery temperature over 60degc) that self-resolved in 10 seconds. there was no preceding battery temperature alarm #48, which occurs when battery temperature is between 50degc and 60degc. battery and power data embedded in ltlogs do not show any temperature rise or elevated battery temperatures, as ltlog data is recorded every 60s. the console¿s batteries and power management system were returned and tested, and no issues were observed. the battery level low alarm was due to a momentary communication glitch between the batteries and pbm, where a single sample from battery data (in this case value of battery temperature) was corrupt.
11805 | 1220648-2025-26020 | 21317949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21317949&pc=OZD | 01/21/2025 | 03/14/2025 |  | 02/05/2025 |  |  | 01/21/2025 | 02/05/2025 |  |  |  |  | 02/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573059 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient had an attempted impella 5.5 implant and during the implant, impella plug to controller cracked due to table movement. staff immediate switch to 2nd impella. there was no harm to the patient. | [5 empty] | The investigation of delivery issues has been completed. the root cause of the delivery issue was most likely use related since they accidentally cracked the impella plug to the automated impella controller. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12650 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12650. h.10 instructions for use were added as they were inadvertently omitted from the initial manufacturer device report number 1220648-2024-12650. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11806 | 1220648-2025-25966 | 21302357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21302357&pc=OZD | 01/15/2025 | 03/03/2025 |  | 02/04/2025 |  |  | 01/15/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597000 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | Hemolysis | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while the pump was being placed into left ventricle (lv), the pump appeared to kink or bend inside heart due to small lv anatomy. as the pump was being withdrawn to straighten out, it popped out of the lv into the aorta. attempts were made to try and place it back in lv without success including using a pigtail to pin leaflet of valve open. however, because of the initial kink/bend in the catheter, the pump would not advance back into lv. the patient was hypotensive at the time and the decision was made to try a new pump due to patient instability. a new pump was successfully placed and patient was supported. | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and the next day a positioning issue was encountered followed by hemolysis. a patient was admitted in as a transfer from the community hospital and the impella cp device was placed as care escalation from an intra-aortic balloon pump. the left ventricle was noted to be small. the pump kinked and came out of position. the following day there were signs of hemolysis. dark red urine was noted after foley placement. central venous pressure was eight, and after rechecking, it was down to six. a 500 milliliter bolus was added. the performance level was reduced to p-7 to reduce hemolysis risk. the echocardiogram position of the impella cp was checked, and it appeared to be resting on the mitral apparatus. multiple attempts were made to adjust positioning with no success at placing the pump mid ventricle. the impella cp was weaned further to performance level p-5 to help with hemolysis. two days later, the impella cp device was weaned and explanted successfully. the patient was reported to be stable following the event. |  |  |  |  | The investigation has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. based on clinical description, the root cause of positioning issue was most likely patient condition related. the root cause of kinked cannula was most likely patient condition related. | B.5 event description was updated since initial manufacturer device report number 1220648-2025-25966 was submitted no longer include hemolysis as a failure mode. the complaint associated with impella cp pump with serial number (b)(6) includes hypotension. d. 4 revised serial number, lot number and unique identifier (udi) # since initial manufacturer device report number 1220648-2025-25966 was submitted. h. 4 revised device manufacture date since initial manufacturer device report number 1220648-2025-25966 was submitted due to change in serial number. h.6 health effect - clinical code was changed from 1886 reported on the initial manufacturer device report number 1220648-2025-25966 to 1914. health effect - impact code was changed from 4641 reported on the initial manufacturer device report number 1220648-2025-25966 to 4629. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11807 | 1220648-2025-25967 | 21302430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21302430&pc=OZD | 12/05/2024 | 03/07/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1215713 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | During impella cp mechanical circulatory support for a patient undergoing high-risk percutaneous coronary intervention, the user facility reported placement signal unreliable alarms were encountered with the pump. the audio was disabled and support continued without interruption. abiomed's investigation engineer requested return of the automated impella controller (aic) to further investigate if this issue could have been a result of the aic. there was no report or indication of patient harm. after five days of support, the impella cp was successfully weaned and explanted with the patient outcome as survived, awake and alert. | [5 empty] | The device was received for investigation and it was determined that it is not a reportable malfunction and a regulatory report is not necessary. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11808 | 1220648-2025-25968 | 21302487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21302487&pc=OZD | 01/16/2025 | 02/25/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573260 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed limb ischemia. after completion of the procedure, the assessment of right lower extremity (rle) was mottled and absent of distal pulses. recommendation was to pull back on repositioning sheath which was completed. however, this did not result in improvement in leg ischemia. slow impella wean per physician orders resulted in a hemodynamically stable patient. the impella cp was turned off, explanted and within minutes, the rle mottling improved and warmed up with return of distal pulses. the patient remained hemodynamically stable and was taken to the unit for further management. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined.
11809 | 1220648-2025-25969 | 21302574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21302574&pc=OZD | 01/16/2025 | 03/05/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579131 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced a pump stop at the beginning of support. the patient had non-st-segment elevation myocardial infarction with elevated troponin while inpatient at a prior facility. cardiac work up showed multi-vessel coronary artery disease and severe aortic stenosis. the patient was transferred in with plans for impella assisted pci and transcatheter aortic valve replacement (tavr). for the impella cp implant, a 14 french long impella sheath was inserted but kinked due to tortuous anatomy. therefore, a 16 french cook sheath was then inserted without some deformation. balloon aortic valvuloplasty was performed. the impella was inserted through a cook sheath with light resistance and crossed into the left ventricle. shortly after the pump started, it stopped with a motor current high alarm. several more start attempts were made with controller failure and motor current high alarms. the decision was made to move to tavr procedure. a new sheath was placed, but the physician was unable to deliver the tavr through the sheath. therefore, the decision was made to move the procedure to the contralateral femoral artery. the tavr was successfully delivered with the decision to do minimal pci without impella mcs. there were no adverse effects to the patient. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation has been completed. the device was inserted but insertion sheaths experienced kinking due to tortuous anatomy. the data logs were reviewed and it was shown that the pump was not started as the motor current spiked immediately upon starting with no pump flow was achieved. each spike corresponded with a motor current high alarm, even though the purge was connected and flowing with adequate pressure. the pump was returned for investigation. a visual inspection found a catheter kink at the 46 centimeter mark, with no other abnormalities found. it was then run on p-9 for 10 minutes where no adverse events were reported and the pump stop/could not start could not be reproduced. the cause of the pump did not start was not determined since the failure mode was unable to be reproduced. 14 french insertion sheaths experienced kinking due to tortuous anatomy and a 16 french cook sheath had to be used instead. no sheaths were returned. the cause of the introducer damage - mechanical was most likely patient condition related. h.6 codes were added to medical device problem code and component code as they were inadvertently omitted from initial manufacturer device report number 1220648-2025-25969.
11810 | 1220648-2025-25970 | 21302677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21302677&pc=OZD | 01/16/2025 | 02/28/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533843 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported attempts to implant an impella cp via left axillary graft were unsuccessful. due to noted resistance from the patient's anatomy, the decision was made to abort the implant. the patient was on venoarterial extracorporeal membrane oxygenation (va ecmo) and impella cp mechanical circulatory support prior to the event. the patient was taken to the operating room for emergent ecmo reconfiguration with an attempt to wean off va ecmo with full impella cp support, but the patient did not tolerate it well. the previous impella cp device was explanted, and an arterial return cannula was placed. the new impella cp device to left axillary graft (old arterial return ecmo site) was attempted but was unsuccessful due to anatomy. the case was therefore aborted. there was no report or indication of harm to the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11811 | 1220648-2025-25971 | 21302985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21302985&pc=OZD | 12/29/2024 | 03/17/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564635 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported in preparation of an impella 5.5 device implant for mechanical circulatory support to a patient with cardiomyopathy, it was noted the graft insertion kit had poor water passage and, therefore, was replaced. there was no harm to the patient. | [5 empty] | The investigation of introducer damage ¿ mechanical has been completed. the product was not returned for investigation, and a data log review was not required for this case. according to the clinical notes, a defect was observed in the short 23fr introducer due to poor water passage. additionally, no clinical image was provided for the reported issue. the cause of the introducer damage issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11812 | 21303109 | 21303109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303109&pc=OZD | 09/24/2024 | 01/15/2025 | 01/15/2025 | 02/04/2025 | 01/15/2025 | 02/04/2025 |  | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [8 empty] | US | [5 empty] | No answer provided |  |  | Not Available | [6 empty] | User Facility report | Yes | Nurse | [6 empty] | NA |  | Unknown |  |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Impella 5.5 leaking solution when added.
11813 | 1220648-2025-25972 | 21303164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303164&pc=OZD | 01/09/2025 | 03/20/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2025564965 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed via the 14 french sheath. the vessel was damaged at the placement of the sheath, but the damage was unseen for hours til in the intensive care unit. the damaged vessel was treated by a covered stent. the procedural outcome was the patient survived the surgery. additionally, the team gave blood products of at least two units. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
11814 | 1220648-2025-25973 | 21303191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303191&pc=OZD | 01/16/2025 | 03/14/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557333 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. during support, the tuohy-borst kept staying open. the staff repeatedly attempted to close it unsuccessfully. additionally, the patient's foot became caught on the sterile sleeve of the pump and tore it away from the tuohy-borst. the pump was explanted from the patient and discarded. no patient harm was reported, and the patient was reported to be stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device is not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of product damage (sterile sleeve damage) and product damage (bond failure) has been completed. the pump was not returned for investigation, and a data log review was deemed unnecessary for this case. according to the clinical report, the tuohy borst was found open several times during the pump use. additionally, the sterile sleeve was torn away from the tuohy borst after the patient caught her foot in it. no clinical image was provided for the sterile sleeve damage. the product damage (bond failure) failure mode applies to the loose bond on the kink protector bond (catheter, patient cable, or purge lumen), which allows for unimpeded translation and rotation. the failure mode of "product damage (bond failure)" has been removed, as the loose tuohy borst employs a snap-fit mechanism and is part of the sterile sleeve. this issue has been addressed under the failure mode "product damage (sterile sleeve damage)". the cause of the sterile sleeve damage issue was most likely use-related, as the patient accidentally tore the sterile sleeve when his/her foot became caught in it.
11815 | 1220648-2025-25974 | 21303306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303306&pc=OZD | 01/17/2025 | 03/19/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548368 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Death | Yes | Yes | Infusion or Flow Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, there was a pressure saturation low alarm. instructions were given to optimize position and parallel during call could see saturated flow and increasing mc >1050ma. the doctor was advised to escalate to impella 5.5 support due to a risk of a pump stop. the patient subsequently passed away on support the following day. it is unknown if the complication contributed to the patient's passing. | [5 empty] | The investigation of saturated flow has been completed. the product was not returned for evaluation. data logs were reviewed and the logs of full case show at time of ¿placement signal low¿ alarms, motor current is variable at time of alarms with multiple spikes in placement signal. at beginning of case, saturated flow is observed with a max flow of approximately 4.3 which is higher than the set p-level range (p-9) at that time. real time (rt) log at time of issue shows multiple spikes in placement signal with ¿impella in aorta¿ and ¿impella in ventricle¿ alarms prior to ¿placement signal low¿ alarms. at beginning of alarms occurring, pump flows became low with pump running at around p-3. two large motor current spikes with speed deviations are observed at time with low pump flow. low pump flows seems to resolve slightly and once flows are within a normal range for the p-level, the flows become fully saturated for a few instances with no pulsatility. clinical details noted that saturated flows were observed and an increase in motor current (mc) to greater than 1050ma was seen. pump was repositioned for optimal positioning. the root cause of the saturated flow was most likely due to biomaterial. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11816 | 1220648-2025-25975 | 21303321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303321&pc=PYX | 01/16/2025 | 10/07/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025585358 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient post cardiothoracic surgery was implanted with an impella rp flex for mechanical circulatory support (mcs) and experienced a pump stop that resulted in explant with replacement of the device. approximately six days after start of support, it was noted the impella rp flex placement signals waveforms changed and flow at support level p-8 went from 3.0l/min to 0. reportedly, this issue immediately happened after the swan catheter was repositioned. the patient remained stable and hemodynamics were unchanged. the intensivist had the team switch automated impella controller (aic) and when the rp flex was plugged into the other aic, the pump immediately alarmed impella defective. the patient was transported to the catheterization lab to replace the impella rp flex which was successfully completed. the rp flex was removed over a guidewire, and a new rp flex was placed via the right internal jugular without difficulty. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-25975 b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The investigation for the pump stop has been completed. clinical details stated there was increased placement signal values and decreased flow values. additionally, there was an impella defective alarm when pump was plugged into new console. no product was returned for analysis. the data logs show a sudden placement signal shift before placement signal started drifting upwards with flow drifting downwards. clinical mention that this happened immediately after swan repositioning. there was alarm #119 pump stop right before explant. the data logs show impella defective alarm when pump was plugged into a different console. the imc logs of the second console were unable to be retrieved. the root cause of the placement signal issue was most likely due to sensor drift. the root cause of the pump not detected likely from the pump stop was not determined as no product was returned and the imc logs for replacement console were not provided.
11817 | 1220648-2025-25976 | 21303382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303382&pc=OZD | 01/16/2025 | 02/28/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594754 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during the implant procedure, the introducer was inserted into the graft and then the physician placed both graft locks. however, there was blood coming around the top of the introducer. a new introducer was taken from another impella 5.5 box and the implant continued without further issues. there was no patient harm. | [5 empty] | The investigation for the introducer damage has been completed. product was not returned for investigation. a data log review was not required for this case. according to the clinical report, bleeding occurred from top the introducer. the issue was resolved after changing the introducer. clinical details state that the bleeding resulted from the damaged introducer. the cause of the access site bleeding issue was not determined since introducer was not returned for investigation. | The impella device was not received from the customer as the device remains implanted supporting the patient at the time of this report being submitted and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
11818 | 1220648-2025-25977 | 21303488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303488&pc=OZD | 01/16/2025 | 03/03/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025596384 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella 5.5 device for circulatory support and experienced increased bleeding from a right subclavian nick. a right axillary pocket cut down was performed with some difficulty. increased bleeding occurred from a nick of the right subclavian. an eight millimeter hemashield was anastomosed to the right axillary artery with a 60-degree beveled edge, tunneled one inch by one inch. the impella 5.5 was backloaded over an impella .018 guidewire into the left ventricle. the guidewire was pulled, and the pump torqued into position. performance level was slowly ramped up to p-9 with good flows. the pocket was closed after full reversal with protamine. the patient was transferred to the unit on support with upcoming plans for pci. four days later, a slight kink in the catheter closet to the lock under the sleeve was noted. there was no change in purge flow or issues with placement signal. support was continued with the protected pci being completed the following day. the patient remained hemodynamically stable. the impella was successfully weaned and explanted without incident. | [5 empty] | The investigation for the vascular damage and the product damage has been completed. pump was not returned for investigation. data log review was not required for this case. as per the clinical, the right axillary pocket cut down was performed with some difficulty, and increased bleeding was observed due to a nick in the right subclavian artery. 8 mm hemashield graft was placed and bleeding was resolved after closing the pocket. there was no issue related to kinking of the device. the cause of the vascular damage issue was most likely use related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11819 | 1220648-2025-25978 | 21303563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303563&pc=PYX | 01/17/2025 | 02/27/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  | 01/31/2026 |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025462199 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support. the patient was already on continuous renal replacement therapy (crrt). however , during support the patient's lactate dehydrogenase levels were rising. therefore, the impella rp flex was explanted. the patient outcome is unknown. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
11820 | 1220648-2025-25979 | 21303665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303665&pc=OZD | 12/16/2024 | 02/10/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568013 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | Malposition of Device |  |  |  |  | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race) with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and during support positioning issue and high purge pressure were encountered. the patient would eventually expire after six days of support. the patient with an extensive medical history was deemed heart failure with reduced ejection fraction of 20%. an echocardiogram showed the left ventricle severely dilated, global hypokinesis, and small mural mass. left ventricular end-diastolic diameter was 6.6cm. the patient with ischemic cardiomyopathy complicated by cardiogenic shock and inotrope dependent with class iv heart failure. the patient was planned for an impella 5.5 device as a bridge to heart transplant or bridge to left ventricular assist device. the impella 5.5 device was implanted, and poor positioning and high purge pressure alarm were noted to have occurred approximately six days after start of support. there were also persistent intermittent suction alarms and motor current was flat, and aorta/left ventricle decoupled. the patient was sliding and in more shock. the patient was taken for an impella 5.5 device exchanged and venoarterial extracorporeal membrane oxygenation (va ecmo) placement. however, the patient continued doing poorly, and support was removed the same day of the new 5.5 insertion. the cause of death is unknown. there is not enough evidence to exclude the impella as a possible contributing factor in the overall outcome. however, it should be noted the patient had a "poor" heart function, less options available (not a candidate for percutaneous coronary intervention or coronary artery bypass graft surgery), and did not fully recover from the cardiac event. | [5 empty] | The investigation of positioning issues and high purge pressure has been completed. the failure mode "positioning issues" was changed to "low pump flows" because low pump flow was the higher severity issue of the two resulting pump performance issues of the same root cause. the pump was returned and inspected. there was biomaterial on the impeller and in the outflow cage. the pump was soaked and the biomaterial was cleared. the pump was run in the lab and the low pump flow did not reproduce. the data logs showed a motor current spike corresponding with a speed deviation, suction alarms, and a shift in placement signal. the logs showed a shift in purge pressure at the time of the spike before the purge pressure values are elevated later on and high purge pressure alarm is triggered. the high purge pressure resolved after about 40 minutes. the positioning issues reported by clinical are most likely due to biomaterial impacting optical sensor and causing false impella position in aorta alarms. the suction alarms are also a result of ingested biomaterial. the root cause of the low pump flow was determined to be due to biomaterial. the root cause of the high purge pressure was determined to be biomaterial. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ h.6 a new medical device problem code was added to reflect failure mode changes noted in the investigation results. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2025-25979 in accordance with updated procedures. | The impella pump with the data stick and purge cassette has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11821 | 1220648-2025-25980 | 21303718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303718&pc=OZD | 10/03/2024 | 05/13/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1663431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | During impella 5.5 mechanical circulatory support to a patient with acute myocardial infarction/cardiogenic shock, there were decreased flows, abnormal waveform, and suspected clot with the pump per the physician, and the abiomed's investigation engineer has now requested return of the automated impella controller to further evaluate instances of low signal to noise ratio. | [5 empty] | The investigation of automated impella controller (aic) - loss of opm data has been completed. the data analysis showed there was an alleged abnormal waveform (placement signal) that was confirmed in the lt logs with periods of placement signal not reliable and raw optical sensor dropping to -3276.8 mmhg the console was returned for investigation and was investigated and was found to have failing 5s parameters with low contrast and a marginally passing signal not reliable (snr). the analog outputs from the ccd array showed no large abnormalities though there was a sharper slope on the left side compared to the right. root cause: the cause of the optical signal issue was most likely the optical bench. d9 date device returned to manufacturer added. | The automated impella controller has been requested for return but has not been received from the customer at the time of this mdr writing. therefore, an evaluation of the device was not possible at this time. upon investigation closure, a supplemental mdr will be filed. the medical device report for the user facility's complaint on the pump is captured under manufacturing report number 1220648-2024-20971. | The cause of the issue was not determined since product was not returned. the cause of the issue was not determined since product was not returned. b3 date of event revised to reflect correct event date on manufacturer device report number 1220648-2025-25980 e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25980 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11822 | 1220648-2025-25981 | 21303723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303723&pc=OZD | 01/16/2025 | 02/18/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585631 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an impella cp device was implanted to support a patient admitted in for a high-risk percutaneous coronary intervention (pci). it was noted the patient experienced significant hemodynamic drifts throughout the procedure; the protected pci of the left anterior descending artery was successful. however, due to body habitus, hemostasis was not achieved, and the access site continued to bleed. the decision was made to obtain a retrograde access and place a covered stent to control the bleeding. the patient was noted to have survived and reported to be in critical condition following the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely use issue since both the perclose failed. e4 should have been left blank on manufacturer device report 1220648-2025-25981. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11823 | 1220648-2025-25982 | 21303726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303726&pc=OZD | 01/16/2025 | 02/26/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593229 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 95 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an attempt to implant an impella cp device pump for mechanical circulatory support in a patient experiencing acute myocardial infarction/cardiogenic shock was aborted due to the patient's anatomy. the patient had a heavily calcified aorta. this pump became damaged when attempting to move pass aortic arch. the guidewire was bent inside the pump not allowing for removal; therefore, the pump and guidewire were removed as a unit and discarded. the was reported to be in stable condition following the event. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, patient with heavily calcified aorta and pump became damaged when attempting to move pass aortic arch. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ corrections to the initial mfr report: added d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11824 | 1220648-2025-25983 | 21303731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303731&pc=OZD | 01/16/2025 | 03/03/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594956 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device and experienced pump positioning issues. two days after start of support, the pump was repeatedly found malpositioned at the mitral apparatus. the physician attempted to re-adjust the pump position as the patient was suffering from signs of hemolysis. the following day, the patient was weaned to support level p-3 in the unit and taken to catheterization lab for removal without incident. the patient was reported to be stable following the event. | [5 empty] | The investigation of positioning issues has been completed. no product was returned. clinical information notes the pump appeared to be caught in the mitral apparatus and was repositioned a little bit. the cause of the positioning issue was not determined since the product was not returned and insufficient clinical details were provided. b.1 revised to reflect product problem only as adverse event was reported incorrectly on the initial manufacturer device report number 1220648-2025-25983. b.2 revised to reflect as blank for all selections as adverse event with required intervention was reported incorrectly on the initial manufacturer device report number 1220648-2025-25983. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25983 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.1 revised to reflect malfunction as serious injury was reported incorrectly on the initial manufacturer device report number 1220648-2025-25983. h.6 codes 1886, 4629 and 4627 were reported incorrectly on initial manufacturer device report number 1220648-2025-25983. new codes were added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11825 | 1220648-2025-25984 | 21303732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303732&pc=OZD | 01/16/2025 | 03/03/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1325609 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported when automated impella controller (aic) was turned on, a message was displayed indicating that the controller was faulty. aic was replaced and there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. on the complaint date, immediately after the initial console boot-up, an "optical sensor lamp defective - controller error" occurred, confirming the reported failure mode. the console was then shut down and restarted twice, and the event reappeared each time. analysis of the long term log and real time log data showed that, during the event, contrast values fluctuated between -3000 and 3000, and the signal-to-noise ratio was near zero. this indicated that light was not reaching the optical bench detector. this console was tested. they were unable to reproduce the reported failure by power cycling the console. voltage from the power battery manager at connector (j3) was 4.9 volt (v). the voltage on the optical bench lamp connector (p4) was 4.8v, and the lamp resistance was 5.1 ohm as well as the 5s testing showed the light output passing the required specification and the lamp was visibly outputting light. the root cause for the controller error (defective optical sensor lamp) was most likely due to the defective optical bench or the lamp assembly. d.9 revised date as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25984. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25984 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact fax number as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25984.
11826 | 1220648-2025-25985 | 21303847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21303847&pc=OZD | 01/16/2025 | 02/25/2025 |  | 02/04/2025 |  |  | 01/16/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558883 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during an attempt to place an impella cp device for mechanical circulatory support to a patient undergoing high-risk percutaneous coronary intervention (pci), the physician was unable to advance the motor into distal aorta. impella was removed and a 7mm balloon angioplasty was performed. the impella cp device was reattempted post percutaneous transluminal coronary angioplasty without success. an 8mm shockwave of right iliac was performed, and the impella cp device was attempted again without the ability to fully advance the pump into aorta. consequently, the impella cp device implant was aborted. access was maintained for percutaneous coronary intervention with a companion sheath and successful pci of left main and left anterior descending arteries followed. the patient showed signs of brief hemodynamic instability; however, the patient recovered quickly without bailout intervention. there was no report or indication of patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, pump inserted but unable to advance motor into distal aorta.
11827 | 1220648-2025-25986 | 21304087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21304087&pc=OZD | 01/17/2025 | 02/12/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1271681 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while attempting a double-barrel impella cp to impella 5.5 escalation, the automated impella controller (aic) powered on with a red battery failure alarm. the aic remained plugged in for 30 minutes, but the alarm remained with a 0 percent battery charge. the aic was subsequently replaced for the planned support. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) battery failure-red alarm has been completed. console logs from the reported complaint date were consistent with what was reported in the complaint. there were two instances of battery failure (transport connection blown/missing) alarms [event set #360] observed in the imc logs. console was returned for evaluation finding the reported battery failure alarm was reproduced while testing. batttest results revealed that the cell stacks in both batteries had voltages of less than 2.5v which resulted in battery lockout. the battery failure alarm was resolved after testing this console with two known good batteries. the battery failure issue was determined to be caused by depleted batteries resulting in battery lockout.
11828 | 1220648-2025-25987 | 21304276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21304276&pc=OZD | 01/16/2025 | 03/06/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485286 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional details stated the patient is bleeding at impella site and discussed redressing and angle matching. | The complainant reported the patient was on impella cp support. while on support, labs were hemolyzing and urine output decreased and coca-cola in color. patient¿s systemic vascular resistance was high, and pump was on auto. recommend repositioning to eliminate potential unseen issue within the heart however it is unclear if repositioning took place. the procedural outcome was the patient survived surgery. |  |  |  |  | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the data log was reviewed: quick resolved position alarms in ventricle and aorta observed. no mc spike outside of p-level changes observed. no suctions seen. no placement signal issue found. pump was repositioned. patient¿s systemic vascular resistance (svr) was high, and pump was on auto. unclear if the hemolysis issue resolved after repositioning or medical treatment during the case. the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. access site bleeding -major added since on 1/18: patient was bleeding at impella site and discussed redressing and angle matching without providing details. the cause of the access site bleeding -major issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. b2 revised to reflect additional information that was not included on initial manufacturer device report number 1220648-2025-25987. b5 revised to reflect additional information that was not included on initial manufacturer device report number 1220648-2025-25987. h6 revised to reflect additional information that was not included on initial manufacturer device report number 1220648-2025-25987. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11829 | 1220648-2025-25988 | 21304954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21304954&pc=OZD | 01/17/2025 | 03/07/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600546 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, left iliac angiogram was then performed via right femoral artery (rfa), showing dissection of left iliac artery. rfa access was exchanged for 45 centimeter destination sheath over guidewire. the doctor made multiple attempts at crossing dissection with various guidewires, but was unsuccessful. the decision was made to transfer the patient directly to the operating room for open vascular repair. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. due to lack of sufficient clinical information and no product returned, the root cause of the vascular damage - great vessel could not be determined. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2025-25988. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11830 | 1220648-2025-25989 | 21305012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21305012&pc=OZD | 01/17/2025 | 03/07/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597123 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support received one unit of packed red blood cells due to gastrointestinal bleed. heparin was put on hold. the patient remained stable and on support. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11831 | 1220648-2025-25990 | 21306475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21306475&pc=OZD | 01/17/2025 | 02/26/2025 |  | 02/04/2025 |  |  | 01/17/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589787 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed ischemia in their left lower extremity following impella cp implant. their leg became mottled with an absence of pulse. intervention was discussed, but the patient's family eventually made the decision to withdraw care and the patient passed away. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11832 | 1220648-2025-25991 | 21307114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21307114&pc=OZD | 01/05/2025 | 03/12/2025 |  | 02/04/2025 |  |  | 01/10/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587498 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp for mechanical circulatory support and escalated to an impella 5.5 four days later due to the expected prolongation of support. approximately after 11 days on impella 5.5 support, leaking from purge cassette was observed. per the clinical engineer liquid was leaking from the purge cassette on the fourth day after starting use. consequently, the purge set was replaced. however, 16 days after start of support, abnormal noise was heard around the purge cassette insertion area of the automated impella controller (used with both the impella cp and impella 5.5), which prompted replacement of the purge set. the abnormal noise was resolved after replacing the purge housing. at some point the patient developed sepsis as case wrap-up noted ¿cardiac function recovers poorly, and eventually sepsis disappears.¿ the timeframe for this adverse event and its relationship to the impella pump is unclear. however, the patient would eventually expire based on reflected case note ¿expired on support,¿ which occurred on the 25th day of support. reportedly the patient went into multiple organ failure caused by sepsis. antibiotics were administered. no detail information surrounding the patient's condition, the sepsis and death were provided, and with limited information, there is not enough information to exclude the impella device as a contributing factor in the patient¿s outcome. | [5 empty] | The investigation of infection/sepsis has been completed. not enough clinical information was provided to determine the root cause of this failure. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. this is one of two devices associated with the event. refer to medical device report for the automated impella controller serial number (b)(6) with date of event 05-jan-2025 and identifier ending in (b)(6). section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11833 | 1220648-2025-25992 | 21308653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21308653&pc=OZD | 01/19/2025 | 03/07/2025 |  | 02/04/2025 |  |  | 01/19/2025 | 02/04/2025 |  |  |  |  | 02/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582254 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing for a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was placed along with extracorporeal membrane oxygenation. the limbs were bilaterally ischemic with mottled appearance. the support remained despite this for a week. the pump was explanted after 7 days and at the time of explant the vascular team had to intervene for a retroperitoneal (rp) bleed. the rp bleed prompted vascular to surgically perform a common femoral artery. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of limb ischemia and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25992 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11834 | 1220648-2025-25958 | 21291019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21291019&pc=OZD | 12/31/2024 | 02/19/2025 |  | 02/03/2025 |  |  | 01/21/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564007 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | Renal Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry, the patient was on impella 5.5 support. while on support, the patient endured atrial arrhythmia and was treated with intravenous amiodarone. further the patient had a cerebrovascular accident. a cat scan on the head showed lacunar infarcts and cerebellar infarct. it was also reported that the patient endured renal failure and treated with creatinine however later the patient passed away before it was resolved. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of vf/vt/af/at, stroke/cva, and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11835 | 1220648-2025-25952 | 21291193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21291193&pc=OZD | 01/15/2025 | 02/20/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED INPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119392 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during impella mechanical circulatory support, the automated impella controller (aic) had a controller error (yellow alarm). the aic was replaced with no report or indication of interruption in support and no harm to the patient. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the automated impella controller (aic) controller error-yellow alarm has been completed. data logs were reviewed which stated the console displayed ¿controller error (loss of comm (pbm1)) ¿ alarm¿. this confirms the reported failure mode. continuous powermod_1 communication issues were recorded in the logs. the console was returned for evaluation and confirmed power battery manager (pbm) corrosion. it was found that te front panel optical connector was contaminated and they were unable to remove the debris upon cleaning. the root cause for the controller error and the pbm1 communication issue was pbm corrosion. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25952 and a new code was added.
11836 | 1220648-2025-25950 | 21291536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21291536&pc=OZD | 01/15/2025 | 03/06/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585510 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Fluid/Blood Leak | Difficult to Remove | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) for bridge to transplant and experienced removal difficulties. the impella insertion was successful without complication and patient on was support until bridged to transplant after approximately 16 days of mcs. the impella pump was cut out with native heart, and the catheter was left in the ascending aorta for post-transplant removal. blood was noted leaking at the red impella plug. clamping impella external catheter would not stop the blood from leaking. graft pulled out from cutdown slightly, the impella catheter was fully removed from locking mechanism, and graft was clamped. site closed post successful bridge to transplant. | [5 empty] | The impella pump has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for access site bleeding has been completed. the data log review was not required for this case. the pump was returned for investigation. the motor side catheter was cut at the kink protector red handle and was missing from the return package. the red handle was opened, and no blood was observed inside it. no further investigation could be performed on the returned product. according to the clinical report, the pump was turned off and left in place inside the patient, where bleeding was observed from the open lumen of the catheter. clamping did not resolve the bleeding. the catheter was fully removed and the site was closed. the reported failure mode of product damage (hole in catheter) was applicable for the blood leaking inside the red handle caused by the hole in the catheter line. however, in this case, no blood was observed inside the red handle. hence, this failure mode was removed and replaced by the accessible bleeding failure mode, since the open end of the catheter was leaking while the motor end was still inside the patient. the cause of the access site bleeding issue was most likely use related, based on returned product investigation and given clinal details. b.1 revised to reflect adverse event only, based on the investigation findings since the initial manufacturer device report number 1220648-2025-25950 was submitted. b.2 revised based on the investigation findings since the initial manufacturer device report number 1220648-2025-25950 was submitted. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25950 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.1 revised to serious injury based on the investigation findings since the initial manufacturer device report number 1220648-2025-25950 was submitted. h.6 revised codes to reflect adverse event only with intervention, based on the investigation findings since the initial manufacturer device report number 1220648-2025-25950 was submitted.
11837 | 1220648-2025-25948 | 21291602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21291602&pc=OZD | 01/15/2025 | 03/11/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274731 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the three-pronged plug came off from the electrical cable of the automated impella controller. there was no patient involved with the reported event. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - cabling issue has been completed. the data logs are not relevant to this complaint. the console was returned for investigation. upon arrival, the console underwent testing, which revealed an issue during visual inspection. internal cables within the plug were found to be disconnected. the issue was most likely the insufficient length of the wire covering, which was noticeably shorter than that of a properly functioning unit. the cables outer sheath/jacket length was inadequate to prevent pinching at the plug, which is why the cord broke loose and defective wire. the reported complaint was most likely due to an insufficient length of the wire outer sheath or covering.
11838 | 1220648-2025-25954 | 21291672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21291672&pc=OZD | 01/15/2025 | 03/05/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576324 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient came into the hospital with what appeared to be sepsis and additionally was with cardiogenic shock heart failure. the patient was taken to the catheterization lab for an impella cp device placement. upon connecting the impella purge line to impella cp, purge fluid was noted at the purge fluid reservoir area. there appeared to be a leak. reportedly, the pump was never implanted in the patient. a new impella cp device was obtained and implanted; the patient was stabilized and transferred to the intensive care unit on support. latest update noted the patient was on support for approximately six days was successfully weaned, and the device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak-pump has been completed. clinical information reported that the purge was leaking right out of the box. the pump was returned for investigation. upon pressurizing the purge line, fluid began to leak out of the filter housing, where further investigation revealed a crack along its length. no data logs were returned as the pump was not started before the leak was discovered in the field. the cause of the purge leak - pump was determined to be sidearm filter damage. the cause of the crack is unknown.
11839 | 1220648-2025-25960 | 21291829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21291829&pc=OZD | 01/15/2025 | 02/25/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580896 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient in cardiogenic shock from an acute myocardial infarction was attempted to be escalated to an impella 5.5 from an impella cp but was unable due to poor vasculature per the cardiothoracic surgeon. once patient was induced with anesthesia, the patient decompensated requiring inotropes and pressor. the physician performed cut down of right axillary artery and attached 10mm graft. the wire and diagnostic catheter were used to cross aortic valve. multiple attempts with wires and catheters were used to cross. once across the impella 5.5 back loaded over .018 guidewire. was unable to get impella 5.5 pass the innominate artery with multiple attempts. a .035 stiff buddy guidewire was attempted without success. due to extreme tortuosity, the physician aborted the right and went to the left axillary artery. left atria artery, cut down and 10 mm graft was attached. the same issues occurred on the left side with wire and impella 5.5. the physician attempted using lunderquist wire but was unable to get past the tortuosity. the physician decided to leave in the impella cp device and noted the limb ischemia was "better." the patient would eventually expire five days later. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two devices associated with the patient implant experience. refer to manufacturing report number 1220648-2025-25924 for the impella cp device events with patient's demise. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of delivery issue is determined to be patient condition because pump was unable to get past tortuosity.
11840 | 1220648-2025-25955 | 21291890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21291890&pc=OZD | 01/16/2025 | 03/05/2025 |  | 02/03/2025 |  |  | 01/16/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592248 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a fluid leak was discovered at the purge sidearm during impella 5.5 mechanical circulatory support to a patient with cardiomyopathy. it was noted that dextrose was building up in the purge reservoir as a result of the leak. both the purge flow and purge pressure remained unaffected by the leak. the following day the plan was for the patient to go for an orthotopic heart transplant, and the impella 5.5 device was explanted. the was no report or indication of patient harm. | [5 empty] | The investigation of purge leak ¿ pump has been completed. no product was returned for evaluation. clinical details noted that purge was leaking from purge sidearm and a build-up of dextrose was observed. additionally, no changes to purge pressure or purge flow were observed when leaking. the root cause of the purge leak - pump was most likely due to a damaged sidearm. however, the exact location of the leak could not be confirmed as no product was returned. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ d.6a/d.6b revised as dates were missing from initial manufacturer device report number 1220648-2025-25955. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25955 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2025-25955. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11841 | 1220648-2025-25949 | 21291928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21291928&pc=OZD | 01/11/2025 | 02/27/2025 |  | 02/03/2025 |  |  | 01/11/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534629 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient who was admitted in cardiogenic shock and acute myocardial infarction. the pump was placed but explanted after a day due to purge/red plug leak. the team had attempted first to change the purge cassette but the low purge remained and the pump was explanted and replaced by an intra-aortic balloon pump. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the data logs were reviewed and the purge leak was confirmed. the returned product was connected to a console and purged. once the pump was connected to the console, the impella defective alarm triggered. the red pump plug was opened and there were signs of blood and purge fluid inside. when the pump was actively purged, it was found that the source of the leak was from a hole in the catheter at the 70 cm marking. to confirm this was the only source of the leak, a syringe with colored fluid was connected to the sidearm and the fluid was pushed through. again, fluid only leaked from the location of the hole in the catheter. in the field, whatever broke the hole in the catheter most likely also punctured the purge lumen allowing for purge fluid to leak out into the catheter and back into the red pump plug. there is potential that during the time the antegrade sheath was being placed, the hole was punctured. with diminished purge pressure blood was also able to ingress back into the pump plug. based on returned product the root cause of the purge leak was determined to be damaged purge tubing. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h3 device not returned to manufacturer for evaluation was inadvertently selected "yes", however the device was returned for investigation at the time of initial manufacturer device report number.
11842 | 1220648-2025-25953 | 21292037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21292037&pc=OZD | 01/15/2025 | 02/27/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108315 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the biomed department had an impella automated controller (aic) repeatedly alarm for controller failures during a procedure. the team removed the aic from service. no further details were made available. the outcome of the patient was indicated as "unknown." however, there was no report or indication of interruption in support or patient harm. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console logs from the reported event showed three unexpected shutdowns, with the console¿s software automatically restarting after each shutdown. the first reboot took place after two hours of the console being powered on. this reboot followed a partial reset of the watchdog and a missed acknowledgment from the process osslopsens. the second reboot occurred after another partial reset of the watchdog and another missed acknowledgment from the process osslopsens. after the three system reboot, the "unexpected shutdown" alarm was triggered, and the watchdog restarted. the console was tested, but the issue could not be reproduced, and no unexpected shutdowns or performance irregularities were observed. additionally, the console was manually power-cycled 10 times without issue. no power supply interruptions were detected between the power battery manager and the 4-in-1 assembly. the most likely cause of the aic unexpected shutdown was a software anomaly. the root cause of the unexpected reboot was due to a software process failure. the cause of thesoftware process failure was not determined. a.3 revised as information was inadvertently omitted in the initial manufacturer device report number 1220648-2025-25953. d.9 revised device returned to manufacturer as selection was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25953.
11843 | 1220648-2025-25957 | 21292322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21292322&pc=OZD | 01/15/2025 | 02/27/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557322 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, some blood was noted to be leaking from the red impella plug of the 5.5. during implant, the doctor went back to fix the graft anastamosis after the 5.5 was already placed. it was assumed that they potentially nicked the catheter when re-suturing the graft, causing blood to leak into the catheter down to the red impella plug. knowing that the microchip within the plug is on the underside of the plug, the plug was inverted with the serial number towards the ceiling and the logo towards the patient's skin, tilted at an angle to help the plug drain and to protect the microchip. blood stopped oozing from the plug on the same day about an hour after this was noticed. it is assumed that the catheter has since clotted off. there has been no blood noted since and the device has been working perfectly. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (hole in catheter) has been completed. the impella device was not returned for evaluation. log review showed no abnormalities related to the hole in catheter. the root cause of the product damage was not determined since the product/images were not returned. a.1 revised patient identifier as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25957. e.1 revised first and last name as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25957. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25957.
11844 | 1220648-2025-25951 | 21292485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21292485&pc=OZD | 01/15/2025 | 02/18/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594135 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and developed low pump flow. impella placed via right femoral artery. notedly, the patient had a small aortic arch and left ventricle. during delivery, the impella cp kinked near the outlet cage. the impella cp device was removed and reinserted which resulted in the cannula kinking a second time. despite the noted issue, the physician decided to proceed with the case. due to the kink, the impella cp device was only able to provide flow around 2 l/min. the pci was completed with drug-eluting stents placed in the left anterior descending and circumflex arteries. the pci concluded and the impella device was explanted without incident. the patient was reported to be stable following the event. | [5 empty] | The impella pump with the cassette has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. a cannula kink was found. no biomaterial around impeller was found. no other damages or abnormalities. the data logs showed low flows from the case start. suction alarms at start. no motor current spikes outside of p-level changes. the cause of the low pump flow was most likely a kinked cannula. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address section as it was previously entered incorrectly on the initial manufacturer device report number.
11845 | 1220648-2025-25947 | 21292545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21292545&pc=OZD | 01/14/2025 | 03/06/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025530469 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella 5.5 device for mechanical circulatory support and experienced a cerebral vascular accident (cva). the patient had an impella 5.5 device placed and then pci and stenting of both the first diagonal, mid and proximal left anterior descending arteries. the patient lost pulsatility multiple times but maintained 3.1-3.5 flows and adequate mean arterial pressure. at the end of the procedure, it was decided that the patient would benefit from continued support until extubated and medications were weaned off. the patient was transferred to the unit on support. later that evening, it was noted the patient experienced a stroke. while decreasing sedation to extubate the patient, it was noted that the patient was unable to move their left arm or leg. the patient underwent a series of computed tomography (ct) scans, and medications was administered per stroke protocol. the next day it was noted that the patient was able to respond to commands, twitch toes on their left foot, and could feel painful stimuli to the left arm although they were unable to move with purpose yet. medication treatment was continued. the following day, the patient was maintaining and making improvements. ct was run and showed no progression of bleeding or increase in the cva area. the patient was able to move their left leg and toes but still no movement of the left arm as of yet. on day three and final day of support, it was reported that hemodynamics were stable, and the patient continued to make progress neurologically. the impella 5.5 device was successfully explanted with a patient outcome as survived. | [5 empty] | The investigation for the stroke has been completed. the returned product was visually inspected and biomaterial was observed on the outside of the motor housing which doesn't affect the results of the investigation. clinical details were not sufficient to determine root cause. therefore the root cause of the stroke could not be determined. added d9 device return date. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11846 | 1220648-2025-25959 | 21292671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21292671&pc=OZD | 01/15/2025 | 02/26/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544851 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with the impella 5.5 support. the physician attempted to insert the impella 5.5. however, there was significant resistance in aorta due to patient anatomy. the procedure was aborted and there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the delivery issue was most likely patient condition related. during attempted impella insertion there was significant resistance in the aorta due to patient anatomy preventing insertion despite trouble shooting and multiple attempts. d.3 revised fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25959. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11847 | 1220648-2025-25956 | 21292959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21292959&pc=OZD | 01/15/2025 | 02/27/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562965 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. after multiple failed attempts, the pump became kinked and would not make the turn down into the aorta. a new pump was subsequently prepped and placed for ongoing support. there was no noted patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the pump was returned for inspection. upon inspection it was noted that the pump was kinked and it was also noted that catheter near cannula also got kinked. the root cause of delivery issue is determined to be patient condition because of the small vessels and tortuous condition of the vessels. d.9 date device returned/information was added.
11848 | 1220648-2025-25962 | 21293157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21293157&pc=OZD | 01/15/2025 | 02/19/2025 |  | 02/03/2025 |  |  | 01/15/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547329 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient admitted in with poor underlying condition of hypotension, st-segment elevation myocardial infarction, and other cardiac and systemic comorbidities was implanted with an impella cp device for mechanical circulatory support (mcs) despite the limb being known to have no flows and ischemia resulted. the pump was not bypassed with any perfusion sheaths. additionally, the team assessed the neurological status to be poor so not intervention was done; rather the patient was transferred to tertiary center. the family made the patient a do not resuscitate (dnr) status due to lack of neuro response. the physicians decided to leave the impella cp in place while awaiting neuro status. no pulses in opposite leg with central line in place. later, the patient woke up and followed commands. the dnr was removed. the patient was escalated to an impella 5.5, and after removal of the impella cp it was noted, the patient had pulses in the leg. the latest update noted after seven additional days on an impella mcs, the patient was successfully weaned, and the device explanted with a survived outcome. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25962. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25962. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11849 | 1220648-2025-25961 | 21293514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21293514&pc=OZD | 09/09/2024 | 02/26/2025 |  | 02/03/2025 |  |  | 01/16/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528156 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured retroperitoneal bleed with a "probable" relationship to the impella 5.5 device used: ¿pci on (b)(6) 2024, hh at 9.6 prior to pci pel. dropped to 7.0 immediately after 1-unit prbc given. hh still dropped. ct done showing rp hematoma. femstop applied and 3 units of prbc given + 1 unit of plts¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11850 | 1220648-2025-25963 | 21297409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21297409&pc=OZD | 01/16/2025 | 03/03/2025 |  | 02/03/2025 |  |  | 01/17/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573966 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support had a hematoma at the incision site. the patient was taken to the operation room, and a pocket washout was done for the hematoma. the patient remained stable and on support. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11851 | 1220648-2025-25964 | 21298655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21298655&pc=OZD | 06/10/2024 | 02/19/2025 |  | 02/03/2025 |  |  | 01/16/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025483551 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular fibrillation with a "possible" relationship to the impella device used: "pt bared down, converted into vt. maps in 40s. intensivist at bedside, initial shock delivered in chair, temporarily converted into paced rhythm. quickly returning to vt multiple times, for which several lidocaine and amiodarone boluses were given along with sedation, mag sulfate, and several successive defibrillations, as soon as safe, pt was lifted back to bed and intubated by intensivist. total of 6 external shocks and 2 internal shocks (each shock was successful but quickly returned to vt). dr. westerdhal and ep team to bedside. internal icd therapies turned off and ppm interrogated/ av pacing optimizes by ep. currently pacing at 100 bpm w occasional pvc's.¿ the patient recovered with no sequelae. | [5 empty] | The investigation of ventricular fibrillation (vf) and ventricular tachycardia (vt) has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the vt/vf was not determined. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25964. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25964 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2130 was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25964. code 4641 was reported incorrectly on initial manufacturer device report number 1220648-2025-25964. a new health effect - impact code was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11852 | 1220648-2025-25965 | 21298727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21298727&pc=OZD | 01/16/2025 | 02/12/2025 |  | 02/03/2025 |  |  | 01/17/2025 | 02/03/2025 |  |  |  |  | 02/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786119 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a automated impella controller (aic) with a broken purge disc. the aic was removed from use and there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the console logs did not contain any data relevant to this complaint. the console was examined and a broken purge disc retainer was identified. the root cause of this issue was a broken purge disc retainer. e4 should have been left blank on manufacturer device report.
11853 | 1220648-2025-25928 | 21282076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282076&pc=OZD | 01/10/2025 | 02/28/2025 |  | 01/31/2025 |  |  | 01/10/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571276 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Handling Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella cp being prepped for support of a patient admitted in acute myocardial infarction/cardiogenic shock was inadvertently turned on outside the patient's body with the easy guide lumen in place. reportedly, the said impella cp device was discarded. a new impella cp device was obtained and implanted for mechanical circulatory support. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore. an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the pump did not start issue has been completed. no product or data logs were returned for evaluation. clinical details noted that the pump was initiated outside of the body with the easy guide lumen in place. the root cause of the pump did not start was most likely a red lumen issue as the pump was initiated outside of the body with the red lumen still in place.
11854 | 1220648-2025-25929 | 21282304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282304&pc=OZD | 01/14/2025 | 02/28/2025 |  | 01/31/2025 |  |  | 01/14/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594297 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. under fluoroscopy, the guidewire within the catheter looked "strange" and as if it may have a deformity within the catheter. reportedly, the pump "flawlessly" performed during the high-risk pci. the impella cp was weaned and successfully explanted. the patient was transferred to the intensive care unit in stable condition to recover overnight. there was no report of adverse effects to the patient. | [5 empty] | The investigation for the product damage-kink has been completed. the impella device was returned and was visually inspected. cannula kink near the outlet cage was observed. no other abnormality observed. data logs were not returned for review. clinical stated the wire within the catheter looked strange and it may have a deformity within the catheter which mostly occurred during insertion and caused cannula kink due to improper technique. the cause of the delivery issue most likely was use related due to improper technique resulting in a kinked guidewire. the cause of the product damage (cannula kink) was cannula kink due to improper technique. | The impella pump, purge cassette and data stick have been received from the customer, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11855 | 1220648-2025-25930 | 21282328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282328&pc=OZD | 11/20/2024 | 02/19/2025 |  | 01/31/2025 |  |  | 01/14/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485982 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Cardiogenic Shock | Respiratory Failure | [10 empty] | Required Intervention | Hospitalization | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced cardiogenic shock, acute respiratory failure and anemia with a "possible" relationship to the impella device used. the patient was discharged from one hospital after a length admission for critical limb ischemia. later this day, the patient felt week and had low pressure, ems was called and the patient experienced syncopal episodes with ventricular fibrillation arrest. cardiopulmonary resuscitation was performed with one shock and return of spontaneous circulation (rosc) was achieved. the patient had inferior st-segment elevation myocardial infarction (stemi) and was urgently taken back to the catheterization lab for stemi. percutaneous coronary intervention (pci) was started but the physician was unable to wire the left circumflex artery with the available equipment. patient started on neo drip for hypotension in the catheterization lab. the vessel remained occluded, and the decision was made to implant an impella cp device for mechanical circulatory support, proactively intubate and transfer the patient to an area hospital. automated impella controller (aic) to aic transfer was completed. impella flows were turned from auto due to impella flows reduced alarm. device noting suction above support level p-5. echocardiogram was completed and float swan-ganz to assess the right ventricle failure. an attempt to open the left circumflex artery was made. a transthoracic echocardiogram (tte) was performed showing impella position shallow and empty left ventricle with an empty right ventricle. the physician was able to cross the left circumflex artery and completed the pci. post pci, the patient transferred to the unit to rest overnight with plans for impella explant the following day. this next day, the impella remained at p-4 due to suction. the patient received two units of packed red blood cells for low hemoglobin/hematocrit. during support the patient had experienced cardiogenic shock, acute respiratory failure, and anemia. additional information from the registry stated the patient became increasingly hypotensive, rhythm ventricular tachycardia with flat arterial and pulmonary artery lines. the patient cardiac arrested, cpr was completed with return of rosc after five minutes. the patient continued to decline despite max inotropes and pressors, arrested again with cpr performed and rosc achieved. the family requested to make the patient a do not resuscitate. the patient was weaned, and the impella cp device was successfully explanted with a survived patient outcome. | [5 empty] | The investigation of the reported issues has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause was not determined. b.2 revised selections as required intervention to prevent permanent impairment/damage was inadvertently omitted from initial manufacturer device report number 1220648-2025-25930 and hospitalization - initial or prolonged was entered incorrectly on the initial manufacturer device report number 1220648-2025-25930. h.6 code 4617 was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25930. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11856 | 1220648-2025-25931 | 21282417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282417&pc=OZD | 01/14/2025 | 03/05/2025 |  | 01/31/2025 |  |  | 01/14/2025 | 01/31/2025 |  |  | 02/28/2026 |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478463 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the attempted implant of an impella 5.5 device for mechanical circulatory support (mcs) to a patient with cardiomyopathy was unsuccessful. preoperative computed tomography (ct) scan was obtained morning of surgery to evaluate anatomy. surgeon expressed concern about tortuosity of vessel and the impella's ability to be safely passed through vessel . after collaboration with cardiology, decision was made to proceed with impella placement due to elevated risk during scheduled robotic procedure. patient was prepped and draped in a normal fashion. incision was made over right axillary artery. artery was determined to be of good caliber and quality. a 10mm gelweave graft was sewn to the artery with approximately 30-degree bevel. at this point attention was turned to obtaining ventricular access. an .035 guidewire was inserted to the impella access sheath with difficulty passing through innominate artery into the ascending aorta. a pigtail catheter was used to aid wire but attempts to access ascending aorta were still unsuccessful. the .035 wire was exchanged for the impella .027 wire without successful access. the pigtail catheter was exchanged for an al-1 catheter, but again this was unsuccessful due to a sharp curve in patient's innominate artery. at this point, interventional cardiology was called in. a bern catheter in combination with impella .027 wire was able to successfully access the left ventricle by cardiology. attention was then turned to placement of the impella device. it became immediately apparent that it was going to be very difficult to pass the impella through the tortuous anatomy. multiple attempts to place different buddy wires were made with various stiff wires including stiff guidewires and a lunderquist but were unable to pass through the curvatures. discussion was made regarding straightening or repositioning the patient's arm, but the surgeon was concerned about too much continued manipulation of the artery. discussion was made regarding alternative access sites. it was noted the patient already had a sternotomy and will require another for left ventricular assist device placement; therefore, to redo sternotomy for direct placement was not an option. the left axillary artery appeared to be feasible on ct scan, however surgeon did not want to attempt as the pump and left axillary insertion site would be in the way for the left robotic diaphragm plication. the patient had remained stable throughout the entirety of the case in which the decision was made to abort the case. additional comments from the physician noted it was not the fault of the impella pumps and was an expected outcome given the tortuosity of the vessel. | [5 empty] | The impella pump has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the returned product was visually inspected and a cannula kink was observed. the root cause of the delivery issue was most likely patient condition related, surgeon expressed concern about tortuosity of vessel and impella ability to be safely passed through vessel.
11857 | 1220648-2025-25932 | 21282555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282555&pc=OZD | 01/14/2025 | 03/17/2025 |  | 01/31/2025 |  |  | 01/14/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533103 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  |  | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the impella implant was noted as uneventful and the patient was additionally placed on extracorporeal membrane oxygenation due to a weak right ventricle and a weak lung. while on support the patient experienced access site bleeding and had a purge fluid leak. approximately, four days after start of the report is when the patient had an access site bleed and was administered "at least" two units of red blood cells. hemostasis was not achieved, however. additional action taken as a result of the event is unknown. however, it was noted that purge fluid was leaking from the red impella plug (not at the filter and not at the yellow luerlock), and purge values were within the normal rang. it was explained, however, that it was impossible to understand whether and how much purge fluid is still flowing through the motor gap, and the pump can stop at any time. pump replacement as soon as possible was recommended and to closely monitor the motor current until then. extensive discussion about escalation was made, as the patient still had a relatively high catecholamine requirement. the decision on escalation was to be made the following day. the following day, however, the patient would expire. the patient was in multi-organ failure (mof) and care was withdrawn. the patient required extracorporeal life support along with impella mcs and would pass away nine days after start of support it appears because of mof. no further details regarding the final outcome of the bleeding event and purge fluid leak were provided. given the impella-related events, there is not enough information to exclude the impella device as a possible contributing factor of the patient's outcome. | [5 empty] | The investigation of access site bleeding and purge leak ¿ pump has been completed. the pump was returned for investigation. a data log was not provided for the case run. no physical abnormalities were observed on the returned product. the repo sheath was found sutured through the catheter, and this is where the leak was reproduced. the catheter was further cut down, and the damaged purge line was observed leaking. the purge pressure was within specifications during testing due to the slow leak. the leak inside the catheter could reach the red handle, as mentioned in clinical details. the cause of the purge leak was damaged purge tubing at the 75 centimeter mark due to the application of inappropriate sutures. the cause of the access site bleeding issue was patient condition related to coagulation disorder. no heparin was given in the purge fluid at the time of bleed. d.9 date device returned/information was added. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25932 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was reported incorrectly on initial manufacturer device report number 1220648-2025-25932. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2025-25932 and were added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11858 | 1220648-2025-25933 | 21282559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282559&pc=OZD | 01/15/2025 | 03/13/2025 |  | 01/31/2025 |  |  | 01/16/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1497915 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during patient transfer to another hospital, a controller error alarm appeared on the automated impella controller. it was noted that the battery level was displayed at 50% at the time of the alarm. the alarm resolved without troubleshooting and support continued uninterrupted. there was no noted patient harm. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console had five controller errors for power batter manager (pbm) voltage out of specification that resolved before the next error triggered. these varied in pbm1 and pbm2 and affecting either or both vpump and vpurge. after the last error, the console operated for thirty minutes without error before the pump was removed from the console. the error did not manifest during additional boot ups in the field with no pump running. the 50% battery power was due to pbm glitch in communication and was not related to the batteries' functionality. during the console investigation the reported failure mode was not reproduce during testing. root cause: the controller error was most likely caused by a transiently malfunctioning pbm that was unable to power multiple sub-circuitries in the console. d9 date device returned to manufacture was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25933. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25933 as it is unknown if the initial reporter also sent the report to the food and drug administration. h3 the device was returned at the time the initial manufacturer device report number 1220648-2025-25933 was submitted, "device not returned" was inadvertently select yes.
11859 | 1220648-2025-25909 | 21282663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282663&pc=OZD | 01/15/2025 | 02/28/2025 |  | 01/31/2025 |  |  | 01/15/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563959 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was admitted in with cardiogenic shock and was on impella 5.5 support for 19 days. during support, there was a pump/purge sidearm leak. the leak did not prompt early pump explant or other forms of intervention. the pump was explanted and a left ventricular assist device was placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. there was no product returned. the purge pressure remained stable throughout case duration around 550 mmhg while purge flow fluctuated around 10 milliters per hour. there were no alarms associated with a leak. the root cause of the purge leak - pump could not be determined since the product was not returned.
11860 | 1220648-2025-25934 | 21282693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282693&pc=OZD | 11/23/2024 | 03/06/2025 |  | 01/31/2025 |  |  | 01/16/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573220 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella 5.5 device used: ¿concern for hemolysis noted (b)(6) 2024 with ldh 400 -> 700 and red tinged urine. impella was pulled back 0.5 cm under echo guidance and pulled back 1 cm (b)(6) 2024. no noted hemolysis as of (b)(6) 2024.¿ the patient recovered with no sequelae. loqi additionally reported, that polymorphic vt was "possibly related¿. patient with episodes of non sustained, self-terminating polymorphic vt starting (b)(6) 2024, longest 20 beats. ongoing short runs of polymorphic vt on (b)(6) 2024. etiology thought to be due to impella touching posterior wall of lv. lidocaine gtt started. (b)(6) 2024 impella decreased to p6 and epi turned off with resolution of nsvt, however overnight recurrent stacked polymorphic nsvt, pacing at 110 to suppress. impella was pulled back 1 cm on (b)(6) 2024. received stellate ganglion block (b)(6) 2024 with significant improvement in vt. no significant vt noted as of am of (b)(6) 2024." the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the ventricular tachycardia (vt) and hemolysis has been completed. data logs showed the pump ran without issue during support. there were placement signal not reliable, positioning and suction alarms triggered during the case but were resolved quickly. the product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the vt was unable to be determined due to insufficient clinical details. g.2 added selection as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25934. h.6 code 1721 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25934 and a new code was added. additional health effect - impact code was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25934. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11861 | 1220648-2025-25935 | 21282779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282779&pc=OZD | 01/18/2025 | 03/07/2025 |  | 01/31/2025 |  |  | 01/18/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594744 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, there was ongoing high pressure alarms. tissue plasminogen activator protocol (tpa) was ordered through the purge. the purge pressures normalized back to 400-500, and purge flow increased to four cubic centimeters per hour. the impella stopped and displayed impella defective alarm flashing bright and dark yellow. the soft buttons were unresponsive. no patient harm has been reported. | [5 empty] | The investigation of high purge pressure (hpp) has been completed. no product was returned for evaluation. data logs were reviewed and the logs showed a rise in purge pressure with drop in purge flow over approximately 30 minutes before initial ¿high purge pressure¿ alarms are triggered. purge pressure remains high but slightly variable for approximately 3 hours before ¿high purge pressure¿ alarms are triggered again and were sustained for approximately 6 hours before resolving. no motor current spikes prior to purge pressure rise were observed. clinical details noted that on day 2 of support, patient received 500 of heparin. additionally, tissue plasminogen activator (tpa) was added to purge solution and was able to resolve hpp alarms. the root cause of the high purge pressure was most likely biomaterial buildup. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11862 | 1220648-2025-25936 | 21282795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282795&pc=OZD | 01/18/2025 | 03/28/2025 |  | 01/31/2025 |  |  | 01/18/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573062 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, the physician reported difficulty advancing the catheter as it has become deformed. the catheter cover has contracted and a thickening has occurred. it cannot be pushed through the repositioning sheath. no patient harm has been reported, and the patient remains on impella support. | Additional information was provided that there was high purge pressure. tissue plasma activator was added to the purge. |  |  |  |  | The investigation was completed. the pump was not returned. the data logs were not recovered. 26 days into care, the medical doctor reported purge pressure high alarms. the discussion with medical team about recombinant tissue plasminogen activator -lyse and exchanging the pump occurred with no confirmation. the pump was used two more days successfully before the patient was weaned off. the pump was not returned. the data logs were not recovered. the cause of the high purge pressure was not determined since the product was not returned and lack of clinical details. the medical doctor reported difficulty advancing the catheter as it had become deformed. it was reported to be contracted and thickened. no harm was reported. the pump was not returned. the data logs were not recovered. the cause of the delivery issue was not determined since the product was not returned and lack of clinical details. b.5 additional information was received. d.6b explant date was added. h.6 code was added due to new information received. d.7a revised as selection was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25936. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
11863 | 1220648-2025-25937 | 21282939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21282939&pc=OZD | 01/18/2025 | 03/03/2025 |  | 01/31/2025 |  |  | 01/18/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590779 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, there was little to no distal limb perfusion in the patient's right femoral artery. an external femoral/femoral bypass was performed in response to the condition and distal flow was subsequently confirmed. support with the implanted pump was uninterrupted. the patient survived the event. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
11864 | 1220648-2025-25938 | 21283049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21283049&pc=OZD | 01/18/2025 | 02/20/2025 |  | 01/31/2025 |  |  | 01/18/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590506 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had a hematoma at access site and continuous oozing. dressing was removed and it was noted that the repositioning sheath was sutured down rather than using forward tension sutures. the sutures were removed and the site stopped oozing immediately. forward tension sutures placed. staff monitored. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since the bleed was resolved by angle matching. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11865 | 1220648-2025-25939 | 21283439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21283439&pc=OZD | 01/19/2025 | 03/07/2025 |  | 01/31/2025 |  |  | 01/19/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573386 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 11 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Gastrointestinal Hemorrhage | No Clinical Signs, Symptoms or Conditions | [10 empty] | Death | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 while on patient support had high purge pressure and purge system blocked alarms. no kinks were observed in the tubing. impella 5.5 was decreased to p-3 and p-2 where flows were appropriate. transesophageal echocardiography was performed to confirm impella 5.5 outlet position related to centrally cannulated aortic cannula. the impella 5.5 measured five centimeters in the left ventricle (lv) and was properly positioned above the aortic valve. heparin purge was started along with systemic heparin drip. purge flow, purge pressure, and lv waveform was spontaneously returned to normal along with normal flows at p-4. the patient remains on impella support. | It was further reported that the patient¿s labs were hemolyzing and that the patient developed gastrointestinal (gi) bleed. bival was discontinued. the next day the gi bleed was stabilized. the family decided to withdrew care and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | B.1 added adverse event based on the additional information received. b.2 added death and date of death due to new information received. b.5 additional information was received. d.6b explant date was added. h.1 revised to death based on the additional information received. h.6 added additional codes for the new information received. h.10 additional instructions for use were added based on the additional information received. the investigation has been completed. no product was returned for analysis. the data logs confirmed purge pressure high alarms. the logs showed a motor current spike with speed deviation and a drop in purge flow right before high purge pressure occurrence. it was unknown if tissue plasminogen activator was added to the purge. there was high purge pressure for about two hours after which values started decreasing towards normal purge pressure values. the root cause of the high purge pressure was most likely biomaterial. the root cause of the hemolysis was most likely biomaterial as evidenced by motor current spike in the logs before hemolysis occurrence. the root cause of the vascular damage - peripheral was not determined as insufficient clinical information was provided as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11866 | 1220648-2025-25940 | 21283699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21283699&pc=OZD | 01/19/2025 | 03/05/2025 |  | 01/31/2025 |  |  | 01/19/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525060 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during delivery of an impella cp pump for mechanical circulatory support. it was documented that while delivering the device over the guidewire, the technician powered on the pump. due to the resulting complication, the decision was made to replace the devices. there was no patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues and product damage has been completed. impella cp was returned for evaluation. upon visual inspection, no physical damage to the pump was observed. data logs were reviewed and rt log of case shows an alarm #13 motor current high pump stop occurred. placement signal throughout rt log was at approximately 0mmhg indicating pump was never inserted into the body. pump flow and motor speed remained at 0 for entirety of log. clinical details noted that pump was turned on by tech when wire was still inserted. the root cause of the delivery issue was most likely use related as the pump was initiated before the pump was inserted into the body with the guidewire still inserted.
11867 | 1220648-2025-25941 | 21283870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21283870&pc=OZD | 01/29/2025 | 02/20/2025 |  | 01/31/2025 |  |  | 01/01/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570042 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was a suspected thrombus in the cannula. the staff were no longer able to achieve/see left ventricle waveforms over p4 and flow did not equate to p-level. the pump was scheduled to be explanted however no other information was provided at this time. | [5 empty] | The investigation of thrombus and low pump flow has been completed. the product and data were not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. thrombus failure mode was removed as no thrombus was observed, only suspected. d.6a implant date was added. b.3 revised date of event as it was reported incorrectly on initial manufacturer device report number 1220648-2025-25941. e.1 revised facility name as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25941. g.1 revised reporting contact address line 1, reporting contact us city, reporting contact state, reporting contact us zip and reporting contact country as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-25941. revised manufacturing site name and address section as it was reported incorrectly on initial manufacturer device report number 1220648-2025-25941. g.3 date received by manufacturer was reported incorrectly on initial manufacturer device report number 1220648-2025-25941. the date should of reflected 29-jan-2025. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d6a and d6b implant/explant info was requested but not yet verified a supplemental mdr will be filed should any new information be received. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
11868 | 1220648-2025-25942 | 21284699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21284699&pc=OZD | 12/01/2024 | 03/21/2025 |  | 01/31/2025 |  |  | 01/20/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575020 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that the patient on impella cp support developed hemolysis. the support level of p-8 could not be reduced so the impella was exchanged to a 5.5. the patient is stable and survived the event. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11869 | 1220648-2025-25943 | 21285573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21285573&pc=OZD | 01/13/2025 | 02/18/2025 |  | 01/31/2025 |  |  | 01/13/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578457 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device and experienced access site bleeding with blood administration. care was eventually withdrawn, and the patient met their demise. prior to the day of starting mechanical circulatory support, the patient underwent a left heart catheterization which revealed left anterior descending artery (lad) lesion which was treated with one drug-eluting stent. the patient was stable and sent to the unit for recovery. however, subsequently, the patient decompensated with drop in blood pressure requiring three amps of epinephrine. levophed and dopamine drips were started, and the patient was taken back to the catheterization lab. imaging revealed non-obstructive coronary artery disease with patent lad stent. right heart catheterization was completed with elevated pressures. the patient was in cardiogenic shock stage e. the decision was made to implant an impella cp device. the first impella cp being prepped for support was inadvertently turned on outside the patient's body with the easy guide lumen in place. a new impella cp device was obtained, and the patient was successfully placed on impella cp support in the setting of cardiogenic shock after going into heart block approximately three hours prior. due to high mean arterial pressure in the 140s, performance level was decreased to p-7 as dopamine and levophed were titrated. dopamine was discontinued. suction alarms began and performance level were adjusted to resolve. the impella cp placement was confirmed and volume evaluation was continued. the patient was transferred back to the unit in critical condition and soon after arriving, the patient's blood pressure "significantly" dropped. echocardiogram confirmed the impella cp depth and right heart dysfunction. dobutamine was ordered, and the patient required an amp of epinephrine with an epinephrine drip that followed. the team discussed the patient's status with family and no escalation was expected at such said time. the following day it was noted that the patient had multiple suction events overnight requiring 500 milliliters of normal saline bolus. suction events resolved after the normal saline bolus, and the impella cp support level was weaned from support level p-5 to p-8. epiphrine was weaned down, and it was noted that the site was clear of bleeding. the patient continued on support and two days later, however, access site oozing was now noted. blood administration was started with one unit of red blood cells and one unit of fresh frozen plasma. the impella cp's depth was adjusted to 3.66 centimeters from 4.23 centimeters with no plans to escalate at such said time. additionally, it was noted there was concern for bleeding in the lungs based on computed tomography scan earlier in the day. however, a bronchoscopy was performed showing no active bleeding. heparin remained on hold. the patient experienced intermittent episodes of non pulsatility requiring some titration of drips. the following day care was withdrawn per the family's wishes. the patient was critically ill in stage e (extremis) cardiogenic shock and never fully recovered from the cardiac event. however, the impella cannot be excluded as a possible contributing factor to the overall event outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this is one of two devices related to the patient's implant experience. refer to manufacturing report number 1220648-2025-25928 for the initial impella cp device. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely anticoagulation management since bleeding was resolved by withholding the heparin.
11870 | 1220648-2025-25944 | 21288810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21288810&pc=OZD | 01/18/2025 | 03/03/2025 |  | 01/31/2025 |  |  | 01/18/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1540253 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant in slovenia/eu had an automated impella controller (aic)-5.5 pump support ongoing when after 8 days the aic shut down and was replaced. the instructions for use were followed and a 2nd aic was used. the pump support continued another 2 days without harm or injury. patient survived. | [5 empty] | The investigation for the console (aic) shutdown restart issue has been completed. data logs were reviewed which showed 8 days after the pump was started, the console was shut down unexpectedly. then, the console was manually restarted after 2 minutes 51 seconds. after restart, the console displayed ¿unexpected controller shutdown ¿ alarm,¿ event # 603. this confirms the reported failure mode. the pump resumed running at the same p-level p5 as it was before the unexpected shutdown. the support continued for 55 minutes with ic8426 before swapping the console. the console was returned for evaluation. during testing, it was unable to reproduce the reported failure mode upon running the optical test pump. the root cause for the unexpected controller shutdown was not determined due to the inability to reproduce. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11871 | 1220648-2025-25945 | 21288818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21288818&pc=PYX | 01/18/2025 | 09/22/2025 |  | 01/31/2025 |  |  | 01/18/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590629 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella rp support developed bleeding at the impella access site. pressure was held and suture was added. additionally, there was a placement signal alarm due to the impella sensor malfunctioning. the team monitored motor current. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. data logs are not relevant to the complication and the product was not returned. clinical details reported that the patient was bleeding from the access site and required an extra suture to resolve it. however, additional bleeding was reported two days later which was treated with blood products, another suture, and a sandbag. it is reported that the patient's partial thromboplastin time was greater than 200 at the time of the impella rp site bleed, though it is not reported if this is within the hospital's therapeutic range. additionally, the intra-aortic balloon pump access site also had oozing two days prior which required several days of blood products. bleeding at the impella site had resolved at the time of explant. based on the clinical details provided, the root cause of the access site bleed was most likely patient condition related. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25945 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed and there were no placement signal not reliable alarms and all optical signal metrics were stable during the case. that being said, there were abnormalities in the placement signal calculations around the reported date. the pump was not returned for investigation. the failure mode was changed from placement signal to optical signal issue based on the fact that data log review an optical signal metric. based on the data log review and clinical details, the root cause of the optical signal issue was determined to most likely be patient condition. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-25945.
11872 | 1220648-2025-25946 | 21288821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21288821&pc=OZD | 01/19/2025 | 03/05/2025 |  | 01/31/2025 |  |  | 01/19/2025 | 01/31/2025 |  |  |  |  | 01/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1275229 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | After 69.48 hours of support the decision was made by the family to withdraw care place patient on palliative care and the patient expired. the exact date of death is unknown. per the medical safety officer review it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported that during purge cassette replacement, the disc holder on the automated impella controller (aic) was inadvertently torn off. the aic was replaced to resolve the damage. there was no patient harm and support continued with the implanted pump. |  |  |  |  | B1 revised to reflect both adverse event and product problem on manufacturer device report 1220648-2025-25946. b2 revised to reflect death, the exact date of death is unknown 1220648-2025-25946. b5 revised to include patient outcome and mso statement which inadvertently was not included on manufacturer device report 1220648-2025-25946-1. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25946 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 manufacturing site name and address revised as was inadvertently reported incorrectly on manufacturer device report number 1220648-2025-25946-1. h1 type of report revised to reflect the patient outcome of death on manufacturer device report number 1220648-2025-25946. h6 revised to reflect the patient outcome of death on manufacturer device report number 1220648-2025-25946. | Additional information provided in d9, h3, h6 and h10. the investigation into the reported issue has been completed. console logs did not contain any data relevant to this complaint. the console was evaluated, and the reported broken purge disc retainer was confirmed. the root cause of this issue was a broken purge disc retainer. before sending this console back to the customer, field service was instructed to perform the following: replace the blue purge disc retainer, purge disc detect flag, and purge flag spring, perform preventative maintenance and all functional checks on the console and optical system. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
11873 | 1220648-2025-25894 | 21272919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21272919&pc=OZD | 01/14/2025 | 02/18/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574849 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 was attempted to be placed in a patient with an acute myocardial infarction. however, due to the patient¿s anatomy the case was aborted. the patient's outcome as a result of this event is unknown. reportedly, an impella cp device was eventually implanted to support the patient and at end of support, the patient was successfully weaned and device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy.
11874 | 1220648-2025-25895 | 21272992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21272992&pc=OZD | 01/14/2025 | 02/26/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561429 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with the impella 5.5 support. it was reported that impella 5.5 was backloaded and there was difficulty passing the cannula through the artery. contrast was infused through the impella introducer. the medical doctor noted the vessel to be clamped down and was small. the decision was made to remove the impella 5.5 and an impella cp was placed via the axillary artery graft. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella 5.5 backloaded, difficulty passing cannula through the artery. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿.
11875 | 1220648-2025-25896 | 21273055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273055&pc=OZD | 01/14/2025 | 02/28/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  | 02/28/2026 |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480011 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information included that groin site bleeding was noted post patient movement when the sutures were noted to be loosened and the dressing needed reinforced. manual pressure held and groin site redressed achieving hemostasis. | The complainant had an impella cp pump placed to support a patient admitted in with acute myocardial infarction and cardiogenic shock. post-insertion and initiation of the pump, a kink of impella cannula near the outlet appeared, leading to continuous suctions alarms. despite troubleshooting, the physician requested a new pump for support. there was no reported harm to the patient. |  |  |  |  | The investigation of product damage (cannula kink) and access site bleeding has been completed. the failure mode product damage (cannula kink) changed to low pump flow since the clinical stated post-insertion and initiation of pump, kink of impella cannula near outlet appeared leading to continuous suctions alarms, despite troubleshooting. returned compliant device visually inspected. cannula kink near outlet cage observed. no biomaterial or broken blade observed. no other abnormality observed. pump connected to aic and run for around 10 minute to reproduce low pump flow issue. no low flow issue reproduced and pump flow was within specified limit. data log not available to review. the cause of the low pump flow issue most likely was cannula kink related due to improper technique. failure mode access site bleeding - major added since groin site bleeding post patient movement when sutures were noted to be loosened and dressing needed reinforced. manual pressure held and groin site redressed achieving hemostasis. the cause of access site bleeding - major most likely issue was due to patient movement. instructions for use as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b1 revised to reflect both adverse event and product problem on manufacturer device report number 1220648-2025-25896. b5 revised to include the additional information that was not included on initial manufacturer device report number 1220648-2025-25896. h1 revised to serious injury to reflect the additional information that was not included on initial manufacturer device report number 1220648-2025-25896. h6 revised to reflect the additional information that was not included on initial manufacturer device report number 1220648-2025-25896. h10 instructions for use added to reflect access site bleeding. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11876 | 1220648-2025-25897 | 21273176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273176&pc=OZD | 01/08/2025 | 02/19/2025 |  | 01/30/2025 |  |  | 01/08/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580794 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during the impella cp insertion, pre-high-risk percutaneous coronary intervention (pci), an unsuccessful attempt was made utilizing a long 14 fr impella introducer sheath. upon removal of dilator, the introducer sheath kinked in two places; the wire was able to pass, and a second introducer sheath was unsuccessfully attempted as it kinked as well. the decision was made to utilize a competitor sheath, which passed without issues. the impella cp pump was placed and had good flows for the duration of the pci. after pci completion, the impella cp was successfully weaned and explanted. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two devices associated with this event. refer to medical device report for the introducer sheath lot/serial number unknown with date of event 08-jan-2025 and identifier ending in (b)(6). | The investigation of introducer damage ¿ mechanical has been completed. data logs are not relevant to the complication and the product was not returned for investigation. clinical details reported that each time the team attempted to remove the dilator from the 14 french long sheath, it would kink. the patient's vessels were reportedly tortuous. the complication was resolved by using a 16 french gore sheath. based on the clinical details provided, the root cause of the introducer damage was determined to most likely be patient condition.
11877 | 1220648-2025-25898 | 21273223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273223&pc=OZD | 01/14/2025 | 02/14/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533779 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the patient developed some bleeding at the impella cp access site during support. the systemic heparin drip was suspended due to the bleed and femostop was placed to manage the condition. support continued uninterrupted following the intervention. the outcome of the patient following the event is unknown. however, two days following the event the patient was successfully weaned and the impella cp device was explanted with a survived patient outcome. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue could not be determined since the product was not returned and insufficient clinical details were provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25898 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25898. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11878 | 1220648-2025-25899 | 21273228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273228&pc=PYX | 01/14/2025 | 02/27/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025599015 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella rp flex placed to support a patient admitted and being considered for a transplant. the patient was additionally supported by a impella 5.5 pump. on day 5 of the rp flex support there was concerns of clot entrainment and hemolysis. the rp was explanted and biomaterial was observed on the inflow of the explant. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of hemolysis and thrombus has been completed. the returned pump was inspected and biomaterial was found around the impeller in the inlet. there were no other abnormalities. the data logs show suction alarms on 01/09 when increased hgb values were reported and show 2 motor current spikes with speed deviations on 01/11. the complaint intake mentioned increased hgb values on 01/09/2025 and after with highest hgb levels on 01/13. no other information was provided about patient support or management. it is likely that the hemolysis worsened with biomaterial ingestion. the returned pump was hemolysis tested after biomaterial was cleared and it passed. the root cause of the hemolysis was most likely due to biomaterial. as the necessary information was not provided, the root cause of the thrombus was not determined e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25899 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source was revised to reflect study as it was not included on the initial manufacturer device report number 1220648-2025-25899. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11879 | 1220648-2025-25901 | 21273235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273235&pc=OZD | 11/13/2024 | 02/25/2025 |  | 01/30/2025 |  |  | 01/09/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558800 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular tachycardia with a "possible" relationship to the impella device used: ¿patient was placed on cpap however at that time he went into afib w/rvr which progressed to vt with a pulse. amio bolus given. patient re-sedated and shocked at 360. following vt episode patient had increase back on pressor requirements followed by multiple long pauses.¿ the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source was revised; study was missing from manufacturer device report 1220648-2025-25901.
11880 | 1220648-2025-25900 | 21273236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273236&pc=OZD | 01/14/2025 | 03/06/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571444 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | Unspecified Vascular Problem | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp for care escalation from an intra-aortic balloon pump (iabp) for the patient admitted in with acute myocardial infarction and cardiogenic shock. the iabp site was utilized and the sheath exchange was done for the cp but, during this time vascular damage and bleeding occurred. the bleed led to a hematoma. the site then suffered from limb ischemia. the ischemia prompted pump explant and replacement in the operating suite, where additionally femoral repair was done. the team eventually escalated to the 5.5 pump as femoral vessels clamped down. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding, limb ischemia, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the cause of the access site bleeding issue was most likely use related since best practices were not followed. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of vascular damage was unable to be determined as limited clinical details were provided and no product was returned for review. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25900 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11881 | 1220648-2025-25902 | 21273251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273251&pc=OZD | 01/14/2025 | 02/25/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503251 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, an echocardiogram was done after transfer which showed signs of left ventricle thrombus. the patient remained on support and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25902 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11882 | 1220648-2025-25903 | 21273253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273253&pc=OZD | 01/14/2025 | 03/14/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516875 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that insertion of an impella 5.5 implant attempt to a patient off-pump coronary artery bypass surgery was aborted due to the patient's anatomy. the outcome as a result of the event is unknown. however, there is no report or indication of harm to the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was determined to be patient condition because the insertion was aborted due to patient anatomy. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d.6a/d.6b implant and explant dates were added as they was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25903. e.1 added facility name as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25903. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25903 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25903. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25903 and a new code was added. h.10 instructions for use were added as they were inadvertently omitted on initial manufacturer device report number 1220648-2025-25903. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11883 | 1220648-2025-25904 | 21273344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273344&pc=OZD | 01/08/2025 | 03/13/2025 |  | 01/30/2025 |  |  | 01/08/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580794 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | UNIDENTIFIED INTRODUCER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during the impella cp insertion, pre-high-risk percutaneous coronary intervention (pci), an unsuccessful attempt was made utilizing a long 14 fr impella introducer sheath. upon removal of dilator, the introducer sheath kinked in two places; the wire was able to pass, and a second introducer sheath was unsuccessfully attempted as it kinked as well. the decision was made to utilize a competitor sheath, which passed without issues. the impella cp pump was placed and had good flows for the duration of the pci. after pci completion, the impella cp was successfully weaned and explanted. there was no report of harm to the patient. | [5 empty] | Product-specific additional information was requested and is unavailable. per the abiomed clinical coordinator the introducer sheath was not from an impella cp box but rather a standalone introducer sheath that was ordered by the catheterization lab, which was discarded after use before it could be retained for return. therefore, an evaluation of the device is not possible. a no product-specific information is available, respective fields have been marked as "unknown" or left blank accordingly. the device brand name and procode are based on the 14french introducer kit used with an impella cp device. should updated, further information become available and/or upon investigation closure, a supplemental mdr will be filed accordingly. this is one of two devices associated with this event. refer to medical device report for the impella cp serial number (b)(6) with date of event 08-jan-2025 and identifier ending in (b)(6). | The investigation for the introducer damage has been completed. the product was not returned. the event logs were not applicable to the reported event. clinical details stated that two 14french long sheaths were kinked during this case. the root cause of the mechanical introducer damage was most likely a sheath kink related to the patient's vessel tortuosity. corrections to the initial mfr report: d1, d2 brand name, product code, and common device corrected to reflect impella cp. the initial mfr report incorrectly reported out on the unidentified introducer. d4-corrected model number, catalog number, serial number, lot number, expiration date, and udi number to reflect the reported impella cp d6a/d6b added implant and explant date of impella cp d10-concomitant device listed the unidentified introducer g1- updated manufacturer email address to current h4-added device manufacturer date h6-component code 3038 changed to 4742 and 925.
11884 | 1220648-2025-25905 | 21273505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273505&pc=OZD | 01/05/2025 | 03/06/2025 |  | 01/30/2025 |  |  | 01/12/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1581781 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Use of Device Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp for mechanical circulatory support and escalated to an impella 5.5 four days later due to the expected prolongation of support. approximately after 11 days on impella 5.5 support, leaking from purge cassette was observed. per the clinical engineer liquid was leaking from the purge cassette on the fourth day after starting use. consequently, the purge set was replaced. however, 16 days after start of support, abnormal noise was heard around the purge cassette insertion area of the automated impella controller (used with both the impella cp and impella 5.5), which prompted replacement of the purge set. the abnormal noise was resolved after replacing the purge housing. at some point the patient developed sepsis as case wrap-up noted ¿cardiac function recovers poorly, and eventually sepsis disappears.¿ the timeframe for this adverse event and its relationship to the impella pump is unclear. however, the patient would eventually expire based on reflected case note ¿expired on support,¿ which occurred on the 25th day of support. reportedly the patient went into multiple organ failure caused by sepsis. antibiotics were administered. no detail information surrounding the patient's condition, the sepsis and death were provided, and with limited information, there is not enough information to exclude the impella device as a contributing factor in the patient¿s outcome. | [5 empty] | The investigation for the console (aic) purge issues has been completed. console logs were reviewed as a precaution but did not contain any data relevant to the complaint. console was tested after arriving in field service. the reported abnormal noise could not be reproduced during testing. the console operated without any issues and no noise was observed from the purge assembly. the root cause of the reported issue could not be determined due to the inability to reproduce. console operated normally and no abnormal noises were observed while testing the console. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. this is one of two devices associated with the event. refer to medical device report for the impella 5.5 serial number (b)(6) with date of event 05-jan-2025 and identifier ending in (b)(6).
11885 | 1220648-2025-25907 | 21273753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273753&pc=OZD | 01/14/2025 | 02/26/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582993 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that one day following impella cp implant, the patient experienced oozing from the access site. femostop was applied and one unit of packed red blood cells were given to the patient to treat the condition. support continued uninterrupted. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided.
11886 | 1220648-2025-25906 | 21273931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21273931&pc=OZD | 01/13/2025 | 02/18/2025 |  | 01/30/2025 |  |  | 01/13/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573978 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed a hematoma. heparin was held and hematoma improved. additionally, the patient was having non-sustained ventricular tachycardia runs. impella was repositioned. the patient survived the event. | [5 empty] | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely anticoagulation management since bleeding was resolved by withholding the heparin. the root cause of the vt/vf cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11887 | 1220648-2025-25908 | 21274069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21274069&pc=OZD | 01/14/2025 | 02/25/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  | 04/30/2025 |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371684 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported post impella cp device implant, same day, the patient experienced oozing from the access site. surgiseal was applied to treat the condition, which was successfully resolved. support continued uninterrupted. the patient was reported to be stable following the event. | [5 empty] | The investigation of access site bleeding has been completed. the product was not returned. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details were provided. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25908. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
11888 | 1220648-2025-25910 | 21274315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21274315&pc=OZD | 01/15/2025 | 02/18/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582164 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, oozing/bleeding along with hematoma was noted at impella cp site after exchanging for the repositioning sheath. manual pressure was held but bleeding persisted. pump was explanted and 14fr sheath inserted in its place. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the root cause of the access site bleeding issue cannot be determined since insufficient clinical details were provided. d6a and d6b were inadvertently reported incorrectly on manufacturer device report 1220648-2025-25910. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11889 | 1220648-2025-25911 | 21274330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21274330&pc=OZD | 01/15/2025 | 02/25/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp being placed emergently in a transcatheter aortic valve replacement patient. the anatomy precluded placement and delivery was aborted. the patient¿s outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy.
11890 | 1220648-2025-25912 | 21274401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21274401&pc=OZD | 01/14/2025 | 02/26/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592117 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a patient admitted into community hospital and then sent to a tertiary center. the impella cp was supporting but during the support, the pump was repositioned due to persistent and recurring runs of ventricular tachycardia. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the vt was not determined.
11891 | 1220648-2025-25913 | 21274561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21274561&pc=OZD | 12/22/2024 | 03/20/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567912 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hypoxia | Unspecified Infection | Sepsis | Respiratory Failure | [9 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured respiratory failure with a "possible" relationship to the impella device used: ¿patient lost pulsatility with the impella. patient was extubated on (b)(6) 2024 after impella procedure but reintubated on (b)(6) 2024. on (b)(6) 2024, patient had worsening hypoxia with secretion burden. rml and rll collapsed with mucous plugging. bal was positive for pseudomonas aeruginosa and +gram cocci. patient was extubated on (b)(6) 2024.¿. the patient recovered with no sequelae. another notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding the patient as they endured sepsis with a "possible" relationship to the impella device used: ¿on (b)(6) 2024, patient had worsening hypoxia with secretion burden. rml and rll collapsed with mucous plugging. bal on (b)(6) 2024 was positive for +gram cocci and pseudomonas aeruginosa. patient was had ecmo placed on (b)(6) 2024. patient developed elevated temps after. patient was treated with zosyn and cefepime. id consulted and followed patient. patient also had c.diff on (b)(6) 2024 and was treated with 10d of vancomycin. team determined patient was septic though cause of infection was unknown.¿. the patient recovered with no sequelae | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the failure mode "hypoxia/desaturation" was removed because the hypoxia is unrelated to impella since impella is not a device that oxygenates the blood per medical safety officer review. b.2 revised as selection on initial manufacturer device report number 1220648-2025-25913 was reported incorrectly. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25913. h.6 code 1918 was reported on the initial manufacturer device report number 1220648-2025-25913 and should be removed per investigation results. code 2484 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25913.
11892 | 1220648-2025-25914 | 21274719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21274719&pc=OZD | 01/13/2025 | 03/04/2025 |  | 01/30/2025 |  |  | 01/13/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596968 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device followed by the implant of an impella rp flex device the next day for bipella mechanical circulatory support and would experience hemolysis and thrombus of the device respectively. the patient was initially implanted with the impella cp device and post implant of the impella cp device, three to four hours later approximately, the patient's labs began to hemolyze. an echocardiogram was completed to assess the impella cp positioning and after two hours of attempting to visualize impella placement via echocardiogram, the physician decided to pull the impella cp back 1 centimeter and drop the flow to p-7 support level. the next day, the impella cp position was checked on echocardiogram. the device was repositioned and the support level was able to be increased to p-9 with no suction. however, severe right ventricular dysfunction was noted on echocardiogram and the patient was taken to the catheterization lab for an impella rp flex implant. thereafter, the patient was transferred to an outside hospital per plan later this day of the impella rp flex implant. the patient arrived combative and on bilevel positive airway pressure. flows on the impella rp flex were noted to be 0.5 liters per minute. the impella rp flex was clotted. a plan was made to discontinue use of the impella rp flex, in which the impella rp flex was explanted at bedside and cardiology was consulted the following morning for percutaneous coronary intervention (pci). additionally, this day, it was noted the patient continued to hemolyze. echocardigram showed the impella cp at 4.1 centimeters from the valve with pigtail in papillary muscle and mitral leaflets were hitting the inflow cage. epinephrine and levoohed were started. labs came back hemolyzed again. continuous renal replacement therapy was initiated. no blood products were given overnight. hemoglobin and hematocrit were noted to be stable. heparin drip was started. echocardiogram overnight showed an ejection fraction of 12%. the following day, the physician was awaiting plans for a pci. the patient was a turn down for surgery. the next day, the high-risk pci was completed and the patient received four stents. weaning of the impella was started. support level was down to p-4. eventually a successful wean off and explant of the impella cp device occurred. the patient was noted to have survived at explant with no further updates. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this is one of two devices associated with the patient's implant experiences. refer to manufacturing report number 1220648-2025-25871 for the rp flex serial number (b)(6) with date of event 14-jan-2025 and identifier ending in (b)(6).
11893 | 1220648-2025-25915 | 21274954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21274954&pc=OZD | 01/15/2025 | 02/12/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a battery failure occurred on the automated impella controller (aic). the aic was removed from service. there was no patient involvement. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) battery failure-red alarm has been completed. console logs from the reported complaint date are consistent with the complaint. there were multiple battery failure (transport connection blown/missing) alarms observed from the reported complaint date. console was returned for evaluation finding the reported battery failure issue was unable to be reproduced. no battery failures or charging issues were present while testing this console. additionally, results of running a batttest on this console showed that there were no issues with the batteries. reported battery failure issue was determined to be use related. it is likely that the console was booted up while the console was connected to ac power and the battery switch was disengaged since the issue occurred upon first boot of the console after it came back from pm.
11894 | 1220648-2025-25916 | 21274958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21274958&pc=OZD | 12/08/2024 | 03/10/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1765953 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, placement signal not reliable occurred. during investigation of the impella cp, the engineer determined that the issue may be due to an optical signal issue with the automated impella controller and requested that it be returned for further evaluation. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) loss of opm data has been completed. data logs were reviewed which showed that while running pump, the console had consistently low snr throughout. additionally, there was a portion of the case where optical sensor intermittently went to -3276.8mmhg as which triggered a placement signal not reliable alarm as seen in the complaint. on the previous case, there was additional evidence of low snr. the console was not returned for investigation. the cause of the placement signal not reliable alarm was not determined since product was not returned.
11895 | 1220648-2025-25917 | 21275231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21275231&pc=OZD | 01/15/2025 | 02/24/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570466 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, there was oozing and a hematoma at the access site. manual pressure was held, and one unit of packed red blood cells was given to the patient. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. h.6 code 4641 was reported incorrectly on the initial manuacturer device report number 1220648-2025-25917.
11896 | 1220648-2025-25918 | 21275397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21275397&pc=OZD | 01/15/2025 | 02/19/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  | 10/31/2025 |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409232 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp pump for the support of a cardiogenic shock and acute myocardial infarction patient. the pump was placed but the team had high purge pressure that warranted troubleshooting with tissue plasminogen activator added to the purge and conversations with the central serous chorioretinopathy. there was no patient harm. | [5 empty] | The investigation has been completed. the pump was not returned for investigation. total case run time was 30 hours. a data log was not provided for the first three hours of pump operation. the provided data log showed no signs related to high purge pressure. the cause of the high purge pressure issue was not determined since the product was not returned and sufficient clinical information, including data logs, was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11897 | 1220648-2025-25919 | 21275840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21275840&pc=OZD | 01/14/2025 | 03/13/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597142 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. at the end of the 6-day support the patient was successfully weaned and the device was explanted with some difficulty. during explant, removal issues were noted as the physician identified an approximately 6-centimeter-long tissue formation at the impella inlet. reportedly, the patient survived the explant and there was no harm to the patient as a result of the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. during the explant of the device, the medical doctor identified an approximately six centimeter long unusual tissue formation at the impella inlet. no product nor data logs were returned. the failure mode was changed from removal issue to thrombus as although the issue was found during removal of the device, there were no issues related to specifically to the removal of the device. as the necessary information was not provided, the root cause of the thrombus was not determined. b.1 revised since the initial manufacturer device report number 1220648-2025-25919 was submitted based on the investigation results. b.2 revised to add a selection since the initial manufacturer device report number 1220648-2025-25919 was submitted based on the investigation results. h.1 revised since the initial manufacturer device report number 1220648-2025-25919 was submitted based on the investigation results. h.6 based on the investigation results, health effect - clinical code, health effect - impact code and medical device problem code were revised since the initial manufacturer device report 1220648-2025-25919 was submitted.
11898 | 1220648-2025-25920 | 21276313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21276313&pc=OZD | 01/15/2025 | 02/26/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568606 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the impella access site had an infection. cultures grew back yeast. additional information on this event was provided from loqi database reporting that the patient was put on empiric treatment with daptomycin, piperacillin, tazobactam, and fluconazole. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source was revised on manufacturer device report 1220648-2025-25920; study was inadvertently omitted.
11899 | 1220648-2025-25921 | 21278009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21278009&pc=OZD | 01/13/2025 | 03/17/2025 |  | 01/30/2025 |  |  | 01/14/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581468 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Complainant in japan reported the patient was on impella cp support and during support, there was a purge blockage alarm that occurred. the pump stopped after the alarm. it resumed immediately. the pump was explanted the following day. it was noted there was no significant change in purge flow or pressure. the patient survived the event. | [5 empty] | The investigation of pump stop has been completed. the product was returned and evaluated. there was significant biomaterial found wrapped around the impeller. no other abnormalities were found. based on the data logs, about an hour after a regular bag change, there was a spike in purge pressure and "purge system blocked" alarm that resolved within a few minutes. about 4 hours later, data logs show a large motor current (mc) spike prior to the "impella stopped - mch" alarms before the pump restarts shortly after. when the pump restarted, the motor current remained higher than expected. the pump continued to run for 12 hours before being successfully weaned and explanted. the root cause of the temporary pump stop was biomaterial since there was significant biomaterial wrapped around the impeller and the data logs indicate ingestion. e4 should have been left blank on the initial manufacturer device report number (b)(4) as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
11900 | 1220648-2025-25922 | 21278028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21278028&pc=OZD | 01/15/2025 | 03/03/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572053 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, impella was malposition but not having alarms. the physician attempted to reposition and came back across the valve. at that point, the decision was made to escalate to axillary 5.5. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
11901 | 1220648-2025-25923 | 21278065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21278065&pc=OZD | 01/15/2025 | 02/20/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1136692 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller¿s (aic) on button was pushed, and the button fell inside the console. the aic was removed. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - kbd/rotary knob issues has been completed. the data logs were not provided. inspection of the console¿s underside revealed a missing actuator plate from the battery transport switch. upon opening the console¿s housing, both the actuator plate and spacer plate were found inside the console, not attached. the transport switch was still operational. it is most likely that the battery transport switch was erroneously disassembled during aic battery pack replacement during prior preventative maintenance (pm) on 2023-12-14, which went undetected during later pm on 2025-01-06. after the aic was returned from the pm, the issue was observed during first use at the site, presumably preventing staff from powering on the console. the root cause of the aic issue was a work instruction / technician error. d.4 revised model number since the initial manufacturer device report number 1220648-2025-25923 was submitted in accordance with updated procedures. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25923 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25923 and a new code was added.
11902 | 1220648-2025-25924 | 21278233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21278233&pc=OZD | 01/15/2025 | 02/18/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599127 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had impella cp support and during support, the patient had access site bleeding and limb ischemia. the intervention for the access site bleeding is unknown. staff monitored the patient's limb ischemia and support continued. the patient expired after 129.08 hours of impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding and limb ischemia could not be determined since the product was not returned and insufficient clinical details were provided. h.6 code 1884 was inadverently omitted from the initial manufacturer device report number 1220648-2025-25924 and was added. code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25924.
11903 | 1220648-2025-25925 | 21278325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21278325&pc=OZD | 01/14/2025 | 04/02/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567219 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The us complainant had placed a 5.5 pump to support a patient admitted in with multiple comorbidities while awaiting transplant. during the support there was a cerebrovascular accident. the team questioned clot ingestion as the heparin had been stopped systemically in preparation for the transplant. the impella remains on for support (with sodium bicarb in purge) despite the harm. | Additional information was received regarding the patient's outcome. the family decided to withdraw care and the patient expired. it is unknown if the use of the impella was related to the patient's death. thus, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25925. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Additional information was received regarding the patients outcome. the family decided to withdraw care and the patient expired. it is unknown if the use of the impella was related to the patients death. thus, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. b.2 revised from the initial manufacturer device report number 1220648-2025-25925 to reflect the patient outcome of death and date of death was added. b.5 revised from the initial manufacturer device report number 1220648-2025-25925 to add the patient outcome of death. d.6b added explant date as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25925. e.1 added fax number as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25925. h.1 type of report was revised from the initial manufacturer device report number 1220648-2025-25925 to reflect the patient outcome of death. h.6 health effect - impact code was revised from the initial manufacturer device report number 1220648-2025-25925 to reflect the patient outcome of death. h.10 added information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25925.
11904 | 1220648-2025-25926 | 21278421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21278421&pc=OZD | 01/15/2025 | 03/07/2025 |  | 01/30/2025 |  |  | 01/15/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589627 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the eu had a 5.5 pump being delivered to support a post aortic valve replacement patient. the team attempted many times and methods but the delivery across the valve failed. the 5.5 implant was aborted. no patient harm has been reported. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the team attempted insertion many times and methods but the delivery across the valve failed. g.1 revised manufacturing site address line 1 as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25926.
11905 | 1220648-2025-25927 | 21278979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21278979&pc=OZD | 01/16/2025 | 03/14/2025 |  | 01/30/2025 |  |  | 01/16/2025 | 01/30/2025 |  |  |  |  | 01/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584073 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Eu complainant reported the patient had an attempted impella cp implant, however it was aborted due to patient anatomy. no further details were provided. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the delivery is aborted due to patient anatomy.
11906 | 1220648-2025-25866 | 21263282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21263282&pc=OZD | 01/13/2025 | 02/26/2025 |  | 01/29/2025 |  |  | 01/13/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573218 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. approximately three days after start of support, bleeding at the access was noted and resolved by the in-house physician placing a suture. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the access site bleeding was determined to be use issue because the bleeding was resolved by adding a suture at the bleeding site.
11907 | 1220648-2025-25867 | 21263381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21263381&pc=OZD | 11/05/2024 | 02/26/2025 |  | 01/29/2025 |  |  | 01/12/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554535 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Cardiac Perforation | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient experienced a left ventricular apex perforation with a possible relationship to the impella device used and also experienced limb ischemia with a probable relationship. additionally, another event was experienced. the patient had an emergent angiogram during which the patient had ventricular fibrillation and brief cardiopulmonary resuscitation with five shocks and bipella¿s placed for hemodynamic support. the impella cp was placed in left groin and impella rp was placed in right leg for hemodynamic support. then patient was then brought emergently to the operating room for triple bypass grafting surgery. during this procedure, a perforation at the left ventricular apex was noted, likely from the wire used while placing the impella cp device. the tissues were quite friable. the perforation was repaired with two, 2-0 prolene sutures with pledgets. following after the start of support, same day of the implant, the patient experienced another event, however. the impella cp leg groin was mottled and cold with no signals at the foot and critical limb ischemia secondary to impella cp placement in profunda femoral artery on the left. additionally, the patient had hemorrhaging from the chest. mass transfusion was started with emergent operating room takeback for chest exploration, washout and removal of left femoral artery impella cp. right ventricular tear was found and repaired. the left femoral artery impella cp was removed, arteriotomy repaired, and 4-compartment fasciotomy was performed due to prolonged lower limb extremity ischemia. the impella cp was successfully removed and the impella rp was left in place. hemostasis was achieved as well as good perfusion of lower left extremity. it was dopplerable distal signal at the posterior tibial artery. excellent pulsatility at present. the left leg fasciotomy closed two days later, after start of support, after surgery for washout and repacking of the chest. the patient recovered from the events with no sequelae. | [5 empty] | The investigation for the limb ischemia and the ventricle perforation has been completed. the product was not returned for investigation. data logs were reviewed and there were no positioning alarms during the entirety of support. clinical details stated that the suspected cause of the ventricle perforation was a guidewire during impella cp access. however, there is no evidence reported to support this. there were no reports of difficulties with access, and the timing of the injury is unknown. due to the product not being returned and limited clinical details, the root cause of the ventricle perforation could not be determined. the root cause of the limb ischemia could not be determined without additional information. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
11908 | 1220648-2025-25868 | 21263444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21263444&pc=OZD | 01/13/2025 | 02/19/2025 |  | 01/29/2025 |  |  | 01/13/2025 | 01/29/2025 |  |  | 08/31/2025 |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. however, on the day of implant, echocardiogram showed left ventricle thrombosis and plan was to monitor and medically treat. in the following two days the patient's condition deteriorated. continuous renal replacement therapy was started, and the patient was shocked due to supraventricular tachycardia. and atrial fibrillation/rapid ventricular response. the patient would eventually expire. the family decided not to escalate care and ultimately withdrew care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
11909 | 1220648-2025-25869 | 21263572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21263572&pc=OZD | 01/13/2025 | 02/19/2025 |  | 01/29/2025 |  |  | 01/13/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580300 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced ectopy and limb ischemia with a demised outcome. the patient was a recent admission for acute systolic and diastolic heart failure (baseline ejection fraction of 25%) and underwent 3-vessel coronary artery bypass graft surgery in the morning. the patient arrived at the unit on two inotropes/vasopressors and throughout the day had increased needs for medical support. cardiothoracic surgery consulted with cardiology for an intra-aorta balloon pump placement, but cardiology opted to implant an impella cp device which was successfully completed. after start of support, some ectopy was experienced and was mostly resolved with minor positioning adjustments. the patient had anti arrhythmic medications prior to impella due to premature ventricular contractions post-operation. additionally, the next day, the patient was noted to have a popliteal pulse. however, pulses were absent in left foot. the patient had a history of raynauds and the team therefore was evaluating such as the cause because this primarily has affected the patient's left foot in the past. also, there was discussion regarding, if needed, placing a new impella cp device from an axillary approach. however, the patient would expire the next day. the family decided to withdraw care. although an alternative cause for the limb ischemia was being considered, the cause of death is unknown, and there is not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | G6 type of report; 30 day was inadvertently not selected manufacturer device report 1220648-2025-25869-1. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation has been completed. the product was not returned. as the necessary information was not provided, the root cause of the ectopy and limb ischemia was not determined.
11910 | 1220648-2025-25870 | 21263615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21263615&pc=OZD | 01/14/2025 | 02/17/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613678 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient developed a hematoma around the access site. manual pressure was held and femstop was applied to the site to help achieve hemostasis. anticoagulants were put on hold until the hematoma resolved. support continued following the intervention and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient movement since the hematoma occurred after the patient was moved to icu. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11911 | 1220648-2025-25871 | 21263637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21263637&pc=PYX | 01/14/2025 | 02/18/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590532 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device followed by the implant of an impella rp flex device the next day for bipella mechanical circulatory support and would experience hemolysis and thrombus of the device respectively. the patient was initially implanted with the impella cp device and post implant of the impella cp device, three to four hours later approximately, the patient's labs began to hemolyze. an echocardiogram was completed to assess the impella cp positioning and after two hours of attempting to visualize impella placement via echocardiogram, the physician decided to pull the impella cp back 1cm and drop the flow to p-7 support level. the next day the impella cp position was checked on echocardiogram, the device was repositioned, and the support level was able to be increased to p-9 with no suction. however, severe right ventricular dysfunction was noted on echocardiogram and the patient was taken to the catheterization lab for an impella rp flex implant. thereafter, the patient was transferred to an outside hospital per plan later this day of the impella rp flex implant. the patient arrived combative and on bilevel positive airway pressure. flows on the impella rp flex noted to be 0.5l/min. the impella rp flex was clotted. a plan was made to discontinue use of the impella rp flex, in which the rp flex was explanted at bedside, and consult cardiology the following morning for percutaneous coronary intervention (pci). additionally, this day, it was noted the patient continued to hemolyze. echo showed the impella cp at 4.1cm from valve with pigtail in papillary muscle and mitral leaflets was hitting inflow cage. epi and levo were started. labs came back hemolyzed again. continuous renal replacement therapy was initiated. no blood products were given overnight, hemoglobin and hematocrit were noted to be stable. heparin drip was started. echocardiogram overnight showed an ejection fraction of 12%. the following day the physician was awaiting plans for the pci. the patient was a turn down for surgery. the next day the high-risk pci was completed, the patient received four stents. weaning of the impella started; support level down to p-4 to eventually a successful wean off and explant of the impella cp device. the patient was noted to have survived at explant with no further updates. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two devices associated with the patient's implant experiences. refer to initial medical device report for impella cp serial number 557877 with date of event 13-jan-2025 and identifier ending in (b)(6).
11912 | 1220648-2025-25872 | 21264138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264138&pc=OZD | 01/13/2025 | 02/18/2025 |  | 01/29/2025 |  |  | 01/13/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025523250 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 device placed via the right axillary artery for a patient admitted in with cardiomyopathy and bridging to transplant. the impella 5.5 was placed as care escalation from an intra-aortic balloon pump. approximately three days after start of support an access site hematoma was noted, and the patient underwent a hematoma evacuation surgical procedure. the pump remains implanted supporting the patient who was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the root cause of the access site bleeding issue could not be determined since the product was not returned and insufficient clinical details were provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25872 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 1888 was reported incorrectly on initial manufacturer device report number 1220648-2025-25872. a new code was addded. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-25872 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
11913 | 1220648-2025-25873 | 21264177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264177&pc=OZD | 01/13/2025 | 03/03/2025 |  | 01/29/2025 |  |  | 01/13/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582391 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support (mcs) followed by extracorporeal membrane oxygenation (ecmo) support seven days later. while on impella mcs, approximately 20 days after start of support, the nursed called to report that the distal end of the sterile sleeve had become detached and was wrapped with tegaderm. the nursed raised the question regarding if there is any risk of infection/contamination with the sterile sleeve being separated and was informed to consult with their provider as abiomed has no recommendations per the instructions for use for a separated sterile sleeve. further discussion noted that the nurse and team were advised to be conscience of the potential risk of infection when/if the impella needs repositioned. the patient would expire four days later. the family made the decision to withdraw care. the cause of death was not provided, and there is not enough information to exclude the impella device used as a contributing factor to the patient's outcome. | [5 empty] | The investigation of product damage (sterile sleeve damage) has been completed. due to limited clinical details and no product return, the root cause of the sterile sleeve damage could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25873 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11914 | 1220648-2025-25874 | 21264179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264179&pc=PYX | 01/14/2025 | 02/18/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025483289 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | LEFT FEMORAL SWAN. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient was implanted with an impella rp flex pump for mechanical circulatory support. following successful implant, a swann was inserted and the pump suddenly stopped working unexpectedly. both the swann and rp flex were explanted as a result of the failure. there was no known patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. the product was not returned for analysis. the data logs confirm pump stop alarms. the logs show motor current spike with speed deviation and a drop in pump flow. it also shows unsuccessful restart attempts. the clinical details mention interventionalist placed another left femoral swan, after placing swann there was pump stoppage. the root cause of the pump stop was most likely due to an interaction with another device. e4 should have been left blank on manufacturer device report. d10 concomitant medical products was inadvertently omitted from manufacturer device report.
11915 | 1220648-2025-25875 | 21264556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264556&pc=OZD | 01/16/2025 | 03/03/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1538795 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the automated impella controller (aic) had a yellow controller error. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console logs from the reported day of event are consistent with complaint and show controller error alarm triggered during the clinical procedure. the real time (rt) logs confirmed that all signals received from optical bench. the controller error and associated with it momentary loss of placement signal was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. d9 date device returned to manufacturer added. d4 model number corrected on manufacturer device report number. d4 udi number revised as it was inadvertently omitted on the initial manufacturer device report number. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
11916 | 1220648-2025-25876 | 21264572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264572&pc=OZD | 01/16/2025 | 03/05/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592237 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support placed to support a patient admitted in and escalated from cp and extra corporeal membrane oxygenation. the pump had high purge pressures which prompted troubleshooting. the purge cassette was replaced but the high purge pressure alarms were not resolved. the team used tissue plasminogen activator in the purge and the pump remained on for support after 41 days now as care/comfort versus transplant being discussed. the patient expired when the patient was made care/comfort after day 42. | [5 empty] | The investigation for the high purge pressure has been completed. no product was returned for evaluation. data logs were reviewed and long term log at time of ¿high purge pressure¿ alarms showed a rise in purge pressure over a period of approximately 30 minutes prior to alarms. no changes in motor current were observed at time of alarms. full long term logs of the case showed no motor current spikes prior to high purge pressure alarms. purge pressure remains high for the remainder of the case. clinical details noted that the patient was on support for 43 days before "high purge pressure" alarms were triggered. the root cause of the high purge pressure was most likely due to biomaterial since there was no motor current spike preceding the purge pressure rise. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25876. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25876 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11917 | 1220648-2025-25877 | 21264637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264637&pc=OZD | 12/06/2024 | 10/20/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1505149 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant had placed a 5.5 pump to support a patient as escalation from the cp. the pump was placed but lost optical signal. the loss of signal did not prompt explant. the pump remained on for another day until the patient had care withdrawal and expired. during the investigation of the pump, the team determined the automated impella controller warranted investigation. | The us complainant had placed a 5.5 pump to support a patient as escalation from the cp. the pump was placed but lost optical signal. the lost of signal did not prompt explant. the pump remained on for another day until care was withdrawn and the patient and expired. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | After further review there is currently no allegation of a malfunction or serious injury associated with the automated impella controller (aic). a regulatory report is no longer necessary. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. console logs from the reported day of event were analyzed within pump complaint (b)(4), and indicated that the issue was most likely caused by the console. console had not been returned for investigation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
11918 | 1220648-2025-25878 | 21264655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264655&pc=OZD | 01/15/2025 | 03/14/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583844 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, bleeding was noted at the access site. the impella cp was explanted and was replaced with a new impella. the patient survived the event. | [5 empty] | The investigation for the access site bleed has been completed. the returned product was visually inspected and the inflow seems to be slightly deformed and no other issues are alleged against the impella. the root cause of the access site bleeding issue was not determined since insufficient clinical details provided. added d9 device return date. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11919 | 1220648-2025-25879 | 21264720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264720&pc=OZD | 01/16/2025 | 02/26/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591639 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient that was on impella cp support lost pulses in the impella leg with a mottled appearance. the staff discussed escalating the patient to an impella 5.5. the staff decided to possibly wean extracorporeal membrane oxygenation, place a jump graft, and monitor. the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11920 | 1220648-2025-25880 | 21264828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264828&pc=OZD | 01/14/2025 | 02/28/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503016 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during the explant, the operating room table was adjusted and damaged the connection of the impella 5.5 to the automated impella controller. the pump stopped. the patient was urgently placed on bypass and the impella 5.5 was removed. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. data logs were reviewed, and leading up to (b)(6) 2025, there were no abnormalities in any of the pump metrics. on the reported date, an alarm suddenly was triggered, which is a signature of an electrical failure, but there was no motor current spike along with it. additionally, "impella disconnected (while running)" triggered within the same second. real time logs at the time of the event were not available. logs showed that seconds after the pump stop alarm, a pump was plugged in and unplugged within a second, not enough time for the console to register the pump's serial number or the impella defective alarm to trigger. clinical details reported that while the operating room table was adjusted, the connection between the pump and the console was damaged. based on data log review and the clinical details provided, the pump stop was determined to most likely be an electrical open. the location of the damage, however, could not be determined without the returned product.
11921 | 1220648-2025-25881 | 21264853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21264853&pc=OZD | 01/16/2025 | 03/18/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  | 01/31/2026 |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468926 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the impella 5.5 was primed, but there was no placement signal available. the pump was replaced. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. after priming of the pump and insertion into position, there was no ao placement signal despite position confirmed by echo. pump was returned. data logs could not be recovered. no abnormalities were found upon initial inspection. pump was run for 10 minutes at nominal bench and operating parameters, without reproducing placement signal unreliable alarms. the cause of the optical signal issue was not determined since this issue could not be reproduced and no data logs were returned. d9 date device returned to manufacturer added.
11922 | 1220648-2025-25884 | 21265176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21265176&pc=OZD | 01/14/2025 | 02/24/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  | 02/28/2026 |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465658 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed oozing from the access site. the perclose was pushed and the bleeding resolved. the next day, the site was oozing with drainage. a sandbag was placed, and heparin was held. bleeding was resolved and the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the root cause of the access site bleeding is determined to be anti coagulation management because the partial thromboplastin time (ptt) values were very high(200) than the therapeutic range during the time of bleeding and the heparin was kept on hold to bleeding was resolved later.
11923 | 1220648-2025-25882 | 21265495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21265495&pc=OZD | 01/16/2025 | 03/10/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597008 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following initial placement of the impella device, guidewire removal was met with resistance. both the pump and the guidewire were pulled back and a kink/twist was visible near the motor housing. the .018 guidewire was also curled up. as a result of the damage, the decision was made to replace the devices. there was no patient harm. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and product damage (guidewire damage - great vessel) has been completed. the data log review was not required for this case. product was returned for investigation. the returned pump had a catheter kink at the motor housing. additionally, bending damage was noted on the guidewire, but it was neither broken nor stretched. no further investigation was required. the pump operated as expected in a bucket of water. according to the given clinical report, some resistance was encountered during the wire removal after the pump was placed in the ventricle. both the 5.5 mm wire was pulled back, and the kink was visible on the catheter. the product damage (guidewire damage - great vessel) failure mode was removed since the guidewire was only slightly bent without a full or partial fracture. additionally, the product damage (cannula kink) failure mode was added to address the catheter kink damage on the returned product. the cause of the product damage issue was most likely use related. the cause of the delivery issue was most likely use related.
11924 | 1220648-2025-25885 | 21265731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21265731&pc=PYX | 01/16/2025 | 02/18/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025503251 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella rp flex. the wire was unable to be advanced via the right ventricular failure. the decision was made to manually insert the impella via the right atrium. manipulation was unable to be done. the case was aborted. the patient remains stable. | [5 empty] | The investigation of delivery issues has been completed. no product was returned for analysis. the clinical details mention that the insertion was aborted due to patient anatomy. the root cause of the delivery issue was most likely due to patient condition. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp flex with smartassist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ impella rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
11925 | 1220648-2025-25886 | 21265978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21265978&pc=OZD | 01/16/2025 | 02/26/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566770 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had attempted to place an impella cp for support of a patient admitted in for surgery and was found to have a need for high-risk percutaneous coronary intervention. the native anatomy was accessed, via single access, but the pump was unable to be delivered. the implant was aborted. the patient¿s outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy.
11926 | 1220648-2025-25887 | 21266031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21266031&pc=OZD | 01/14/2025 | 02/25/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572146 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was inserted, access site bleeding was noted. the physician controlled the bleeding with angle matching, tightening of perclose sutures, and forward tension sutures. the access site bled through three pressure dressings overnight. small sandbag was applied. anticoagulation held for 4 hours and now hemostatic. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed has been completed. the impella as not returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the ptt values were also in therapeutic range during the time of bleed and the hemostasis was achieved by manual pressure. hence, the root cause is not determined.
11927 | 1220648-2025-25883 | 21266497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21266497&pc=OZD | 01/14/2025 | 03/05/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568534 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Vascular Dissection | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during utilization of an impella 5.5 pump for mechanical circulatory support. delivery of the device proved complicated for which multiple attempts were required to ultimately place the pump. during the complicated delivery, a contrast injection showed a potential isolated dissection in the axillary artery for which an altered approach was required. no intervention was necessary in regards to the dissection. upon completion of support, the staff attempted to explant the pump at which point the patient coded. multiple chest compressions were performed in an attempt at resuscitation, but ultimately the patient expired due to peri-procedural myocardial infarction. | [5 empty] | The investigation of vf/vt/af/at and vascular damage - great vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
11928 | 1220648-2025-25888 | 21266568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21266568&pc=OZD | 01/14/2025 | 03/05/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593445 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following successful implant of an impella cp pump, the device experienced persistent suction alarms. an x-ray noted a kink in the pump cannula just below the outlet cage. attempts to reposition the device failed to straighten the kink and the decision was made to replace the device. there was no resulting patient harm associated with the failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the data logs confirmed low pump flow when pump started & remained to the rest of the case. suction alarms were seen. during the visual inspection it was discovered there was a kink on can about 15 mm from the bonding site of can & of cage. the root cause of low flow was due to kinked cannula.
11929 | 1220648-2025-25889 | 21266747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21266747&pc=OZD | 01/14/2025 | 03/03/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547326 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had impella cp support and after the implant, the cannula was found kinked. impella flows dropped and suction alarms on automated impella controller displayed, despite multiple attempts to unkink cannula, decision was made to remove impella and place new catheter. on removal of this device, impella kink clearly visible under fluoroscopy. there was no reported harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the low pump flow has been completed. clinical details state that the device experienced kinking once the peel away sheath was removed. this was reported by the physician and confirmed on the returned device during investigation. per the data logs, the kink caused the pump to flow in the field at around 1l/min at p-4, less than the expected 2.0-2.5l/min, while suction alarms presented. they also show no motor current spike. the cause of the low flow was most likely a kinked cannula. added: d9 device return date.
11930 | 1220648-2025-25890 | 21267416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21267416&pc=OZD | 01/12/2025 | 02/28/2025 |  | 01/29/2025 |  |  | 01/12/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584929 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced high purge pressure and positioning issues during support. the impella was inserted via the right axillary artery, started on support level p-2 and slowly increased to p-6 with flows at 3.4lpm. the patient was transferred to the unit on support in stable condition awaiting heart and liver transplants. approximately 19 days after start of support, purge pressure high/purge system blocked alarm triggered. consequently, the purge fluid was changed to tissue plasminogen activator (tpa), the team was provided steps to disable the audio alarm and advised to closely monitor the motor current for upward trends. following, two days later, the physician called for assistance with repositioning the impella due to aortic placement signal waveform occasionally ventricular with suction alarms at p-7 support level. the impella initially measured deep at approximately 7 cm and was pulled back just under 2 cm with echocardiogram guidance. the flow returned to p-7 with all waveforms and numbers within normal limits. there was no report or indication of adverse effects to the patient as a result the events, and the patient remains on impella 5.5 mcs. | [5 empty] | The investigation for the high purge pressure and the positioning issues has been completed. no product was returned for evaluation. data logs were reviewed and lt log at time of ¿high purge pressure¿ alarms shows multiple instances of purge pressure increases over a period of multiple days with step ups in purge pressure and some being able to resolve to lower purge pressure values. prior to start of purge pressure increases, a prolonged bag/cassette change of 2 minutes and 12 seconds with a drop in purge pressure at time of bag/cassette change. ¿high purge pressure¿ alarms are triggered with last rise in purge pressure over a period of approximately 6 hours. purge pressure remains high for approximately 4 hours before slowly resolving. clinical details noted that tpa was not used and no heparin was added to purge solution. the root cause of the high purge pressure was most likely a use related issue as a prolonged bag change caused multiple instances of purge pressure step ups prior to purge pressure alarms being triggered. data logs were reviewed and lt logs of full case show suction alarms occurring towards end of case. ¿impella in ventricle¿ alarms also triggered while suction alarms are occurring. pump flow is lower than expected for p-level. clinical details noted that pump was experiencing suction alarms while at p-6 and echo showed that pump was interfering with the mitral apparatus. it was not noted if pump was properly fixated at time of alarms or if patient's condition was suboptimal at time of suction alarms. the root cause of the positioning issues cannot be definitively determined as limited clinical details were provided. added d5-operator of device was blank should have been health care professional. | The impella device was not received from the customer as it appears the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11931 | 1220648-2025-25892 | 21268274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21268274&pc=PYX | 01/16/2025 | 03/17/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025570387 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a rp flex placed to support a patient admitted in post heart transplant. the patient was additionally supported by extra corporeal membrane oxygenation. the pump was supporting for 4 days when the console alarmed and purge pressure / pump flows prompted the team to troubleshoot with tissue plasminogen in the purge. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of high purge pressure has been completed. the device was returned for evaluation. the data logs showed a high purge pressure event occurred four days into care that was resolved on the same day a tissue plasminogen activator (tpa) purge solution was advised. no evidence of biomaterial on the impella, cannula, nor purge gap was found. the device was run in a bucket at p-9 for 10 minutes with the high purge pressure unable to be reproduced. the cause of the high purge pressure was most likely biomaterial due to data log review and the clinical use of tpa that was used the same day the high purge pressure resolved. low pump flow was added as additional failure mode. the data logs were reviewed and it was found that the mean flow was trending down and the placement signal was trending up. these trends appear to occur independently from the high purge pressure event as they begin to trend before the event and continue after that event was resolved. no visual abnormalities were found on the returned device. the device was run in a continuous flow loop for 60 hours but the drift was unable to be reproduced, indicating the dp sensor to be functioning properly and low flow not reproduced in the lab. the cause of the low pump flow was not determined due to lack of clinical details. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25892 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11932 | 1220648-2025-25893 | 21268401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21268401&pc=OZD | 01/16/2025 | 03/17/2025 |  | 01/29/2025 |  |  | 01/16/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556108 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the eu had a cp pump placed to support a patient when the pump had to be explanted due to a kink in the cannula. pump flow then dropped. no harm or injury noted, and pump removed and noted to be successful on the 2nd pump. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of product damage (cannula kink) has been completed. returned complaint pump visually inspected. cannula kink observed near out let. no biomaterial or broken blade found. pump connected to console and run for 10 mins at p-9 to reproduce low pump flow issue. flow was within specified limits and no low flow reproduced. failure mode: product damage (cannula kink) changed to failure mode: low pump flow since cannula kink observed under fluoroscopy and pump flow was 1l on auto mode and suction on p-2. the cause of the low pump flow issue was due to a kink in the cannula which likely was related to calcification vessel conditions, which led to this cannula kink during insertion. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25893 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11933 | 1220648-2025-25891 | 21268899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21268899&pc=OZD | 01/14/2025 | 02/18/2025 |  | 01/29/2025 |  |  | 01/14/2025 | 01/29/2025 |  |  |  |  | 01/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545743 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to needing an high-risk percutaneous coronary intervention. the doctor initially wanted to leave the patient on impella support after the procedure however, during the explant of the repositioning sheath, it was noted that significant hematoma was present after reposition was advanced into position. several attempts were made to reposition the sheath however, the bleeding and hematoma persisted. patient was stable requiring no pressors or ventilatory support. the doctor decided to rewire vessel via reposition sheath, explant impella, and implant long impella peel away sheath. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the root cause of the access site bleeding issue cannot be determined since the product was not returned and insufficient clinical details were provided. e4 should have been left blank on manufacturer device report 1220648-2025-25891.
11934 | 1220648-2025-25837 | 21251408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251408&pc=OZD | 01/10/2025 | 02/18/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573265 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support and experienced hematoma and bleeding around the axillary site. the patient had prolonged cardiac arrest and route to the hospital reportedly received continuous chest compressions by lucas device and was intubated. the patient underwent a right heart catheterization and was found to have severe triple-vessel disease with chronic total occlusion of the right coronary artery and subacute complete occlusion of the proximal left anterior descending artery. an intra-aortic balloon pump (iabp) was placed to provide additional support due to high pressor requirements. however, given the lack of improvement, it was decided to escalate to an impella. an impella 5.5 device was successfully implanted and the iabp was removed. the patient was transferred to the unit noted to be in stable condition. while on support, two days later, hematoma and bleeding were noted around the axillary site, around the jackson-pratt drain. consequently, heparin was held, dressing changed, and fresh frozen plasma was administered, and the issued resolved. support continued and 12 days later it was noted the patient was planned for a durable left ventricular assist device, and the impella 5.5 was explanted. the patient was noted to have survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely anticoagulation management since bleeding was resolved by withholding the heparin.
11935 | 1220648-2025-25838 | 21251467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251467&pc=OZD | 09/22/2023 | 02/24/2025 |  | 01/28/2025 |  |  | 01/09/2025 | 01/28/2025 |  |  | 07/31/2025 |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377890 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 46 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via the impact study a patient experienced hemolysis one day following the implant of an impella 5.5 device for mechanical circulatory support while undergoing a high-risk coronary artery bypass graft surgery. the patient had a lactate dehydrogenase level of 608 units per liter and was anemic. consequently, two units of red blood cells were administered. per the study report, the patient recovered from the event and the impella was explanted six days later with a survival outcome. the event was deemed as a probable relationship to the impella device used. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. based on the clinical details, the cause of the hemolysis issue most likely was patient condition since blood volume resolved the hemolysis.
11936 | 1220648-2025-25839 | 21251505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251505&pc=OZD | 10/19/2024 | 02/17/2025 |  | 01/28/2025 |  |  | 01/07/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025516160 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient experienced an atrial fibrillation with a rapid ventricular response during impella cp support. a 150mg amiodarone bolus was given to the patient in response to the condition. the event was initially deemed possibly related to the impella device. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25839. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
11937 | 1220648-2025-25840 | 21251520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251520&pc=OZD | 01/09/2025 | 02/24/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1086076 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while connecting an impella cp for a patient support, that a defective impella catheter alarm occurred. they unplugged and attended to connect the pump to the console several times with the same result. they grabbed a different console and the pump connected and primed for normal operation. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Data analysis: console logs from the reported day of event are consistent with complaint and show "impella defective" alarm (#3, error reading eeprom) when the pump was connected on 9th jan 2025 at 19:06:43 with in 17 seconds. additionally, after which the pump was connected twice and similar alarms were triggered both times. device analysis: the console (b)(4) was tested upon arrival at fs, and the reported failure mode was successfully reproduced when the pump was connected. when the impellatronic was replaced with a known good unit, the aic functioned properly, with both the pump and purge cassette connected and tested for several hours without issue. a thorough visual inspection of the internal components was conducted and no other issues were found. root cause: the root cause for the pump not detected "impella defective (error reading eeprom) - alarm issued" was identified as a malfunction in the impellatronic board.
11938 | 1220648-2025-25841 | 21251637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251637&pc=OZD | 01/10/2025 | 01/31/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577611 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of an impella cp device in a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful due to the patient's cardiac anatomy. the device was unable to pass the aortic arch after several attempts and the case was aborted. additionally, once the impella cp device was removed, a crack in the peel-away sheath was noted. the hemostatic valve was leaking and upon removal, the sheath had split down the center. the access site was closed with two perclose, manual pressure and reversed with protamine. there were no adverse effects to the patient. the physician noted a reconsult with the team on the patient's options as the patient is a surgical turndown as well. the patient was reported to be stable following the event. | [5 empty] | The introducer was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The introducer was received from the customer and the evaluation of the device is underway. upon investigation closure, a supplemental manufacturer device report will be filed. h.10 in manufacturer device report number stated the introducer was not received. however, the introducer was received for investigation.
11939 | 1220648-2025-25842 | 21251755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251755&pc=OZD | 01/09/2025 | 03/14/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575400 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp device was being placed for the support of a patient admitted with acute myocardial infraction. the team had a failure at the setup that prompted the team to replace the pump. the lumen failed to be removed which would have prevented the pump to start. there was no patient harm as a result of the event. | [5 empty] | The investigation for the pump did not start issue has been completed. clinical details state that during delivery, red lumen could not be removed after the wire was inserted. pump was returned and inspected. no data logs were provided. investigation revealed the red lumen was kinked inside the cannula, just before the opening of the outflow cage. the cause of the delivery issue was a red lumen removal issue due to red lumen found kinked in device and provided clinical details. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
11940 | 1220648-2025-25843 | 21251790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251790&pc=OZD | 01/10/2025 | 02/28/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579014 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the impella was inserted further, and the patient went into ventricular tachycardia (vt). the patient was shocked and support continued. additionally, the patient had access site bleeding with a hematoma at the access site. pressure was applied and sutures were placed. support continued. furthermore, the pigtail of the pump was caught in pap/chordae tendinea. the doctor decided to keep placement where it is at. the pump was eventually explanted and replaced with an impella 5.5. the patient survived the event. | [5 empty] | The investigation for the ventricular tachycardia (vt), access site bleed, and positioning issue has been completed. product was not returned for review. the root cause of the access site bleeding issue was most likely use related per clinical review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the vt could not be determined without additional information. corrections to the initial mfr report: d4 updated model number. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11941 | 1220648-2025-25844 | 21251814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251814&pc=OZD | 01/10/2025 | 09/25/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571164 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device. once implanted, suction alarms occurred, and the performance level was reduced to p-2. at this point an impella in position in aorta red alarm triggered. performance level was increased to p4 to assess the left ventricular waveform and at p-4, the red alarms resolved but suction was initiated. at p-3, there was intermittent suction. the plan was to give intravenous fluids and assess with echocardiogram. the echocardiogram was completed and despite no significant change in motor current pulsatility and "good" position confirmed with an echocardiogram, the impella was giving the position alarm "position in aorta" when turned from p-3 to p-2. consequently, the impella ran in auto mode for duration of procedure providing 3.6 l/min of support. patient sent to unit for further recovery with tentative plan to explant the following day. the following day, however, approximately 24 hours later, the nurse reported a groin site bleed. no further information was provided. therefore, intervention for the bleeding is unknown. the clinical representative was unable to confirm if blood products were given or if heparin was withheld. the patient was successfully weaned and the impella cp device was explanted this day. the patient was noted to have survived following the explant. | [5 empty] | The investigation for the access site bleed and low pump flow has been completed. the returned product was inspected and there were no physical abnormalities. the optical parameters were within specification. the data logs show a motor current spike with speed deviation characteristic of biomaterial ingestion. moments after there was a purge pressure spike. the "impella position in aorta" alarms and "impella in ventricle" alarms occur after the motor current spike indicative of the biomaterial possibly interacting with the optical sensor causing false positioning alarms. the logs show suction alarms at p2 and p3. clinical details mention that pump position was confirmed with echo. the root cause of the optical signal issue was most likely due to ingested biomaterial. the returned pump was run in the lab and low pump flow did not reproduce. the root cause of the low pump flow issue was not determined. as limited clinical information was provided, the root cause of he access site bleeding - major was not determined. corrections to the initial mfr report: g1 mfr site name and address updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | H6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-25844. d10 concomitant medical products and therapy dates updated. e4 should have been left blank in the initial report.
11942 | 1220648-2025-25845 | 21251818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251818&pc=OZD | 01/10/2025 | 02/28/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591364 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient with respiratory distress and acute myocardial infarction / cardiogenic shock. the pump was removed and replaced after minutes due to low pump flows. the pump was replaced and support proceeded without harm. the patient survived the event. | [5 empty] | The investigation of low pump flow has been completed. data logs confirmed low pump flow when pump started and remained thru the rest of the case. suction alarms and placement signal declining trend were seen during the case. logs also showed pump was in improper position near the end of the case. the product was not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25845 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
11943 | 1220648-2025-25846 | 21251842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21251842&pc=OZD | 01/11/2025 | 03/10/2025 |  | 01/28/2025 |  |  | 01/11/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563616 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a partially-mobile thrombus was discovered via echocardiogram during impella cp support. due to the noted clot, plans were made for clot removal and pump escalation to impella 5.5 support. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date was revised as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25846. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25846 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11944 | 1220648-2025-25847 | 21252092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21252092&pc=OZD | 01/10/2025 | 02/24/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591658 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and post implant while in the procedure area, it was noted the patient developed a hematoma at the impella insertion site which was noted to be "easily" managed without intervention. a fem-stop was applied, and two units of packed red blood cells were administered. the following day, it was noted the patient's hemodynamics drastically improved overnight with volume replacement. fem-stop off and continued to require no iv drips; the impella cp device was successfully weaned and explanted the next day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The root cause of access site bleeding is determined to be anti coagulation management because the act values were higher than 250 during the time of bleed.
11945 | 1220648-2025-25848 | 21252529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21252529&pc=OZD | 08/29/2024 | 03/07/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508922 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via the impact study a patient experienced an infection approximately three days after the study device, impella 5.5, was explanted; medication was administered. per the study report there was an increased infection parameter and no signs of cause yet known. urine sample was tested and most likely pleura empyema infection. however, before a diagnosis was determined, the patient (unknown age, race and ethnicity) had discharged himself from the facility. no further information was provided. the event was deemed possibly related to the impella and surgical procedures. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 manufacturing site name and address was revised to reflect the correct site on manufacturer device report number 1220648-2025-25848.
11946 | 1220648-2025-25849 | 21252587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21252587&pc=OZD | 01/11/2025 | 02/18/2025 |  | 01/28/2025 |  |  | 01/11/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592342 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was escalated from impella cp to impella 5.5 due to limb ischemia. the patient was on impella 5.5 support and despite the impella cp being removed, the leg continued to be ischemic due to clots throughout the leg. the patient was taken emergently to the operation room for femoral-popliteal bypass. after, the patient became increasingly acidotic, went into rhabdomyolysis, and care was withdrawn, the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11947 | 1220648-2025-25850 | 21252619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21252619&pc=OZD | 01/11/2025 | 02/25/2025 |  | 01/28/2025 |  |  | 01/11/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551827 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient received a septic workup completed with the addition of broad-spectrum antibiotics being added. additionally, suction was noted on impella. correction positioning verified trans-esophageal echocardiography (tee) and no clot visualized. volume status was evaluated, and additional volume was given. multiple repositioning attempts were made with no success. the patient¿s jackson-pratt drain had multiple clots when being emptied by nursing post-operation back on the unit. decision was made to leave impella at p3, access for hemolysis. pump was explanted. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the thrombus, low pump flow, and infection has been completed. returned complaint cp pump was visually inspected. cannula kink was near outlet observed. no biomaterial was observed around impeller. no broken blade observed. returned complaint pump connected to aic to check on low pump flow. no low pump flow reproduced and flow was within specified limits. data logs were reviewed finding no motor current (mc) spike outside of p-level changes observed. many suction alarms observed throughout the case. flow was stable and within specified limits. pump turned down for repositioning, suction alarms still occur after with flows on low end but in range. positioning alarms occurred after then which is not resolved the reported suction alarms after repositioning based on clinical data. the root cause of the low pump flow was not determined due to insufficient data logs and low flow issue not reproduced in lab. the root cause of the thrombus and infection could not be determined without additional information.
11948 | 1220648-2025-25851 | 21252690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21252690&pc=OZD | 01/13/2025 | 02/19/2025 |  | 01/28/2025 |  |  | 01/13/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537966 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient developed ischemia in their right lower extremity. the condition was accompanied by a pale limb that was cool to the touch with no distal flow noted on fluoroscopy. the patient was taken to the operating room where the pump was explanted and a clot was removed from the superficial femoral artery. distal flow was returned to the leg following the intervention. the patient remained supported via extracorporeal membrane oxygenation and survived the event. | [5 empty] | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the thrombus and limb ischemia was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
11949 | 1220648-2025-25852 | 21252748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21252748&pc=OZD | 01/12/2025 | 02/25/2025 |  | 01/28/2025 |  |  | 01/12/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568614 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump being delivered to escalate care from a cp when the 5.5 pump failed to deliver. the acute myocardial infarction / cardiogenic shock and covid positive patient instead was placed on extra corporeal membrane oxygenation when the 5.5 failed. the patient remains stable. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the graft anastomosis.
11950 | 1220648-2025-25853 | 21253180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21253180&pc=OZD | 01/14/2025 | 02/27/2025 |  | 01/28/2025 |  |  | 01/14/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597127 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed oozing from the impella access site. femstop was placed and anticoagulant was held. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be anticoagulation management because anticoagulant was held and hemostasis was achieved. d.6b revised as explant date was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25853. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11951 | 1220648-2025-25854 | 21253551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21253551&pc=OZD | 01/08/2025 | 02/19/2025 |  | 01/28/2025 |  |  | 01/08/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589905 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced oozing from impella cp access site due to a crack in the introducer. the introducer was replaced with a shorter sheath to resolve the issue. support remained uninterrupted and there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review was performed. as per quality notification (qn) introducer from batch dor9273740 failed the post sterile inspection check for unit has a stain on tyvek and was reworked and accepted. bounding is complete and is specific to 1 unit of batch dor9273740. no additional containment was required because each receipt requires 100% visual inspection per the quality inspection plan for the material. material review board (mrb) confirmed this batch was 100% inspected. qn has been put in-process. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
11952 | 1220648-2025-25855 | 21253875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21253875&pc=OZD | 01/09/2025 | 03/12/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575312 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was unable to put the impella cp 0.018 guidewire in the easy guide lumen. the support was aborted and there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the pump was returned for inspection and the guidewire was not returned. no abnormalities were noted with the pump upon inspection. the root cause of delivery issue is not determined because of insufficient clinical details and neither the guidewire or the red lumen were returned for inspection.
11953 | 1220648-2025-25856 | 21254056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21254056&pc=OZD | 01/10/2025 | 02/28/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528095 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and the pump would not pass over the innominate artery. due to patient anatomy, staff were unable to advanced the catheter despite multiple attempts including balloon angioplasty to affected area. the implant was aborted and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The pump was returned for inspection and no abnormalities were not with the pump upon inspection. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the aorta due to porcelain condition of the aorta. d9 date device returned to manufacturer added.
11954 | 1220648-2025-25857 | 21254429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21254429&pc=OZD | 01/11/2025 | 02/26/2025 |  | 01/28/2025 |  |  | 01/11/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786113 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella support when the automated impella controller (aic) turned off. the aic was replaced and impella support continued. the patient was on impella support for 9.08 hours before expiring. | [5 empty] | The impella device has received from the customer, upon completion of the investigation, a supplemental mdr will be filed. | The investigation for the console (aic) shutdown issue has been completed. console logs from the reported day of the event are consistent with the complaint, showing an unexpected shutdown followed by "unexpected shutdown" alarm after the boot-up. there were no warning signs or precursors prior to the unexpected shutdown, and no software processes showed signs of failure. after the shutdown, the console¿s software automatically restarted, with no interruption in pump operation. the console was returned for evaluation. during inspection the issue could not be reproduced during testing with the pump. no unexpected shutdowns or performance issues were observed. a thorough visual inspection of the internal components was conducted, but no physical problems were found. there was no interruption in the supply were detected between the pbm and the 4-in-1 assembly. the root cause of the intermittent, unexpected controller shutdown could not be determined, as there was no clear indication of the cause, and the issue could not be reproduced. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11955 | 1220648-2025-25858 | 21255386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21255386&pc=OZD | 01/16/2025 | 02/19/2025 |  | 01/28/2025 |  |  | 01/16/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586759 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Life Threatening | Disability |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was escalated to an impella 5.5 from an impella cp and during impella 5.5 support the patient experienced an ischemic cerebral vascular accident (cva) with a demised outcome. the patient had chest pain for one week, was taken to the emergency department, and diagnosed with st-segment elevation myocardial infarction. the patient was transported for an emergent percutaneous coronary intervention (pci). one drug-eluting stent (des) was placed in the right coronary artery (rca), and the patient decompensated leading to cardiac arrest. in between resuscitation attempts, the decision was made to implant the impella cp. once on mcs, the patient stabilized. the pci continued with a second des to the rca and two des to the left circumflex artery. the impella cp site was dressed with minimal oozing and no hematoma present. the patient continued on impella support and was prepped for airevac to an outside hospital. a couple of hours later, however, a "large" hematoma around the impella access site was noted. heparin rate was reduced to manage the condition, and the site was monitored. following, approximately five days later, the patient was escalated to an impella 5.5. the patient was continued on impella and extracorporeal membrane oxygenation supports. three days after start of impella 5.5 support, a computed tomography scan was completed and revealed an ischemic cva. an angiogram confirmed a right common carotid occlusion of unknown age. the patient experienced a neuro status decompensation as a result of the cva and care was withdrawn. the patient subsequently passed away. a cause for the cva is unknown; the nurse reported the hospital will perform an autopsy to help determine the cva cause. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two devices associated with the patient implant experience. refer to initial medical device report for impella cp serial number (b)(6) with date of event 08-jan-2025 and identifier ending in (b)(6). | The investigation of stroke/cerebrovascular accident (cva) has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2025-25858 and were added.
11956 | 1220648-2025-25859 | 21256915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21256915&pc=OZD | 07/06/2024 | 02/26/2025 |  | 01/28/2025 |  |  | 01/09/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501658 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported via impact documentation that a patient on impella 5.5 support experienced acute blood loss anemia. it was alleged that the condition had a probable relationship to the implant procedure. the patient was given one unit of packed red bleed cells to treat the condition. the patient's condition was listed as stable following the intervention. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. data logs are not relevant to the complication and the product was not returned for investigation. vascular damage was reported through the impact study. the source of the bleed was unknown, but it was believed to likely be in relation to the impella. the patient had an aortic valve replacement procedure just before impella insertion. due to limited clinical details, the root cause of the vascular damage could not be determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25859 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25859. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25859. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
11957 | 1220648-2025-25860 | 21256929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21256929&pc=OZD | 01/10/2025 | 02/28/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025596339 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant due to being escalated from an impella cp. the patient had a torturous anatomy and the catheter kinked. it was a difficult insertion, unable to implant first 5.5 and case aborted. physician altered the technique and new 5.5 implanted. there was a period of time (failed implant) when patient was not on support and patient placed back on extracorporeal membrane oxygenation as the patient had significant hemodynamic change. the patient would expire 6 days later. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the product was received for investigation and the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25860 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11958 | 1220648-2025-25861 | 21256941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21256941&pc=OZD | 01/10/2024 | 03/06/2025 |  | 01/28/2025 |  |  | 01/10/2025 | 01/28/2025 |  |  | 11/30/2025 |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430825 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pneumonia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a 7 day support with a impella 5.5 pump for a patient admitted in for a high risk surgery. the support was completed and the impact study team notified of a adverse event with "possible" relationship to the impella procedure. "pneumonia as result of prolonged ventilation after rethoracotomy." the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event inadvertently reported incorrectly on manufacturer device report number 1220648-2025-25861. d6a/d6b implant/explant dates inadvertently reported incorrectly on manufacturer device report number 1220648-2025-25861. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25861 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11959 | 1220648-2025-25862 | 21256944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21256944&pc=OZD | 01/11/2025 | 02/28/2025 |  | 01/28/2025 |  |  | 01/11/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580374 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the flows were suboptimal. position was checked and the pump was caught in papillary muscle. the doctor attempted to reposition and pump popped across aortic valve. the pump was explanted. no patient harm has been reported. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the positioning issues and low pump flow has been completed. product and data logs were not returned for review. based on clinical description, the root cause of low flow was most likely due to pump being in improper position. the root cause of the positioning issue was most likely patient condition related per clinical review.
11960 | 1220648-2025-25863 | 21256954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21256954&pc=OZD | 01/16/2025 | 02/24/2025 |  | 01/28/2025 |  |  | 01/12/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592863 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, heparin drip paused due to small hematoma at right subclavian access site. computed tomography done to assess hematoma and it decreased in size. patient has intermittent atrial fibrillation (afib). the patient survived the event. | [5 empty] | The investigation of access site bleeding and atrial fibrillation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25863 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.3 date of awareness was reported incorrectly on the initial manufacturer device report number 1220648-2025-25863. date should have been 16-jan-2025. h.6 code 1888 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25863. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
11961 | 1220648-2025-25864 | 21257051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21257051&pc=OZD | 01/14/2025 | 02/12/2025 |  | 01/28/2025 |  |  | 01/14/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557936 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the eu had a console with grey buttons getting stuck. the console was utilized for the 5.5 pump support until the patient expired after 24 hours. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - kbd/rotary knob issues has been completed. unable to confirm the reported failure with the log evidence. during the device analysis was unable to reproduce the reported issue. the root cause of the grey button getting stuck was not determined due to the inability to reproduce. e4 should have been left blank on manufacturer device report 1220648-2025-25864.
11962 | 1220648-2025-25865 | 21257052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21257052&pc=OZD | 01/09/2025 | 03/17/2025 |  | 01/28/2025 |  |  | 01/14/2025 | 01/28/2025 |  |  |  |  | 01/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584242 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 91 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in japan reported that the patient on impella cp support had a cardiac perforation requiring thoracotomy. the impella caused/contributed to the cardiac perforation. additionally, it is believed that the patient expired from hemorrhagic shock due to liver rupture. liver rupture is not impella related. there is not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of ventricle perforation issue has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25865 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
11963 | 1220648-2025-25787 | 21240090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240090&pc=OZD | 01/08/2025 | 12/02/2025 |  | 01/27/2025 |  |  | 01/08/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218783 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility the patient was transferred in from an outside hospital for a coronary artery bypass graft surgery and elective placement of an impella 5.5 device. the impella was successfully implanted and soon after, still in the procedure area, a controller failure with pump stopped occurred. consequently, the impella 5.5 pump was replaced with a new impella 5.5. reportedly, the patient did well thereafter with no complications. | [5 empty] | The investigation for the reported issue has been completed. data analysis: based on the logs, the console displayed ¿impella stopped (motor current high) [detected by epc] ¿ alarm,¿ event #13. although the user restarted the pump several times, the impella stopped each. this confirms the reported failure mode. even after the pump was moved from (b)(6) to (b)(6) as per the constellation, it was still unable to run, and the (b)(6) displayed event #13. a new pump (sn: (b)(6)) was then used in (b)(6), and no pump stop issues occurred. the support was successfully provided for 4 days and 17 hours. device analysis: not applicable. although ic10564 was requested, it was not returned for investigation. root cause: the root cause of the pump stop was not determined as the (b)(6) was not returned for the investigation. there was no evidence in the log to confirm the reported failure mode, ¿controller error.¿ it is most likely that the pump stop was misreported as a ¿controller error.¿ h6 investigation type, findings and conclusions. | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two devices associated with this event. refer to manufacturing report number 1220648-2025-25736.
11964 | 1220648-2025-25788 | 21240283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240283&pc=OZD | 01/09/2025 | 02/24/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575792 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pseudoaneurysm | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp for care escalation for a transferred patient who had an intra-aortic balloon pump. the cp site was bleeding which prompted the team to assess site. the team observed a pseudoaneurysm and vascular surgeon involved. the cp was explanted and escalated to a impella 5.5 pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and vascular damage has been completed. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely anticoagulation management related as act was 345 when bleeding was reported. the root cause of vascular damage was unable to be determined as limited clinical details were provided and no product was returned for review.
11965 | 1220648-2025-25789 | 21240322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240322&pc=OZD | 01/09/2025 | 02/25/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was inserted to support a patient in acute myocardial infarction and cardiogenic shock. impella cp support was for 10 hours as the patient urine turned brown. the care team believed that the patient did not need impella support anymore, so impella was removed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs showed pump ran without issue during support. there were positioning alarm triggered in the end of the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
11966 | 1220648-2025-25790 | 21240383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240383&pc=OZD | 01/03/2025 | 02/07/2025 |  | 01/27/2025 |  |  | 01/03/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584942 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with a history of cerebrovascular accident in 2023 with left sided residual weakness was admitted for non-st-segment elevation myocardial infarction and acute systolic heart failure. while on impella 5.5 support, the patient developed left middle cerebral artery stroke with weakness on right side. the patient has been recovering and the right-side function is improving. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed the pump passed all post inspection checks. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
11967 | 1220648-2025-25791 | 21240410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240410&pc=OZD | 01/09/2025 | 02/18/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563313 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, an echocardiogram revealed a thrombus in the patient's left ventricle. attempts were made to swap the pump for intra-aortic balloon pump support, but the patient decompensated and the removal was aborted. the patient remained on support for another day while the thrombus was being monitored. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. e4 should have been left blank on manufacturer device report 1220648-2025-25791.
11968 | 1220648-2025-25792 | 21240411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240411&pc=OZD | 01/09/2025 | 02/24/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584576 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, extensive clot noted from aortic arch to descending aorta, encapsulating impella cp catheter. decision was made to abort explant and aggressively gave an anticoagulant. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the thrombus has been completed. the returned product was visually inspected and no abnormalities were found. without additional clinical details, the root cause of the thrombus could not be determined. d.9 date device returned/information was added. h.6 code 4641 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25792 and was added.
11969 | 1220648-2025-25793 | 21240427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240427&pc=OZD | 01/06/2025 | 02/18/2025 |  | 01/27/2025 |  |  | 01/06/2025 | 01/27/2025 |  |  | 06/30/2025 |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364869 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Swelling/ Edema | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with impella 5.5 device for mechanical circulatory support. approximately three days after start of support, the patient experienced premature ventricular contractions/runs of ventricular tachycardia which were resolved with repositioning of the pump. four days later, the patient had swelling at impella site and down right arm, dressing changed with bedside, no active bleeding. however, there was leakage from insertion site with turns due to swelling. doppler ultrasound was completed on all four extremities, and patient received one unit of packed red blood cells. the next day the patient went for a wash out of the site; there was no significant bleeding, but the staff found a surgical device left in the pocket. the device was removed, and the pocket was washed out, dried and site reclosed. physician commented it is believed the incident of the device occurred at an outside facility. the patient was reported to be in stable condition following the event. latest follow up noted 13 days later, the patient was received a left ventricular assist device; impella 5.5 was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because the surgicel was left in the pocket and hemostasis was obtained by removing it and applying the pressure the site. as the necessary clinical information was not provided, the root cause of the ventricular tachycardia was not determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25793 as it is unknown if the initial reporter also sent the report to the food and drug administration.
11970 | 1220648-2025-25794 | 21240443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240443&pc=OZD | 01/08/2025 | 01/30/2025 |  | 01/27/2025 |  |  | 01/08/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1484655 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Select Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) was replaced due to inability of the purge disc to be detected by the aic. there was no report of harm to a patient. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated imepella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm. the console was examined and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected issue. h10 investigation results were inadvertently omitted from manufacturer device report 1220648-2025-25794.
11971 | 1220648-2025-25795 | 21240447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240447&pc=OZD | 01/09/2025 | 02/27/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589154 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support a patient admitted in with cardiomyopathy. the patient was being discussed for heart transplant. the pump supported for 23 hours til the pump came out of position with torque and patient rolling. the pump was attempted to be re-delivered and was found to kink. the pump was replaced. there was no reported patient harm. | [5 empty] | The investigation of positioning issues and product damage (cannula kink) has been completed. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely patient movement related. the root cause of kinked cannula was most likely use related. h1 type of reportable event was revised as it was inadvertently selected incorrectly on initial manufacturer device report 1220648-2025-25795. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
11972 | 1220648-2025-25796 | 21240571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240571&pc=OZD | 01/05/2025 | 02/18/2025 |  | 01/27/2025 |  |  | 01/06/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576596 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | [12 empty] | Death | Life Threatening | Disability | Hospitalization | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received from abiomed's long-term outcome and quality indicator impella registry regarding a patient who endured cardiogenic shock with an unknown/undetermined relationship to the device used. the patient continued to decompensate after impella placement, due to critical illness and unlikelihood of recovery, the family decided on comfort care measures. the patient would expire approximately 24 hours after start of this support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided, the root cause of the vt/vf was not determined. e4 should have been left blank on manufacturer device report 1220648-2025-25796.
11973 | 1220648-2025-25797 | 21240577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240577&pc=OZD | 01/08/2025 | 02/25/2025 |  | 01/27/2025 |  |  | 01/08/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519319 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with st-segment elevation myocardial infarction was found to have an 100 percent left anterior descending artery occlusion. after opening the vessel, the patient coded requiring cardiopulmonary resuscitation and was defibrillated ten times. decision was made to implant an impella cp device to the right femoral artery, which was successfully completed. however, during support, after implant, it was noted there was a vascular injury and significant bleeding at insertion site requiring vascular cutdown and pump removal. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. clinical details mentioned are very limited. the location and cause of the vascular damage is unknown. without this information, the failure could not be investigated further. therefore, the root cause of the vascular damage issue was not determined since the product was not returned and insufficient clinical information was provided.
11974 | 1220648-2025-25798 | 21240604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240604&pc=OZD | 01/07/2025 | 03/13/2025 |  | 01/27/2025 |  |  | 01/07/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592499 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and two days after start of support, a leakage between impella filter and pressure reservoir was observed. protocol for purge filter bypass was requested. reportedly the following day medical staff decided to end therapy due to the patient's "poor" general condition. the relationship between the patient's outcome and the impella device used is unknown. no further details were provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of purge leak ¿ pump has been completed. no product returned and no data logs. clinical details reported a leakage between the filter and pressure reservoir. bypass was performed. the cause of the purge leak - pump was most likely a leak from the sidearm but the exact location of the leak was unknown. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ a.1 revised patient identifier as it was reported incorrectly on the initial manufacturer device report number 1220648-2025-25798. d.3 revised us zip code as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25798. e.1 revised last name as it was reported incorrectly on the initial manufacturer device report number 1220648-2025-25798. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25798 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.4 revised device manufacture date as it was reported incorrectly on the initial manufacturer device report number 1220648-2025-25798. h.10 instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2025-25798.
11975 | 1220648-2025-25799 | 21240629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240629&pc=OZD | 01/06/2025 | 03/14/2025 |  | 01/27/2025 |  |  | 01/07/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103169 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient was on an impella 5.5 mechanical circulatory support, and the automated impella controller (aic) had a controller error. the top part of screen (ao and lv numbers/waveforms) went black while lower motor current, flows, and other views remained visible. the issue self-resolved in approximately one to two minutes. it was reported the patient remained stable throughout, the aic was exchanged with no interruption in support. there was harm to the patient as a result of the event. | [5 empty] | The investigation for the reported issue has been completed. the impella device was returned for evaluation. the reported failure mode was confirmed to have occurred with the automated impella controller (console) event logs. the console was tested and after running the console for 18 hours with a test cassette and pump at power level p9, the reported failure mode was unable to be reproduced. the root cause for the communication issues and the controller errors were not determined due to the inability to reproduce. h8 usage of device corrected to reuse. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
11976 | 1220648-2025-25800 | 21240637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240637&pc=OZD | 01/07/2025 | 02/19/2025 |  | 01/27/2025 |  |  | 01/07/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552575 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported during an impella 5.5 implant, the cannula kinked and kept kinking in same spot. the impella 5.5 device was explanted, tissue tract enlarged, and the second impella 5.5 device was placed without difficulty. reportedly there was no harm to the patient as a result of the event. | [5 empty] | The investigation of product damage (cannula kink) and delivery issues has been completed. the pump was returned. cannula kink was found. catheter kink was found under motor housing. as the cannula kink occurred during delivery issues, failure mode (fm) product damage cannula kink was captured under fm delivery issues. there was difficulty inserting the first pump, causing cannula kinking. the tissue tract was enlarged and the second pump was able to be inserted with no issues despite the possible tissue tract obstruction. the cause of the delivery issue was most likely use related as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2199 was entered incorrectly on the initial manufacturer device report number. a new code was added. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number and were added. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
11977 | 1220648-2025-25801 | 21240642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240642&pc=OZD | 01/09/2025 | 02/25/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512720 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the impella was experiencing suction. the patient was given three units of blood products and albumin. additionally, the peel away sheath was left in place and an external bypass was placed due to the patient have occluded flow. this helped the patient flow but flow was still limited until explant. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the low pump flow and the limb ischemia has been completed. data logs were reviewed which showed flow was fluctuated after pump started and remained for 3 hours corresponded with the time the patient coded. after that, flow was in normal range to the rest of the case. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely patient condition related as the low flow was resolved after fluid compensation.
11978 | 1220648-2025-25802 | 21240647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240647&pc=OZD | 01/08/2025 | 02/18/2025 |  | 01/27/2025 |  |  | 01/08/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531174 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and developed a hematoma at the access site that was successfully managed. the patient had chest pain for one week, was taken to the emergency department, and diagnosed with st-segment elevation myocardial infarction. the patient was transported for an emergent percutaneous coronary intervention (pci). one drug-eluting stent (des) was placed in the right coronary artery (rca), and the patient decompensated leading to cardiac arrest. in between resuscitation attempts, the decision was made to implant the impella cp. once on mcs, the patient stabilized. the pci continued with a second des to the rca and two des to the left circumflex artery. the impella cp site was dressed with minimal oozing and no hematoma present. the patient continued on impella support and was prepped for airevac to an outside hospital. a couple of hours later, however, a "large" hematoma around the impella access site was noted. heparin rate was reduced to manage the condition, and the site was monitored. following, approximately five days later, the patient was escalated to an impella 5.5. the impella cp was weaned down and explanted prior to insertion of the impella 5.5 device. access side port was accessed with a 0.35 wire, and the impella cp was discontinued and successfully removed. a manta closure device was inserted over the wire into the common femoral artery and closure was successful with deployment. latest patient update for status noted the patient was continued on impella and extracorporeal membrane oxygenation supports, experienced an ischemic cerebral vascular accident, was moved to comfort care and expired. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement as the patient developed a hematoma around insertion site after being transferred by flight to new facility. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25802 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). this is one of two devices associated with the patient implant experience. refer to initial medical device report for impella 5.5 serial number (b)(6) with date of event 16-jan-2025 and identifier ending in (b)(6).
11979 | 1220648-2025-25803 | 21240679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240679&pc=OZD | 01/08/2025 | 02/14/2025 |  | 01/27/2025 |  |  | 01/08/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545748 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and developed limb ischemia the following day. the patient would later expire. a patient with st-segment elevation myocardial infarction coded, return of spontaneous circulation (rosc) was achieved, and was transported to the catheterization lab intubated and unresponsive. the patient was prepped on the table. radial access was obtained for cardiac catheterization. the patient coded again, and cardiopulmonary resuscitation commenced with rosc attained and maintained. repeat arterial blood gases and activated clotting time (act) were completed every few minutes as ordered. after rosc, coronary angiogram resumed with a tiger catheter via right radial access. prior to the percutaneous coronary intervention (pci), the impella cp device was prepped for cardiac support. the left femoral access obtained for impella access, insertion was successful and cardiac flow was maintained at 3.2 -3.3 flow throughout the pci to the left circumflex, mid left anterior descending and diagonal branch arteries. act was maintained at therapeutic range between 230 -250 range. the patient was transferred to the intensive care unit with stable vital signs and the access site was clean dry and intact (no hematoma or bleeding). the following day, however, the patient experienced limb ischemia and a plan was made to explant the pump. the patient remained intubated and sedated and pressor demands remained unchanged. the next day, the patient was noted be in multiorgan failure. goals were discussed with the family and the patient was made do not resuscitate/intubate and would expire. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11980 | 1220648-2025-25804 | 21240683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240683&pc=OZD | 01/08/2025 | 02/18/2025 |  | 01/27/2025 |  |  | 01/08/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562945 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Pain | Swelling/ Edema | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and approximately four days after start of support developed a "large" hematoma at the axillary site. consequently, the patient experienced arm swelling and pain and was successfully weaned from impella 5.5 to native heart function. the patient underwent a subclavian artery repair and an axillary washout. the patient was noted to be doing better and stable post the event. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the root cause of the access site bleeding cannot be determined since insufficient clinical details were provided. e4 should have been left blank on manufacturer device report 1220648-2025-25804.
11981 | 1220648-2025-25805 | 21240758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240758&pc=OZD | 01/06/2025 | 02/14/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  | 09/30/2025 |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399718 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [5 empty] | VV EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a pediatric patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and 25 days after start of support, a computed tomography head scan reviewed a cerebellar infract. the patient was on venovenous extracorporeal membrane oxygenation as well as impella 5.5 mcs at the time of the finding. follow-ups will be completed as the patient progresses; currently, the team is unable to determine the patient's deficit due to the patient being intubated and sedated. support with the implanted pump continues as per the latest plan of care update 19 days following the finding. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of stroke/cerebrovascular accident (cva) has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25805 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4614 was entered incorrectly on the previously submitted manufacturer device report number 1220648-2025-25805 and a new code has been added. h.10 revised as a portion of the additional manufacturer narrative instructions for use were omitted on the initial submission of manufacturer device report number 1220648-2025-25805.
11982 | 1220648-2025-25806 | 21240771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240771&pc=OZD | 01/08/2025 | 02/07/2025 |  | 01/27/2025 |  |  | 01/08/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573267 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing an st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support and extracorporeal membrane oxygenation (ecmo). the next day after start of support, a large left ventricular thrombus was seen on a transesophageal echocardiogram (tee) (which was noted not previously seen on a prior tee completed in the catheterization lab). consequently, the impella cp device was removed. the patient was escalated to an impella 5.5 and still continues on ecmo. the patient was noted to have survived and in stable condition following the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
11983 | 1220648-2025-25807 | 21240779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240779&pc=OZD | 01/09/2025 | 02/07/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563974 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support during an ablation procedure. after the ablation procedure, the patient experienced some oozing from the impella insertion site following the explant. the physician opened the pocket to evaluate the nature of the oozing, manage the condition. the bleed resolved, and the patient was considered to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. the device history record review confirmed that the pump passed all post sterility checks. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
11984 | 1220648-2025-25808 | 21240848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240848&pc=OZD | 01/09/2025 | 02/18/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584998 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an attempt to insert an impella cp pump was unsuccessful. upon exchanging an indwelling impella cp for a new one, wire position was lost during explant and the new impella cp device could not be delivered. manual pressure was applied until hemostasis was achieved. there was no groin complication noted; the patient remained hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely use issue since wire position was lost during explant and new device could not be delivered. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a and d6b were inadvertently omitted from manufacturer device report. e1 facility name was inadvertently reported incorrectly on manufacturer device report. e4 should have been left blank on manufacturer device report. h10 instructions for use were inadvertently omitted from manufacturer device report.
11985 | 1220648-2025-25809 | 21240863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240863&pc=OZD | 12/31/2024 | 03/26/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573326 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information received abiomed¿s long-term outcome and quality indicator (loqi) impella registry reporting the patient having a stroke during bipella support. | Us complainant reported the patient was on impella 5.5 support and during support, the impella 5.5 had suction until dobutamine was turned off. 5.5 increased to p8. support continued. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right axillary hematoma with a definite relationship to the impella device used: patient with known right axillary hematoma from (b)(6)2024 at impella insertion site. hematoma does not seem to be improving. patient booked for impella axillary site exploration and washout (b)(6)2025 exploration and washout, placed drain and closed wound. the patient recovered with no sequelae. the patient remains on impella support. |  |  |  |  | As noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.5 adittional information was received and was inadvertently omitted on follow up 1 manufacturer device report number 1220648-2025-25809. b.7 added information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25809. g.2 added a selection as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25809. g.3 date of awareness on follow up 1 manufacturer device report number 1220648-2025-25809 was reported incorrectly as the date should have been 24-feb-2025. h.6 added code to health effect - clinical code as it was inadvertently omitted on follow up 1 manufacturer device report number 1220648-2025-25809. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the initial submitted manufacturer device report number 1220648-2025-25809 incorrectly. new ifus have been added. | The investigation for the low pump flow and the access site bleed has been completed. data logs confirmed low pump flow when pump started that triggered multiple suction alarms. this event remained to the rest of the case. product was not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11986 | 1220648-2025-25810 | 21240873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240873&pc=OZD | 01/12/2025 | 02/27/2025 |  | 01/27/2025 |  |  | 01/12/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491148 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had setup a cp pump for the support of a patient without any abiomed support onsite. the system failed to prime or do automated impella controller self checks appropriately. as there was a risk, the team replaced and placed new pump for support. no patient harm has been reported. | [5 empty] | The investigation of priming issue has been completed. the root cause of the priming issue was not determined since the product and logs were not returned and insufficient clinical details were provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25810 as it is unknown if the initial reporter also sent the report to the food and drug administration. d4 expiration date was revised as it was not included on the initial manufacturer device report number 1220648-2025-25810. d4 unique identifier (udi) was revised as it was not included on the initial manufacturer device report number 1220648-2025-25810. h4 device manufacture date was revised to reflect the corrected date on manufacturer device report 1220648-2025-25810. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
11987 | 1220648-2025-25811 | 21240886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240886&pc=OZD | 01/13/2025 | 03/03/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597182 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the left groin was noted to have a small hematoma around the impella insertion site. staff held manual pressure for around 20 minutes and the hematoma resolved. upon reassessment, another small hematoma was noted so staff once again held pressure until hematoma was resolved. staff again assessed groin and noted there was a hematoma forming in addition to notable swelling of the scrotum. cangrelor infusion was discontinued and staff continued to hold pressure over impella sheath in intensive care unit. additionally, the patient was given two units of packed red blood cells. impella was explanted. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be patient movement because left groin was noted to have a small hematoma after taken to the cath lab recovery. added h6 impact code 4643. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11988 | 1220648-2025-25812 | 21240896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240896&pc=OZD | 01/13/2025 | 02/20/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168282 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the automated impella controller (aic) had a controller error alarm. the alarm did resolve spontaneously but team decided to switch to a new console for patient safety. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) controller failure -yellow alarm has been completed. console logs confirmed that as case start was completing, the console was removed from ac power with 8 seconds later triggering a voltage out of spec condition on both the pump and purge line on just pbm1. this triggered a controller error that was resolved immediately after ac power was reconnected. the console was returned for evaluation. despite many iterations of removing ac power with the console performing case start and a pump and purge plugged in, the failure mode did not reproduce. since both the pump and purge subsystems had a voltage out of spec condition, the commonality was the pbm most likely transiently malfunctioned causing the controller error. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11989 | 1220648-2025-25813 | 21240909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240909&pc=OZD | 01/13/2025 | 02/28/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579057 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient pulled impella back 10cm. unable to safely reposition, the surgeon made the decision to replace device contralaterally. the patient survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was due to patient movement since the patient pulled impella back 10cm and it was unable to safely reposition. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
11990 | 1220648-2025-25814 | 21240915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21240915&pc=OZD | 01/12/2025 | 02/28/2025 |  | 01/27/2025 |  |  | 01/12/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589179 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed for a patient admitted in acute myocardial infarction/cardiogenic shock. after three days, the purge block alarms began. the team discussed troubleshooting with pharmacology and the pump remained on for support with tissue plasminogen activator running in the purge. the patient survived the event. | [5 empty] | The investigation of high purge pressure has been completed. the pump was returned for evaluation. upon visual inspection, biomaterial was present on returned pump around the purge gap and in the outflow cage. pump was purged in lab for approximately 10 minutes and run for an additional 5 minutes at p-9. "high purge pressure" alarms did not reproduce and purge pressure remained stable and in normal range during lab run. data logs were reviewed and the logs at time of issue show step up in motor current over a period of approximately 7 hours. initial step up in purge pressure occurred immediately after bag change. clinical details noted that purge pressure alarms occurred on 4th day of support. purge cassette was changed and tissue plasminogen activator (tpa) was added to purge and did not resolve alarms. the root cause of the high purge pressure was most likely due to biomaterial. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
11991 | 1220648-2025-25815 | 21241070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241070&pc=OZD | 01/12/2025 | 02/19/2025 |  | 01/27/2025 |  |  | 01/12/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, there was intermittent loss of pulsatility. the patient¿s right lower extremity (rle) was mottled and without a dorsalis pedis/ posterior tibial pulse - has a popliteal. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
11992 | 1220648-2025-25816 | 21241123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241123&pc=OZD | 12/06/2024 | 02/26/2025 |  | 01/27/2025 |  |  | 12/06/2024 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568004 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs). approximately eight days after start of support, heparin-induced thrombocytopenia was suspected. bicarbonate was added to the purge and thr patient started on an angiomax drip. the patient continued on impella mcs. however, approximately 38 days later, it was reported the patient had runs of ventricular tachycardia with positioning change. the patient maintained a blood pressure during the ventricular tachycardia runs. additionally noted, the patient was evaluated for gastrointestinal bleeding. systemic anticoagulation was discontinued, and a gastrointestinal consult was completed. no further information regarding the bleeding was noted. the patient continues on the impella mcs waiting for a heart transplant. | [5 empty] | The investigation for the thrombocytopenia, ventricular tachycardia (vt), and gastrointestinal bleed (gi) has been completed. the impella device was not returned for evaluation. clinical details were not sufficient to determine the root cause. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the cause of the gi bleed and vt was not determined since product was not returned and there was insufficient clinical information. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11993 | 1220648-2025-25817 | 21241155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241155&pc=OZD | 01/01/2025 | 03/03/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582793 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant had a cp pump placed to support a patient admitted in with st elevated myocardial infarction. the pump had supported for under a day when explanted. abiomed¿s long-term outcome and quality indicator (loqi) impella registry database later reported upon hemolysis and pump losing left ventricle positioning. the loqi puts allegation of relationship as definite. later the loqi reported upon limb ischemia which was of a definite relationship. the foot was dusky and the limb was cold. the patient survived. | [5 empty] | G.2 revised as study selection was missing from previously submitted manufacturer device report number 1220648-2025-25817. | The investigation for the limb ischemia, hemolysis, and positioning has been completed. data log was reviewed and suctions occurred followed by impella position in aorta alarms at the date hemolysis issue reported, but looks hematuria was observed pre-occurring positioning alarms. the cause of the hemolysis issue could not be determined since the complaint device was not returned for analysis and insufficient clinical details were provided. the cause of the positioning issue could not be determined due to insufficient clinical data. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. h.6 component code 3038 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25817 and was added. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
11994 | 1220648-2025-25818 | 21241168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241168&pc=OZD | 01/13/2025 | 03/03/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570637 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a short sheath was initially placed and while going up with pigtail, bleed-back occurred around the catheter. the issue was not improved by removing and reseating the hemostatic valve. the short sheath was exchanged for a long. no patient harm has been reported. | [5 empty] | Product was not returned for investigation. a data log review was not required for this case. according to the clinical report, bleeding occurred from the hemostatic valve after the insertion of the catheter. the issue was resolved after changing the introducer. the reported introducer damage failure mode was changed to the access site bleeding failure mode since the bleeding resulted from the damaged introducer. the cause of the access site bleeding issue was not determined since introducer was not returned for investigation. according to the investigation findings the following was revised on manufacturer device report 1220648-2025-25818. b1 revised to adverse event. h1 revised to serious injury. h6 revised to reflect investigation findings. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
11995 | 1220648-2025-25819 | 21241264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241264&pc=OZD | 01/10/2025 | 04/24/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  | 09/30/2025 |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403342 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complaints team received an impact study that found a possible relationship between the impella device to the patient's bleeding. the patient had a blood coagulation disorder that was possibly caused by impella therapy. the patient needed a transfusion for the blood loss. the patient survived the event. | The complaint team received new information from an impact study that reported the patient went into cardiac arrest, went asystole and expired. the study found there was a remote relationship between the impella and the patient's death and a possible relationship between the impella procedure and the patient's death. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vt/vf has been completed. the root cause of the vt/vf cannot be determined since the product was not returned and insufficient clinical details were provided. | B.2 death and date of death were added. b.5 new information was received. h.1 revised due to new information received. h.6 health effect - clinical code and health effect - impact code were addded due to new information received. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-25819 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
11996 | 1220648-2025-25820 | 21241291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241291&pc=OZD | 01/10/2025 | 02/19/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084303 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller failure alarm twice. the aic was removed. there was no patient involvement. | [5 empty] | The investigation of automated imeplla controller (aic) - controller failure (red alarm) has been completed. the logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the cause of the aic issue was a software issue. e2 was inadvertently reported incorrectly on manufacturer device report. e3 was inadvertently reported incorrectly on manufacturer device report. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
11997 | 1220648-2025-25821 | 21241463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241463&pc=OZD | 01/13/2025 | 02/28/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585630 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [5 empty] | CONTINUOUS RENAL REPLACEMENT THERAPY | EXTRACORPOREAL MEMBRANE OXYGENATION | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient's lactate dehydrogenase levels were 2200. the patient had acute kidney injury and continuous renal replacement therapy was needed. the patient was making minimal urine so the staff could not confirm hemolysis. the team explanted the pump and implanted and impella 5.5. the patient survived the event. | [5 empty] | The investigation of hemolysis and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
11998 | 1220648-2025-25822 | 21241563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241563&pc=OZD | 01/13/2025 | 02/18/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592993 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support an acute myocardial infarction/cardiogenic shock patient with cardiac history. the limb accessed became ischemic and distal perfusion was placed. the impella cp was then explanted and escalated to an impella 5.5 pump. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
11999 | 1220648-2025-25823 | 21241673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241673&pc=OZD | 05/30/2024 | 03/12/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  | 02/28/2026 |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025480808 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Unspecified Vascular Problem | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella 5.5 support and during support, the patient had tachycardia related to the impella. the impella was explanted and the tachycardia stopped. additionally, an impact study found a probable relationship between the impella procedure and the patient's pain in their pectoral. patient was admitted for a revision of the subclavian artery, right side. the patient had pain in the thorax right side, suspected distress from the impella goretex graft, surgery. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of tachycardia and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the tachycardia was not determined. a probable relationship between the impella procedure and the patient's pain in their pectoral pain was reported. the patient was admitted for a revision of the aortic subclavia, right side. the patient had pain in the thorax right side, suspected distress from the impella goretex graft, surgery. the cause of the vascular damage peripheral vessel was not determined due to lack of clinical information and no product returned. b.3 revised date of event as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-25823. e.1 revised first name and address line 1 as they were previously reported incorrectly on the initial manufacturer device report number 1220648-2025-25823. g.1 revised manufacturing site address as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-25823. g.2 revised as health professional was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25823. g.6 30-day was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25823. h.6 code 1721 was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-25823 and a new code was added. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-25823 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
12000 | 1220648-2025-25824 | 21241827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21241827&pc=OZD | 01/13/2025 | 02/28/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592224 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, a motor current spike occurred and then impella stop alarm appeared. issues were unable to resolve so the pump was explanted. the patient¿s outcome is unknown. | [5 empty] | The investigation of pump stop has been completed. the root cause of the pump did not start was most likely due to an electrical open issue, but the reason for electrical open error could not be determined as no product was returned and insufficient information related to failure was provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12001 | 1220648-2025-25825 | 21242166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21242166&pc=OZD | 01/13/2025 | 02/19/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576994 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, distal pulses could not be found in impella leg. impella was explanted due to limb ischemia. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
12002 | 1220648-2025-25826 | 21245911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21245911&pc=OZD | 01/10/2025 | 03/04/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580378 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was inserted to support a percutaneous coronary intervention (pci) procedure. at some point during the case, blood was noted on the floor. the physician noted bleeding from the impella sheath hemostatic valve. wires and companion sheath were removed and bleeding stopped. the companion sheath was reinserted and no issues with bleeding were reported after. pci procedure was completed and the impella was wean and explanted. the patient received one unit of packed red blood cells and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the pump and the introducers were returned for inspection and no issues or abnormalities were noted upon inspection. the root cause of access site bleeding is determined to be use issue because there were no issue when they have tried again later the same day. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12003 | 1220648-2025-25827 | 21248118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248118&pc=OZD | 01/09/2025 | 02/25/2025 |  | 01/27/2025 |  |  | 01/09/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558066 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. patient developed a hematoma in the right groin due to excessive movement and agitation overnight. the decision was made to switch to the left groin for cp insertion. additionally, the patient developed hemolysis. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding - major most likely was due to patient's excessive movement and agitation overnight. the data logs were reviewed and there was no motor current spike observed. quick resolved alarms occurred in the aorta and ventricle. placement signal low alarms were observed on the date of hemolysis issue reported. based on the clinical details it was noted the impella was deep on the echocardiogaram with accompanying placement signal low alarms but the decision was made to decrease power level rather than reposition. the root cause of hemolysis issue most likely was due improper positioning related. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25827 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12004 | 1220648-2025-25828 | 21248127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248127&pc=OZD | 01/10/2025 | 02/26/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584553 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced active bleeding from their impella cp access site during support. a purse string suture was placed coinciding with a femstop to manage the condition. four units of packed red blood cells were given to counteract the blood loss and support continued for one more day until planned escalation. upon removal of the device, a clot was seen on the tip of the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the access site bleed and thrombus has been completed. the pump was returned for investigation with the broken catheter between the repo unit and the pump. the returned product was upon investigation, biomaterial was noted on the cannula near the catheter and the cannula was noted be slightly kinked. the ptt value is200 during the time of bleeding after insertion of gwru in the cath lab. the root cause of access site breeding is determined to be use issue because successful hemostasis achieved by applying purse string suture. added d9 device return date. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12005 | 1220648-2025-25829 | 21248170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248170&pc=OZD | 01/10/2025 | 02/26/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, bleeding was noted at the impella insertion site. bleeding successfully managed by manual pressure and applying a fem-stop. heparin was off due to bleeding at impella insertion site. large hematoma remained. care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the product was not returned. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details were provided. b.2 revised as required intervention was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25829. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25829 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12006 | 1220648-2025-25830 | 21248188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248188&pc=OZD | 01/10/2025 | 06/23/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  | 02/28/2025 |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338392 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella cp was inserted pre-percutaneous coronary intervention (pci). initial flows were within range but after repositioning, flows noted to be decreased and suction occurring. kink noted in impella cannula on fluoro imaging. attempts to straighten cannula unsuccessful. decision was made to complete pci procedure and pump was explanted after. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E4 should have been left blank on the initial manufacturer device report number 1220648-2025-25830 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The investigation for the product damage has been completed. returned complaint cp pump visually inspected. cannula kink near outlet observed. no biomaterial, no broken blade observed. the cause of the low pump flow issue was cannula kink due to improper repositioning technique.
12007 | 1220648-2025-25831 | 21248223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248223&pc=OZD | 01/10/2025 | 02/26/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573347 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was placed the patient was taken to intensive care unit and a small hematoma was noted. pressure was applied and hematoma resolved. heparin was held temporarily. the patient remained stable until successful weaning of the pump. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement, and the anticoagulation management as the heparin was stopped to achieve hemostasis as per the clinical description. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25831 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12008 | 1220648-2025-25832 | 21248248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248248&pc=OZD | 01/10/2025 | 02/24/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555085 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to deliver an impella cp device for the support of a patient undergoing high-risk surgery. an impella 5.5 device was initially placed into right axillary artery but due to patient's anatomy and inability to push the impella 5.5 device passed innominate, the impella 5.5 device was removed and the impella cp was attempted, which was placed into axillary artery after impella 5.5 was attempted and removed. however, impella cp device could not get passed innominate as well. impella cp device was removed, and the surgeon decided to perform the coronary artery bypass graft (cabg) surgery without impella mechanical circulatory support. reportedly, the patient did well after the cabg surgery. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. instructions for use: delivery issues impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a and d6b were inadvertently omitted from manufacturer device report 1220648-2025-25832. e4 should have been left blank on manufacturer device report 1220648-2025-25832. h10 instructions for use were inadvertently omitted from manufacturer device report 1220648-2025-25832. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two devices associated with the patient's implant experience. refer to manufacturing report number 220648-2025-25778.
12009 | 1220648-2025-25833 | 21248281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248281&pc=OZD | 01/10/2025 | 02/13/2025 |  | 01/27/2025 |  |  | 01/10/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580347 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported during implant of an impella 5.5 device to a patient (unknown race) for mechanical circulatory support (mcs), the sterile sleeve came apart. the patient was being escalated to an impella 5.5 from an impella cp. the cp was pull and placed in surgical mode while implanting the impella 5.5. once the impella 5.5 was placed, removal of the impella cp from the right groin without any complications was completed. while placing the impella 5.5 and trying to optimize position, the contamination sleeve came apart from the distal end. the was sleeve pulled forward and a tegaderm was used to secure sleeve to hub. a 3-point fixation was completed, and the patient was returned to the unit continuing with impella mcs. there was no report or indication of harm to the patient as a result of this event. | [5 empty] | The investigation for the product damage has been completed. the pump was not returned for investigation. data log review was not required for this case. as per the given clinical description, while advancing the locking hub into the hp 10 millimeter graft, the contamination sleeve detached. it was repositioned and secured with a tegaderm. the cause of the product damage issue was not determined since the product was not returned and sufficient clinical information was not provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25833 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12010 | 1220648-2025-25834 | 21248292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248292&pc=OZD | 01/11/2025 | 04/10/2025 |  | 01/27/2025 |  |  | 01/11/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561845 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient developed a severe hematoma in their groin and scrotum following impella cp implant. following a successful high-risk coronary intervention, explant was planned for later in the day due to the hematoma. the patient survived the event. | [5 empty] | The investigation access site bleeding has been completed. the returned product was visually inspected and no abnormalities were found .the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. d9 date device returned to manufacturer added. e1 initial reporter first and last name was inadvertently omitted from the initial manufacturer device report 1220648-2025-25834 e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25834 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the access site bleeding issue was not determined since insufficient clinical details were provided. corrected information provided in h10. erroneous investigation results were entered on the supplemental manufacturer device report number 1220648-2025-25834-1.
12011 | 1220648-2025-25835 | 21248293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248293&pc=OZD | 12/17/2024 | 02/18/2025 |  | 01/27/2025 |  |  | 01/12/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571152 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry, the patient was on impella cp support. while on support, the patient endured ventricular arrhythmia. once pigtail crossed into left ventricle, the patient had premature ventricular contractions (pvcs) that led to sustained monomorphic ventricular tachycardia (mmvt) requiring single direct current cardioversion (dccv) at 200j with restoration to sinus rhythm. the patient survived the event. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 was inadvertently left blank on manufacturer device report 1220648-2025-25835. e4 should have been left blank on manufacturer device report 1220648-2025-25835. g1 reporting contact fax number was inadvertently omitted from manufacturer device report 1220648-2025-25835. g2 report source; study was missing from manufacturer device report 1220648-2025-25835. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12012 | 1220648-2025-25836 | 21248296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21248296&pc=OZD | 01/13/2025 | 02/28/2025 |  | 01/27/2025 |  |  | 01/13/2025 | 01/27/2025 |  |  |  |  | 01/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553754 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had access site bleeding. manual pressure was held and a femstop was applied. additionally, the patient had pink urine. the p-level was reduced but staff wanted to explant due to patient's injuries. before explanting, the pump alarmed, position in aorta. waveforms looked like position was too deep. the pump was explanted. the patient survived the event. | [5 empty] | The investigation of access site bleeding, hemolysis, and positioning issues has been completed. the root cause of access site bleeding, hemolysis, and positioning issues was unable to be determined as limited clinical details were provided and no product was returned for review. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25836 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12013 | 1220648-2025-25760 | 21228592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21228592&pc=OZD | 01/05/2025 | 02/14/2025 |  | 01/24/2025 |  |  | 01/05/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584944 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported upon attempted delivery of an impella 5.5 pump for mechanical circulatory support to a patient with cardiomyopathy, the physician was unable to advance the pump due to the size of the patient's vessel. after multiple failed attempts, the impella 5.5 pump was removed and an impella cp device was placed for ongoing support (for approximately 8 days until successfully weaned and explanted). the patient was noted to have survived with no resulting patient harm as a result of the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist f section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use were inadvertently omitted from manufacturer device report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12014 | 1220648-2025-25761 | 21228764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21228764&pc=OZD | 01/06/2025 | 02/13/2025 |  | 01/24/2025 |  |  | 01/06/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hypoxia | Vascular Dissection | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was being implanted with an impella cp via single access when there was a vessel dissection. the patient would later expire. upon arrival to catheterization lab the patient required cardiopulmonary resuscitation (cpr) and achieved return of spontaneous circulation (rosc). first, impella prepped, inserted and started as recommended, and single access attempted for angiogram, but right femoral artery ended up having small dissection in femoral artery so left femoral artery was accessed using a companion sheath. angiogram revealed culprit lesion of left circumflex (lcx) artery with left anterior descending (lad) artery total occlusion. a total of two stents was deployed in lcx and two stents in lad. during percutaneous coronary intervention, the patient remained with a mean arterial pressure over 80. however, due to hypoxia, the cardiothoracic surgery consulted for ecmo, but surgeons refused. the patient continued to decompensate and required 12 shocks and cpr. rosc achieved and additional drips added. team decided to further stabilize patient on the unit. a swan catheter was swapped to a dialysis catheter, and the patient transferred to unit. however, the patient would later expire the morning after a total of seven hours on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage has been completed. pump was not returned for investigation. data log review was not required for this case run. as per the given clinical, during the procedure, a small dissection (tear) occurred in the right femoral artery (rfa) while attempting single access for the angiogram. due to this complication, the team switched to the left femoral artery (lfa), using a companion sheath for access to continue the procedure. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided.
12015 | 1220648-2025-25762 | 21228825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21228825&pc=OZD | 12/28/2024 | 02/27/2025 |  | 01/24/2025 |  |  | 12/28/2024 | 01/24/2025 |  |  | 10/31/2025 |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466946 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant placed a impella cp to support a patient admitted in with chest pain and nausea/vomiting. the cp was placed (as well as extracorporeal membrane oxygenation (ecmo)) and after four days ischemia was noted in the impella limb, and a fasciotomy was performed. then 13 days later a above-the-knee amputation was performed. the patient remained on support for a total of over 28 days and was then explanted. the patient remained on ecmo and survived. | [5 empty] | The device was not returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12016 | 1220648-2025-25763 | 21228905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21228905&pc=OZD | 01/07/2025 | 02/24/2025 |  | 01/24/2025 |  |  | 01/07/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593471 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after the impella was placed, the peel-away sheath was removed and repositioning sheath placed. forward tension sutures were placed at the site, but groin still had some bleeding. an additional stitch was placed and pressure held with resolution. additionally, same day post procedure it was noted the impella was sitting shallow but unable to be advanced further due to frequent ectopy. support continued. three days later the impella was weaned and removed with the patient continuing on extracorporeal membrane oxygenation therapy. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary information was not provided, the root cause of the vt/vf was not determined.
12017 | 1220648-2025-25764 | 21229333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21229333&pc=OZD | 01/07/2025 | 03/20/2025 |  | 01/24/2025 |  |  | 01/07/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536775 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and same day of implant had heavy bleeding at the access site requiring two units red blood cells. status of the patient as a result of the event is unknown. no further details were provided. the patient continues on impella mcs. | It was further reported that the patient expired. there is not enough information to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 revised from the initial manufacturer device report number 1220648-2025-25764 to reflect the patient outcome; death, date of death was unknown. b5 revised from the initial manufacturer device report number 1220648-2025-25764 to reflect the details of the patient outcome. d3 manufacture zip code was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25764. d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25764. h1 revised from the initial manufacturer device report number 1220648-2025-25764 to reflect the details of the patient outcome. h6 revised from the initial manufacturer device report number 1220648-2025-25764 to reflect the details of the patient outcome.
12018 | 1220648-2025-25765 | 21229462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21229462&pc=OZD | 12/30/2024 | 09/22/2025 |  | 01/24/2025 |  |  | 12/30/2024 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547840 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that during support, there were placement signal issues. the placement signal was not reliable. motor current, purge flow, and pressure have all remained stable. support continued. | The complainant reported the patient was on impella 5.5 support and during support, a leak at the yellow to yellow connection was observed. staff changed the purge cassette however this did not resolve the issue. the pressures and flows were in range so support continued and staff monitored. there was no patient harm. |  |  |  |  | The investigation of purge leak ¿ pump has been completed. the root cause of the purge leak issue was not determined since product was not returned. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs were investigated and all the 5s parameters were found to be in specification. no motor current spikes, suction alarms observed. no positioning issues were observed. impella position unknown alarms were observed prior to the placement signal not reliable (psnr) alarm. placement signal trend was seen to be trending down into negative value of 80mmhg and then going out of specification on 89th day of support. therefore, the root cause of the optical signal issue was sensor silicone wear which happened after the design life of the pump and is therefore misuse of the pump. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial. manufacturer device report number 1220648-2025-25765. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.". ¿clean the connector cable with 70% isopropyl alcohol.¿.
12019 | 1220648-2025-25766 | 21230036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21230036&pc=OZD | 01/07/2025 | 02/26/2025 |  | 01/24/2025 |  |  | 01/07/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577085 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction was implanted with an impella cp device and developed low pump flow with concerns of a clot. the impella was start in auto and there was very dampened motor current. flows were less than 1l. the position was optimized, but there was no improvement. concern was expressed for clot being a possible cause. impella device was explanted and decision made to proceed with a new pump. the patient was reported to be stable following the event. | [5 empty] | The device was not returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of low pump flow and thrombus has been completed. the data log were not returned for review. the pump was returned for analysis and visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. the root cause of low flow was unable to be determined as the failure mode could not be reproduced and no problem was found on the pump. as the necessary information was not provided, the root cause of the thrombus was not determined. d9 date device returned to manufacturer added.
12020 | 1220648-2025-25767 | 21230177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21230177&pc=OZD | 01/07/2025 | 02/24/2025 |  | 01/24/2025 |  |  | 01/07/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552573 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) was implanted with an impella cp device for mechanical circulatory support and approximately 11 days after start of the support, the patient had bloody stool. a gastrointestinal bleed was confirmed. the patient was successfully weaned, and the device was explanted. heparin was running in the impella 5.5 pump purge line and may have exacerbated the bleeding event. the patient was noted to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Pump was not returned for investigation. as per the clinical information provided, the gi bleed was found by the bloody stool of the patient but the location of the bleed is unknown. therefore the failure cannot be investigated further. therefore, the root cause of the vascular damage issue was not determined since product was not returned and the location of the bleed is unknown.
12021 | 1220648-2025-25768 | 21230361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21230361&pc=OZD | 01/22/2025 | 03/07/2025 |  | 01/24/2025 |  |  | 01/01/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576327 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the doctor attempted to reposition the pump without success. the pump was then explanted and replaced. there was no harm tot he patient. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. b3 date of event was revised as it was inadvertently reported incorrectly on the initial manufacturer device report number. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12022 | 1220648-2025-25769 | 21230698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21230698&pc=OZD | 11/25/2024 | 03/12/2025 |  | 01/24/2025 |  |  | 01/08/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547476 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Complainant in japan reported the patient was on impella cp smart assist support and during support, the patient had systemic bleeding. the patient's chest was reopened and hemostasis was achieved. the cause of the bleeding was not from the impella, however the doctor stated that the impella exacerbated the bleeding. the patient had a replacement pump and expired on support of the replacement pump. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. as per the clinical information provided, the bleeding was occurring prior to the impella insertion. after pump insertion, the bleeding was possibly exacerbated and hemostasis was achieved by reopening the chest. no other information regarding the vascular damage was provided. therefore, the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and the relation to impella is unknown. b.2 revised as required intervention was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25769. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25769 and a new code was added. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-25769 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
12023 | 1220648-2025-25770 | 21230703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21230703&pc=OZD | 01/04/2025 | 02/24/2025 |  | 01/24/2025 |  |  | 01/04/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565717 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | Ventricular Fibrillation | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient went hypotensive and then ventricular tachycardia (vt)/ventricular fibrillation (vf) multiple times with ability to shock out of it. k 2.9 treated, amio bolus and drip started. this kept patient out of vf/vt. additionally, there were suction event with likely cause thrombus ingestion. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus, vf/vt/af/at, and low pump flow has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the vt/vf was not determined. data logs show motor current spike followed by a decrease in flow and suction alarms. there was motor current pulsatility loss. low flow resolved after 36 hours. the cause of the low flow was most likely biomaterial. thrombus failure mode (fm) was removed since it is captured under low flow fm with the root cause of biomaterial. e4 should have been left blank on manufacturer device report 1220648-2025-25770. h6 code 2993 was reported incorrectly on manufacturer device report 1220648-2025-25770. new code was added to medical device problem code.
12024 | 1220648-2025-25771 | 21230717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21230717&pc=OZD | 01/06/2025 | 02/18/2025 |  | 01/24/2025 |  |  | 01/06/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556083 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp explant, the patient continued to bleed from their groin site despite tightening of the existing perclose sutures. an additional suture was placed and manual pressure was held to achieve hemostasis. the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, additional percloses placed over wire to resolve the bleed. e4 should have been left blank on manufacturer device report 1220648-2025-25771.
12025 | 1220648-2025-25772 | 21230810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21230810&pc=PYX | 01/09/2025 | 02/26/2025 |  | 01/24/2025 |  |  | 01/09/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025508353 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and during support, there was a high current alarm and with a motor off alarm. under fluoro the rp flex looked to be bent. the pump was restarted and motor current was still high. the pump was repositioned and once the pump was back in appropriate position, the pump was turned back on to p-2 and was increased appropriately without any further issues. after that incident, the doctor noted that they were unable to pull or push the rp flex for any repositioning confirming the tuohy burst lock was unlocked. after reviewing the position was appropriate and no need to reposition. motor current caused pump to stop, brought p-level to 0 and started again. the pump continued to support the patient for an additional 6 days with no patient harm. | [5 empty] | The investigation of temporary pump stop has been completed. the product was returned and evaluated. no abnormalities were found related to the temporary pump stop. there was no biomaterial or cannula kinks found. log review showed motor current spikes near the beginning of the case followed by a motor current high pump stop and alarms. the pump was able to be restarted as noted in the clinical details. there were no abnormalities seen with the logs after the temporary pump stop. the root cause of the temporary pump stop was most likely use-related since the pump stop occurred after the pump became incorrectly positioned after pulling out the peel-away sheath for sheath exchange. the pump was able to be restarted after repositioning the pump. instructions for use: impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25772 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact fax was revised as it was not included on the initial manufacturer device report number 1220648-2025-25772. h10 was revised to include instructions for use not included on the initial manufacturer device report number 1220648-2025-25772. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12026 | 1220648-2025-25773 | 21230865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21230865&pc=OZD | 01/09/2025 | 02/18/2025 |  | 01/24/2025 |  |  | 01/09/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585955 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding from the impella access site. the patient groin site previously stable had begun oozing despite proper patient positioning and dressing techniques. the patient was switched to bival and is off heparin. the patient required transfusion of packed red blood cells and platelets. it was further reported that the impella was explanted for patient to go on comfort measures and the patient expired after withdrawal. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
12027 | 1220648-2025-25774 | 21231091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21231091&pc=OZD | 01/02/2025 | 05/30/2025 |  | 01/24/2025 |  |  | 01/10/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1407754 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) as a bridge to transplant and approximately 14 days after start of support, the automated impella controller (aic) had a controller failure. the issue self-resolved. the staff were instructed to switch the aic, however support continued. there was no report or indication of interruption in support and there was no report of harm to the patient as a result of this event. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis revealed there were significant number of invalid placement signal log entries leading up to the 1/6 placement signal complaint. the lt logs show during these times that the signal nor reliable (snr) was low for an extended period of time corresponding with placement signal railing around 3000 mmhg. later in the case, around the 1/10 "aic failure" alert complaint, there were multiple instances in the aic logs of event triggering and resolving minutes later. event 1045 triggers a controller error not a controller failure as described in the complaint. the lt logs at this point show 5s parameters represented by -16384 as placement signal goes to 3000 again. device analysis: neither of the failures were reproduced during testing. the 5s was within specification as outlined in preventative maintenance and the analog voltages of the ccd array looked as expected with and without a pump. root cause: there was no controller failure, instead it was a controller error. the cause of the optical signal issue could not be determined as the failure mode was not reproduced and the optical bench was within specification. the controller error and loss of placement signal described is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. d9 device returned to manufacturer date added. | The investigation for the console (aic) controller failure (red alarm) has been completed. logs from the complaint date show placement signal, light, contrast all spiked as the snr dropped significantly. these triggered an unreliable placement alarm that was what the complaint was about. the pump was transferred due to the alarm. the cause of the abnormal optical signals could not be determined since the console was not returned. | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports related to the patient implant experience. refer to initial medical device report for the impella 5.5 device serial number (b)(6) with date of event 02-jan-205 and identifier ending in (b)(6).
12028 | 1220648-2025-25775 | 21231218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21231218&pc=OZD | 08/05/2024 | 03/04/2025 |  | 01/24/2025 |  |  | 01/10/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508642 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu in (b)(6) 2023 had a 5.5 pump support for 6 days for a high risk surgical patient. the pump was explanted and patient survived. later in (b)(6) 2025 the impact study forms documented an infection with possible relationship to the impella. wound healing disorder was treated with piperacillin/tazobactam. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. corrections to the initial mfr report: b1-adverse event should have been selected.
12029 | 1220648-2025-25776 | 21231476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21231476&pc=PYX | 01/10/2025 | 02/21/2025 |  | 01/24/2025 |  |  | 01/10/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590421 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella rp while on patient support flows decreased and purge pressure was high. the impella rp was exchanged. the patient survived the event. | [5 empty] | The investigation for the low pump flow has been completed. returned complaint device visually inspected. cannula kink near inflow cage observed. big chunk of biomaterial around impeller, purge gap and in the outflow cage observed. diagnostic catheter attached through the rp flex cannula and dried biomaterial accumulated inside the cannula. it looks the catheter passed through the cannula during explant in order to gain wire access to the access site to put another pump in to clamp the cannula to do this. cannula kink may occurred during removal. data logs were reviewed finding suctions occurred. high purge pressure observed. motor current (mc) spike outside of p-level changes occurred with drop in flows per reported event. no positioning issue observed. the root cause of the low pump flow issue most likely was biomaterial ingestion. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12030 | 1220648-2025-25777 | 21231594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21231594&pc=PYX | 01/05/2025 | 09/24/2025 |  | 01/24/2025 |  |  | 01/03/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025503925 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella rp device for mechanical circulatory support and began bleeding from an unknown source. the patient would eventually expire. the patient presented to the emergency department after having a ventricular fibrillation arrest at home. in emergency department the patient coded multiple times, intubated, and was transferred to catheterization lab for an angiogram. after 10% ejection fraction on left ventricular gram, the right coronary artery (rca) disease from angiogram, and cardiogenic shock with unstable hemodynamic status, the decision was made to implant an impella cp prior to percutaneous coronary intervention (pci). after the impella cp inserted, the pci was performed on rca, the patient lost pulsatility and continued to decompensate hemodynamically. a right heart catheterization was completed and pulmonary artery pulsatility index revealed as 0.3. the decision was made to escalate to an additional impella rp device (with heparin in the purge system). the patient's pulsatility came back as soon as the impella rp was started, and the mean arterial pressure was maintained above 55. the impella was secured, the sheath was secured and dressed, and the patient transferred to unit on bipella support. following the implant, pulses in the right leg (impella cp side) became weak and the right foot appeared mottled. the leg began to swell, and a fasciotomy was subsequently completed to restore pulse to the leg. four days later, pulse was once again absent in the patient's leg (impella cp side). the patient had been given ten units of blood as a result of the condition and intervention. coinciding with the ischemia, the patient also developed hemolysis with four units of bleed being given to help manage the hemolysis. the support level was lowered to reduce suction that may have been contributing to the hemolysis. as a result of the prior interventions, the patient experienced bleeding complications from the fasciotomy sites. multiple blood products were given to counteract the bleed and heparin levels were reduced. the patient later passed away after seven days of support. it is unknown if the complications contributed to the patient's passing. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. placement signal issue: the fm "optical signal issue" was updated to "placement signal issue" since the complaint involves the dp sensor. no product was returned. log review shows the placement signal drifting up before going out of spec and intermittently coming back in and out throughout the rest of the case. the pump flow drifted down before going out of spec (to zero). the motor current was stable and followed the changes in p-level. the cvp was also stable until going out of spec with the placement signal. psnr alarms related to the dp sensor were also triggered. the optical parameters were steady indicating that the optical sensor was not affected. the root cause of the placement signal issue was most likely sensor drift since data logs indicate the pattern. vascular damage - peripheral vessel: the bleeding was from the right leg fasciotomy procedures, which were performed to treat the limb ischemia in the right leg (cp site). per mso review, if heparin is used in the purge, the pump cannot be dissociated from any bleeding. heparin was used in the purge solution of the rpsa, so the rpsa cannot be dissociated from the right leg vascular damage. please refer to 25-40478-1 for further investigation into the right leg vascular damage. the root cause of the vascular damage was not determined since insufficient clinical details were provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | The event date reflects the same as the associated medical device report accordingly (best estimate) given totally of the information for the event. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. note: as heparin was in the impella rp purge, the impact/exacerbation to the bleeding is unknown. therefore, this is one of two medical devices associated with this event. refer to manufacturing report number 1220648-2025-25587. use of the repositioning sheath and the 23 fr peel-away introducer: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 manufacturing site name and address was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2025-25777.
12031 | 1220648-2025-25778 | 21232447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21232447&pc=OZD | 01/10/2025 | 02/26/2025 |  | 01/24/2025 |  |  | 01/10/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025572812 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump being placed to support a patient in need of coronary artery bypass graft. the pump was unable to be delivered and so instead a cp was placed. the patient survived the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the delivery issue was most likely patient condition related. the impella 5.5 could not pass innominate. d.4 revised model number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25778. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12032 | 1220648-2025-25779 | 21232558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21232558&pc=OZD | 01/11/2025 | 03/06/2025 |  | 01/24/2025 |  |  | 01/11/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606127 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preparation for patient support, an impella cp pump experienced a fluid leak from the purge filter. a new pump was subsequently prepped and implanted for patient support. there was no patient harm. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the pump was returned for evaluation. upon visual inspection, no damage to sidearm filter was observed. pump was attempted to be purged and a leak was observed from the sidearm filter air holes. clinical details noted that pump was leaking from the purge filter when attempting to prime pump. the root cause of the purge leak - pump was due to a damaged sidearm filter. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
12033 | 1220648-2025-25780 | 21232758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21232758&pc=OZD | 01/11/2025 | 02/24/2025 |  | 01/24/2025 |  |  | 01/11/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575711 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding at the impella access site. femstop was placed and one unit of red blood cells was administered. the patient is stable and survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. h.6: code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25780.
12034 | 1220648-2025-25781 | 21232912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21232912&pc=OZD | 01/13/2025 | 02/28/2025 |  | 01/24/2025 |  |  | 01/13/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484148 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, purge pressure high/purge system blocked alarms were present. the nurse reported tissue plasminogen activator protocol (tpa) purge fluid running since 2100. audio alarm was disabled. the patient remains on impella support. | [5 empty] | The investigation for the high purge pressure has been completed. the pump was not returned for investigation. the data log shows a small gradual rise and drop in purge pressure before a 10-minute gradual rise in purge pressure, culminating in a high purge pressure alarm. the purge pressure resolved with a gradual downward trend, and the clinical team suggests that the tpa helped to resolve the high purge pressure. the cause of the high purge pressure issue was most likely biomaterial, based on data log review and clinical detail. added h6 impact code 4644. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12035 | 1220648-2025-25782 | 21232980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21232980&pc=OZD | 01/13/2025 | 02/20/2025 |  | 01/24/2025 |  |  | 01/13/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129495 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) has a broken power cord. the entrance of the cord has been crushed. the aic was removed and has been returned for investigation. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) cabling issue has been completed. console logs from the reported day of event do not contain any information relevant to external condition of the console. console was inspected in and a cut in the ac cord insulation was confirmed. there was also damage to the rear and front housing of the console, consistent with equipment being dropped to the floor. console¿s electronics operated within specification during testing. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12036 | 1220648-2025-25783 | 21233341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21233341&pc=OZD | 01/13/2025 | 10/01/2025 |  | 01/24/2025 |  |  | 01/13/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576499 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella 5.5 was having an abnormal aortic waveform/placement signal. additionally, there were suction alarms with higher p-level. echocardiography was performed and impella position was acceptable. the patient survived the event. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the investigation for optical signal issue has been completed. data logs were not available and product was not returned for investigation. clinical details report that placement signal not reliable (psnr) did not trigger, though the patient update the day after the complication states that the placement signal was not displaying, which only occurs when psnr is active. position unknown alarms were reported, but pump position had been verified with echo. clinical details also report that they considered administering epinephrine and the pump flows and suction alarms had improved by the next day, but there is no confirmation. since product was not returned and clinical details are limited and unclear, the root cause of the optical signal issue could not be determine. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. corrections to the mfr report: h6 med dev prob code 2917 added. | The investigation for the low pump flow has been completed. data logs were not available and product was not returned for investigation. clinical details report that pump flows calculations had become inaccurate and suction alarms were triggering above p-5. the following day, it was reported that pump flows had improved and suction alarms had resolved, but the only troubleshooting method mentioned was that they considered starting the patient on epinephrin. with product not being returned and limited clinical details being provided, the root cause of the low pump flows could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12037 | 1220648-2025-25784 | 21233997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21233997&pc=OZD | 01/09/2025 | 02/27/2025 |  | 01/24/2025 |  |  | 01/09/2025 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577766 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was additionally reported that the patient had bleeding at the access site and the heparin was withheld. | The complainant reported the patient was on impella cp support and during support, there were low pump flows. staff implanted an impella rp and the suction alarms on the cp continued. the impella cp was removed and there was no patient harm. |  |  |  |  | The investigation has been completed. data logs showed steady placement signal, suction alarms and low pump flow with no motor current spikes. blood volume was given and the pump was repositioned but did not resolve the pump flow issues. right heart failure was noted. the cause of the low flow was not determined since the product was not returned and there was insufficient clinical information. | The cause of the access site bleeding - major was not determined since the product was not returned and there was insufficient clinical information. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.1 revised to reflect both adverse event and product problem since the first followup manufacturer device report number 1220648-2025-25784 was submitted. b.2 add selection due to information added in b.5 since the initial manufacturer device report 1220648-2025-25784 was submitted. b.5 information was added that was not included on the initial manufacturer device report 1220648-2025-25784. h.1 revised type of report since the initial manufacturer device report 1220648-2025-25784 was submitted due to new information in b.5. h.6 added codes since the initial manufacturer device report 1220648-2025-25784 was submitted due to additional information in b.5 h.10 added investigation results for the adverse event and instructions for use for the adverse event as they were not included on previously submitted manufacturer device report number 1220648-2025-25784. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12038 | 1220648-2025-25785 | 21234849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21234849&pc=OZD | 02/23/2024 | 02/11/2025 |  | 01/24/2025 |  |  | 02/23/2024 | 01/24/2025 |  |  | 11/30/2025 |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435568 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 65 year old male patient for mechanical circulatory support for myocardial infarction. it was reported that upon explanting the pump, the patient had a hematoma evacuation and washout on their right groin. the site was then surgically closed. however, further details states that due to the patient's rapid deterioration, the patient continued to code and a new pump was placed. the patient eventually expired while on support. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the cause of the access site bleeding was not determined since the product was not returned and not enough clinical information was provided. e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25785. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25785 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact facility name and reporting contact fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-25785. h.2 should have been left blank on the initial manufacturer device report number 1220648-2025-25785 as it is was an initial manufacturer device report.
12039 | 1220648-2025-25786 | 21235590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21235590&pc=OZD | 12/25/2024 | 01/24/2025 |  | 01/24/2025 |  |  | 12/25/2024 | 01/24/2025 |  |  |  |  | 01/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558933 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Death |  |  |  |  | ECMO AND PROTEK DUO | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported an impella 5.5, along with ecmo and protek duo, in a 72 year old female patient for a bridge to recovery for cardiomyopathy. the team observed a left ventricle (lv) clot and were unable to increase impella flows. however the pump remained on for support for over 32 days. care was eventually withdrawn and the patient expired. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
12040 | 1220648-2025-25720 | 21216797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21216797&pc=OZD | 01/04/2025 | 02/28/2025 |  | 01/23/2025 |  |  | 01/06/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579368 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and approximately 11 days after start of the support the pump stopped. subsequent conversation with the physician noted the pump stopped due to high motor current, and there was no attempt performed to restart the pump. the automated impella controller was shut off, and the pump was pulled back in the aorta. the patient passed away one day later referred to his cardiac disease. however, there is not enough evidence to exclude the impella cp device as an associated factor in the patient's outcome. | [5 empty] | The investigation for the pump stop has been completed. the returned product was inspected and an abnormally large purge gap was seen and the bearing was exposed. the data logs confirmed that the pump was run for 290 hours (approximately 12 days) longer than the design life of the impella which is 8 days. the root cause of the pump stop was determined to be due to bearing wear. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.9 device return date/information was added. b.2 revised as there was an incorrect selection on the initial manufacturer device report number 1220648-2025-25720. e.1 revised as facility name was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25720. g.1 revised manufacturing site name and address section was it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25720. h.10 revised instructions for use as they were reported incorrectly on the initial manufacturer device report number 1220648-2025-25720. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12041 | 1220648-2025-25721 | 21216952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21216952&pc=OZD | 01/06/2025 | 02/12/2025 |  | 01/23/2025 |  |  | 01/06/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814730 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | Appropriate Term / Code Not Available | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) did not recognize the purge disc during preparation for use with an impella 5.5 device. the aic was exchanged, and there was no report. reportable due to potential patient harm. mdr and pmda report required. | [5 empty] | The investigation for the console (aic) disc purge/flag issues has been completed. data logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) twice. device was returned for evaluation finding a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not recognized issue. added d9 device return date, h6 impact code 4629. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12042 | 1220648-2025-25722 | 21217017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21217017&pc=OZD | 11/20/2024 | 01/28/2025 |  | 01/23/2025 |  |  | 01/02/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573291 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was a purge leak, high purge pressure and saturated flow issues. the purge side arm of the impella 5.5 gen 2 was noted to be leaking inside the clear housing. it was confirmed that it looked like it was leaking at about 1-2 droplets per hour. the purge solution was sodium bicarbonate and the purge parameters were within normal limits. additionally, high purge pressure was encountered and a tissue plasminogen activator (tpa) bag was planned to be used. the pump also began to show intermittent signs of saturated flows with the left ventricle tracing above the placement signal and the difference in min/max flows was decreasing. eventually, the impella flow was turned down and which cleared the purge and no further leaking was encountered. | [5 empty] | The investigation of purge leak ¿ pump, high purge pressure, and saturated flow has been completed. the data logs were returned and investigated. the root cause of the purge leak issue was a leak from the sidearm but the exact location of the leak was unknown. the root cause of the high purge pressure issue was most likely biomaterial buildup. the following have been reported incorrectly or missing from manufacturer device report 1220648-2025-25722: b3 date of event incorrect. d4 model number incorrect. d6b explant date incorrect. e4 should have been left blank. g1 reporting contact first name incorrect. g1 reporting contact last name incorrect. g1 reporting contact email revised in accordance with updated procedures. h10 investigation was omitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12043 | 1220648-2025-25723 | 21217100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21217100&pc=OZD | 01/07/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/07/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582883 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported companion sheath was contaminated. no patient harm has been reported. | [5 empty] | The investigation of packaging issues- sterility has been completed. neither companion sheath was not returned for investigation nor the pictures were provided. clinical information is very little and gives no information about how the sterility was compromised. therefore, the root cause of the packaging issue was not determined since product was not returned. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12044 | 1220648-2025-25724 | 21217276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21217276&pc=OZD | 01/06/2025 | 02/14/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593395 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 14fr introducer distal end kinked. the long sheath was exchanged for the short peel-away sheath. there was no harm to the patient. | [5 empty] | The investigation of introducer damage ¿ mechanical has been completed. the product was received for investigation and a sheath kink was found. the scoreline was not split. the cause of the sheath kink was not determined since there is insufficient clinical information. d9 date device returned to manufacturer added. b3 date of event was inadvertently reported incorrectly on manufacturer device report. g2 report source; health professional was missing from manufacturer device report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12045 | 1220648-2025-25725 | 21217532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21217532&pc=OZD | 01/08/2025 | 03/07/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571334 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump supporting a patient admitted in with acute myocardial infarction / cardiogenic shock. the pump was supporting but was replaced after just 25 hours due to blood loss at the repositioning sheath. the cp was replaced by a 5.5 pump. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. h.6 code 4641 was reported incorrectly on initial manufacturer device report number 1220648-2025-25725.
12046 | 1220648-2025-25726 | 21217833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21217833&pc=OZD | 01/07/2025 | 03/13/2025 |  | 01/23/2025 |  |  | 01/07/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025549014 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that roughly eleven days into impella 5.5 support, a leak was discovered inside the sidearm retainer. as a result of the leak, crystalized material was beginning to form in the retainer. the leak had no effect on the purge pressure, purge flow rate, or functionality of the pump. as a result, the decision was made to monitor the issue with no further action required. there were no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. a leak was reported inside the side arm retainer during use. no product nor data logs were returned for evaluation. the cause of the purge leak - pump was a leak from the sidearm but the exact location of the leak was unknown as product was not returned.
12047 | 1220648-2025-25727 | 21218038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218038&pc=OZD | 01/08/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | Death |  |  |  |  | CONTINUOUS RENAL REPLACEMENT THERAPY. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had access site bleeding. pressure was held and angle matching was performed. heparin was withheld until labs were ran. additionally, while on support, the patient had hemolysis. dialysis was started and the patient became hypotensive so continuous renal replacement therapy (crrt) was started. the patient did have acute kidney injury prior to impella support, however it is unknown if the impella contributed to the patient's kidney issues. care was withdrawn and the patient expired after 68.47 hours of impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the hemolysis, access site bleed, renal failure has been completed. pump was not returned for investigation. as per the clinical information provided, hematoma and access site was bleeding on day 2 and was addressed with manual dressing and angle matching. no other information was provided about acts and heparin was withheld. therefore, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. data logs were investigated and around the time of reported failure there were no positioning issues, suction alarms, motor current issues, or placement signal trends. clinical mentioned that the pump was seated deep and repositioned at 4.1cm mark but data logs don't support that. therefore, the root cause of the hemolysis issue was not determined since product was not returned. as per the clinical information provided, the patient has an acute kidney injury prior to impella support. on day 3 of impella support when patient was going through dialysis, the patient was switched to crrt due to patient being hypotensive. it is unclear if the impella pump contributed to the renal issues. therefore, the root cause of the renal issue was not determined since product was not returned and insufficient clinical information was provided.
12048 | 1220648-2025-25728 | 21218081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218081&pc=PYX | 12/27/2024 | 09/22/2025 |  | 01/23/2025 |  |  | 12/27/2024 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560599 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Cardiac Arrest | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp flex device for mechanical circulatory support (mcs) and developed high purge pressure. the patient had a central right ventricular assist device (rvad) for four days, and was taken for rvad decannulation two days prior to the impella rp flex implant. during this procedure large amount of clots in chest was noted, decannulation was aborted and the chest was left open. the rvad decannulation and impella rp flex implant was completed two days later. two days after start of impella mcs, it was noted patient was resting on the impella rp flex support while assessed for right ventricular recovery post durable left ventricular assist device procedure. now 16 days after start of impella mcs, high purge pressure and purge system blocked alarms were encountered. resolution and action taken are unclear. approximately eleven days later the patient expired. care was withdrawn; however, there is not enough information to exclude the event with impella rp flex as an associated factor with the patient's outcome. | The complainant reported that the impella rp while on patient support had high purge pressure and placement signal issue. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. no product was returned for analysis. the logs showed no motor current spike before high purge pressure alarms. the logs showed a gradual rise in purge pressure before alarms were triggered. there was a motor current spike right before high purge pressure was resolved likely as a result of biomaterial being pushed off the impeller. it is unknown if tissue plasminogen was added to the purge. the root cause of high purge pressure was most likely biomaterial e.1 facility name and us phone number were inadvertently omitted from initial manufacturer device report number 1220648-2025-25728 and were added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25728 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4614 and 1762 were reported incorrectly on initial manufacturer device report number 1220648-2025-25728. a new code was added. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the failure mode "optical signal issue" was changed to "placement signal issue" as issue is related to dp sensor of the pump. the failure mode "high purge pressure" was included as high purge pressure was reported in complaint. no product was returned for analysis. the data logs confirm placement signal not reliable alarms #1051 dp sensor out of range. the logs show placement signal drifting downwards, with flows drifting upwards before placement signal went out of range. the placement signal recovers for some hours before drifting and going out of range again. the cvp is out of specification and flow calculation is disabled. the logs show no motor current spike before high purge pressure alarms. the logs show a gradual rise in purge pressure before alarms were triggered. there was a motor current spike right before high purge pressure was resolved likely as a result of biomaterial being pushed off impeller. it is unknown if tpa was added to the purge. the root cause of the placement signal issue was most likely sensor drift. the root cause of high purge pressure was most likely biomaterial. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
12049 | 1220648-2025-25729 | 21218088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218088&pc=OZD | 01/03/2025 | 02/26/2025 |  | 01/23/2025 |  |  | 01/03/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580396 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, oozing was noted from the companion sheath diaphragm during procedure while 7fr guide was inserted. treatment of the bleeding is unknown and there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the companion sheath and introducer were returned for analysis. the companion sheath had a damaged sheath tip and leak was reproduced in the lab. the leak occurred at the valve at about 120mmhg. the introducer was tested and the leak did not reproduce till 290-300mmhg. the root cause of the access site bleeding - major was most likely due to a damaged valve of the companion sheath. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25729 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12050 | 1220648-2025-25730 | 21218093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218093&pc=OZD | 01/06/2025 | 02/05/2025 |  | 01/23/2025 |  |  | 01/06/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1654996 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that during inspection outside of patient use, an automated impella controller (aic) experienced a system self-check failure and would not start up. the aic was rebooted twice; however, the issue persisted. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the reported failure mode was reproduced during testing at the field service. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. a.1 and a.3 revised as selections were inadvertently omitted from the initial manufacturer device report number 1220648-2025-25730. d.5 revised section as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25730. e.1 revised facility name and address line 2 as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-25730. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25730 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
12051 | 1220648-2025-25731 | 21218099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218099&pc=OZD | 01/06/2025 | 02/17/2025 |  | 01/23/2025 |  |  | 01/06/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. following the initial insertion, the physician was unable advance the device due to iliac stenosis. the pump was removed and a balloon angioplasty of the left iliac was performed. however, upon reinsertion, the device was still unable to advance due to the patient's anatomy and the decision was made to abort the implant. an intra-aortic balloon pump was placed via the right groin post percutaneous coronary intervention for ongoing support. the patient was reported to be stable following the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient condition since unable to advance impella due to iliac stenosis and calcification of the vessel. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ e4 should have been left blank on manufacturer device report. d6a and d6b were inadvertently omitted from manufacturer device report. h10 instructions for use were inadvertently omitted from manufacturer device report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12052 | 1220648-2025-25732 | 21218120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218120&pc=OZD | 01/08/2025 | 02/17/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537950 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the pump shifted out of position and triggered supraventricular tachycardia. cardioversion was performed to stabilize the patient and planned support continued uninterrupted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided, the root cause of the vt/vf was not determined. e4 should have been left blank on manufacturer device report 1220648-2025-25732.
12053 | 1220648-2025-25733 | 21218138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218138&pc=PYX | 01/09/2025 | 02/25/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  | 04/30/2025 |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024334331 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Malfunction | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp support and during support, the pump came out of position and the doctor was unable to reposition the pump. during this event, a large hematoma was noted with bleeding. the decision was made to stop heparin and remove the pump. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp with smart assist system with automated impella controller: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation for the positioning issues and access site bleed has been completed. product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. added h6 clinical code 4582 corrections to the initial mfr report: b1-deselected adverse event b2-deselected required intervention h6 clinical code 1888 and 1884 removed *access site bleeding failure mode was removed because bleeding was reported from the cp right femoral artery site. access site bleeding failure mode has been reported in mdr mfr report #1220648-2025-25784.
12054 | 1220648-2025-25734 | 21218147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218147&pc=OZD | 01/09/2025 | 02/12/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1615855 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported damage to the automated impella controller (aic). the purge cassette is loose, it does not hold in the latch. the cassette does not hold on to the center piece in the back. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) damage before therapy has been completed. console logs from the complaint date were analyzed as a precaution. numerous imc-ahp error prompts were observed in the logs which indicates that there may have been some intermittent communication issues between the rfid and the purge cassette. console was returned for evaluation. however, the reported loose purge cassette issue was unable to be reproduced. the console was able to function normally while running with a purge cassette kit and pump. while testing this console, the purge insert had no problem fitting a known-good purge cassette. the purge cassette functioned normally and stayed in the purge insert when the purge door was closed which is how it is intended to be used. this loose purge cassette issue was likely use related. it is likely that the purge door was not closed after insertion of the purge cassette into the consoles purge assembly. this ensures that the purge cassette stays in place during use. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12055 | 1220648-2025-25735 | 21218152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218152&pc=OZD | 01/10/2025 | 02/26/2025 |  | 01/23/2025 |  |  | 01/10/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567092 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following successful placement, bleeding was noticed from the hemostatic valve of the introducer. the companion sheath was adjusted with some improvement, but the bleed persisted. additional intervention was not required and the supported procedure was completed without complication. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the introducer was pressure tested with the companion sheath and it passed the spec of180 mmhg per specifications without leaking. the root cause of the access site bleeding issue was most likely use related, unable to cross aortic valve over abiomed 0.018 wire so different company wire was used. d9 was revised on manufacturer device report number 1220648-2025-25735 to reflect the device was returned to the manufacturer e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25735 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12056 | 1220648-2025-25736 | 21218419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218419&pc=OZD | 01/08/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505440 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an impella 5.5 implant and there was a controller failure with a pump stop. the pump was explanted and replaced. the patient survived the event. | [5 empty] | The investigation of pump stop has been completed. no product was returned for evaluation. data logs were reviewed and real time log of the pump run showed pump stop occurred 21 minutes after the pump was started. ¿impella stopped ¿ motor current high¿ was triggered at the time of pump stop. a motor current spike was observed at time of pump stop. multiple attempts were made to restart the pump. however, the pump was not able to remain on when attempting to restart. clinical details noted that pump stop occurred on the first day of support. the root cause of the pump stop could not be definitively determined as no product was returned and limited clinical details were provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12057 | 1220648-2025-25737 | 21218690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21218690&pc=PYX | 01/08/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  | 08/31/2025 |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395738 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp support. while on support, the patient had renal failure. creatinine jumped from 1.1 to 4.0 and urine output dropped significantly; patient was anuric. additionally, the patient had hemolysis. patient was given two units of red blood cells. impella was removed due to hemolysis. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of hemolysis and renal failure has been completed. the cut pump was returned for investigation. biomaterial was found in the inlet cage. no further investigation was performed due to the cut product. the data log shows a small spike in motor current, with an observed rise in placement signal and a drop in flow on the same day of pump use. the cvp dropped, and a rising trend in purge pressure was seen during the impeller interaction event. additionally, pump flow was slightly lower than expected during the rest of the case, with some trends noted on the placement signal and motor current. the cvp was positive in the meantime. no abnormalities were reported on 5s optical signal parameters during case run. the cause of the hemolysis issue was most likely biomaterial, based on data log review and returned product investigation. the cause of the renal failure issue was not determined due to insufficient clinical details. d4 unique identifier (udi) # was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25737.
12058 | 1220648-2025-25738 | 21219218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21219218&pc=OZD | 01/08/2025 | 02/17/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513191 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient endured premature ventricular contractions (pvc)/ ventricular tachycardia (vtach) arrhythmias which required defibrillation. device was repositioned due to being close to the septum. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12059 | 1220648-2025-25739 | 21219461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21219461&pc=OZD | 01/23/2024 | 02/10/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  | 11/30/2025 |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024424083 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Impact study reported that the patient on impella 5.5 support experienced a subacute infarct. the patient was examined in the morning on (b)(6) 2024 and noted to be following commands and moving the right upper extremity (rue). in the afternoon, sedation was held, and patient was re-examined. at this time, the patient was not following commands and with very minimal movement of rue. computed tomography showed subacute infarcts in the left parietooccipital region and right occipital lobe. the patient recovered on the same day of the event. per the study, this event is possible related to the study device and procedure. the patient survived the event. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact fax number was inadvertently omitted from manufacturer device report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12060 | 1220648-2025-25740 | 21219579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21219579&pc=OZD | 01/09/2025 | 02/14/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  | 02/28/2026 |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474130 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported delivery issues with an impella 5.5. unable to pass 5.5 due to subclavian disease. procedure was aborted, and a cp was placed instead. no patient harm has been reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related as they were unable to pass the impella 5.5 due to subclavian disease. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12061 | 1220648-2025-25741 | 21219632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21219632&pc=OZD | 01/09/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579750 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, a placement signal alarm occurred. the patient also had labs that were hemolyzed. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and optical signal issue has been completed. the pump was returned for investigation. no physical abnormalities were reported on the returned product. the pump passed the laser continuity test, and all the parameters were found to be within the specified range. the pump was run according to specifications in a bucket of water. the data log shows several placement signal low alarms while running on a higher p-level, with a drop in snr and a trending placement signal. the pressure level was reduced to p4 after reporting a suction alarm and low pump flow, which continued for a day of case run, accompanied by a trending placement signal and motor current. a recovery trend was observed in the motor current and pump flow towards the end of the case run. the cause of the hemolysis issue was patient condition related, based on data log review and returned product investigation. b5-the description of the event incorrected cited that the impella 5.5 was supporting the patient. however, this should have stated that the patient was on impella cp support.
12062 | 1220648-2025-25742 | 21220137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21220137&pc=OZD | 01/09/2025 | 02/07/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566570 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump inserted for an acute myocardial infarction / cardiogenic shock patient when, after a day of support, the unresolving low flows/suction prompted pump explant and replacement. the patient survived the event. | [5 empty] | The investigation into the reported issue has been completed. pump passed all post sterile inspection checks. data logs confirmed the failure and clinical narrative. logs showed after pump started and run for 17 hours, motor current suddenly spiked to 1500 mmhg followed low flow and suction. this event remained to the rest of the case. there is a brief position alarm which resolved; clinical states repositioning pump did not resolve issue. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely due to biomaterial ingestion. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12063 | 1220648-2025-25743 | 21221862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21221862&pc=OZD | 01/23/2024 | 05/28/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  | 09/30/2025 |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024414988 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complaints team received an impact study that found a culture from the device following explant was positive for traces of e. faecium. the patient was treated with vancomycin but switched to linezolid per infectious disease recommendation. the patient did not have a fever or elevated white blood cell count. the patient survived the event. | Additional information received from the impact study stating the patient had an axillary artery injury with a possible relationship to the impella device. no further details were provided. |  |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 was inadvertently omitted from manufacturer device report 1220648-2025-25743. g2 report source; study was missing from manufacturer device report 1220648-2025-25743. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. b5 updated with additional information received from the clinical site h6 revised to included additional information received from the clinical site.
12064 | 1220648-2025-25744 | 21222991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21222991&pc=OZD | 01/05/2025 | 04/04/2025 |  | 01/23/2025 |  |  | 01/06/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573130 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and approximately 16 hours after start of support, it was reported the pump's position was completely in the aorta and repositioning was impossible. patient critical was taken back to the catheterization lab, and the pump was explanted and replaced with new impella cp device. the patient was returned to the intensive care unit on impella mcs. however, within a couple of hours on the second impella cp device, it was noted to be dislocated; this pump slipped into the aorta as well. according to the physician, the tuohy borst valve was fixed. consultation was made with the chief physician and the decision was made to explant and place another new (now third implant) impella cp device as it was noted the patient was in ¿very¿ critical condition. the third replacement pump was successfully placed. however, the following day within some hours the patient expired. | [5 empty] | After further review the complainant reported the patient was on impella cp support and during support, the pump was completely in the aorta. repositioning was not possible and a new pump was implanted. the patient had a replacement pump and expired on support of the replacement pump. the patient's death has been documented in manufacturer device report 1220648-2025-25747 b1 revised from manufacturer device report number 1220648-2025-25744 to reflect mso review from death to product problem. b2 revised from manufacturer device report number 1220648-2025-25744 to reflect mso review from death to product problem was left blank. h1 type of reportable event revised from manufacturer device report number 1220648-2025-25744 to reflect mso review from death to malfunction. h6 health effect - impact code 1802 omitted from manufacturer device report number 1220648-2025-25744 to reflect mso review from death to product problem. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports related to the patient implant experience. refer to initial medical device report for impella cp serial number (b)(6) with date of event 05-jan-2025 and identifier ending in (b)(6). | The investigation of positioning issues has been completed. the pump was not returned. the impella cp had to be replaced due to positioning issues. a second pump was placed, also having positioning challenges. the patient conditions or any anatomical abnormalities are unknown. the cause of the positioning issue was unable to be determined due to lack of clinical details. a.3 revised as selection was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25744. d.3 revised as us zip code was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25744. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25744. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25744 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12065 | 1220648-2025-25745 | 21223010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21223010&pc=OZD | 01/08/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593776 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had hemolysis. the patient had dark cherry urine. the pump was repositioned and fluid volume was given. additionally, neither dorsalis pedis or posterior tibial were able to be found by doppler bilaterally. bilateral lower extremities were reported as warm to touch with cool feet. the pump was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the hemolysis and limb ischemia has been completed. the returned complaint device was visually inspected. biomaterial was observed around the purge gap and impeller. no sharp edge or broken blade were found. no other abnormalities found. clinical details stated that the plasma free hgb down to 40 after impella repositioning and fluid bolus, urine clear and yellow. the root cause of the hemolysis issue most likely was improper positioning related since plasma free hgb down to 40 after impella repositioning. biomaterial collected around impella during support. no logs available to review. the root cause of the limb ischemia could not be determined without additional information. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12066 | 1220648-2025-25746 | 21223356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21223356&pc=OZD | 01/09/2025 | 02/18/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579794 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in with acute myocardial infarction / cardiogenic shock and history of atrial fibrillation. the patient was transferred on support from community to tertiary center for possible care escalation. the access site oozed and the bleed was treated by manual pressure, tightening of the sutures, and applied a hemostatic dressing. the pump remained on until the escalation to the 5.5 pump as prior discussed. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access issue was most likely use related, tightening pre-close sutures, and applying hemostatic dressing stopped the bleeding at the impella site.
12067 | 1220648-2025-25748 | 21224204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224204&pc=OZD | 01/06/2025 | 02/20/2025 |  | 01/23/2025 |  |  | 01/06/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582191 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the pump continuously ran for approximately 12 days. the patient was taken to the catheterization lab for a pump exchange and during weaning, the pump stopped. the pump was unable to be restarted, and the patient was medically managed until a new impella pump was implanted. it was reported there was no harm to the patient as a result of the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the pump stop has been completed. the impella cp was returned for evaluation. no physical abnormalities were found on the returned product, and the pump passed electrical testing. upon further inspection, a large purge gap was found which is due to bearing wear. data logs were reviewed, and a high motor current pump stop with alarm #13 occurred. multiple attempts were made to restart the pump without success. the root cause of the pump stop was due to bearing wear as pump was run past indicated design life. the instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿.
12068 | 1220648-2025-25749 | 21224267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224267&pc=OZD | 01/07/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/07/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536234 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Th complainant reported the patient was on impella cp support. while on support, persistent suction and low flow present despite adequate preload and position. the pump was repositioned, lvedp 20s, and good right ventricle function was evaluated before removing device and replacing with new impella. the patient never showed any improvement in neuro status. brain flow study completed and indicative of brain death. the decision was made to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the low pump flow has been completed. data logs were reviewed finding no (motor current) mc spike observed. no positioning issue alarm. no placement signal occurred. the cause of the low pump flow issue cannot be determined since the complaint device not returned for analysis and lack of clinical details.
12069 | 1220648-2025-25750 | 21224292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224292&pc=OZD | 01/02/2025 | 01/23/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577563 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support had a computed tomography scan that showed a cerebral hemorrhage. activated clotting time was within the usual control range. the patient will be monitored without additional treatment. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿