30635 | 1220648-2024-07167 | 18799760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18799760&pc=OZD | 09/19/2022 | 03/12/2025 |  | 02/28/2024 |  |  | 09/19/2022 | 02/28/2024 |  |  | 05/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230096 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2022 |  |  |  |  | The user facility reported a 78-year-old female patient was being implanted with an impella cp device for mechanical circulatory support. after the procedure was completed, it was noted that the patent had multiple access site bleeds. forward sutures and angle matched were used to resolve the oozing, however it was noted that the sites were still oozing. therefore, the impella was explanted. additional sutures were deployed and the access site was cleaned without any reported issues. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30636 | 1220648-2024-07168 | 18800717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18800717&pc=OZD | 09/15/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 09/15/2022 | 02/28/2024 |  |  | 05/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023223652 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/15/2022 |  |  |  |  | The user facility reported a 82-year-old female noted as a high risk percutaneous coronary insertion was being placed on an impella 5.5 device for mechanical circulatory support. it was reported that the medical staff was unable to deliver the impella pump due to the patient's anatomy. all attempts were unsuccessful and the procedure was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30637 | 1220648-2024-07169 | 18800859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18800859&pc=OZD | 09/15/2022 | 03/11/2025 |  | 02/28/2024 |  |  | 09/15/2022 | 02/28/2024 |  |  | 12/31/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171827 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2022 |  |  |  |  | The user facility reported a 77-year-old male in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, it was reported that the patient pulled their central line out, and bleeding around the impella cp access site was observed. the physician repositioned the impella cp device and re-stitched the blue suture hub to resolve the bleeding at the access site. | [5 empty] | The impella cp device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30638 | 1220648-2024-07170 | 18801969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18801969&pc=OZD | 09/16/2022 | 03/11/2025 |  | 02/28/2024 |  |  | 09/16/2022 | 02/28/2024 |  |  | 04/30/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223292 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The user facility reported a 21-year-old male was implanted with an impella cp device requiring mechanical circulatory support. it was reported that the patient had bleeding from multiple sites and was in disseminated intravascular coagulation (dic). there was no information provided on the treatment of the bleeding issues. additionally, the impella cp device had to be repositioned under trans-esophageal echocardiography (tee). the device was weaned and explanted but the patient continued on extracorporeal membrane oxygenation (ecmo) support. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30639 | 1220648-2024-07171 | 18802979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18802979&pc=OZD | 10/03/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 10/03/2022 | 02/28/2024 |  |  | 11/30/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169541 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/03/2022 |  |  |  |  | The user facility reported a 83-year-old male patient needing mechanical circulatory support. it was reported that after being placed on impella cp support, the patient developed a bleed at the groin site. the impella was removed from the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30640 | 1220648-2024-07172 | 18803892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18803892&pc=OZD | 04/21/2022 | 03/10/2025 |  | 02/28/2024 |  |  |  | 02/28/2024 |  |  | 11/30/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023169734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 60 year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced oozing from the insertion site. epinephrine was injected at the site to gain hemostasis. the patient also experienced diminished pulses in the right foot. ultrasound was performed. a small occlusion was detected near the impella site in the groin; however, the foot appeared to be warmer and was not getting any worse. no treatment was required. the patient continued on support. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia ¿ reversible and access site bleeding - major issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30641 | 1220648-2024-07173 | 18803915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18803915&pc=OZD | 04/26/2022 | 02/28/2024 |  | 02/28/2024 |  |  |  | 02/28/2024 |  |  | 12/31/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023170342 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/26/2022 |  |  |  |  | The complainant reported that the 60-year-old male patient was implanted the impella 5.5 pump for mechanical circulatory support. it was reported that the pump was unable to be advanced past the level of the innominate artery due to angle and heavy calcification. the impella 5.5 device was removed and a impella cp was placed without issue. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issues was most likely due to the patient condition, as patient anatomy had heavy calcification. this report is being filed as part of a retrospective review of historical records.
30642 | 1220648-2024-07174 | 18803985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18803985&pc=OZD | 04/26/2022 | 09/21/2024 |  | 02/28/2024 |  |  |  | 02/28/2024 |  |  | 06/30/2022 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1492674 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2022 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported the 45-year-old male patient was on impella 5.0 support for mechanical circulatory support. it was reported that the impella 5.0 had purge pressure high alarms and flows were less than 1. the purge was changed to tissue plasminogen activator (tpa) and this resolved the issue. additionally, the patient went into ventricular fibrillation and automatic implantable cardioverter-defibrillator (aicd) shocked the patient. the patient was shocked more than 50 times in a 24 hour time span. care was eventually withdrawn the patient expired. no allegations were made towards the impella for cause of death. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation into the reported ¿high purge pressure¿ and ¿ventricular fibrillation¿ has been completed. no products were returned for investigation. however, data logs were returned and investigated. there was a gradual purge pressure buildup and purge system block alarms suggesting biomaterial buildup. the most likely root cause of the high purge pressure issue was most likely biomaterial. additionally, any device contacting the heart wall in conjunction with a poor patient condition could result in vf/vt but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the vf/vt is unable to be determined. this report is being filed as part of a retrospective review of historical records.
30643 | 1220648-2024-07175 | 18803992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18803992&pc=OZD | 05/02/2022 | 02/28/2024 |  | 02/28/2024 |  |  |  | 02/28/2024 |  |  | 01/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183124 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported that a 32-year-old female patient had attempted delivery of an impella 5.5 pump for mechanical circulatory support. it was reported that the device could not be advanced due to the size of the patient's vessel. as a result, the implant was aborted and an impella cp was subsequently placed for patient support. there was no patient harm. | [5 empty] | The investigation into the reported issue has been completed. no device was returned for evaluation. the root cause of the reported issue was most likely due to the patient condition of axillary vessel size. this report is being filed as part of a retrospective review of historical records.
30644 | 1220648-2024-07176 | 18804008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18804008&pc=OZD | 05/23/2022 | 05/09/2024 |  | 02/28/2024 |  |  |  | 02/28/2024 |  |  |  |  | 02/28/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1652655 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The complainant reported that the 40-year-old male patient underwent successful impella 5.5 implant. the automated impella controller (aic) failed to recognize an installed purge cassette. the aic was replaced to resolve the noted issue and support continued. there was no patient harm. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the console was tested finding that high purge pressure, low purge flow, and purge fluid not detected issues were unable to be reproduced but were likely due a use-related cause. console was able to detect the purge cassette/disc normally and the purge sensor assembly was visibly sound. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. data logs were not available for the date of occurrence. the root cause of this issue was a broken purge retainer. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30645 | 1220648-2024-07177 | 18804025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18804025&pc=OZD | 09/23/2022 | 02/28/2024 |  | 02/28/2024 |  |  |  | 02/28/2024 |  |  | 04/30/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224069 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported a 65-year-old male patient was implanted with an impella cp urgently during high risk percutaneous coronary interventions (hrpci). while being on impella cp support, the patient had some bleeding at the insertion site. the physician removed the pump at beside due to the concerns with the bleeding. no further information was provided concerning the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30646 | 1220648-2024-07178 | 18804202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18804202&pc=OZD | 01/29/2024 | 03/22/2024 |  | 02/28/2024 |  |  |  | 02/28/2024 |  |  | 11/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024432631 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 47-year-old male patient presenting for acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circularoty support. following implant, the impella cp pump was started with the guidewire still in place. the pump stopped and was subsequently replaced. there was no patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. the device nor data logs were returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the pump stop was user-related due to improper technique of starting the pump while the guidewire was still in place. d4-correct model number. | The impella device was not returned by the customer and therefore, evaluation of the device was not possible. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results.
30647 | 1220648-2024-07080 | 18788494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18788494&pc=OZD | 04/20/2022 | 02/27/2024 |  | 02/27/2024 |  |  | 04/20/2022 | 02/27/2024 |  |  | 10/31/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED |  |  | US |  | 2022152028 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The user facility reported a 60-year-old female patient on impella cp for mechanical circulatory support needed to be escalated to an impella 5.5. during insertion of the impella 5.5, the patient's anatomy would not allow the device to advance past the anastomosis. the physician decided to continue use of the current impella cp. no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature due to patient's anatomy. this report is being filed as part of a retrospective review of historical records.
30648 | 1220648-2024-07081 | 18788705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18788705&pc=OZD | 10/17/2022 | 03/11/2025 |  | 02/27/2024 |  |  | 10/17/2022 | 02/27/2024 |  |  | 06/30/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The complainant reported that a 45-year-old male patient was on impella cp support. the patient lost distal pulses on the impella side hemolysis was noted overnight. urine was scant red. distal pulses could be obtained the following day but peripheral sat on impella side had been trending down from 78% to 50%.the patient continued to have hemolysis and the physician was unable to the optimize the pump¿s position as the pump was posterior, behind the papillary muscle. the patient has distal perfusion cannula (dpc) from venoarterial extracorporeal membrane oxygenation (va ecmo). the patient was started on continuous renal replacement therapy (crrt). the impella cp was subsequently removed, and the patient remained on va ecmo for support. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30649 | 1220648-2024-07082 | 18788753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18788753&pc=OZD | 10/18/2022 | 03/11/2025 |  | 02/27/2024 |  |  | 10/18/2022 | 02/27/2024 |  |  | 06/30/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247260 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 85-year-old female patient for high risk percutaneous coronary insertion (hrpci). it was reported that the impella cp easily passed through the introducer and was placed in the left ventricle (lv). the pump was started in auto with average flows of 1-1.5 l with suction. however, the position of the pump within the patient appeared to be the culprit and the impella was removed and rewired. the impella cp position was optimal and the pump speed slowly increased to auto. no patient harm was detected. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
30650 | 1220648-2024-07083 | 18788770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18788770&pc=OZD | 10/21/2022 | 04/15/2025 |  | 02/27/2024 |  |  | 10/21/2022 | 02/27/2024 |  |  | 06/30/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247263 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported that a 53-year-old male patient, while on being placed on impella cp support experienced oozing at the groin site around the peel away sheath. the patient continued oozing from the groin site that was not amenable to manual pressure, surgical, or local anesthetics. the patient then received a total of three units of packed red blood cells. groin issues were complicated by high vagal tone resulting in decreased blood pressure and heart rate when manipulating the groin site. the pump was eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30651 | 1220648-2024-07084 | 18788955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18788955&pc=OZD | 04/21/2022 | 05/20/2025 |  | 02/27/2024 |  |  | 04/21/2022 | 02/27/2024 |  |  | 12/31/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED |  |  | US |  | 2023180074 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 70-year-old male patient on impella cp for mechanical circulatory support. the echocardiogram images were reviewed and it appeared that the impella cp was under the pap muscle. the physician was unable to torque free without pulling it back across valve. the physician made the decision to leave and monitor. the device was later repositioned. no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30652 | 1220648-2024-07085 | 18789313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18789313&pc=OZD | 06/20/2022 | 02/27/2024 |  | 02/27/2024 |  |  | 06/20/2022 | 02/27/2024 |  |  | 10/31/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154706 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a (b)(6)year-old male patient needed mechanical circulatory support. the impella 5.5 pump was unable to advance through the aortic arch during attempted delivery due to the patient's anatomy. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The investigation into the reported delivery issues- anatomy has been completed. no products were returned for investigation. the root cause of the delivery issue - anatomy was most likely due to the patient condition of not being able to pass the aortic arch. this report is being filed as part of a retrospective review of historical records.
30653 | 1220648-2024-07086 | 18789498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18789498&pc=OZD | 06/26/2022 | 02/27/2024 |  | 02/27/2024 |  |  | 06/26/2022 | 02/27/2024 |  |  | 03/31/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023208431 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/26/2022 |  |  |  |  | The user facility reported a 85-year-old female requiring impella 5.5 for mechanical circulatory support. it was reported that after multiple attempts, the 5.5 would not pass through the innominate turndown. the procedure was aborted and the intra-aortic balloon pump (iabp) was left in place. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient anatomy related. no product damage noted. this report is being filed as part of a retrospective review of historical records.
30654 | 1220648-2024-07087 | 18789957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18789957&pc=OZD | 02/01/2024 | 04/22/2024 |  | 02/27/2024 |  |  | 02/16/2024 | 02/27/2024 |  |  | 02/29/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2023191006 |  | 005066 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella 5.0 support and there was an infection noted at the access site. the access site was previously used by the impella cp twice, and this time, with the 5.0, the wound became dirty, pus came out, and there was a tendency for infection, so the doctor stated that they explanted the pump and switched to left ventricular assist device early. | [5 empty] | The investigation into the infection/sepsis issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information has not been provided, the root cause of the infection cannot be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿
30655 | 1220648-2024-07088 | 18791165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791165&pc=OZD | 01/29/2024 | 03/19/2024 |  | 02/27/2024 |  |  | 01/29/2024 | 02/27/2024 |  |  | 12/31/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318158 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient had to have a washout of the axillary due to increased jp drain output. the patient remained on support for two additional days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the product was not returned for investigation. the root cause of the access site bleeding was most likely due to patient condition since hematoma and bleeding was occurring at multiple sites.
30656 | 1220648-2024-07089 | 18791314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791314&pc=OZD | 01/30/2024 | 04/03/2024 |  | 02/27/2024 |  |  | 01/30/2024 | 02/27/2024 |  |  | 10/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | GM |  | 2024421036 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported an unknown gender and age patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during impella support for patient of unknown age and gender, there was a pump stop. the pump was noted to have been restarted several times without success. the console was exchanged but the pump stopped again. therefore, the pump was explanted. there was no reported patient harm. | [5 empty] | The investigation into the pump stop issue has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the pump stop issue was not determined since product and data logs were not returned for investigation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
30657 | 1220648-2024-07091 | 18791352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791352&pc=OZD | 02/09/2024 | 03/29/2024 |  | 02/27/2024 |  |  | 02/09/2024 | 02/27/2024 |  |  | 10/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025439355 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 26-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella support there was compromised blood flow to the left lower extremity, the pedal and posterior tibial (pt) pulses were absent. the decision was made to take the patient to the cardiac catheterization lab for external bypass. it was further reported that the patient had a steady ooze from the impella access site. bleeding was stabilized and heparin was started. the patient was reported as stable. | [5 empty] | The investigation into the limb ischemia ¿ reversible and the access site bleeding ¿ major has been completed since the initial report was submitted. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the access site bleeding issue was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30658 | 1220648-2024-07090 | 18791355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791355&pc=OZD | 01/30/2024 | 04/03/2024 |  | 02/27/2024 |  |  | 01/30/2024 | 02/27/2024 |  |  | 11/30/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | GM |  | 2024429785 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | No | No | Break | Defective Component | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 75 -year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during the implant the peel away sheath was removed from the vessel, but it would not peel open and one wing of the sheath broke off. the peel away sheath had to be cut open leading to severe bleeding. therefore, the impella cp was explanted. the patient was transferred to the intensive care unit with no further complications. | [5 empty] | The investigation into the introducer damage ¿ mechanical and the access site bleeding ¿ major issues has been completed since the initial report was submitted. the introducer was not returned for investigation. the root cause of the introducer damage is undetermined. the cause of the access site bleeding issue was most likely damaged valve of introducer. the failure type is unknown without return product. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: bleeding ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30659 | 1220648-2024-07092 | 18791455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791455&pc=PYX | 02/12/2024 | 03/29/2024 |  | 02/27/2024 |  |  | 02/12/2024 | 02/27/2024 | 02/27/2024 |  | 11/30/2025 | 11/30/2025 | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  | IMPELLA RP | IMPELLA RP |  | PYX | PYX |  | IMPELLA RP | IMPELLA RP |  | ABIOMED, INC. | ABIOMED, INC. |  | US | US | 2024434393 | 2024434393 | 1000323 | 1000323 | No | No |  | No | No | Yes | No | Health Professional | Health Professional |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported an 88-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the physician elected to use single access technique to place an 8f sheath next to the 15f repositioning sheath of the impella. despite multiple purse string sutures, impella access site continued to bleed. manual pressure was applied to the area and eventually hemostasis was achieved. blood products were required. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d1-updated product code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
30660 | 1220648-2024-07093 | 18791522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791522&pc=OZD | 10/04/2021 | 02/27/2024 |  | 02/27/2024 |  |  | 10/04/2021 | 02/27/2024 |  |  | 06/30/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022104948 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/04/2021 |  |  |  |  | The user facility reported a (b)(6)year-old female patient was attempting to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 pump was unable to advance beyond the innominate artery due to high calcification within the patient's vasculature. therefore, the impella 5.5 implant was aborted and an impella cp was subsequently placed for patient support. there was no patient harm reported. | [5 empty] | The investigation into the reported event has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30661 | 1220648-2024-07095 | 18791561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791561&pc=OZD | 09/13/2022 | 03/11/2025 |  | 02/27/2024 |  |  | 09/13/2022 | 02/27/2024 |  |  | 05/31/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228127 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 72-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced some mild oozing at the groin access site. the physician elected to explant the device. no further information has been provided. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30662 | 1220648-2024-07094 | 18791562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791562&pc=OZD | 11/16/2021 | 02/27/2024 |  | 02/27/2024 |  |  | 11/16/2021 | 02/27/2024 |  |  | 07/31/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022118267 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2021 |  |  |  |  | The user facility reported a 65-year-old female patient was attempting to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an impella 5.5 pump was unable to advance beyond the innominate artery due to high calcification within the patient's vasculature. the impella 5.5 implant was aborted and an impella cp was subsequently placed for patient support. there was no patient harm reported. | [5 empty] | The investigation into the reported event has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue: anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30663 | 1220648-2024-07096 | 18791584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791584&pc=OZD | 12/01/2021 | 02/27/2024 |  | 02/27/2024 |  |  | 12/01/2021 | 02/27/2024 |  |  | 06/30/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022109083 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2021 |  |  |  |  | The user facility reported a 67-year-old male patient was being implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 pump lost access into the ventricle during attempted delivery and was unable to re-cross the valve. the impella 5.5 was removed, and an impella cp pump was implanted without issue for support. there was no patient harm reported. | [5 empty] | The investigation into the reported event has been completed. no product was returned. the root cause of the delivery issue was use related per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30664 | 1220648-2024-07097 | 18791915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18791915&pc=PYX | 02/13/2024 | 03/29/2024 |  | 02/27/2024 |  |  | 02/13/2024 | 02/27/2024 |  |  | 11/30/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024286753 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during the sheath introduction from the left femoral vein (lfv). despite the noted damage, the rp pump was successfully placed for patient support. there was no patient harm. | [5 empty] | The investigation for the reported product damage has been completed. the impella was not returned for evaluation. clinical details note that patient did have peripheral artery disease/peripheral vascular disease (pad/pvd). the first 23fr introducer sheath split at the tip of the introducer. the second 23fr introducer used also split at the distal tip and kinked. a third introducer was able to successfully place the pump. the root cause of the introducer damage was most likely due to the introducer sheath tip being split on insertion as the patient had pad/pvd. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30665 | 1220648-2024-07098 | 18792085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18792085&pc=OZD | 02/14/2024 | 04/01/2024 |  | 02/27/2024 |  |  | 02/14/2024 | 02/27/2024 |  |  | 11/30/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024425255 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to systolic function to be severely reduced with a dilated heart. while on support, there was a large apical clot found. the plan was to switch to bival. | [5 empty] | The investigation into the thrombus issue has been completed since the initial report was submitted. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30666 | 1220648-2024-07099 | 18792328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18792328&pc=OZD | 02/10/2024 | 04/18/2024 |  | 02/27/2024 |  |  | 02/16/2024 | 02/27/2024 |  |  | 10/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024417433 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 57-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that the patient close to two months on impella support has a break mark on the shaft. the patient was moving and when lying down on the side, the shaft may have been crushed, resulting in the fracture. purge flow (pf) decreased slightly compared to me, but purge pressure (pp) remained unchanged. the pump is currently working without any problems. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the product damage was catheter kink due to patient movement.
30667 | 1220648-2024-07101 | 18792562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18792562&pc=OZD | 12/08/2021 | 02/27/2024 |  | 02/27/2024 |  |  | 12/08/2021 | 02/27/2024 |  |  | 09/30/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143110 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2021 |  |  |  |  | The user facility reported that the medical team experienced difficulty implanting an impella 5.5 in a 67-year-old male patient for mechanical circulatory support due to the patient's anatomy. after multiple failed attempts, the impella 5.5 implant was aborted and an impella 5.0 was subsequently placed for patient support. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. make sure that another report is not scheduled and that there is no next report due date as the investigation was completed and an smdr is not needed.
30668 | 1220648-2024-07100 | 18792577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18792577&pc=OZD | 02/16/2024 | 04/05/2024 |  | 02/27/2024 |  |  | 02/16/2024 | 02/27/2024 |  |  |  |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025438851 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 65-year-old female patient of unknown race and ethnicity was noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed an access site hematoma, manual pressure and one unit of red blood cells were required as treatment | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of access site bleeding was not determined because of insufficient clinical details and the product was not returned for investigation.
30669 | 1220648-2024-07102 | 18792961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18792961&pc=OZD | 02/17/2024 | 09/05/2024 |  | 02/27/2024 |  |  | 02/17/2024 | 02/27/2024 |  |  |  |  | 02/27/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1687316 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a 37-year-old female patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while the purge cassette was being changed the automated impella controller (aic) alarmed and failed. the aic and the purge cassette were replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Upon further review the returned product investigation concludes that the reported failure mode was most likely a defective purge cassette and not an issue with the automated impella controller (aic). a regulatory report is no longer necessary.
30670 | 1220648-2024-07103 | 18793020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793020&pc=OZD | 02/17/2024 | 04/05/2024 |  | 02/27/2024 |  |  | 02/17/2024 | 02/27/2024 |  |  | 06/30/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023237227 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a 69-year-old female patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that there was bleeding at the access site that required the patient to received blood products. it was reported the patient developed ischemia which prompted the pump to be explanted and had to undergo a penumbra thrombectomy, to restore flow to the limb. | [5 empty] | The investigation into the limb ischemia ¿ reversible and access site bleeding ¿ major issues has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30671 | 1220648-2024-07104 | 18793028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793028&pc=OZD | 08/04/2021 | 02/27/2024 |  | 02/27/2024 |  |  | 08/04/2021 | 02/27/2024 |  |  | 05/31/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022086838 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/04/2021 |  |  |  |  | The user facility reported an impella 5.5 was implanted in 57-year-old male patient for mechanical circulatory support. it was reported that the medical team experienced difficulty implanting impella due to patient's anatomy. ultimately the medical team aborted the procedure. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30672 | 1220648-2024-07105 | 18793082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793082&pc=OZD | 02/19/2024 | 04/08/2024 |  | 02/27/2024 |  |  | 02/19/2024 | 02/27/2024 |  |  | 06/30/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024427232 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to being biventricular non pulsatile systole. while on support, the nurse repositioned the patient and the impella cable became caught on the bed railing. the impella was dislodged and pulled out of the aortic valve. the impella was repositioned, however the pump was unable to flow. the physician decided to remove the impella. | [5 empty] | The investigation into the positioning issues and the pump stop issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the root cause of the pump stop was due to an electrical open of the motor cable lines as they were broken due to excessive force. the root cause of the positioning issue was due to patient movement as the impella cable got caught on the bed railing when the nurse was rolling the patient. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
30673 | 1220648-2024-07106 | 18793098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793098&pc=OZD | 06/28/2022 | 02/27/2024 |  | 02/27/2024 |  |  | 06/28/2022 | 02/27/2024 |  |  | 03/31/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023203041 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 67-year-old male was being placed on an impella 5.5 for mechanical circulatory support. the complainant reported that an impella 5.5 pump implant attempt had to be aborted due to the patient's anatomy and an impella cp was successfully placed. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella 5.5 device has not been returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
30674 | 1220648-2024-07107 | 18793103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793103&pc=OZD | 02/19/2024 | 04/10/2024 |  | 02/27/2024 |  |  | 02/19/2024 | 02/27/2024 |  |  | 07/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025437383 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 62-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support the patinet had a slow flow down the right leg, requiring multiple pressors, and inotropes. the patient was taken to the operation room for emergent below knee amputation (bka) of ischemic right lower extremity (rle) due to progressive acidosis and increasing pressor requirements. the impella cp was exchanged for a 5.5. | [5 empty] | The investigation into the limb ischemia ¿ irreversible issue has been completed since the original report was submitted. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
30675 | 1220648-2024-07108 | 18793144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793144&pc=OZD | 02/18/2024 | 03/08/2024 |  | 02/27/2024 |  |  | 02/20/2024 | 02/27/2024 |  |  |  |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1100033 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Component Incompatible | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a ¿purge disc not detected¿ alarm occurred. the aic (automated impella controller) was exchanged. there was no harm to the patient. | [5 empty] | The impella device was received from the customer on 26-feb-2024 and therefore, an evaluation of the device was is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. d4-updated model number
30676 | 1220648-2024-07109 | 18793539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793539&pc=OZD | 09/14/2022 | 04/16/2025 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 04/30/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2023226532 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 66-year-old female requiring mechanical circulatory support. during the implantation of an impella cp, the physician had difficulty cannulating the right coronary artery with the impella cp due to chronic total occlusion (cto). the physician was able to cannulate it with the fr4, but due to lack of support, the physician was unable to advance the coronary wire. after multiple attempts to seat the catheter in the artery and multiple attempts to reposition, the physician elected to remove the impella cp and perform the percutaneous coronary intervention with vasopressor support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30677 | 1220648-2024-07110 | 18793548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793548&pc=OZD | 08/18/2021 | 02/27/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 02/28/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022046272 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/18/2021 |  |  |  |  | The user facility reported an impella 5.5 was attempted to be implanted in a 58-year-old female patient for mechanical circulatory support. it was reported that the patient had a short ascending aorta and the graft was just short of 7. the medical doctor decided to place the pump. the pump was placed and started at p2 up to p5. there was positioning issues due to patient anatomy and placement of graft. the decision was made to remove the pump. | [5 empty] | The investigation into the reported position issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the delivery issue was patient condition related because the impella was unable to pass through vasculature. this report is being filed as part of a retrospective review of historical records.
30678 | 1220648-2024-07111 | 18793558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793558&pc=OZD | 09/09/2021 | 02/27/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 05/31/2023 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022085242 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/09/2021 |  |  |  |  | The user facility reported an impella 5.5 was implanted in 70-year-old male patient for mechanical circulatory support. it was reported that the medical team was having difficulty delivering the pump due to the patient's anatomy and the procedure was ultimately aborted. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30679 | 1220648-2024-07112 | 18793574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793574&pc=OZD | 02/05/2024 | 03/25/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 11/30/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024431406 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella 5.5 implant however the staff was unable to implant the pump due to anatomy. the impella 5.5 was aborted and an impella cp was implanted instead. | [5 empty] | The investigation into the delivery issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of delivery issue is determined to be patient condition since they were unbale to pass the pump due to patient anatomy. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30680 | 1220648-2024-07113 | 18793598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793598&pc=OZD | 02/12/2024 | 03/29/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 07/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442212 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, echo was completed at bedside and the catheter was out of position. the pigtail was likely trapped under the chordae. the patient's urine was red due to hemolysis. the doctor declined repositioning due to possibly dislodging the impella. suction alarms were also noted. additionally, the patient's right lower extremity was ischemic. the levo dose was increased and the foot went cold roughly an hour after. the impella was removed and the impella 5.5 was placed. | [5 empty] | The investigation for the reported low pump flow, limb ischemia, positioning issues, and hemolysis has been completed. the impella device was not returned for evaluation. data logs confirmed the low flow and the pump was not in proper position. logs also showed 7 hours after low flow occurred, the pump was not in proper position corresponding with clinical time hemolysis occurred. the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of positioning issue was most likely use related. low flow was confirmed on data logs but without product return, the root cause of the low flow was unable to be determined. without additional clinical information, the true root cause of limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30681 | 1220648-2024-07114 | 18793642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793642&pc=OZD | 02/13/2024 | 03/21/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 08/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382386 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had bleeding from anastomosis of graft on aorta. sutures were placed and the bleeding resolved. the patient remained on support for three days. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleed was most likely use-related since adding additional sutured resolved the bleeding issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30682 | 1220648-2024-07115 | 18793647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793647&pc=OZD | 02/13/2024 | 04/17/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  |  |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435417 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported impella 5.5 implant procedure had to be aborted as the device was unable to be inserted due to patient anatomy. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. added- h6 impact code 4627 d4-updated model number | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30683 | 1220648-2024-07116 | 18793681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793681&pc=OZD | 02/15/2024 | 04/16/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 11/30/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024426791 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support for 67-year-old male patient, the pump stopped. act (activated clotting time) was noted to be in range. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. according to the clinical, on the 17th day of impella cp support a pump stop occurred. product evaluation confirmed the presence of an abnormally large purge gap. the cause of the pump stop was bearing wear after 17 days of use, which is beyond 8-day design life per the device¿s design trace matrix.
30684 | 1220648-2024-07117 | 18793688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793688&pc=OZD | 02/16/2024 | 04/03/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 06/30/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024431076 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there was an impella in aorta alarm. the catheter was pulled into the descending aorta. the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical data provided, the cause of the positioning issue was use related due to improper technique.
30685 | 1220648-2024-07118 | 18793710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793710&pc=OZD | 09/14/2023 | 03/28/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 11/30/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024305207 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported an 81-year-old female patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after impella removal the patient became anemic, and blood products were given. the impella was a probable reason for the blood loss. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that prior to removal, the patient's activated coagulation time (act) was prolonged at 279 . therefore, the root cause of the access site bleed was most likely due to improper anticoagulation management. the instructions for use states to ¿assess access site for bleeding and hematoma.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30686 | 1220648-2024-07119 | 18793727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793727&pc=OZD | 02/16/2024 | 03/28/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 10/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410779 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella placement after removing peel away and advancing repo sheath, bleeding was observed around the repo sheath. the physician could not get the bleeding to stop. decision was made to remove the impella and close the groin site. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30687 | 1220648-2024-07120 | 18793799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18793799&pc=OZD | 02/17/2024 | 04/03/2024 |  | 02/27/2024 |  |  | 02/17/2024 | 02/27/2024 |  |  | 10/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024411510 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 stopped and was not able to restart. there was a tear in the catheter shaft with wires exposed on the proximal end of the catheter just before the red impella plug. the impella was exchanged for another 5.5. there was no patient harm. | [5 empty] | The investigation into the pump stop issue has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical data provided, the most likely cause of the pump stop was an open electrical line. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
30688 | 1220648-2024-07122 | 18794101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18794101&pc=OZD | 02/19/2024 | 03/25/2024 |  | 02/27/2024 |  |  | 02/19/2024 | 02/27/2024 |  |  | 05/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024356212 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 62-year-old female patient was being escalated from impella cp to the impella 5.5 for circulatory support. it was reported that during placement the 5.5 was pulled back into the aorta after peel away sheath was removed. the impella was removed and a new peel away sheath used to re-insert the impella 5.5 into the left ventricle. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. data logs showed about 45 minutes into the case, pump was in the aortic area that triggered impella position in aorta. pump then was disconnected and restarted 27 minutes later. the root cause of the positioning issue was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30689 | 1220648-2024-07123 | 18794130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18794130&pc=OZD | 02/20/2024 | 03/08/2024 |  | 02/27/2024 |  |  | 02/20/2024 | 02/27/2024 |  |  |  |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1606415 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that while the staff was removing the purge disk from the transmitter, the purge transmitter broke off from the aic (automated impella controller). there was no patient harm. | [5 empty] | The impella device was received from the customer on 26-feb-2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge retainer was broken. the cause of the issue was a broken purge disc retainer.
30690 | 1220648-2024-07124 | 18794143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18794143&pc=OZD | 02/20/2024 | 04/09/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 02/29/2024 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197650 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 78-year-old female patient noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp placement was aborted as it was unable to past the aortic arch due to tortuous and calcified aortic vasculature. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. per clinical information provided, the root cause of the delivery issue is determined to be patient condition because it was reported that the patient vessels are calcified and were unable to pass beyond aortic arch and the case was aborted due to patient anatomy.
30691 | 1220648-2024-07125 | 18794164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18794164&pc=OZD | 02/20/2024 | 04/08/2024 |  | 02/27/2024 |  |  |  | 02/27/2024 |  |  | 07/31/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025443413 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was pushed through the infarcted left ventricle wall while positioning. perforation requiring covered stent and/or balloon tamponade. the impella was explanted. the patient is stable. | [5 empty] | The investigation into the ventricle perforation issue has been completed since the initial report was submitted. the device was not returned for investigation. data log analysis confirmed pump position in the ventricle alarms with a ventricular placement signal waveform. the cause of the ventricle perforation was improper positioning. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows:
30692 | 1220648-2024-07126 | 18794171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18794171&pc=OZD | 02/22/2024 | 04/23/2024 |  | 02/27/2024 |  |  | 02/22/2024 | 02/27/2024 |  |  | 11/30/2025 |  | 02/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427062 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/22/2024 |  |  |  |  | In addition to the noted positioning issue, it was reported that the patient was experiencing hemolysis at support level p-4. the support level was lowered to p-1 but the condition persisted. following pump explant, the hemolysis was documented to have resolved. | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the positioning of the implanted impella cp pump was unstable during patient support. it was noted via fluoroscopy that the catheter was moving violently within the left ventricle. it was deemed unreliable to adjust positioning to a shallower depth and the decision was made to explant the device. there was no patient harm. |  |  |  |  | The impella device has not been returned by the customer and therefore, evaluation of the device has not been possible. should any new information be acquired, a supplemental mdr will be filed. | The investigation for the reported hemolysis and positioning issues has been completed. the cp pump was returned and visually inspected. no biomaterial was observed. no broken blades or sharp edges were found. no sharp edge were observed. no abnormalities were observed. hemolysis testing was conducted. no hemolysis was reproduced and mih=4.51. the cause of the hemolysis issue was improper positioning related due to improper technique because the pump was unstable during patient support and was moving violently within the left ventricle. the cause of the positioning issue most likely was use related due to improper technique. added-d9 device return date. d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
30693 | 1220648-2024-07048 | 18777615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18777615&pc=OZD | 10/09/2022 | 03/07/2025 |  | 02/26/2024 |  |  | 10/09/2022 | 02/26/2024 |  |  | 06/30/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023248049 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2022 |  |  |  |  | The user facility reported a 79-year-old female patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced mild bleeding at the insertion site. the patient was given one unit of packed red blood cells (prbcs). the bleeding was resolved and the site was redressed. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
30694 | 1220648-2024-07049 | 18777741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18777741&pc=OZD | 04/18/2022 | 02/26/2024 |  | 02/26/2024 |  |  | 04/18/2022 | 02/26/2024 |  |  | 01/31/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188353 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/18/2022 |  |  |  |  | The user facility reported a 48-year-old male patient needed mechanical circulatory support. it was reported that the while performing an echocardiogram for placement of the impella cp, a large thrombus was noted. patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. as the necessary clinical information was not provided, the root cause of the thrombus was not determined. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use provides details on how prevent thrombus formation in the pump. it states the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
30695 | 1220648-2024-07050 | 18777748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18777748&pc=OZD | 01/31/2024 | 09/21/2024 |  | 02/26/2024 |  |  | 01/31/2024 | 02/26/2024 |  |  | 11/30/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/31/2024 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the team attempted to deliver but were unable to get beyond the iliac. the patient had repeated episodes of ventricular fibrillation and expired in the ccl. no allegation made that impella use caused/contributed to the death. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the device was not returned. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to pass through the vasculature due to stenosis as per the clinical description provided.
30696 | 1220648-2024-07051 | 18777787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18777787&pc=OZD | 02/09/2024 | 04/03/2024 |  | 02/26/2024 |  |  | 02/09/2024 | 02/26/2024 |  |  | 11/30/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429781 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 77-year-old male patient, there was oozing from the repo sheath (from the gap between white and blue part). the patient's hemoglobin dropped from 11.1 to 6.6 however, there was also significant bleeding also from skin transplant in face. the peel away sheath was exchanged and the issue resolved. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30697 | 1220648-2024-07052 | 18777970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18777970&pc=OZD | 02/14/2024 | 04/01/2024 |  | 02/26/2024 |  |  | 02/14/2024 | 02/26/2024 |  |  | 12/31/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024315802 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump began to experience complications during patient support. it was originally noted that the lv (left ventricle) signal was reading higher than the placement signal and that the flows in systole and diastole were reading identical. however, upon placing an echo probe on the patient, it was noted that there was no unloading in the ventricle. the staff was advised that the pump was experiencing saturated flows and the decision was made to explant the pump. there was no patient harm. | [5 empty] | The investigation into the saturated flow issue has been completed since the original report was filed. product and data logs were investigated. data logs show the there was a motor current spike followed by saturated flows. also, heavy biomaterial was found wrapped around the impeller blades during product investigation. therefore, the root cause of the saturated flow issue was biomaterial. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
30698 | 1220648-2024-07053 | 18778008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778008&pc=PYX | 02/14/2024 | 03/29/2024 |  | 02/26/2024 |  |  | 02/14/2024 | 02/26/2024 |  |  | 07/31/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023229876 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 74-year-old male patient in acute myocardial infarction cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that during insertion the 23f sheath was placed and sheath kinked upon removal of the dilator. the dilator reinserted and 0.018 v18 wire advanced through the dilator. v18 removed and replaced with impella 0.027 wire. impella rp loaded onto 0.027 wire. there was difficulty with catheter kinking in right atrial (ra) by the outflow. buddy wire advanced for support. back tension was placed on 0.027 wire, while clockwise torque placed on rp catheter. the impella was successfully placed in the left pulmonary artery (lpa). no patient harm was reported. | [5 empty] | The investigation into the introducer damage ¿ mechanical and the delivery issues has been completed since the initial report was submitted. the product was not returned for investigation. per the clinical information, the root cause of the introducer mechanical damage issue was most likely an introducer sheath kink. the root cause of the delivery issue was use related to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30699 | 1220648-2024-07054 | 18778090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778090&pc=OZD | 02/15/2024 | 03/28/2024 |  | 02/26/2024 |  |  | 02/15/2024 | 02/26/2024 |  |  | 11/30/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427044 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 50-year-old male patient of unknown race and ethnicity in st elevated myocardial infarction in need of coronary artery bypass graft surgery (cabg) was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on day 3 of support the patient developed atrial fibrillation and was cardioverted and was escalated to amiodarone. the pump remained on for a total of 7 days before explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported atrial arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the atrial arrhythmia could not be determined. the instructions for use states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
30700 | 1220648-2024-07056 | 18778122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778122&pc=OZD | 02/14/2024 | 04/04/2024 |  | 02/26/2024 |  |  | 02/16/2024 | 02/26/2024 |  |  | 06/30/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024430952 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella cp insertion was aborted due to the patient¿s anatomy. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because of the patient company. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30701 | 1220648-2024-07057 | 18778126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778126&pc=OZD | 09/28/2022 | 03/07/2025 |  | 02/26/2024 |  |  | 09/28/2022 | 02/26/2024 |  |  | 04/30/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023213500 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | Increase in Pressure | Pumping Stopped |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a male patient for acute myocardial infarction and cardiogenic shock. it was reported that the purge sidearm was cracked at the yellow luer. a purge sidearm bypass was completed but the pump subsequently continued to have high purge pressure alarms. patient was not a candidate for tissue plasminogen activator (tpa) and as a result, a pump stop occurred. the pump was exchanged to a different impella cp. there was no patient harm reported. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30702 | 1220648-2024-07055 | 18778155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778155&pc=OZD | 10/11/2022 | 03/04/2024 |  | 02/26/2024 |  |  | 10/11/2022 | 02/26/2024 |  |  | 04/30/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224442 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 61-year-old male patient was implanted with an impella cp device for mechanical circulatory support. it was reported that after being placed on impella cp support, the patient developed right lower extremity (rle) ischemia. the patient was treated using distal perfusion catheters, 5 fr antegrade for the right femoral artery (rfa) and 6 fr retrograde for the left femoral artery (lfa). | [5 empty] | The impella device has not been returned for evaluation. when the investigation has been completed, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30703 | 1220648-2024-07058 | 18778198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778198&pc=OZD | 02/16/2024 | 03/27/2024 |  | 02/26/2024 |  |  | 02/16/2024 | 02/26/2024 |  |  |  |  | 02/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1504376 |  | 0042-000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The us complainant had an aic (automated impella controller) failure noted by the hospital biomed department. the aic battery failed to charge. the aic was removed from service before any patient use. no patient involvement. | [5 empty] | The investigation into the aic - battery charging/other issues has been completed since the initial report was submitted. the console was returned and tested, and the issue was reproduced. the cause of the aic issue was a defective power supply. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30704 | 1220648-2024-07059 | 18778210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778210&pc=OZD | 09/30/2022 | 02/26/2024 |  | 02/26/2024 |  |  | 09/30/2022 | 02/26/2024 |  |  | 03/31/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Sepsis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/30/2022 |  |  |  |  | The user facility reported that a 65-year-old male patient developed septic shock after being placed on impella cp support. additionally, the patient went into runs of ventricular tachycardia (vt) and atrial fibrillation (afib). the patient was shocked multiple times. the impella cp was explanted due to the patient improving and the need to determine the source of sepsis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
30705 | 1220648-2024-07060 | 18778261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778261&pc=OZD | 02/19/2024 | 03/27/2024 |  | 02/26/2024 |  |  | 02/20/2024 | 02/26/2024 |  |  | 11/30/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024431468 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced bleeding from their axillary access site following impella 5.5 implant. heparin was subsequently held in response to the bleed and one unit of prbc (packed red blood cells) was transfused. support was maintained following the intervention. | [5 empty] | The investigation into the access site bleeding major issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30706 | 1220648-2024-07061 | 18778372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778372&pc=OZD | 10/16/2022 | 09/20/2024 |  | 02/26/2024 |  |  | 10/16/2022 | 02/26/2024 |  |  | 06/30/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248343 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2022 |  |  |  |  | Additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported that an 80-year-old male, while being placed on impella cp support the pump was unable to purge. the impella was discontinued and a new impella was prepped, which was successfully purged and inserted. there was no patient harm reported. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07061 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30707 | 1220648-2024-07062 | 18778379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18778379&pc=OZD | 10/13/2022 | 02/27/2024 |  | 02/26/2024 |  |  | 10/13/2022 | 02/26/2024 |  |  | 09/28/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253547 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported a 78-year-old female patient was being placed on impella cp support for mechanical circulatory support. multiple attempts were made to implant the impella cp into the patient. abdominal aorto-angioplasty and lithotripsy were both performed without success. the impella implantation was aborted due to patient anatomy. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition since calcification of vessels was observed along with tortuous abdominal aorta. this report is being filed as part of a retrospective review of historical records. | D4 revised unique identifier (udi) # as previously entered incorrectly.
30708 | 1220648-2024-07067 | 18779461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18779461&pc=OZD | 02/01/2024 | 03/15/2024 |  | 02/26/2024 |  |  | 02/01/2024 | 02/26/2024 |  |  | 08/31/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385046 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a 68-year-old female (race unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the axillary insertion failed repeatedly and the procedure was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition, as the impella was unable to pass through the vasculature via the axillary artery.
30709 | 1220648-2024-07068 | 18779495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18779495&pc=OZD | 01/20/2024 | 04/11/2024 |  | 02/26/2024 |  |  | 02/09/2024 | 02/26/2024 |  |  | 11/30/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278444 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 67-year-old male patient, there were continuous suction alarms. neither pump repositioning nor fluids helped, therefore, pump was explanted. following explant, there was tissue found in the cannula, which was determined to be "fibrin" thrombus. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of low flow was most likely due to biomaterial ingestion. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
30710 | 1220648-2024-07069 | 18779545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18779545&pc=OZD | 02/13/2024 | 09/03/2024 |  | 02/26/2024 |  |  | 02/13/2024 | 02/26/2024 |  |  |  |  | 02/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 74-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) had a purge cassette failure. the team swapped out the aic with another with no further issues. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Upon further review, the likely root cause is associated with the purge cassette and not an issue with the automated impella controller. a regulatory report is no longer necessary.
30711 | 1220648-2024-07070 | 18779629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18779629&pc=OZD | 02/14/2024 | 03/28/2024 |  | 02/26/2024 |  |  | 02/14/2024 | 02/26/2024 |  |  | 07/31/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431931 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 76-year-old male patient of unknown ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during patient support an echocardiogram revealed a large thrombus in the left ventricle. the decision was made to remove the pump and replace with an intra-aortic balloon pump for continued support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
30712 | 1220648-2024-07071 | 18779951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18779951&pc=OZD | 02/14/2024 | 09/24/2025 |  | 02/26/2024 |  |  | 02/14/2024 | 02/26/2024 |  |  |  |  | 02/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 02/14/2024 |  |  |  |  | Additional information: it was further reported the patient expired. per the abiomed medical safety office, abiomed does not have enough information to associate use of device with outcome. | The user facility reported that a 78-year-old female patient post cardiotomy cardiogenic shock, low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during the preparation for patient support the automated impella controller (aic) failed to recognize a connected impella 5.5 pump. the aic was swapped to resolve the noted issue. there was no patient harm reported. |  |  |  |  | The investigation into the pump not detected issue has been completed since the original report was filed. the pump was returned for investigation and data logs were made available. data analysis: pump 463338 was plugged into ic7567 on 2/14. immediately, calculator placement signal 1036 errors are immediately seen in the logs followed a minute later with the user receiving "case start: optical sensor system error: disconnect cable". the rt logs for the first pump insertion show that raw optical sensor was steady at -3276.8 mmhg which is the value when a pump is not completely connected and there is an air gap. pump 463338 was moved over to ic1229 and the error did not reappear. device analysis: the failure mode was able to be reproduced by not completely plugging in the pump. fully plugging in the pump resolved the controller error. root cause: the cause of the optical signal failure was an air gap from between the optical cable and the pump. additional information b5 added d9 f6 and f8 should have been blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 adverse event updated. b2 updated to death. h1 updated to death. h6 updated to include optical signal coding. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and reporting contact fax number updated.
30713 | 1220648-2024-07072 | 18780028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18780028&pc=PYX | 02/15/2024 | 03/22/2024 |  | 02/26/2024 |  |  | 02/15/2024 | 02/26/2024 |  |  | 02/28/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023199604 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported that a 59-year-old male patient post cardio-thoracic surgery was implanted with an impella rp for mechanical circulatory support. it was reported that there was bleeding noted at the access site, requiring 2 units of packed red blood cells for treatment. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
30714 | 1220648-2024-07073 | 18780118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18780118&pc=OZD | 02/15/2024 | 04/04/2024 |  | 02/26/2024 |  |  | 02/16/2024 | 02/26/2024 |  |  | 05/31/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345979 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported that a 66-year-old female was implanted with an impella 5.5 device for mechanical circulatory support after being escalated from impella cp. it was reported that while on impella 5.5 support, the patient experienced a non-dominate vertebral stoke. the doctor believes there was an embolism at the impella access site, and it migrated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the stroke/cva has been completed since the initial report was submitted. the device was not returned for investigation. as the clinical information was not provided, the root cause of the stroke/cva could not be determined.
30715 | 1220648-2024-07074 | 18780189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18780189&pc=OZD | 02/08/2024 | 04/04/2024 |  | 02/26/2024 |  |  | 02/17/2024 | 02/26/2024 |  |  | 11/30/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024428145 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Fluid/Blood Leak | Material Perforation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that there was bleeding at the impella insertion site that has since decreased in volume and now stabilized. rn will continue to monitor. it was further reported that blood was noticed in the sterile sleeve, so heparin was held. the physician loosened the back tuohy and drained the blood from the sterile sleeve and retightened the tuohy. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ minor and the product damage (sterile sleeve damage) issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding and the product damage - sterile sleeve issues was not determined since product was not returned and limited clinical information was provided.
30716 | 1220648-2024-07075 | 18783957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18783957&pc=OZD | 04/18/2022 | 05/23/2025 |  | 02/26/2024 |  |  |  | 02/26/2024 |  |  | 01/31/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188288 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2022 |  |  |  |  | The user facility reported a 70-year-old male patient needed mechanical circulatory support. it was reported that the patient was on the impella cp and there was bleeding at the impella cp insertion site. additionally, there were two low purge flow alarms which resolved once the patient was switched to heparin. no kinks were noted in the line. the patient continued on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use lists several potential events:¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ it also states to "¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30717 | 1220648-2024-07076 | 18784839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18784839&pc=OZD | 04/19/2022 | 05/23/2025 |  | 02/26/2024 |  |  |  | 02/26/2024 |  |  | 01/31/2024 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023182811 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/19/2022 |  |  |  |  | The user facility reported a 38- year-old male in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was bleeding from everywhere including the pump site. disseminated intravascular coagulation (dic) had become a concern amongst the medical team. heparin was removed from the pump purge and the patient was transfused with 3 units of packed red blood cells (prbcs), 10 units of cryoprecipitated antihemophilic factor (cryo), and 2 units platelets following a cardiac code. the impella cp was re-sutured and manual pressure was held by the staff. a few hours later the device displayed an aorta alarm and the echocardiogram (echo) revealed that the inflow cannula was noted to be in the left ventricular outflow tract (lvot). the device was advanced using ultrasound guidance to 3.1cm and locked at 89cm on the impella cp catheter. the impella cp was eventually explanted. unfortunately, the patient expired while being on support. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07076 was provided. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30718 | 1220648-2024-07077 | 18785386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18785386&pc=OZD | 01/29/2024 | 04/01/2024 |  | 02/26/2024 |  |  | 01/29/2024 | 02/26/2024 |  |  | 10/31/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024417297 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Delivery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implant procedure for a 78-year-old female patient had to be aborted as the impella cp was unable to be delivered due to the vascular anatomy's massive kinking. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the initial report was submitted. the product was not returned for investigation. based on the clinical information provided, the root cause of delivery issue was determined to be patient condition because of the patient¿s vascular anatomy.
30719 | 1220648-2024-07078 | 18785405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18785405&pc=OZD | 02/09/2024 | 04/09/2024 |  | 02/26/2024 |  |  | 02/09/2024 | 02/26/2024 |  |  | 11/30/2025 |  | 02/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442473 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after switching the impella cp in 'auto mode" during support for a 68-year-old male patient, 'suction alarms' and 'impella flow reduced' alarms appeared. the physician noted that the papillary muscles may have been drawn into the pump causing the suction alarms. the pump slipped into the left ventricle leading to the pump having to be repositioned. however, the pump ultimately stopped 15 minutes later. the pump was explanted and another device was used. after the explant, the pump was noted to be kinked. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. according to the clinical, immediately after beginning impella cp support constant suction alarms were recorded. product evaluation confirmed no abnormalities with the pump, but pump stop was able to be reproduced. data log analysis confirmed positioning issues but did not include impella defective alarm or pump stop. the most likely cause of the positioning issues were use related.
30720 | 1220648-2024-07079 | 18785441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18785441&pc=OZD | 02/06/2024 | 03/18/2024 |  | 02/26/2024 |  |  | 02/09/2024 | 02/26/2024 |  |  |  |  | 02/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1586834 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Charge | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella support, the automated impella controller (aic) displayed a battery failure alarm. the battery was not charging. there was no reported patient harm. | [5 empty] | The investigation into the aic - battery failure (red alarm) issue has been completed since the original report was filed. the console was returned and tested after arriving in fs. battery failure alarm issue was able to be reproduced. the issue was determined to be caused by internal battery cell imbalance. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
30721 | 1220648-2024-07042 | 18774809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18774809&pc=OZD | 02/12/2024 | 05/16/2024 |  | 02/25/2024 |  |  | 02/12/2024 | 02/25/2024 |  |  |  |  | 02/25/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1171604 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility's biomedical team reported the automated impella controller (aic) displayed a yellow alarm on the console. there was no patient involvement. | [5 empty] | The investigation into the aic - controller error (yellow alarm) has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the controller error was not reproduced while running an optical pump simulator. there was no issue with the console hardware related to the reported failure mode. however, per data analysis, due to the force shutdown, when the console was booted up on 13 mar 2024, logs recorded ¿unexpected controller shutdown ¿ alarm. the root cause of the controller error was due to an aic software issue. in accordance with updated procedures, section d2 has been revised from the initial submission. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
30722 | 1220648-2024-07043 | 18774810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18774810&pc=OZD | 02/12/2024 | 04/01/2024 |  | 02/25/2024 |  |  | 02/12/2024 | 02/25/2024 |  |  | 09/30/2025 |  | 02/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403009 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported an (b)(6) male patient in st-elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient had continuous oozing around the repositioning sheath. the patient was given one unit of packed red blood cells and femostop was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleed was unable to be determined due to insufficient clinical information that resulted in the cause of oozing and product was not returned for investigation.
30723 | 1220648-2024-07044 | 18774814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18774814&pc=OZD | 02/13/2024 | 04/01/2024 |  | 02/25/2024 |  |  | 02/13/2024 | 02/25/2024 |  |  | 11/30/2025 |  | 02/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426558 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support due to needing a bridge to left ventricle assist device. while on support, the patient developed a hematoma. heparin was paused and protamine was given. the hematoma was under control. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that the hematoma was resolved after they have stopped the heparin administration. hence, the root cause is determined to be anticoagulation management.
30724 | 1220648-2024-07045 | 18774815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18774815&pc=OZD | 02/14/2024 | 04/09/2024 |  | 02/25/2024 |  |  | 02/14/2024 | 02/25/2024 |  |  |  |  | 02/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1699285 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the automated impella controller had a purge issue. when attempting to decrease the p level, it defaults to p-0 and is unable to increase the p level without it dropping to p-0. | [5 empty] | The investigation into the aic, purge issue, other has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure mode was reproduced with the original parts and configuration. the root cause of the issue was an impellatronic malfunction. | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
30725 | 1220648-2024-07046 | 18774816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18774816&pc=OZD | 02/14/2024 | 04/01/2024 |  | 02/25/2024 |  |  | 02/14/2024 | 02/25/2024 |  |  | 11/30/2025 |  | 02/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438239 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 69-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that upon implanting the impella cp, excessive force was used to get through the calcium on the femoral artery. the impella cp was implanted, and a kink was noticed in the cannula. flows were low and there were suction alarms. the impella cp was explanted and reinserted and the same kink in the cannula was noted. the impella cp had low flows again and the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) and the low pump flow issues has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the low pump flow issue was cannula kink due to excessive force.
30726 | 1220648-2024-07047 | 18774819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18774819&pc=OZD | 02/16/2024 | 04/03/2024 |  | 02/25/2024 |  |  | 02/16/2024 | 02/25/2024 |  |  | 10/31/2025 |  | 02/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422320 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on support due to needing additional support. while on support, there were placement signal alarms. an echocardiogram (echo) was used to confirm the position. the imaging showed the inlet measured 3.8 centimeters from the aortic valve. however, the pump continued to alarm ¿placement signal low¿. the pump appeared to have been caught in the papillary muscle. the pump position was pulled back until resolution of the alarm. in attempt to pull back another half centimeter, the pump was unable to be seen on echo. the pump appeared to be in the aorta yet alarm ¿impella position in ventricle¿ was observed. the ultimate decision was made to explant the device as crossing the valve bedside was unsuccessful. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. no kink, biomaterial or other damages were observed. the cause of the optical signal issue was most likely patient condition related per field investigation and log review.
30727 | 1220648-2024-07003 | 18766330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18766330&pc=OZD | 02/09/2024 | 03/29/2024 |  | 02/23/2024 |  |  | 02/09/2024 | 02/23/2024 |  |  | 02/28/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334414 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 29-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed renal failure one day after implant was inserted. there was no definitive diagnosis of hemolysis. the team alleged there was a thrombus probably present that could have caused/contributed to the hemolysis and continuous veno-venous hemofiltration. the patient remained on support for four more days before the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. | The investigation into the hemolysis and the thrombus issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the hemolysis was unable to be determined as no product or logs were returned for analysis. as the necessary information was not provided, the root cause of the thrombus was not determined.
30728 | 1220648-2024-07004 | 18766402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18766402&pc=OZD | 02/11/2024 | 03/29/2024 |  | 02/23/2024 |  |  | 02/11/2024 | 02/23/2024 |  |  | 09/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410549 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2024 |  |  |  |  | The user facility reported a patient of unknown age, gender, and race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion the staff was unable to pass through the tortuous iliac artery. after troubleshooting, the decision was made to remove the impella cp device and place an intra-aortic balloon pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to pass through tortuous iliac artery as per the clinical description provided.
30729 | 1220648-2024-07005 | 18766467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18766467&pc=OZD | 02/12/2024 | 04/01/2024 |  | 02/23/2024 |  |  | 02/12/2024 | 02/23/2024 |  |  | 11/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435580 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported an 84-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp had numerous purge issues upon set up. purge pressure low and purge system open alarms persisted. troubleshooting was unsuccessful. there was no purge fluid exiting the catheter. this impella cp device was removed and a new impella cp was placed successfully. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned for evaluation with red lumen still inserted. upon visual inspection, no defects or damage were noted. the root cause of the priming issue was most likely due to use issue as there was not issues found with the pump and purge cassette was changed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30730 | 1220648-2024-07007 | 18766787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18766787&pc=OZD | 02/12/2024 | 03/29/2024 |  | 02/23/2024 |  |  | 02/12/2024 | 02/23/2024 |  |  | 11/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024436229 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the us complainant had placed an impella cp with single access to support a patient in need of a high risk percutaneous coronary intervention (hrpci). the 14 french x 25 centimeter sheath was observed to malfunction, as the product was leaking at the seam and so the introducer was replaced. the impella cp supported the patient successfully for over 67 hours until weaned and explanted. | [5 empty] | The investigation for the reported introducer damage has been completed. the introducer was returned for evaluation. upon functional testing, a leak was able to be reproduced on only one of the sides of the introducer. the root cause of the introducer damage was determined to be a damaged hub causing the introducer to leak. added-d9 device return date | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30731 | 1220648-2024-07008 | 18767009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18767009&pc=OZD | 09/25/2022 | 03/11/2025 |  | 02/23/2024 |  |  | 09/25/2022 | 02/23/2024 |  |  | 04/30/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219975 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported a 57-year-old male patient was implanted with an impella cp for mechanical circulatory support. it was reported while being on impella cp support, the patient developed hemolysis. the urine output was red and the labs were hemolyzed. pump placement was checked via echocardiogram and impella flow was decreased to p6. intravenous antihypertensives and inotropics were weaned off. labs were no longer being hemolyzed. the patient continued support. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30732 | 1220648-2024-07009 | 18767024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18767024&pc=OZD | 09/29/2022 | 02/23/2024 |  | 02/23/2024 |  |  | 09/29/2022 | 02/23/2024 |  |  | 04/30/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216732 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported an 86-year-old male patient was implanted with an impella cp for mechanical circulatory support. it was reported that the impella access site suddenly developed a hematoma. the hematoma was expressed and manual pressure was held. sutures were removed for replacement to change angle matching to help control bleeding. however, bleeding would not stop. as a result, the device was weaned and explanted. manual pressure was then held and then femstop was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ for access site bleeding, it ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30733 | 1220648-2024-07010 | 18767848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18767848&pc=OZD | 02/05/2024 | 11/18/2025 |  | 02/23/2024 |  |  | 02/05/2024 | 02/23/2024 |  |  | 08/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394636 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/05/2024 |  |  |  |  | The user facility reported a 50-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after 1 hour of support the pump had low flows. the team reviewed the flows and observed either from a clot or mitral valve leaflet issue. the pump was removed and replaced. the 2nd pump was placed and supported until the patient had care withdrawn by the family. | [5 empty] | The pump was returned for investigation. there were no cannula kinks. no impeller damage was seen but it was observed that the blades were interfering with the inside wall of the outlet cage, which prevented the impeller from rotating freely. this resulted in a motor current high alarm when the pump was run in the lab, due to which no low pump flow could be reproduced. no biomaterial or foreign material was seen inside the cannula, purge gap, or around the impeller blades. data logs were returned for analysis. data logs confirmed the failure mode. suction and "impella flow reduced" alarms were seen about 40 minutes into support, after which the pump was disconnected about 1hr into support. purge pressure was trending up slightly with the purge flow trending down which could indicate possible biomaterial buildup but none was seen on the product and therefore could not be confirmed. low pump flow: the root cause of low flow was unable to be determined as limited clinical details were provided and the product returned could not be run to get additional information. the root cause of low flow was unable to be determined as limited clinical details were provided and the product returned could not be run to get additional information. d6a/d6b are unknown for pump serial number (b)(6). d9 device returned to manufacturer date added statement was inadvertently omitted on manufacturer device report number 1220648-2024-07010-1. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-07010 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-07010 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-07010. h6 investigation findings was erroneously entered on the manufacturer device report number 1220648-2024-07010-1. h6 type of investigation code 10, 4121, and 4110 was inadvertently omitted on manufacturer device report number 1220648-2024-07010-1. h10 investigation results were erroneously entered on the manufacturer device report number as the device was received for investigation 1220648-2024-07010-1. | The pump was returned for investigation. there were no cannula kinks. no impeller damage was seen. no biomaterial or foreign material was seen inside the cannula, purge gap, or around the impeller blades. no damaged blades were observed. based on data log review no pump stops were observed and the pump was able to run during support. data logs were returned for analysis. suction and "impella flow reduced" alarms were seen about 40 minutes into support. no mc spikes outside of p-level changes. purge pressure slightly increasing with decrease in purge flow but still in range with no alarms. low pump flow: the root cause of low pump flow was most likely patient condition based on clinical details provided and data log analysis. b7 other relevant history, including preexisting medical conditions updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
30734 | 1220648-2024-07011 | 18767984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18767984&pc=OZD | 02/12/2024 | 03/29/2024 |  | 02/23/2024 |  |  | 02/12/2024 | 02/23/2024 |  |  | 11/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024433018 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 64-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that patient experienced access site bleeding, manual pressure was applied and packed red blood cells were given. the patient was reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30735 | 1220648-2024-07012 | 18768007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18768007&pc=OZD | 02/12/2024 | 03/28/2024 |  | 02/23/2024 |  |  | 02/12/2024 | 02/23/2024 |  |  | 11/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025438263 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 58-year-old female patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that the pump was explanted due to lack of flow to the impella leg. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30736 | 1220648-2024-07013 | 18768137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18768137&pc=OZD | 02/13/2024 | 11/25/2025 |  | 02/23/2024 |  |  | 02/13/2024 | 02/23/2024 |  |  |  |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1576134 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 68-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that automated impella controller (aic) alternated high and low purge flow and purge pressure alarms. the aic was replaced, and no patient harm was reported. | [5 empty] | The impella device was received from the customer and the evaluation has started but not yet completed. upon investigation closure, a supplemental mdr will be filed.
30737 | 1220648-2024-07014 | 18768200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18768200&pc=OZD | 02/14/2024 | 04/11/2024 |  | 02/23/2024 |  |  | 02/14/2024 | 02/23/2024 |  |  | 10/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024419355 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to severe left main artery disease. while on support, the patient sat up and a large hematoma was noted. manual pressure was applied and the decision was made to remove the impella. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely patient movement as there was patient movement observed during the impella support as per the clinical description provided.
30738 | 1220648-2024-07015 | 18768233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18768233&pc=OZD | 02/15/2024 | 03/29/2024 |  | 02/23/2024 |  |  | 02/15/2024 | 02/23/2024 |  |  | 11/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025438786 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Dislodged or Dislocated | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that ecp (expandable percutaneous heart pump) case was done but the patient became hemodynamically unstable and required impella cp placement for continued support post-pci (percutaneous coronary intervention). after support was initiated for the cp, ecp was weaned to p-0 level and explanted. during explant of ecp, the pigtail got caught around the outlet of cp, and kinked cp. suction events were noted. the kink didn¿t resolve with repositioning. impella cp was explanted and new cp placed. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the initial report was submitted. the device was returned, visually inspected, and tested. kink was found on cannula. the root cause of the low flow was a kinked cannula due to interaction with ecp device.
30739 | 1220648-2024-07016 | 18768376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18768376&pc=OZD | 02/15/2024 | 04/02/2024 |  | 02/23/2024 |  |  | 02/15/2024 | 02/23/2024 |  |  | 10/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024414864 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to non-ischemic cardiomyopathy. while on support the patient had a few runs of ventricular tachycardia (vt), the patient's urine was discolored and appeared cola colored, and the patient had a left ventricle thrombus. support was discontinued and the patient expired. there was no relationship between the impella and cause of death. | [5 empty] | The investigation into the hemolysis and the vf/vt/af/at issues has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation and no sufficient clinical data provided. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30740 | 1220648-2024-07017 | 18768414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18768414&pc=OZD | 02/15/2024 | 04/01/2024 |  | 02/23/2024 |  |  | 02/15/2024 | 02/23/2024 |  |  | 06/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024427259 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and while on support the patient had hemolysis. one unit of blood products was provided and the pump was repositioned. once the pump was pulled back, the hemolysis resolved. however, the impella was removed later that day due to the patient being at baseline. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. based on clinical description & data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position as hemolysis was resolved after repositioning of the pump.
30741 | 1220648-2024-07018 | 18768474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18768474&pc=OZD | 02/17/2024 | 04/04/2024 |  | 02/23/2024 |  |  | 02/17/2024 | 02/23/2024 |  |  |  |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433497 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was in poor position. the team attempted to reposition, but labs continued to hemolyze. the impella cp was exchanged for an impella 5.5. in addition, it was reported that the purge cassette was leaking. the cassette was changed out and the issue resolved. | [5 empty] | The investigation into hemolysis and the purge leak ¿ cassette issues has been completed since the initial report was submitted. the device was not returned for investigation; only the data stick was returned. based on the clinical details and data log revision, the cause of the hemolysis issue was most likely improper positioning issue due to improper technique. the cause of the purge cassette leak cannot be determined since the complaint purge cassette was not returned for investigation. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30742 | 1220648-2024-07019 | 18769307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18769307&pc=OZD | 02/05/2024 | 03/26/2024 |  | 02/23/2024 |  |  | 02/05/2024 | 02/23/2024 |  |  | 10/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024418848 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Twisted/Bent | Material Split, Cut or Torn | [5 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 30-year-old male patient was implanted with an impella 5.5 due to needing a bridge for a transplant. while on support, there was a kink noted in the pump catheter. metrics did not change so support was continued. the pump was repositioned, and the sterile sleeve ripped. the staff taped the sleeve. additionally, the patient went into atrial fibrillation during the repositioning. furthermore, the patient had bleeding. it is unknown how the bleeding was resolved. moreover, the patient had renal failure while on support. the staff noted that continuous renal replacement therapy may be started. additionally, the patient had a fever from unknown causes. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported renal failure, access site bleeding, infection, cannula damage and product kink has been completed. the impella device has not be returned for evaluation. clinical details provided were sufficient to determine the root cause. it was noted that the sterile sleeve ripped when repositioning the pump. the root cause of the sterile sleeve product damage was due to use when clinical team was repositioning pump via excessive torqueing. the root cause of the cosmetics defect on the pump was a catheter kink that was most likely due to excessive force applied when repositioning pump as it did not affect the function of the pump. due to a lack of clinical details of where the bleeding occurred and how it was resolved, the root cause of the bleeding cannot be determined. the root cause of the renal failure cannot be determined as insufficient clinical details were provided. the root cause of the infection cannot be determined. the root cause of the vt/vf was not determined.
30743 | 1220648-2024-07020 | 18770675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18770675&pc=OZD | 02/13/2024 | 04/01/2024 |  | 02/23/2024 |  |  | 02/13/2024 | 02/23/2024 |  |  | 11/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435088 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the us complainant had placed a 14 french sheath to deliver an impella cp to support a patient in need of a high risk percutaneous coronary intervention (hrpci). the team first had to treat the impella accessed iliofemoral anatomy with shockwave. the impella cp was successfully delivered. however, the vessel was dissected. the team instead placed a long 45 centimeter sheath to the dissected anatomy. the vessel still has flow with the sheath. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation into the vascular damage - peripheral vessel issue has been completed since the initial report was submitted. the device was returned, visually inspected, and tested. product evaluation confirmed that the peel away sheath had not been split. there were multiple kinks found along the introducer sheath. the most likely cause of the peripheral vessel vascular damage was use related as there were no vessel conditions and the sheath was left in place and kinked.
30744 | 1220648-2024-07022 | 18771147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18771147&pc=OZD | 09/19/2022 | 04/16/2025 |  | 02/23/2024 |  |  | 09/19/2022 | 02/23/2024 |  |  | 06/30/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237145 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2022 |  |  |  |  | The user facility reported that a 55-year-old male patient on impella cp mechanical circulatory support, developed some bleeding around the sheath. the bleeding was treated with applied pressure and an additional figure eight (8) suture was placed. the patient was sent to the unit with fem-stop (femoral compression system). the patient was not receiving systemic heparin at the time. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30745 | 1220648-2024-07021 | 18771156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18771156&pc=OZD | 09/17/2022 | 03/11/2025 |  | 02/23/2024 |  |  | 09/17/2022 | 02/23/2024 |  |  | 11/30/2023 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023164462 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 62 year old female patient requiring mechanical circulatory support. it was reported that the patient experienced ventricular tachycardia (vtach) when the impella cp needed repositioning. the patient required defibrillation and increased vasopressin use. an echocardiogram was performed, and the pump was repositioned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30746 | 1220648-2024-07023 | 18771628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18771628&pc=OZD | 09/21/2022 | 02/23/2024 |  | 02/23/2024 |  |  | 09/21/2022 | 02/23/2024 |  |  | 05/31/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231622 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported that a 52 year old male patient was on impella cp for mechanical circulatory support. there were concerns that inlet position and suction events were above the p4 performance level. an echocardiogram was performed. the physician advanced and rotated the inlet away from the pap muscle. waveforms improved and performance levels were able to increase without suction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30747 | 1220648-2024-07024 | 18771783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18771783&pc=OZD | 09/21/2022 | 02/23/2024 |  | 02/23/2024 |  |  | 09/21/2022 | 02/23/2024 |  |  | 05/31/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227739 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported an impella cp implanted in a 30-year-old female patient for mechanical circulatory support. it was reported after the impella cp was successfully implanted, a light ooze from the impella cp site was noted. the site was stitched. the patient also had a large hematoma in a right axillary artery with manual pressure performed. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30748 | 1220648-2024-07025 | 18771973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18771973&pc=OZD | 09/26/2022 | 09/18/2024 |  | 02/23/2024 |  |  | 09/26/2022 | 02/23/2024 |  |  | 03/31/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204065 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported an impella cp on an 80 year-old-male requiring mechanical circulatory support. it was reported that after placement of the cp, the septal wall was blown. the patient expired within three hours of transport to the unit. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
30749 | 1220648-2024-07026 | 18772288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18772288&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/23/2024 |  |  | 02/08/2024 | 02/23/2024 |  |  | 09/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024399427 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was successfully weaned and explanted but thrombus was noted during closure of impella access site. it was further reported that after impella insertion the patient had a slight bleeding from the groin site that was managed by manual pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus and the access site bleeding ¿ minor has been completed since the initial report was submitted. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus and the minor the access site bleeding was not determined.
30750 | 1220648-2024-07027 | 18772482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18772482&pc=PYX | 02/12/2024 | 04/01/2024 |  | 02/23/2024 |  |  | 02/12/2024 | 02/23/2024 |  |  | 08/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024399174 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 67-year-old male patient of unknown race and ethnicity post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported that the rp was unable to be delivered. the team attempted multiple times and with multiple wires. the rp implant was aborted. the was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30751 | 1220648-2024-07028 | 18772604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18772604&pc=OZD | 02/12/2024 | 03/29/2024 |  | 02/23/2024 |  |  | 02/12/2024 | 02/23/2024 |  |  | 10/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024423071 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 50-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump had experienced suction issues overnight during patient support. an echocardiogram was performed and revealed a thrombus in the left ventricle. due to the noted thrombus, the decision was made to explant the pump. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30752 | 1220648-2024-07029 | 18772717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18772717&pc=OZD | 02/14/2024 | 03/19/2024 |  | 02/23/2024 |  |  | 02/14/2024 | 02/23/2024 |  |  |  |  | 02/23/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1734346 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | No patient involvement: us complainant reported the automated impella controller had a battery failure (red alarm). the aic was replaced. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) battery failure alarm has been completed. the console was returned for evaluation. the aic was evaluated and it was discovered that ac power disconnected alarm would not resolve as it normally would. when moving the power cord around there would be sporadic ac power disconnected alarm resolved notifications but it was not constant. examined the ac plug portion of the power cord and found that one of the prongs was loose. the data logs were reviewed finding frequent instances of ac power disconnected alarms on the reported occurrence date. ran a batttest to check the status of the batteries and found that there were no evident signs of battery cell imbalance in either of the batteries.
30753 | 1220648-2024-07030 | 18772767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18772767&pc=OZD | 02/15/2024 | 03/08/2024 |  | 02/23/2024 |  |  | 02/15/2024 | 02/23/2024 |  |  |  |  | 02/23/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1097384 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Component Incompatible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 52-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support when the automated impella controller (aic) had an alarm stating the purge disc not detected. the staff tried to remove and reinsert the purge cassette multiple times, but they were unable to clear the alarm. staff switched to a backup console and the alarm resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. added-device return date d4-updated model number, udi #.
30754 | 1220648-2024-07031 | 18772873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18772873&pc=OZD | 02/15/2024 | 03/25/2024 |  | 02/23/2024 |  |  | 02/15/2024 | 02/23/2024 |  |  | 11/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025438726 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp to support a patient in ami/cgs (acute myocardial infarction¿/ cardiogenic shock) . the patient was sent to the or for cabg (coronary artery bypass grafting). while in the or the team used soft clamps to cross clamp the cp. the soft clamp had been a discussion for appropriate use. the aorta was however found to be damaged at the site of the cross-clamp. the team had to repair the aorta with a dacron patch and cp removal and escalation to the 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the vascular damage - great vessel has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the vascular damage - great vessel issue was not determined since product was not returned and insufficient clinical information was provided.
30755 | 1220648-2024-07032 | 18773005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773005&pc=OZD | 02/16/2024 | 05/25/2024 |  | 02/23/2024 |  |  | 02/16/2024 | 02/23/2024 |  |  | 07/31/2025 |  | 02/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205390 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | Upon review by abiomed's medical safety officer, the patient expired on support. there is not enough information to exclude the patient's outcome from the use of device. | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella stopped post repositioning on corevalve tavr (transcatheter aortic valve replacement). the aic (automated impella controller) is now reporting failure. |  |  |  |  | The investigation into the aic - pump stop has been completed since the initial report was submitted. the logs showed multiple impella stopped retrograde flow alarms. it could be observed that before and after each pump stop, the console attempted to restart the rotation of the pump rotator. eventually, numerous communication errors from the impellatronic board software were recorded. the attempts to reset the software eventually lead to the issuance of the impella stopped (mechanical/electrical issue) alarm. multiple issuances of this alarm occurred before the pump was unplugged and re-plugged into the console. the accumulation of this communication errors, resets and alarms led the console to reset the whole system but was unable to pass the system self check. the console successfully booted when it was received in the lab. the internal circuitry was inspected and no abnormality was found. when a known good pump and purge line was attached to the system, the failure mode was reproduced after about three hours of operation at the highest p-level. the failure was characterized by buzzing of an alarm as well as the other symptoms observed on the date of occurrence. rebooting the console following the failure led to system self check failure. the failure could be replicated one more time. it should be noted that the console could normally boot when it was turned off for over 24 hours. when the suspected defective impellatronic board was replaced with a know good board, the console was run overnight for more than 12 hours without any related abnormalities confirming that the failure was isolated to the impellatronic board. the root cause for the pump stop was a defective impellatronic board causing an alarm to trigger. the system self check failure occurred after the console was rebooted with a defective impellatronic board. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.2 common device name and d.4 model number were revised from the initial submission in accordance with updated procedures. d.4 revised catalog number as previously entered incorrectly and expiration date was previously entered incorrectly as there is no expiration date. h.3 device not returned to manufacturer for evaluation was previously selected incorrectly as the device was returned to be evaluated. h.6 code 925 was reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. b.2, b.5 and h.1 were revised due to medical safety officer review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
30756 | 1220648-2024-07033 | 18773021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773021&pc=OZD | 02/16/2024 | 04/03/2024 |  | 02/23/2024 |  |  | 02/16/2024 | 02/23/2024 |  |  | 07/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025441205 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that immediately following implant, the impella cp was found to be stuck under the papillary muscle. repositioning was attempted, but the pump subsequently kinked and support levels could not be raised above p2 with persistent suction. due to the ongoing issue, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical data provided, the cause of the positioning issue was use related due to improper technique and led to kink. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
30757 | 1220648-2024-07034 | 18773265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773265&pc=OZD | 10/04/2022 | 05/22/2025 |  | 02/23/2024 |  |  | 10/04/2022 | 02/23/2024 |  |  | 03/31/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | The complainant reported that a 88-year-old male patient requiring mechanical circulatory support was being weaned from the impella cp. the last angiogram showed distal wire perforation of a branch off of the circulatory. the impella cp then ramped back up to auto and a balloon was inserted to slow flow to the perforation. a covered stent was placed at the site and pericardiocentesis was then performed. the physician was able to successfully shutoff perforation site and draw off around 500cc of blood from the pericardium and the impella cp explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30758 | 1220648-2024-07035 | 18773809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773809&pc=OZD | 08/05/2023 | 03/15/2024 |  | 02/23/2024 |  |  | 01/26/2024 | 02/23/2024 |  |  | 03/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343155 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 73 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support had endured multiple episodes of vt/vf during impella support, with a possible relationship to the impella device: treatment included continuation of "iv amiodarone and iv lidocaine, cpr for 2nd episode, shocks for 2nd and 3rd episode." | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. added- h6 impact code 4644 d4-updated model number.
30759 | 1220648-2024-07036 | 18773821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773821&pc=OZD | 12/12/2023 | 02/26/2025 |  | 02/23/2024 |  |  | 01/24/2024 | 02/23/2024 |  |  | 08/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397909 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Pseudoaneurysm | Unspecified Vascular Problem | [10 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | Followup |  |  | 01/24/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured a left femoral artery pseudoaneurysm from the impella cp access site with a "definite" relationship to the impella device. this was discovered during a follow up visit over 1 month after explant. a repair was administered. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿the patient developed hemolysis. admission total hemoglobin 11.3 g/dl, nadir 8.4 g/dl; peak ldh 633 u/l ((b)(6) 2023). urine was noted as being "pink tinted". urinalysis on (b)(6) 2023 reveals blood in urine.¿. the patient recovered with no sequelae. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the pump stop and the vascular damage - peripheral vessel issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. d4-model number updated to impella cp. | B5 revised to reflect new information. h6 revised to reflect new information. instructions for use hemolysis. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ f6/f8-should have been left blank on manufacturer device report 1220648-2024-07036. g1 reporting contact email revised on manufacturer device report 1220648-2024-07036 in accordance with updated procedures. g1 reporting contact fax number was revised on manufacturer device report 1220648-2024-07036 as it was inadvertently omitted. h6 component code revised on manufacturer device report 1220648-2024-07036 from the initial submission in accordance with updated procedures.
30760 | 1220648-2024-07037 | 18773828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773828&pc=OZD | 11/07/2023 | 03/12/2024 |  | 02/23/2024 |  |  | 01/26/2024 | 02/23/2024 |  |  | 06/30/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024356963 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 53 year old patient who endured a descending aortic dissection with a "possible" relationship to the impella device. no other information is available. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the vascular damage - great vessel issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the great vessel vascular damage could not be determined because not enough information was given.
30761 | 1220648-2024-07038 | 18773830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773830&pc=OZD | 11/06/2023 | 04/26/2024 |  | 02/23/2024 |  |  | 01/26/2024 | 02/23/2024 |  |  | 08/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024395659 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/26/2024 |  |  |  |  | It was noted that after the a-line was removed, additional information provided notes "rue color improved" but "there were still ischemic changes to tips of digits." | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured chest wall bleeding and rue and limb ischemia occurred with a possible relationship to the impella cp device used on this patient. intervention was performed via a fasciotomy and delivery of blood products. |  |  |  |  | Additional information added regarding some improvement of ischemia after removal of a-line. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding and the limb ischemia cannot be determined since the product was not returned and insufficient clinical details were provided.
30762 | 1220648-2024-07039 | 18773831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773831&pc=OZD | 10/20/2023 | 03/15/2024 |  | 02/23/2024 |  |  | 01/26/2024 | 02/23/2024 |  |  | 01/31/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182783 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured "significant bleeding and hematoma (right groin)" requiring at least 1 unit of blood as medical intervention. while the patient presented with bilateral groin bleeds prior to placement of the device, it was alleged this event had a "possible" relationship to the impella device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that post procedure, the patient's activated coagulation time (act) was prolonged at 250. therefore, the root cause of the access site bleed was most likely due to improper anticoagulation management.
30763 | 1220648-2024-07040 | 18773832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773832&pc=OZD | 09/29/2023 | 03/15/2024 |  | 02/23/2024 |  |  | 01/26/2024 | 02/23/2024 |  |  | 07/31/2025 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024371307 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured ventricular tachycardia with a "possible" relationship to the impella device: "multiple episodes: (b)(6) 2023." it was reported that "lido and amini" were infused as a result. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the ventricular tachycardia (vt) issue has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
30764 | 1220648-2024-07041 | 18773833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18773833&pc=OZD | 08/21/2023 | 03/12/2024 |  | 02/23/2024 |  |  | 01/26/2024 | 02/23/2024 |  |  | 12/31/2024 |  | 02/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024305828 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured hemolysis with a "possible" relationship to the impella device: patient was given packed red blood cells to address the issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the hemolysis issue has been completed since the original report was filed. the device was not returned for investigation the root cause of the hemolysis cannot be determined as the pump was not returned and there is no indication of any events that may have caused hemolysis.
30765 | 1220648-2024-06967 | 18756661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18756661&pc=PYX | 02/09/2024 | 10/24/2025 |  | 02/22/2024 |  |  | 02/09/2024 | 02/22/2024 |  |  | 02/28/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024346413 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Pressure Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 34-year-old female was implanted with an impella rp device for mechanical circulatory support. while on support, high purge pressure alarms occurred for the past two days without resolving. the clinician already changed the cassette and will try tissue plasminogen activator (tpa) for the purge system. it was further reported that the impella rp flex pump had stopped alarms. numerous attempts to restart were unsuccessful. the patient was hemodynamically stable and the left sided impella 5.5 support was unchanged and operating well. the impella rp flex pump was removed and a new one was inserted. once the device was removed, a clot burden was noted in the outlet cage of the device. a pulmonary artery catheter was attempted in the left internal jugular to assess cardiac need for replacement. unfortunately, a clot was found in the left internal jugular. the right internal jugular was accessed and premature atrial contractions (pac) were successfully placed. the right heart filling pressures were normal. the decision was made to not place a new impella rp flex device. the patient was also noted to be on bivalirudin for systemic anticoagulation due to possible heparin-induced thrombocytopenia (hit) exposure. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
30766 | 1220648-2024-06969 | 18756972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18756972&pc=OZD | 02/09/2024 | 03/28/2024 |  | 02/22/2024 |  |  | 02/09/2024 | 02/22/2024 |  |  | 10/31/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024412204 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 79-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as escalation of care from an impella cp device. the graft was placed. however due to anatomy, the impella 5.5 could never be successfully delivered. the patient was again delivered an impella cp pump. the patient was still being considered for coronary artery bypass grafting or revascularization. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the delivery issue was determined to be patient condition due to the calcified vessel condition of the patient and impella was unable to advance because of their smaller diameter.
30767 | 1220648-2024-06970 | 18757316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18757316&pc=OZD | 02/10/2024 | 03/29/2024 |  | 02/22/2024 |  |  | 02/10/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424895 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Material Deformation | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2024 |  |  |  |  | The user facility reported a 59-year-old male enrolled in the st-segment elevation myocardial infarction door to unloading protocol was in acute myocardial infarction/cardiogenic shock and was implanted with an impella cp device for mechanical circulatory support. the team delivered the impella cp pump via the long 14 french sheath. the long sheath was placed due to body mass index/obesity and depth of tissue tract. the sheath kinked and the patient had access site bleed that prompted the patient to develop a hematoma. it was felt the repo sheath was not long enough due to the patient's size and best course of action would be to stay on catheterization table and manually hold until angiomax wore off in two hours and then pull pump in same setting. the product was discarded. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major, the introducer damage ¿ mechanical and the limb ischemia - reversible issues has been completed since the initial report was submitted. the product was not returned for investigation. the most likely cause of the major access site bleeding was the patient's condition. the most likely cause of the mechanical introducer damage was a sheath kink. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
30768 | 1220648-2024-06971 | 18757477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18757477&pc=OZD | 02/10/2024 | 03/25/2024 |  | 02/22/2024 |  |  | 02/10/2024 | 02/22/2024 |  |  | 10/31/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415864 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | Unintended Movement | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2024 |  |  |  |  | The user facility reported a 66-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support due to needing a bridge to ventricular assist device or transplant. while on support, the patient sat in a chair and there was an, "impella in aorta alarm". additionally, there was high purge pressure and there was pump stop. the impella was removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the pump stop, the positioning issues, and the high purge pressure has been completed since the initial report was submitted. the product was not returned for investigation. according to the data log analysis, the purge pressure could be seen increasing in a stepwise fashion leading to the pump stop. the most likely cause of the pump stop was a kinked purge line. the most likely cause of the positioning issue was patient movement.
30769 | 1220648-2024-06972 | 18757584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18757584&pc=OZD | 02/11/2024 | 03/29/2024 |  | 02/22/2024 |  |  | 02/11/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025439288 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2024 |  |  |  |  | The user facility reported a 68-year-old male patient was implanted with an impella cp as a bridge to transplant. it was reported that while on support the patient developed hemolysis. the impella cp was repositioned and after no improvement, the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined.
30770 | 1220648-2024-06973 | 18757724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18757724&pc=OZD | 02/13/2024 | 03/29/2024 |  | 02/22/2024 |  |  | 02/13/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424740 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support due to having coronary artery disease. while on support, the patient lost pulses in their left leg (impella side). vascular was consulted and the impella cp device was removed and replaced with an impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the limb ischemia ¿ reversible has been completed since the initial report was submitted. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30771 | 1220648-2024-06974 | 18759273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18759273&pc=OZD | 09/25/2022 | 09/14/2024 |  | 02/22/2024 |  |  | 09/25/2022 | 02/22/2024 |  |  | 02/29/2024 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194752 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 75-year-old male patient for mechanical circulatory support. it was reported that the pump exhibited elevated purge pressures. the patient had multiple episodes of ventricular tachycardia (vtach) while on impella cp support which required defibrillations and increased vasopressin use. an echocardiogram was performed, and the pump was repositioned. | It was noted the patient's care was withdrawn and the patient expired while on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06974 was provided.
30772 | 1220648-2024-06975 | 18759872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18759872&pc=OZD | 02/08/2024 | 03/20/2024 |  | 02/22/2024 |  |  | 02/08/2024 | 02/22/2024 |  |  | 10/31/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422879 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 67-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, there was a thrombus observed in the left ventricle. an attempt was made at removal, but the clot remained. the thrombosis was also present in the extra-corporeal membrane oxygenation (ecmo) accessed limb. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
30773 | 1220648-2024-06976 | 18759930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18759930&pc=OZD | 02/09/2024 | 03/22/2024 |  | 02/22/2024 |  |  | 02/09/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422018 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 73-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support, the patient experienced premature ventricular contraction (pvc). a trans-esophageal echocardiogram (tee) was performed to adjust positioning of the pump and it was documented that the pump could not be moved away from the mitral valve structures. as the patient's hemodynamics were more stable, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues and the premature ventricular contraction (pvc) has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. as the necessary information was not provided, the root cause of the pvc was not determined.
30774 | 1220648-2024-06977 | 18759970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18759970&pc=OZD | 02/11/2024 | 03/29/2024 |  | 02/22/2024 |  |  | 02/11/2024 | 02/22/2024 |  |  | 06/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372871 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2024 |  |  |  |  | The user facility reported a 61-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was experiencing bleeding from the groin. fem-stop was applied, heparin was stopped, and two units of packed red blood cells were given. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30775 | 1220648-2024-06978 | 18760109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760109&pc=OZD | 02/06/2024 | 04/01/2024 |  | 02/22/2024 |  |  | 02/06/2024 | 02/22/2024 |  |  | 06/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431036 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 54-year-old male (race unknown) undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the impella cp was delivered but was replaced due to lost placement signal (ps). the loss of ps caused no harm or injury but, the team placed a new 2nd pump after 58 hours of support. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was not returned for evaluation. data logs were reviewed. there was placement signal not reliable alarm found after 6 minutes of connection pump to the aic controller. the placement signal was noted to be consistent at 3000 mmhg. therefore, per the data log review, the root cause of the optical signal issue was most likely a damaged optical fiber line. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30776 | 1220648-2024-06979 | 18760154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760154&pc=OZD | 09/26/2022 | 02/22/2024 |  | 02/22/2024 |  |  | 09/26/2022 | 02/22/2024 |  |  | 03/31/2024 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214667 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 58-year-old male patient for mechanical circulatory support. it was reported that while on impella cp support the patient developed oozing at the groin, but there was no hematoma. additionally, both feet are discolored, but pulses could be heard with the doppler. the purge was switched to bicarb to help with partial thromboplastin time (ptt) levels and both feet have returned to normal color. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30777 | 1220648-2024-06980 | 18760175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760175&pc=OZD | 02/12/2024 | 03/29/2024 |  | 02/22/2024 |  |  | 02/12/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422098 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 57-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. during support, the sterile sleeve on an implanted impella 5.5 pump was found to be torn. the tear was noted by the staff during attempts to reposition the device. support remained ongoing despite the noted damage. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported product damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the product damage could not be determined.
30778 | 1220648-2024-06981 | 18760233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760233&pc=OZD | 02/12/2024 | 03/28/2024 |  | 02/22/2024 |  |  | 02/12/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430702 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 78-year-old female patient of unknown ethnicity in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that after 21 hours of support the pump was explanted via surgical cutdown due to the impella accessed limb became ischemic noted by loss of pulse. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30779 | 1220648-2024-06982 | 18760242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760242&pc=OZD | 09/26/2022 | 05/21/2025 |  | 02/22/2024 |  |  | 09/26/2022 | 02/22/2024 |  |  | 06/30/2024 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported a 76-year-old male patient was implanted with an impella cp for mechanical circulatory support. after being placed on impella cp support, there was very low pulsatility on placement signal with flows 4lpm. one drug eluting stent (des) was placed to the left main (lm) coronary artery, resulting in recovery of pulsatility. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30780 | 1220648-2024-06983 | 18760269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760269&pc=OZD | 02/13/2024 | 04/01/2024 |  | 02/22/2024 |  |  | 02/13/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427757 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump began to experience high purge pressure during patient support. the purge cassette was exchanged to troubleshoot the issue, but the high purge alarm persisted. due to the noted condition, the decision was made to remove the pump, clean the graft, and implant a new impella 5.5 device. upon explant, a clot was found around the motor housing of the device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported high purge pressure has been completed. the impella has not been returned for evaluation. data logs were reviewed and showed flow saturation and a gradual increase in purge pressure over 12 minutes. the root cause of the high purge pressure issue was most likely biomaterial which would have been a result of the use of the previously used graft.
30781 | 1220648-2024-06984 | 18760331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760331&pc=OZD | 02/13/2024 | 03/29/2024 |  | 02/22/2024 |  |  | 02/13/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024428126 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was ambulating and getting pump in aorta alarms. impella 5.5 was repositioned, but the alarms continued. the decision was made to exchange the impella for another 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the initial report was submitted. the product was not returned for investigation. clinical data confirmed the pump pulled back during ambulation and "pump in aorta" alarm occurred. the cause of the positioning issue was patient movement.
30782 | 1220648-2024-06985 | 18760397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760397&pc=OZD | 02/14/2024 | 03/07/2024 |  | 02/22/2024 |  |  | 02/14/2024 | 02/22/2024 |  |  |  |  | 02/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1598659 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | During training with cvor staff, the complainant reported that when placing the purge cassette in the blue purge mount, it was noticed that the disk flange in the aic (automated impella controller) was broken. the aic was replaced. there was no patient involvement. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. d4-updated model number | The impella device was received from the customer on (b)(6)2024 and therefore,an evaluation of the device is under way. upon investigation closure, a supplemental mdr will be filed.
30783 | 1220648-2024-06986 | 18760641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760641&pc=OZD | 02/14/2024 | 03/08/2024 |  | 02/22/2024 |  |  |  | 02/22/2024 |  |  |  |  | 02/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1236846 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The complainant reported that during a set up for an emergency case, it was noticed that the blue purge disc holder was broken. there was no patient involvement. | [5 empty] | The impella device was received from the customer on 20-feb-2024 and therefore, an evaluation of the device was is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. the aic's reported issue with the broken purge retainer was confirmed. purge disc retainer was completely broken off. the cause of the issue was a broken purge retainer.
30784 | 1220648-2024-06987 | 18760988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18760988&pc=OZD | 02/14/2024 | 04/01/2024 |  | 02/22/2024 |  |  | 02/14/2024 | 02/22/2024 |  |  |  |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437337 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having cardiac arrest. while on support, the patient lost distal flow and the decision was made to explant the impella. reportable due to patient harm and necessary intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30785 | 1220648-2024-06988 | 18761212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18761212&pc=OZD | 02/19/2024 | 04/09/2024 |  | 02/22/2024 |  |  | 02/19/2024 | 02/22/2024 |  |  | 10/31/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024424904 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 89-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and upon implanting the impella the abiomed sheath kinked in the mid shaft after the impella was placed. the sheath was exchanged for the 6f single access sheath. no further complications were noted. the sheath was unable to be obtained. | [5 empty] | The investigation into the delivery issues and the access site bleeding ¿ major issue has been completed since the original report was submitted. per clinical information provided, there was introducer damage ¿ mechanical. the root cause of the introducer damage was most likely due sheath kink since kinking of sheath body was observed. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30786 | 1220648-2024-06989 | 18763243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18763243&pc=OZD | 01/24/2024 | 03/07/2024 |  | 02/22/2024 |  |  | 01/24/2024 | 02/22/2024 |  |  | 08/31/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394654 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Positioning Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 66-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient had a subarachnoid hemorrhage. it was unknown if the impella was related to the subarachnoid hemorrhage. additionally, the patient's urine production was dropping. units of packed red blood cells were provided, and continuous renal replacement therapy was started. also, there were suction events and volume was provided. on support, the pigtail of the impella went through the heart tissue and was sitting in the pericardial space. the impella was pulled back and a patch was sewn in place. furthermore, the patient also had a hematoma that was expressed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported issues flow issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding - major is most likely due to patient movement as it was stated that the bleeding occurred after transfer to the icu. the root cause of the ventricle perforation was most likely due to improper positioning. the root cause of the low pump flow is most likely due to the patient¿s condition. as the clinical information was not provided, the root cause of the hemolysis and the stroke/cva could not be determined.
30787 | 1220648-2024-06990 | 18763719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18763719&pc=OZD | 02/11/2024 | 03/21/2024 |  | 02/22/2024 |  |  | 02/11/2024 | 02/22/2024 |  |  | 09/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024406325 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2024 |  |  |  |  | The user facility reported a 55-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was then transferred to the tertiary center. at transfer, the patient was observed to have a access site bleed. the bleeding was treated by a suture and blood products - six units of fresh frozen plasma and two units of red blood cells. the pump remains on for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since the hub was not sutured leading to profusely bleeding.
30788 | 1220648-2024-06991 | 18763731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18763731&pc=OZD | 02/12/2024 | 03/28/2024 |  | 02/22/2024 |  |  | 02/12/2024 | 02/22/2024 |  |  | 09/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024406447 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 53-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient experienced a mottled leg and loss of pulse in the impella side. the impella was explanted, and the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30789 | 1220648-2024-06992 | 18763751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18763751&pc=OZD | 02/12/2024 | 04/01/2024 |  | 02/22/2024 |  |  | 02/12/2024 | 02/22/2024 |  |  | 09/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397212 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 75-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the implanted impella 5.5 pump experienced high purge pressure and triggered purge system blocked alarms during patient support. the purge cassette was swapped and a tissue plasminogen activator protocol was initiated to troubleshoot the issue but the alarms remained ongoing. due to the persistent failure, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was returned, visually inspected, and tested. upon investigation of the pump, a large clot was observed around the impeller and in the purge gap. the root cause of the high purge pressure was determined to be biomaterial.
30790 | 1220648-2024-06993 | 18763820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18763820&pc=OZD | 02/12/2024 | 04/01/2024 |  | 02/22/2024 |  |  | 02/12/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432586 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 64-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed a hematoma on the impella access site. computed tomography showed a small extravasation of the right femoral artery. heparin was discontinued and pressure was applied. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643 apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no abnormalities were found on the product.
30791 | 1220648-2024-06994 | 18764011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764011&pc=OZD | 02/13/2024 | 03/21/2024 |  | 02/22/2024 |  |  | 02/13/2024 | 02/22/2024 |  |  | 10/31/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024421084 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having heart failure. while on support, low pump flows were noted, pump position was verified, and the team elected to replace the pump. a visible clot was noted on the impella blades during removal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the low pump flow issue has been completed since the original report was filed. the device was not returned for investigation. based on clinical description & data review, the root cause of the low flow was most likely biomaterial ingestion.
30792 | 1220648-2024-06995 | 18764025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764025&pc=OZD | 02/13/2024 | 03/29/2024 |  | 02/22/2024 |  |  | 02/13/2024 | 02/22/2024 |  |  | 11/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426582 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Nerve Damage | [11 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 28-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support ongoing for a patient awaiting transplant. the younger patient (28yo male) was admitted with multiple cardiac comorbidities and a bipella support with protek centri-mag pump. the patient support was ongoing for 9 days when the 5.5 site developed a hematoma and nerve pain. the hematoma was not treated, but the nerve pain was by starting gabapentin. the pump was explanted on next day when heart transplant took place. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and hemolysis could not be determined. added- h6 impact code 4627 d4-updated model number. h5-labeled for single use changed to yes.
30793 | 1220648-2024-06996 | 18764032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764032&pc=OZD | 02/14/2024 | 11/20/2025 |  | 02/22/2024 |  |  | 02/14/2024 | 02/22/2024 |  |  |  |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1161463 |  |  |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | No patient involvement: the complainant reported that a display popped up during priming that could not be bypassed. the aic (automated impella controller) was rebooted to resolve the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Data analysis: imc logs revealed that the console (b)(4) had experienced the epm crystal issue, and that the console had then automatically performed steps to mitigate it. when the epm crystal issue occurs, the console is unable to recognize the presence of the pump. this matches the reported complaint. device analysis: console (b)(4) was sent back fs for further investigation. the reported complaint could not be reproduced. the pump detection issue was confirmed to be due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. d4: corrected the common device name d9: added the device return information. g1: corrected the manufacturer email address. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation.
30794 | 1220648-2024-06997 | 18764038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764038&pc=OZD | 02/14/2024 | 03/21/2024 |  | 02/22/2024 |  |  | 02/14/2024 | 02/22/2024 |  |  |  |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401635 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the placement signal stopped appearing on the aic (automated impella controller) and was followed by a controller error alarm during patient support. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation for the reported console (aic) "controller error" alarm has been completed. the aic was returned for evaluation. data logs were reviewed which confirmed the controller error alarms. functional testing of the console did not confirm the event. the alarm is typically caused by a known pattern failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm, and has a low probability of reoccurrence. therefore the root cause of the controller alarm issue could not be determined. | The investigation for the reported issue has not yet been completed. upon its completion, a follow up report will be submitted.
30795 | 1220648-2024-06998 | 18764074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764074&pc=OZD | 02/09/2024 | 03/29/2024 |  | 02/22/2024 |  |  | 02/15/2024 | 02/22/2024 |  |  | 10/31/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024407792 |  | 0042-0000 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support that was ongoing for 3 days when the team observed access site bleeding. the bleed was at the impella, ecmo (extracorporeal membrane oxygenation), and arterial line. the team treated by the placement of the femostop and gave 1 unit rbc (red blood cell), 1 platelet, and 1 cryo. the event did not prompt explant. the pump supported for over 17 days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and limited information is provided in the clinical description.
30796 | 1220648-2024-06999 | 18764352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764352&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/22/2024 |  |  | 02/07/2024 | 02/22/2024 |  |  | 05/31/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354516 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 57-year-old male patient in st-elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient support had no dopplerable pulses on the impella side and faint on the left leg. the impella cp was removed, and the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into limb ischemia ¿ reversible issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30797 | 1220648-2024-07000 | 18764402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764402&pc=OZD | 11/09/2023 | 03/12/2024 |  | 02/22/2024 |  |  | 01/23/2024 | 02/22/2024 |  |  | 09/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395923 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 56 year old male patient presenting with cardiomyopathy for impella support. during support, the patient endured a stroke with a "probable" relationship to the impella device, followed by ventricular tachycardia (vt) with a "possible" relationship to the impella device. a thrombectomy was performed in response to the stroke, and a cardioversion was performed in response to the vt that occurred on the following day. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke, ventricular tachycardia, and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke, ventricular tachycardia, and thrombus could not be determined.
30798 | 1220648-2024-07001 | 18764476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764476&pc=OZD | 11/11/2023 | 03/13/2024 |  | 02/22/2024 |  |  | 01/23/2024 | 02/22/2024 |  |  | 06/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024368900 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 67 year old male patient presenting with cardiomyopathy for impella support. during support, the patient that had endured ventricular tachycardia with a "possible" relationship to the impella device. intervention was required via cardioversion. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the ventricular tachycardia (vt) issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt was not determined.
30799 | 1220648-2024-07002 | 18764538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764538&pc=OZD | 11/14/2023 | 03/12/2024 |  | 02/22/2024 |  |  | 01/23/2024 | 02/22/2024 |  |  | 09/30/2025 |  | 02/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Ischemia Stroke | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry, pertaining to a 72 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient had endured vt, acute limb ischemia, and stroke alleged as a result of using the impella device. a single shock was delivered in response to vt. two days later, the patient developed ischemia resulting from thrombus in the shared femoral artery that prompted an embolectomy. lastly, the patient developed signs of an ischemic stroke another 2 days later, via ct scan. there was no record of intervention for the stroke. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke, limb ischemia, and ventricular tachycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke, limb ischemia, and ventricular tachycardia could not be determined. d4-updated model number.
30800 | 1220648-2024-06936 | 18764542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18764542&pc=OZD | 02/08/2024 | 03/07/2024 |  | 02/22/2024 |  |  | 02/08/2024 | 02/22/2024 |  |  |  |  | 02/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1324936 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported that an automated impella controller (aic) failed to recognize an installed purge cassette. the aic was replaced to resolve the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag.
30801 | 1220648-2024-06932 | 18746502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18746502&pc=OZD | 02/06/2024 | 03/28/2024 |  | 02/21/2024 |  |  | 02/06/2024 | 02/21/2024 |  |  |  |  | 02/21/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1140351 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Charging Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) was not charging when it was plugged in. the orange/green light emitting diode (led) above the selector knob did not illuminate and the charging indicator did not show the unit as plugged in. the aic was replaced successfully. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - battery charging/other issues has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. it was discovered that the battery was not charging, and both 2a fuses were blown. the root cause of the battery not charging was a power supply malfunction.
30802 | 1220648-2024-06933 | 18746549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18746549&pc=OZD | 02/08/2024 | 03/15/2024 |  | 02/21/2024 |  |  | 02/08/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432800 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 23-year-old female patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp pump experienced an acute stoppage during patient support. it was noted that the motor current had spiked immediately prior to the stoppage. attempts to restart the pump were unsuccessful and the decision was made to replace the device. there was no patient harm. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned and tested. product evaluation confirmed a large purge gap consistent with bearing wear. data log analysis confirmed a motor current spike with no other abnormalities. the cause of the pump stop was bearing wear.
30803 | 1220648-2024-06934 | 18746698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18746698&pc=OZD | 02/08/2024 | 03/25/2024 |  | 02/21/2024 |  |  | 02/08/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED |  |  | US |  | 2024430176 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Unknown |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 67-year-old patient of an unknown gender post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that impella was unable to be implanted due to the patient¿s anatomy. the case was aborted and there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the clinical details states that the patient had tortuosity vessels condition. the cause of the delivery issue was patient condition related due to tortuosity vessels condition. h6 impact code updated to 4627, h6 problem code updated to 2682.
30804 | 1220648-2024-06935 | 18746743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18746743&pc=OZD | 02/08/2024 | 05/15/2024 |  | 02/21/2024 |  |  | 02/08/2024 | 02/21/2024 |  |  |  |  | 02/21/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1253203 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported that an unexpected automated impella controller (aic) shutdown alarm came on during patient support. the aic was exchanged. the was no patient harm reported. | [5 empty] | The investigation for the reported console (aic) shutdown has been completed. the console has been returned for evaluation. the console was tested the reported failure mode could not be reproduced naturally upon running the pump for 16 hours. no problem was found with the flash cards. upon force shutdown of the io-graphics via telnet command ¿slay -s sigabrt io-graphics,¿ the console was shut down and rebooted automatically, as happened at the customer site. data logs were reviewed finding the console received the log ¿systemidle:process io-graphics 339995 died abnormally.¿ after that, the console was shut down unexpectedly in 18 seconds and rebooted automatically in 22 seconds. during bootup following, the log indicated ¿io-graphics.core¿ file has been deleted - ¿impellalib: found core dump file /userflash/io-graphics.core with size : 23281664 and last modified on: 2024-02-08, 10:14:18 pm.¿ following that, the console recorded an error in establishing communication with ¿impellaconnect_bridge,¿ as the software is designed to turn off the ¿impellaconnect_bridge¿ communication after an unexpected console shutdown. then, the console received the unexpected controller shutdown ¿ alarm event set #603, which confirmed the reported failure mode. the root cause of intermittent unexpected controller shutdown was most likely the 3rd party hardware/software that runs io-graphics. in accordance with updated procedures, section d2 has been revised from the initial submission. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
30805 | 1220648-2024-06937 | 18746888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18746888&pc=OZD | 01/31/2024 | 04/11/2024 |  | 02/21/2024 |  |  | 02/09/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024425442 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 62-year-old female patient, part of the luerlock broke off due to user error leading to a purge filter bypass to be installed. there was no reported patient harm, and the pump was successfully weaned off the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported purge leak has been completed. no product was returned for evaluation. clinical details noted that the luerlock was broken due to user error and purge filter bypass was performed. patient was supported well on the pump and bypass did not affect pump performance. the root cause of the purge leak - pump was most likely due to the damaged luerlock on the sidearm as it was resolved with a purge sidearm bypass.
30806 | 1220648-2024-06938 | 18748127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748127&pc=OZD | 08/30/2021 | 09/12/2024 |  | 02/21/2024 |  |  | 01/23/2024 | 02/21/2024 |  |  | 09/30/2021 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 - JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444438 |  | 05066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 22 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/23/2024 |  |  |  |  | An abstract was obtained from the academic society. the date of occurrence is the date estimated from the abstract content. the user facility reported a 22-year-old female with acute myeloid leukemia with into remission and with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was covid positive. around 10 days after impella 5.0 removal, a right subdural hematoma developed and a craniotomy was performed to remove the hematoma. it is unknown if the patient's subdural hematoma was caused by the impella. additionally, there was bleeding noted and a blood transfusion was performed. furthermore, the patient had thrombocytopenia while on impella support. it is unknown if the impella was related to the thrombocytopenia. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The following have been reported incorrectly or missing from manufacturer device report (mdr) 1220648-2024-06938. d6a and d6b revised from the initial submission in accordance with updated procedures f6/f8-should have been left blank. h6 code 1888, 4431, and 4643 were reported incorrectly. new code was added to health effect - impact code. h10 investigation results on supplemental mdr included thrombocytopenia, stroke, and access site bleeding, which was reported in error. investigation was for stroke only. | The investigation into the stroke/cva, the access site bleeding ¿ major and the thrombocytopenia issues has been completed since the initial report was submitted. the product was not returned for investigation. as the clinical information was not provided, the root cause of the stroke/cva could not be determined. the root cause of the access site bleeding and the thrombocytopenia issues was not determined since product was not returned and not much information is provided.
30807 | 1220648-2024-06939 | 18748194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748194&pc=OZD | 01/25/2024 | 03/21/2024 |  | 02/21/2024 |  |  | 01/30/2024 | 02/21/2024 |  |  |  |  | 02/21/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1361559 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Shut Off | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp explant procedure of 74-year-old female patient, the automated impella controller (aic) was impossible to shut down after disconnecting pump. the shut down screen was frozen for 37 minutes and there was continuous beeping. there were no alarms visible on the screen. there was no reported patient harm. | [5 empty] | The investigation for the reported console (aic) shutdown/restart issue has been completed. the aic was returned for evaluation. during functional testing, the issue was not reproduced - there were no unusual delays after shutdown, and no failed shutdowns or boot-ups. the data logs reviewed from the reported day of event show ¿unexpected controller shutdown¿ alarm upon boot-up, and the number of log entries after first shutdown command was issued was uncharacteristically small. this confirms a failed shutdown. amount of time that passed between last entry after shutdown, and first entry upon next boot-up, is consistent with the reported delay in the shutdown process. it is likely that the issue was caused by pbm not withdrawing power from aic components. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
30808 | 1220648-2024-06940 | 18748410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748410&pc=OZD | 02/08/2024 | 08/18/2025 |  | 02/21/2024 |  |  | 02/08/2024 | 02/21/2024 |  |  |  |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363622 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The complainant reported that during preventative maintenance, an aic (automated impella controller) failed self check due to a battery comm. failure. there was no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d2 common device name revised on manufacturer device report in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number. e2 health professional revised from the initial submission in accordance with updated procedures. e3 occupation revised from the initial submission in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 investigation conclusions revised from the initial submission in accordance with updated procedures. h10 additional manufacture narrative revised from the initial submission in accordance with updated procedures.
30809 | 1220648-2024-06941 | 18748448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748448&pc=OZD | 01/26/2024 | 03/27/2024 |  | 02/21/2024 |  |  | 01/26/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428247 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 implant procedure for patient of unknown age and gender, the purge arm was damaged by the surgeon. the pump was explanted on the same day. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the purge leak ¿ pump issue has been completed since the initial report was submitted. the product was not returned for investigation. based on the limited clinical information provided, the root cause of the issue was a leak from the sidearm. the surgeon mentioned that the sidearm was damaged during insertion.
30810 | 1220648-2024-06942 | 18748458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748458&pc=OZD | 02/08/2024 | 03/18/2024 |  | 02/21/2024 |  |  | 02/08/2024 | 02/21/2024 |  |  | 05/31/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024354630 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 76-year-old female patient noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump was unable to cross the 23mm tissue valve during attempted delivery. it was noted that while advancing the pump, the wire began to buckle near an existing aneurysm. an aortic dissection was subsequently discovered, and the decision was made to abort the implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury | The investigation for the reported great vessel damage, product damage, and delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. a preexisting aneurysm located in the root of the aorta prevented the pump from passing and caused the guidewire to buckle. the most likely cause of the great vessel vascular damage was the patient's condition aneurism in the aorta. additionally, the most likely cause of the delivery issues was the patient's condition aneurism in the aorta resulting in resistance at that location.
30811 | 1220648-2024-06943 | 18748493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748493&pc=PYX | 02/08/2024 | 11/10/2025 |  | 02/21/2024 |  |  | 02/08/2024 | 02/21/2024 |  |  | 07/31/2024 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2023276906 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | Difficult to Advance |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/08/2024 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella rp support due to having reduced right ventricle function. while on support, the patient developed a deep vein thrombus distal to the rp access site. heparin was not started until 16 hours after the impella implant. heparin was provided and support continued. additionally, the pump was found to be too deep, however the physician did not want to reposition because the pump had good flows. furthermore, the placement signal was no longer reading reliable. support was continued. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the optical signal issue, the thrombus and the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. per the data log review, the root cause of the optical signal issue was determined to be most likely sensor silicone wear. the root cause of the positioning issue was use related to improper technique since the pump was placed deep per the provided information. per the data log investigation, the root cause of the low pump flow issue was biomaterial due to improper anticoagulation and maintaining low act's. | A5 ethnicity updated. b1 updated with product problem. d4 model number updated. d6a/d6b updated with additional information. d9 device returned to manufacturer date added. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. f6 and f8 should have been left blank. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2024-06943 h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-06943.
30812 | 1220648-2024-06945 | 18748611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748611&pc=OZD | 09/22/2022 | 05/22/2025 |  | 02/21/2024 |  |  | 09/22/2022 | 02/21/2024 |  |  | 04/30/2024 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227007 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 79-year-old male patient for mechanical circulatory support. it was reported that the patient was placed on impella cp support after being turned down surgically and the impella was placed without difficulty. the impella created an occlusive flow with the 14 french peel away sheath before and after being peeled away. the impella was removed. there was no patient harm. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30813 | 1220648-2024-06944 | 18748618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748618&pc=OZD | 02/08/2024 | 03/28/2024 |  | 02/21/2024 |  |  | 02/09/2024 | 02/21/2024 |  |  | 06/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024427863 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 39-year-old male patient of unknown ethnicity and race with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that the pump was explanted due to the patient developing hemolysis. the patient was reported as stable. | [5 empty] | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. based on the clinical information provided, the cause of the hemolysis issue most likely was improper positioning due to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30814 | 1220648-2024-06946 | 18748681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748681&pc=OZD | 02/08/2024 | 09/22/2024 |  | 02/21/2024 |  |  | 02/08/2024 | 02/21/2024 |  |  | 09/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024407084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Failure to Pump | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/08/2024 |  |  |  |  | The user facility reported an 83-year-old male patient in acute myocardial infarction/cardiogenic shock and the impella cp was attempted to be used as a bail out. it was reported that the was not done priming as the doctor continued to place impella into ventricle. the doctor succeeded, but the pump tried to come on and immediately stopped alarm appeared approximately three times. the pump was explanted. | Additional information noted that the procedure outcome was that patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr1220648-2024-06946 was provided in sections b2, b5, d6, h1, and h6. | The investigation into the pump did not start issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the pump unable to start cannot be determined without pump being returned for analysis.
30815 | 1220648-2024-06947 | 18748711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748711&pc=OZD | 02/09/2024 | 04/19/2024 |  | 02/21/2024 |  |  | 02/09/2024 | 02/21/2024 |  |  |  |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1239161 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support that was successful for over 2 weeks despite aic (automated impella controller) alarms. the aic buzzer malfunction did not cause harm or injury. | [5 empty] | The investigation into the aic - buzzer/ alarm malfunction issue has been completed since the initial report was submitted. the console was returned and tested. the reported failures were unable to reproduce upon initial bootup in the lab; however, there was an intermittent and faint static-like audio observed from the speaker. the root cause of the issue was not determined. the console was not making an old telephone ringing audio; however, a faint static background sound would intermittently be observed. the cause of the aic issue was a defective impellatronics board. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30816 | 1220648-2024-06948 | 18748778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748778&pc=OZD | 02/09/2024 | 03/28/2024 |  | 02/21/2024 |  |  | 02/09/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 68-year-old male patient noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that following start up, the pump stopped working. the pump was able to start back up after two restart attempts. the scheduled procedure was completed without further issue. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the initial report was submitted. the device was returned, visually inspected, and tested. biomaterial was observed in the purge gap, and a kink was found on the catheter near the motor. the cause of the pump stop issue was biomaterial based on return product and data log review.
30817 | 1220648-2024-06949 | 18748839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748839&pc=OZD | 02/09/2024 | 03/28/2024 |  | 02/21/2024 |  |  | 02/09/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427498 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient experienced hemolysis. the physician reported the impella might be too mobile causing hemolysis. the pump was repositioned and one unit of packed red blood cells was given. the patient is too unstable to exchange the impella. the plan is to discuss options with family members. continue support until kidney/liver recovery. | [5 empty] | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. based on the clinical information provided, the cause of the hemolysis issue most likely was improper positioning due to improper technique. the cause of the positioning issue was patient movement which was solved on its own. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30818 | 1220648-2024-06950 | 18748872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748872&pc=OZD | 02/10/2024 | 03/29/2024 |  | 02/21/2024 |  |  | 02/10/2024 | 02/21/2024 |  |  | 09/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024408886 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Difficult or Delayed Positioning | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2024 |  |  |  |  | The user facility reported a 22-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp insertion the cp kinked in the aorta. a new impella was used successfully. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through due to patient¿s small vasculature as per the clinical description provided.
30819 | 1220648-2024-06951 | 18748908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748908&pc=OZD | 02/11/2024 | 03/29/2024 |  | 02/21/2024 |  |  | 02/11/2024 | 02/21/2024 |  |  | 07/31/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432456 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2024 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was experiencing bleeding around the repo sheath of the impella. forward tension suture and angle matching attempted, but the patient was still bleeding. manual compression applied. after 30 min of manual pressure, still bleeding so figure 8 suture added at impella access site, followed by fem stop. antixa came back greater than 2, so protamine 25mg was given. the bleeding then improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleed was most likely due to anticoagulation management, as the bleeding improved when the anticoagulation was reduced.
30820 | 1220648-2024-06952 | 18748986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18748986&pc=OZD | 02/12/2024 | 03/29/2024 |  | 02/21/2024 |  |  | 02/12/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430172 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump stopped functioning during patient support. a controller failure alarm coincided with the stop and the decision was made to explant the pump. upon explant, calcium buildup was noted in the device. there was no patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. clinical information states calcium was observed on the pump after implant. however, no product was returned for evaluation. data logs show that the pump was weaned and turned to p-0 and removed for 50 minutes. the pump was then reinserted and ramped up over 8 minutes to p-8 but then had alarm 13 high motor current pump stop. the pump was unable to be restarted. the root cause of the pump stop was most likely biomaterial as seen by calcium observed on the pump upon explant. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
30821 | 1220648-2024-06953 | 18749275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18749275&pc=OZD | 02/09/2024 | 03/28/2024 |  | 02/21/2024 |  |  | 02/09/2024 | 02/21/2024 |  |  | 10/31/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024415339 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 62-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that following pump placement bleeding was noted at the insertion site, after the peel- away sheath was removed and repositioning sheath was inserted the site continued to bleed. the bleeding was managed with a fem-stop over the site, but the patient had no flow down the leg and no pulse in the foot. even with the fem-stop removed the patient continued to have no flow down the leg past the sheath. the decision was made to go to the operating room and do a surgical cutdown and remove the impella and repair the artery. once the artery was repaired the pulse in the foot returned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ischemia: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed since the initial report was submitted. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding and the limb ischemia was not determined.
30822 | 1220648-2024-06954 | 18749353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18749353&pc=OZD | 02/10/2024 | 03/21/2024 |  | 02/21/2024 |  |  | 02/10/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024428626 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2024 |  |  |  |  | The user facility reported a 64-year-old male cancer patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed heavy bleeding from their access site during impella cp support. upon noting the bleed, it was discovered that the repositioning sheath had been pulled back. the repositioning sheath was reinserted, and the site was redressed and sutured to stop the bleed. to counteract the blood loss, ten units of red blood cells, seven units of fresh frozen plasma, one unit of platelets, and one unit of cryo were transfused. support was maintained following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding - major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely a use issue, as it was reported that it was difficult to reposition and was left in place. severe bleeding from the access site occurred when the repo sheath was drawn back, and hemostasis was obtained by reinserting the repo, re-suturing the site, and redressing.
30823 | 1220648-2024-06955 | 18749916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18749916&pc=OZD | 02/10/2024 | 03/28/2024 |  | 02/21/2024 |  |  | 02/10/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426113 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2024 |  |  |  |  | The user facility reported a 78-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, an impella in aorta alarm occurred. the patient was taken back to the operation room for impella replacement. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the initial report was submitted. the product was not returned for investigation. as per clinical, pump was pulled out due to aorta alarm. the cause of the positioning issue was use issue related due to improper technique.
30824 | 1220648-2024-06956 | 18750082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18750082&pc=PYX | 02/11/2024 | 03/28/2024 |  | 02/21/2024 |  |  | 02/11/2024 | 02/21/2024 |  |  | 08/31/2024 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023258328 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2024 |  |  |  |  | The user facility reported a 41-year-old male with pulmonary hypertension was implanted with an impella rp device for mechanical circulatory support. the rp was supporting but on 2nd day of support the team observed hemolysis with hematuria. the md did not infuse back blood nor start dialysis. the pump later stopped but did restart. the pump remains on for support, now on day three. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the pump stop and the hemolysis issues has been completed since the initial report was submitted. the product was not returned for investigation. based on the clinical information provided, the root cause of the temporary pump stop and the hemolysis was most likely biomaterial.
30825 | 1220648-2024-06957 | 18750208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18750208&pc=OZD | 02/12/2024 | 03/29/2024 |  | 02/21/2024 |  |  | 02/12/2024 | 02/21/2024 |  |  | 11/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434927 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed to support a stemi patient. the position was lost and the team was unable to re-deliver the cp. the team explanted and replaced the pump, despite lack of injury or harm from interruption in the support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was due to a patient condition, stenosis vessels condition. d4-updated model number h6-impact code updated to 4629.
30826 | 1220648-2024-06958 | 18750302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18750302&pc=OZD | 09/23/2022 | 04/16/2025 |  | 02/21/2024 |  |  | 09/23/2022 | 02/21/2024 |  |  | 06/30/2024 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237498 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 71-year-old female patient for mechanical circulatory support. it was reported that there was oozing at the insertion site. an extra suture and angle matching dressing was placed at site. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30827 | 1220648-2024-06960 | 18752945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18752945&pc=OZD | 01/26/2024 | 03/19/2024 |  | 02/21/2024 |  |  | 01/26/2024 | 02/21/2024 |  |  | 08/31/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389813 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella cp sensor was damaged by the surgeon leading to the pump being explanted. a new impella cp was placed to provide support for the patient of unknown age and gender. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical issue has been completed. the device was not returned for evaluation. data logs were reviewed and all waveforms were within normal range. although the site reported abnormal ps waveforms, there was no cause found for the optical signal since the device functioned/operated to specification.
30828 | 1220648-2024-06961 | 18753390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18753390&pc=OZD | 02/03/2024 | 10/30/2025 |  | 02/21/2024 |  |  | 02/03/2024 | 02/21/2024 |  |  | 09/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024400989 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant reported the patient was on impella 5.5 support and while on support, there were bleeding concerns and systematic anticoagulants were paused. it is unknown if the bleeding was impella related. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleed was due to patient condition related.
30829 | 1220648-2024-06962 | 18753656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18753656&pc=OZD | 02/08/2024 | 03/26/2024 |  | 02/21/2024 |  |  | 02/08/2024 | 02/21/2024 |  |  | 10/31/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024425267 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support due to having cardiomyopathy. while on support, there was high purge pressure and purge system blocked alarms. tissue plasminogen activator (tpa) was started and the alarms resolved. tpa had to be discontinued. there were suction issues later while on support. two units of packed red blood cells and albumin was given to the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure and the low pump flow issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the high purge pressure issue and the low pump flow was gradual biomaterial buildup due to no heparin in the purge.
30830 | 1220648-2024-06963 | 18753805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18753805&pc=OZD | 09/24/2022 | 04/16/2025 |  | 02/21/2024 |  |  | 09/24/2022 | 02/21/2024 |  |  | 06/30/2024 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235215 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2022 |  |  |  |  | The user facility reported a 69-year-old male patient experienced slight oozing at the groin site after being placed on impella cp for mechanical circulatory support. all systemic heparin was removed. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30831 | 1220648-2024-06964 | 18753831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18753831&pc=OZD | 09/25/2022 | 03/05/2024 |  | 02/21/2024 |  |  | 09/25/2022 | 02/21/2024 |  |  | 06/30/2024 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247609 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2022 |  |  |  |  | The user facility reported an impella cp in a 45-year-old male patient for mechanical circulatory support. it was reported while being placed on impella cp support, the patient developed ischemic leg on impella side. the impella cp pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30832 | 1220648-2024-06965 | 18753917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18753917&pc=OZD | 09/25/2022 | 03/26/2024 |  | 02/21/2024 |  |  | 09/25/2022 | 02/21/2024 |  |  | 06/30/2024 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2022 |  |  |  |  | The user facility reported a 61-year-old female implanted with an impella cp device for mechanical circulatory support, the patient¿s right foot was not moving. nivus scan indicated abnormal monophasic flow in the superficial femoral artery (sfa) and profunda. an external femoral bypass was performed. the leg had a strong popliteal pulse. the pulse in the foot could not be found, but the whole leg was warm to the touch. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the limb ischemia was completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. h6 type of investigation code 4111 was inadvertently not selected on the previous manufacture device report and was added to this report.
30833 | 1220648-2024-06966 | 18754676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18754676&pc=OZD | 11/22/2023 | 03/07/2024 |  | 02/21/2024 |  |  | 01/22/2024 | 02/21/2024 |  |  | 09/30/2025 |  | 02/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402437 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/22/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 69 year old patient, presenting with acute myocardial infarction and cardiogenic shock, had endured acute kidney injury on chronic kidney disease with a "possible" relationship to the impella device: as a result, hemodialysis was given and a bumex drip was initiated. | Upon review of the clinical information received on this complaint event, it appears the patient endured an impella related access site bleed, post procedure. application of 2 tourniquets was successful in hemostasis. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported access site bleed and renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and renal failure could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Investigation still in progress. a supplemental will be filed at the completion of our analysis.
30834 | 18739318 | 18739318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18739318&pc=OZD | 11/15/2023 | 01/02/2024 | 01/02/2024 | 02/20/2024 | 01/02/2024 | 02/20/2024 |  | 02/20/2024 |  |  |  |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0550-0008 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Yes |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 60 YR |  | Female | 86 KG |  | White | Injury | Yes | Yes | [7 empty] | Arrhythmia | Perforation | [11 empty] | Required Intervention | Hospitalization | [23 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | This impella was originally placed early in november. days after, patient had coronary artery bypass graft x 3. during surgery, the impella was turned off and clamped; but unfortunately the impella had perforated the wall of the left ventricle. the impella had to be repositioned in the aortic arch to allow for repair of the ventricle. then, the impella was again repositioned down to the left ventricle, at 4 cm, and it was working well. because of the patient having significant arrhythmias and intermittent malfunctioning of the impella with inadequate drainage of the heart, marginal cardiac outputs, and ventricular arrhythmias, the patient was taken to or for a repositioning of the impella. they used te echo and fluoroscopy, and "attempted to pull back the device and replace it and rotate it so that it would be in a more favorable position. we were having a hard time getting it to advance, however, because the steal rail for the catheter had become broken. we eventually decided we could not get the catheter positioned adequately and decided to remove the existing 5.5 impella." a new impella device was placed. a nurse described the issue with the impella in this way: "there was a kink in the catheter that prevented the impella to be repositioned."
30835 | 1220648-2024-06887 | 18740663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740663&pc=PYX | 01/22/2024 | 03/12/2024 |  | 02/20/2024 |  |  | 01/22/2024 | 02/20/2024 |  |  | 01/31/2024 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023175896 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 65-year-old male patient was on impella rp support due to acute severe right ventricle failure. the impella rp was inserted and ramped up and the impella motor stopped. staff attempted to restart again and the motor stopped again. the impella was repositioned and restarted again, and it was slowly ramped up to p7 then there was a high purge pressure alarm, then alarmed purge blocked. the impella was removed and replaced with a different impella rp. when the second impella was ramped up, there was an immediate motor stop. staff replaced the aic successfully and the second impella that was used was operational. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop and the high purge pressure issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. biomaterial was found wrapped around purge gap and on the impeller blade. pump was tested and did not reproduce pump pause and high purge pressure alarm. the root cause of the pump pause is most likely due to biomaterial found wrapped on impeller. the root cause of the high purge pressure was most likely due to biomaterial wrapped around the purge gap.
30836 | 1220648-2024-06888 | 18740839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740839&pc=OZD | 01/30/2024 | 03/28/2024 |  | 02/20/2024 |  |  | 01/31/2024 | 02/20/2024 |  |  | 06/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024381582 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support of a cgs/ami and covid+ patient. the impella cp was placed when the ccl showed cad and undergoing pci. the cp accessed limb developed ischemia. the team placed bypass to restore flow to the limb. the pulses were worsening and toe color was poor. the pump was not removed and support continues to date. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30837 | 1220648-2024-06889 | 18740860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740860&pc=OZD | 01/31/2024 | 03/21/2024 |  | 02/20/2024 |  |  | 01/31/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024420848 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported patient of unknown age and gender with cardiogenic shock was put on impella cp for mechanical circulatory support following heart attack. however, pump repositioning after x-ray check, led to pump flow dropping to 0.6l. decision was made to explant the pump and implant new impella as the patient had good pressure with catecholamines on the artery. there were no issues reported with the second pump, but patient expired because of cardiogenic shock. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the reported positioning issue has been completed. the impella device has been returned for evaluation. data logs were not returned for review. however, according to clinical details, after the pump was repositioned, the flows immediately dropped. during evaluation of the pump, visual inspection found biomaterial inside the pump housing wrapped around the impeller. therefore, the root cause of the low pump flow was most likely due to biomaterial ingestion. the root cause of the positioning issue could not be determined as sufficient clinical details were not provided.
30838 | 1220648-2024-06890 | 18740876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740876&pc=OZD | 02/01/2024 | 03/04/2024 |  | 02/20/2024 |  |  | 02/01/2024 | 02/20/2024 |  |  | 09/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404392 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for hemodynamic support of a 61-year-old female patient admitted in known critical condition with known carotid artery disease. the patient suffered a cerebral vascular accident (cva) with slurred speech while on the cp support. the stroke was not attributed to the impella, and the physician and extended team were aware the event was a possibility. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva issue has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined.
30839 | 1220648-2024-06891 | 18740884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740884&pc=OZD | 01/30/2024 | 03/13/2024 |  | 02/20/2024 |  |  | 01/30/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024424139 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 45-year-old male with heart failure was implanted with an impella 5.5 device for mechanical circulatory support. on day two of impella support, the patient had ventricular tachycardia (vt). the patient was shocked twice. in the middle of the night, the patient had three episodes of vt. the impella position was verified via echocardiogram. the vt has not repeated so the medical team is uncertain what caused it. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the h the purge cassette failure and the ventricular tachycardia (vt) issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the priming issue was not determined since product and data logs were not returned. as the necessary information was not provided, the root cause of the vt/vf was not determined.
30840 | 1220648-2024-06892 | 18740894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740894&pc=OZD | 02/02/2024 | 04/09/2025 |  | 02/20/2024 |  |  | 02/02/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412708 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 53-year-old patient was scheduled to receive hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). when preparing the impella heart pump, the user noticed that the 0.018" guide wire from the impella kit had a damaged tip. the impella cp heart pump was easily inserted using a different guidewire. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation of product damage (guidewire damage - peripheral has been completed. according to the clinical, before beginning impella cp insertion the guidewire was removed from its sterile packaging and a broken tip was identified. the guidewire was replaced and the no other issues occurred. product evaluation confirmed coil unraveling with no other abnormalities. data log analysis was not required for this complaint. the cause of the out of box guidewire damage is unknown and the product will be returned to the vendor for further analysis. d1 brand name was erroneously entered on the initial manufacturer device report number 1220648-2024-06892. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-06892. d6a and d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2024-06892. g1 reporting contact email revised on manufacturer device report 1220648-2024-06892 in accordance with updated procedures. g1 manufacturing site name, address country, and fax number was erroneously entered on the initial manufacturer device report number 1220648-2024-06892.
30841 | 1220648-2024-06893 | 18740955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740955&pc=OZD | 02/02/2024 | 03/21/2024 |  | 02/20/2024 |  |  | 02/02/2024 | 02/20/2024 |  |  | 09/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410771 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for support of a 72-year-old male patient admitted in sob and nstemi. the impella cp was placed but, along with pressors the patient had signs of limb ischemia. the ischemia was seen by cold limb. the team did not treat the ischemia as condition was poor and care was being withdrawn. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30842 | 1220648-2024-06894 | 18740971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740971&pc=OZD | 02/02/2024 | 03/22/2024 |  | 02/20/2024 |  |  | 02/02/2024 | 02/20/2024 |  |  | 05/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349465 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for support of a 61-year-old male patient admitted with cgs/ami. the pump was supporting the patient and allowed for transfer to tertiary medical center. at the transfer a groin site ooze occurred which caused the team to intervene with 1 unit rbcs and a change of dressing. the pump remains on for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since best practices have clearly not been followed as angle matching was not performed on catheter.
30843 | 1220648-2024-06895 | 18740988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740988&pc=OZD | 02/03/2024 | 03/25/2024 |  | 02/20/2024 |  |  | 02/03/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418832 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump would not advance through the patient's vasculature during attempted delivery. due to the noted resistance, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. h6: impact code: 4627.
30844 | 1220648-2024-06896 | 18740997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740997&pc=OZD | 02/05/2024 | 04/05/2024 |  | 02/20/2024 |  |  | 02/05/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024416589 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for 65-year-old male patient, there was low pump flow <1ml/h and high purge pressure 1060mmhg. there was no reported patient harm and the patient was successfully weaned and recovered well. | [5 empty] | The investigation into the low pump flow and the high purge pressure issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. there was biomaterial found on the impellers. the catheter was also bent near the motor housing end. pump was then run to recreate the failure and there was constant high purge pressure alarm with pressure. the root cause of the high purge pressure issue was biomaterial found on the impeller blades during product investigation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
30845 | 1220648-2024-06897 | 18740999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18740999&pc=OZD | 01/31/2024 | 09/20/2024 |  | 02/20/2024 |  |  | 01/31/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432717 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 58-year-old male noted as high-risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient experienced multiple pump complications during impella cp support. six days into support, it was discovered that the filter on the purge sidearm was cracked and leaking. the staff was walked through the purge bypass process, but an hour later the pump began to experience saturated flows. while on the phone with the staff to discuss troubleshooting, the pump stopped. the pump was restarted temporarily, but stopped again shortly after. the pump was subsequently explanted as a result of the failures. there was no patient harm. | It was noted that the patient expired after explant of the device. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06897 was provided in sections b1, b2, b5, d6, h1, and h6. | The investigation into the purge leak ¿ pump, pump stop, and the saturated flow issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the pump stop was due to the damaged purge filter which led to blood ingress into the purge line and motor, high purge pressure, and the high motor current pump stop.
30846 | 1220648-2024-06898 | 18741007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741007&pc=OZD | 01/30/2024 | 03/27/2024 |  | 02/20/2024 |  |  | 01/30/2024 | 02/20/2024 |  |  |  |  | 02/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1329519 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a controller error alarm. the aic was switched and the new aic had no issues. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - controller error (yellow alarm) issue has been completed since the initial report was submitted. the console was returned and tested. the failure was unable to be reproduced throughout testing. the root cause of the issue could not be determined due to the difficulty in reproducing the failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm with a low probability of reoccurrence.
30847 | 1220648-2024-06899 | 18741029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741029&pc=OZD | 02/07/2023 | 02/20/2024 |  | 02/20/2024 |  |  | 02/07/2023 | 02/20/2024 |  |  |  |  | 02/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1131454 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a console malfunction. the buzzer alarm malfunctioned and sound was not present. there was no patient harm or injury . the ifu was followed and the backup aic used. | [5 empty] | The investigation into the reported automated impella controller (aic) issue has been completed. the suspected console (b)(6) was received for evaluation under mdr 23-13671-1-1 on 08mar2023. additionally, the console logs from the reported day of event were analyzed and were noted to be consistent with the complaint. the logs show staff verifying alarm audio settings, making sure it was enabled. when console was brought for testing, the issue was reproduced, and no audible alarms were present. inspection of console¿s electronics revealed that a cable connecting speakers to the 4-in-1 carrier board was unplugged. after reconnecting the cable - issue was resolved, and alarms were audible. it is most likely the cable was accidentally disconnected during prior pm/fc/repair work on this console. the root cause of the reported event was determined to be a loose connector. the failure modes will be monitored and trended.
30848 | 1220648-2024-06900 | 18741156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741156&pc=OZD | 01/31/2024 | 04/03/2024 |  | 02/20/2024 |  |  | 01/31/2024 | 02/20/2024 |  |  | 09/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024399750 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 72-year-old male patient was to receive hemodynamic support from an impella 5.5 smartassist heart pump as bridge to a durable lvad. immediately after insertion of the impella device, an increased purge pressure and decreased purge flow has been noted. as this could have resulted in a pump stop, the decision has been made to exchange the impella 5.5 heart pump for a new one as a precaution. the patient suffered no harm from the incident. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was returned, visually inspected, and tested. data log analysis confirmed a gradual rise in purge pressure and fall in purge flow immediately after beginning support, which suggests biomaterial occluding the purge gap. the root cause of the high purge pressure could not be reproduced but was most likely due to biomaterial.
30849 | 1220648-2024-06901 | 18741232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741232&pc=OZD | 01/30/2024 | 04/03/2024 |  | 02/20/2024 |  |  | 02/01/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024428551 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 79-year-old male patient was scheduled to undergo coronary artery bypass graft (cabg) under impella support. the impella cp was inserted in the catheterization room, and the patient was transferred to the operating room. echocardiography was performed in the operating room and found a large amount of mitral regurgitation (mr). the impella position was confirmed but it was found that it had been inserted in the direction of the mitral valve. it was determined that the mr had increased due to the problem with the impella position. attempts were made to adjust the impella's position, but it either pulled out or entered the direction of the m valve, making it difficult to adjust. impella was removed and cabg continued. | [5 empty] | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. the cause of the positioning issue was due to patient condition related with pump unable to reposition due to mitral regurgitation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30850 | 1220648-2024-06902 | 18741257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741257&pc=OZD | 01/29/2024 | 04/03/2024 |  | 02/20/2024 |  |  | 01/31/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024433888 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 28 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 28-year-old female patient noted with heart failure/cardiogenic shock and cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that multiple complications were encountered during support. roughly five days into support, bleeding was noted from the graft anastomosis site. hemostatic surgery was performed to manage the condition and an unspecified amount of blood was transfused. intravenous heparin was also stopped. this resulted in an increase in purge pressure. when the pressure reached over 1000 mmhg, the decision was made to increase the heparin concentration in the purge. heparin levels were still held at a lower than desired concentration due to the bleeding issues experienced by the patient. a short time later, the pump stopped and the decision was made to remove the device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the access site bleeding ¿ major, the high purge pressure, and the pump stop issue has been completed since the initial report was submitted. because the product was not returned and there was insufficient clinical data, the root cause of the access site bleeding issue was not determined. data logs showed many high purge pressures and purge system blocked alarms. based on an analysis of the data log, biomaterial was most likely the root cause for the pump stop issue.
30851 | 1220648-2024-06903 | 18741274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741274&pc=OZD | 02/01/2024 | 03/18/2024 |  | 02/20/2024 |  |  | 02/01/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435076 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 40-year-old male patient had poor perfusion to their lower left extremity following impella cp implant. a distal perfusion sheath was subsequently added to the femoral artery to manage the condition. support with the implanted pump was maintained following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the device was not returned. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30852 | 1220648-2024-06904 | 18741380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741380&pc=OZD | 01/26/2024 | 03/19/2024 |  | 02/20/2024 |  |  | 02/02/2024 | 02/20/2024 |  |  |  |  | 02/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1462815 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during prep, the automated impella controller (aic) displayed yellow alarm purge fluid not detected. the console was therefore exchanged. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported console (aic) purge issues has been completed. the aic was returned for evaluation. upon functional testing of the aic, the piston was blocked/dirty and the purge cassette was unable to deair. data logs were reviewed which also confirmed that de-airing step could not be completed at the time of piston blocked log entries. therefore, the root cause of the purge issues was likely due to a block piston cause by glucose ingress. added-h6 impact code 4629.
30853 | 1220648-2024-06905 | 18741426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741426&pc=OZD | 01/14/2024 | 09/20/2024 |  | 02/20/2024 |  |  | 01/14/2024 | 02/20/2024 |  |  | 09/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401525 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Material Separation | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 60-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that when the physician was repositioning the impella, while taking the dressing down the integrity of the sterile sleeve was compromised. the team was advised to use betadine and a sterile dressing. no patient harm was reported. | It was noted that the patient expired after explant of the device. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B2: date of patient death unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06905 was provided in sections b1, b2, b5, d6, h1, and h6. | The investigation for the reported product damage has been completed. the impella device nor sufficient clinical details were provided for evaluation. therefore, the root cause of the reported issue could not be determined.
30854 | 1220648-2024-06906 | 18741570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741570&pc=OZD | 02/02/2024 | 03/21/2024 |  | 02/20/2024 |  |  | 02/02/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427514 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple unsuccessful attempts were made to place impella 5.5, but the case was aborted due to patient¿s vessel tortuosity. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the clinical details states that the patient had calcified and tortuosity vessels condition. the cause of the delivery issue was patient condition related due to calcified and tortuosity vessels condition. added- h6 problem code 2682 h6-updated clinical code 4627.
30855 | 1220648-2024-06907 | 18741600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741600&pc=OZD | 08/20/2023 | 03/27/2024 |  | 02/20/2024 |  |  | 01/31/2024 | 02/20/2024 |  |  | 12/31/2024 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024311755 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella support for a male patient of unknown age and gender, there was blood leakage from red plug. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported product damage has been completed. the impella device was not returned for evaluation. clinical details provided insufficient to determine the root cause. therefore, without the device or clinical details, the root cause of the product damage could not be determined. h.10 the following instructions for use were reported on the mdr and should not have been. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
30856 | 1220648-2024-06908 | 18741647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741647&pc=OZD | 02/03/2024 | 03/21/2024 |  | 02/20/2024 |  |  | 02/03/2024 | 02/20/2024 |  |  | 09/30/2024 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275095 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 66-year-old male patient was on impella cp support due to going asystolic. while on support the patient's urine was dark red/brown in color. labs came back hemolyzed. intervention for the hemolysis is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined.
30857 | 1220648-2024-06909 | 18741735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741735&pc=OZD | 02/06/2024 | 05/02/2024 |  | 02/20/2024 |  |  | 02/06/2024 | 02/20/2024 |  |  |  |  | 02/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | US |  | 1115051 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during prep, the automated impella controller (aic) displayed 'purge system stopped' controller failure alarm leading to console exchange. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the issue was reproduced and the same purge-related alarm was observed upon each boot-up of the console. replacing purge pressure sensor assembly resolved the issue. the data logs from the reported day of event are consistent with complaint and show controller failure alarm presenting upon console boot-up. this was preceded by an error message ¿purgeadmin: purge sensor_defect2 set¿ indicating purge pressure sensor voltage being out of range. the root cause of this issue was a broken purge retainer. added-device return date g1-corrected mfr site name and address
30858 | 1220648-2024-06910 | 18741768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741768&pc=OZD | 02/02/2024 | 04/03/2024 |  | 02/20/2024 |  |  | 02/05/2024 | 02/20/2024 |  |  |  |  | 02/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1250838 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Mechanical Jam | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complaint has contacted cc because the flap behind which the purge cassette is located can no longer be opened. the switch for opening the flap is not working. it is also stuck in the housing. | [5 empty] | The investigation into the aic - purge issue ¿ other has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure mode was confirmed. dextrose deposit was found inside the aic, which was also most likely deposited on the purge door hinge, preventing the opening of the purge cassette assembly door. the root cause of the stuck 'purge cassette assembly' door was most likely the dextrose deposit on the purge door hinge. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
30859 | 1220648-2024-06911 | 18741845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741845&pc=OZD | 02/06/2024 | 03/04/2024 |  | 02/20/2024 |  |  | 02/06/2024 | 02/20/2024 |  |  |  |  | 02/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1154174 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during prep, the automated impella controller (aic) displayed 'purge system stopped' controller failure alarm leading to console exchange. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge issue ¿ other has been completed since the original report was filed. the console was returned and tested after arriving in fs. console inspection revealed significant contamination of the purge insert area, resulting in purge motor shaft being immobilized. after cleaning the purge assembly, the issue was resolved, and console was operating within specification. the cause of the aic issue was contamination.
30860 | 1220648-2024-06912 | 18741891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741891&pc=OZD | 02/07/2024 | 04/09/2024 |  | 02/20/2024 |  |  | 02/07/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024426835 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient of unknown age and gender in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complaint reported that during impella cp support the cp purge sidearm was leaking and broken. there was no patient harm. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found the sidearm joint reservoir to the filter being broken. white discoloration and crack/damage to the sidearm joint was observed due to likely excessive force during support. the cause of the purge leak- pump was determined to be damage at the joint reservoir to filter of the sidearm due to excessive force during support. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
30861 | 1220648-2024-06913 | 18741897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741897&pc=OZD | 02/06/2024 | 04/05/2024 |  | 02/20/2024 |  |  | 02/07/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024433746 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | Company representation |  |  | 77 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | As reported by the protect iv clinical study, a 77-year-old female patient on impella cp pump support had a limb ischemia. a percutaneous transluminal angioplasty (pta) was performed and the event was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30862 | 1220648-2024-06914 | 18741925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741925&pc=OZD | 01/30/2024 | 03/12/2024 |  | 02/20/2024 |  |  | 01/30/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421986 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 76-year-old male patient had a hematoma at their groin during impella cp support. the patient was given one unit of packed red blood cells and one unit of fresh-frozen plasma. a sandbag was placed also to manage the condition. support was continued with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding - major was not determined since product was not returned and there is a lack of clinical details.
30863 | 1220648-2024-06915 | 18741948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18741948&pc=OZD | 02/08/2024 | 03/05/2024 |  | 02/20/2024 |  |  | 02/08/2024 | 02/20/2024 |  |  |  |  | 02/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1314859 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | No patient involvement: us complainant reported the automated impella controller (aic) was powered on and staff did not see a blue cloud light after powering the device on. a variety of troubleshooting was performed unsuccessfully. the aic would not enter maintenance mode or airplane mode. staff tried to factory reset, but continued to see a blinking amber light. device was turned on and off, but was unable to get the device into maintenance mode. the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplementa mdr will be filed. | The investigation for the controller "power on" issues has been completed. the controller and the data logs were returned for evaluation. inspection and testing of the unit revealed remote lan module (rlm) inability to properly boot-up was due to corruption of its microsd (som) card as evidenced by partition switching and an endless boot-up loop. after temporarily replacing microsd card with a known-good one, the unit was able to boot up, connect to a test network and transfer data, while the connection remained stable over time. there were no deficiencies of backstrap cable or aic components discovered.the data logs were reviewed are are consistent with complaint and show persistently failing handshakes (attempts to establish communication between aic and rlm). therefore the root cause of the reported power on issue was due to corrupted microsd in the rlm. however, the cause of the corrupted microsd card could not be determined.
30864 | 1220648-2024-06917 | 18742133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18742133&pc=OZD | 02/08/2024 | 03/25/2024 |  | 02/20/2024 |  |  | 02/08/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438269 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Hispanic | White | Malfunction | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Cardiac Perforation | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported an 85-year-old female patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during manual manipulation of heart, the impella cp pigtail perforated a highly ischemic area of tissue of the lateral wall of the ventricle. after repair of ventricle the physician wanted to shallow out impella. under trans-esophageal echocardiogram (tee) guidance the impella was pulled back however it was pulled back across aortic valve, requiring pump to be replaced. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve. | The investigation into the positioning issues and the ventricle perforation issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the positioning issue was most likely use related as the pump was pulled too far back into the aortic valve requiring explant. the root cause of the ventricle perforation was most likely use related as the issue occurred during manual manipulation of the heart.
30865 | 1220648-2024-06918 | 18742215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18742215&pc=OZD | 02/08/2024 | 03/25/2024 |  | 02/20/2024 |  |  | 02/08/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024424196 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 69-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient experienced a cold leg that started to mottle. the physician did not want to remove the peel away sheath. echocardiogram was done at bedside and pulled back 1cm and the support was continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30866 | 1220648-2024-06919 | 18742278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18742278&pc=OZD | 02/09/2024 | 03/28/2024 |  | 02/20/2024 |  |  | 02/09/2024 | 02/20/2024 |  |  | 09/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395927 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed bleeding and a hematoma at their access site during impella 5.5 support. the patient was taken to the or where an exploration of the hematoma revealed two bleeds that were subsequently repaired. support was maintained with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleed was unable to be determined as insufficient clinical information and product was not returned for the investigation.
30867 | 1220648-2024-06920 | 18742633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18742633&pc=OZD | 02/01/2024 | 03/18/2024 |  | 02/20/2024 |  |  | 02/01/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432244 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support in a 69-year-old male patient admitted with cardiac history and acute myocardial infarction/cardiogenic shock. the pump was placed and support proceeded but, when in the intensive care unit, the patients condition deteriorated. the team determined there was a bleed that was treated with massive transfusion and the vessel perforation was repaired. the team then placed a new pump via the contralateral leg when the left femoral was repaired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported vascular damage has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information regarding the event has been provided. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30868 | 1220648-2024-06921 | 18742665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18742665&pc=OZD | 02/08/2024 | 04/11/2024 |  | 02/20/2024 |  |  | 02/08/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025451868 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported it was impossible to reposition the pump. it was not possible to move the catheter against the repo sheath. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the introducer damage ¿ mechanical has been completed since the original report was submitted. the device was not returned for investigation. it was reported that pump was unable to reposition with the catheter, impossible to slide against the repo sheath. the cause of the positioning issue was not determined since product was not returned and due to insufficient clinical information.
30869 | 1220648-2024-06922 | 18742858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18742858&pc=OZD | 02/07/2024 | 09/25/2025 |  | 02/20/2024 |  |  |  | 02/20/2024 |  |  | 09/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402269 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 76-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the the patient had no radial pulse distal to the impella site. arterial ultrasound ordered to check flow down right arm. a thrombus was observed in the patient. heparin was started the night prior. | The complainant also reported that the patient was on impella cp support when there was a placement signal low alarm. the position was verified, and the impella was 3.5 cm pointing towards the apex. placement showed great position. the placement signal alarm was then disabled. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the optical signal issue, the limb ischemia reversible and the thrombus issues has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the optical signal issue was not determined since product was not returned and due to inconclusive failure trend per data log review. as all the necessary information was not provided, the root cause of the limb ischemia and the thrombus was not determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. data logs were reviewed, and placement signal (ps) low alarm observed at start of case after single occurrence of impella in ventricle alarm. several ps low alarms were seen with no consistent placement signal deteriorating trend. placement monitoring was disabled. 5s parameters had no hard failure but snr and contrast were observed to decrease and later increase back to normal range along with change in p-level. also, no major variable parameters were seen for lamp, gain and light. the cause of the optical signal issue was not determined since product was not returned and due to inconclusive failure trend per data log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. b7 updated. d10 concomitant medical products and therapy dates. f6 and f8 erroneously filled out in the initial report.
30870 | 1220648-2024-06924 | 18743164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743164&pc=OZD | 02/08/2024 | 09/20/2024 |  | 02/20/2024 |  |  | 02/08/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418164 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 63-year-old female patient with anterior wall st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was in icu blood was noted to be leaking from the red impella plug. the physician decided to wait and see how hemoglobin reacted to bleeding and would remove the pump in the morning. no patient harm was reported. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06924 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the product damage (hole in catheter) issue has been completed since the original report was filed. the cause of the hole in the catheter could not be determined because product was not returned.
30871 | 1220648-2024-06923 | 18743179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743179&pc=PYX | 02/08/2024 | 03/20/2024 |  | 02/20/2024 |  |  | 02/08/2024 | 02/20/2024 |  |  | 02/28/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024346412 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during impella insertion when the physician was advancing the repo sheath, the impella stopped. this was due to high motor current without change in pulmonary artery placement signal. the impella was successfully restarted at p2 level after repo sheath was in place and ramped up to p6 with appropriate flow and motor current. there was no patient harm, | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the reported pump stop has been completed. the pump was returned for evaluation finding the pump was cut along catheter and the patient cable/red handle side was not returned. stretching of the catheter was observed on the pump cannula but was not related to pump stop failure. no damage to inflow/outflow cage was observed, and no blood or biomaterial was found in purge gap. data logs were reviewed which confirmed a impella stopped motor current high. the pump was restarted manually and was able to run successfully until the end of the case. the exact root cause of the pump stop could not be determined. added- d9 device return date.
30872 | 1220648-2024-06925 | 18743233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743233&pc=OZD | 02/01/2024 | 03/20/2024 |  | 02/20/2024 |  |  |  | 02/20/2024 |  |  |  |  | 02/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1691229 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for a male patient on unknown age and gender, the automated impella controller (aic) displayed battery failure red alarm leading to console exchange. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported console (aic) battery failure issue has been completed. the aic was returned for evaluation. functional testing was performed on the aic and there no issues found with the console hardware. data logs were reviewed which showed a controller error alarm that aligns with a transient loss of optical data. therefore the root cause of the issue was due to an aic software issue. h6-corrected impact code to 4629.
30873 | 1220648-2024-06926 | 18743272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743272&pc=OZD | 02/08/2024 | 03/22/2024 |  | 02/20/2024 |  |  |  | 02/20/2024 |  |  | 06/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431018 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported an 89-year-old female patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after pump placement, the patient became restless causing access site bleeding; pressure was applied, and (b)(4)unit of packed red blood cells was given as treatment. after 11 hours of support the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- h6 impact codes 4643, 4641.
30874 | 1220648-2024-06927 | 18743345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743345&pc=OZD | 02/09/2024 | 03/22/2024 |  | 02/20/2024 |  |  | 02/09/2024 | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422337 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the peel away sheath was removed, the impella came back into the aorta and inverted itself. it was unable to be repositioned and the impella was removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related.
30875 | 1220648-2024-06928 | 18743429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743429&pc=OZD | 02/02/2024 | 03/22/2024 |  | 02/20/2024 |  |  | 02/02/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428640 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella implant was attempted to be implanted in an 67-year-old male patient. the physician was unable to pass through the patient's vascular. the impella 5.5 implant was aborted and an impella cp was implanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because it was mentioned that the patient had heavily calcified aortic arch which caused the difficulty of inserting the pump.
30876 | 1220648-2024-06929 | 18743592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743592&pc=OZD | 02/02/2024 | 03/21/2024 |  | 02/20/2024 |  |  |  | 02/20/2024 |  |  | 10/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424172 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant placed an impella cp to support a 81-year-old male patient who was admitted with cardiogenic shock and needed care escalated from an intra-aortic balloon pump. the impella cp was repositioned due to concerns/signs of hemolysis. after five hours, the team and family made the choice to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the hemolysis issue was not determined since product was not returned and due to insufficient clinical information.
30877 | 1220648-2024-06930 | 18743771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743771&pc=OZD | 10/26/2023 | 02/20/2024 |  | 02/20/2024 |  |  |  | 02/20/2024 |  |  | 05/31/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024360296 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A female patient of unknown age received hemodynamic support from an impella cp smartassist heart pump and an extracorporeal membrane oxygenation (ecmo) due to de-novo cardiomyopathy. bleeding at the femoral impella cp access site was reported. the bleeding was managed by correcting the repositioning sheath angle and renewing the sutures securing the repositioning sheath and manual pressure. it has been reported that the activated clotting time was higher than recommended. | [5 empty] | The investigation into the reported "access site bleeding" has been completed. no product was returned. the evaluation revealed that the activated clotting time (act) values during the bleeding event was 158 seconds per clinical. all other information was unknown. the cause of the access site bleeding issue was not determined since the product was not returned and due to insufficient clinical information. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
30878 | 1220648-2024-06931 | 18743873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18743873&pc=OZD | 02/05/2024 | 03/21/2024 |  | 02/20/2024 |  |  | 02/05/2024 | 02/20/2024 |  |  | 11/30/2025 |  | 02/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431368 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 67-year-old male patient was attempting to have an impella 5.5 implant and the catheter was unable to pass through the brachial cephalic. the impella 5.5 was removed and the catheter kinked. the team did not feel the pump was able to be reused due to the kink. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. clinical details states that the catheter kinked during insertion and was not used. therefore, the root cause of the delivery issue was user-related.
30879 | 1220648-2024-06886 | 18731155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18731155&pc=OZD | 09/22/2022 | 04/16/2025 |  | 02/19/2024 |  |  | 09/22/2022 | 02/19/2024 |  |  | 06/30/2024 |  | 02/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247282 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 52-year-old male patient for mechanical circulatory support. it was reported while on impella cp support, the patient was experiencing a hematoma at the access site. pressure was held to the site and heparin was stopped. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury". | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30880 | 1220648-2024-06882 | 18728258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18728258&pc=OZD | 12/13/2023 | 02/26/2024 |  | 02/17/2024 |  |  | 01/19/2024 | 02/17/2024 |  |  | 08/31/2025 |  | 02/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024392543 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 49 year old male patient presenting with cardiomyopathy for impella support had endured pericardial effusion with a possible relationship to this device. a pericardiocentesis was administered, as a result | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the reported ventricle perforation has been completed since the original report was filed. according to the clinical, 14 days after explant a pericardial effusion with a possible relationship to the impella device occurred. the timing of this injury could not be verified. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the cause of the ventricle perforation was not determined because product was not returned and not enough information was given.
30881 | 1220648-2024-06883 | 18728259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18728259&pc=OZD | 07/08/2023 | 03/08/2024 |  | 02/17/2024 |  |  | 01/19/2024 | 02/17/2024 |  |  | 11/30/2024 |  | 02/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023282417 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/19/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 60 year old male patient presenting with acute myocardial infarction and cardiogenic shock endured ventricular tachycardia with a "possible" relationship to this impella device. | An additional notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured "thrombus at origin of subclavian artery (limb ischemia)" with a "possible" relationship to this impella device: "action taken: low-dose heparin drip then eliquis." |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported limb ischemia and ventricular tachycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and ventricular tachycardia could not be determined. | B(5): additional information added to narrative. g(4): date of awareness added for new information. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30882 | 1220648-2024-06884 | 18728263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18728263&pc=OZD | 06/28/2023 | 05/01/2024 |  | 02/17/2024 |  |  | 01/19/2024 | 02/17/2024 |  |  | 05/31/2025 |  | 02/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352563 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Respiratory Failure | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 01/19/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 50 year old male patient presenting with cardiomyopathy to have endured atrial arrhythmia with a "possible" relationship to the impella device. a direct cardioversion was performed. | Additional information was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured acute kidney injury and respiratory failure with a "possible" relationship to the impella device used. the patient was intubated and ventilator settings were adjusted. the patient's condition improved by discharge. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported atrial arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the atrial arrhythmia could not be determined. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. the investigation has been reopened and, upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the renal failure could not be determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information.
30883 | 1220648-2024-06885 | 18728265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18728265&pc=PYX | 12/07/2023 | 03/15/2024 |  | 02/17/2024 |  |  | 01/19/2024 | 02/17/2024 |  |  | 10/31/2025 |  | 02/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024418597 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 65 year old impella patient to have endured a cva with a "possible" relationship to the impella device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the stroke/cva issue has been completed since the original report was filed. the device was not returned for investigation. as no clinical information was provided, the root cause of the stroke/cva could not be determined.
30884 | 1220648-2024-06848 | 18721802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18721802&pc=OZD | 01/30/2024 | 04/02/2024 |  | 02/16/2024 |  |  | 01/30/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423580 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an implanted impella cp pump kinked during implant as a result of the 77-year-old male patient's calcified aorta and small arch. the pump experienced multiple suction alarms due to the kink and the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation into the low pump flow has been completed since the original report was filed. the pump was returned, tested, and evaluated. low pump flow. visual inspection found a kink on the cannula. the root cause of low flow was a cannula kink that occurred during delivery through a small calcified aortic arch. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | H.6 health effect - impact code 2199 was inadvertently reported on the original manufacturer device report (mdr). the device was replaced and this code has been added to h.6 in this supplemental mdr.
30885 | 1220648-2024-06849 | 18721977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18721977&pc=OZD | 02/02/2024 | 03/07/2024 |  | 02/16/2024 |  |  | 02/02/2024 | 02/16/2024 |  |  |  |  | 02/16/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1152334 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/03/2024 |  |  |  |  | It was reported that the automated impella controller (aic) had purge disc not detected alarms. the purge cassette and tubing were replaced and this did not resolve the alarms. the aic was replaced successfully. no patient was involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. after analysis, reported issue with the aic not being able to detect the purge disc was confirmed. purge disc not detected alarms were determined to be caused by a broken purge flag. the cause of the issue was a broken purge flag.
30886 | 1220648-2024-06850 | 18722210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722210&pc=OZD | 02/06/2024 | 03/21/2024 |  | 02/16/2024 |  |  | 02/06/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430100 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. this patient was admitted for cardiogenic shock and is on impella 5.5 support and was awaiting a transplant. after multiple weeks of support, the team observed the need to have the access site hematoma washed out in the operating room. the pump remained on for support while awaiting a transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30887 | 1220648-2024-06851 | 18722308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722308&pc=OZD | 01/30/2024 | 09/20/2024 |  | 02/16/2024 |  |  | 01/30/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423069 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the 68-years-old male patient was on impella cp support due to having st elevated myocardial infarction. after the impella implant, left ventricle wave forms and flows dropped. there was concern for clot ingestion despite heparin administration. the impella was taken out of auto and there were suction alarms. after 0.75 hours of support, the patient expired. there were no allegations made towards the impella for cause of death. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06851 was provided in sections b1, b2, b5, d6, h1, and h6. | The investigation into the low pump flow issue has been completed since the original report was filed. the device was not returned for investigation. per the data logs, placement signal was trending downwards along with suction alarms and few positioning issue alarms after low flow issue was seen. the cause of the low pump flow issue most likely was patient condition related per data log review.
30888 | 1220648-2024-06852 | 18722340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722340&pc=OZD | 01/31/2024 | 03/19/2024 |  | 02/16/2024 |  |  | 01/31/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435920 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported a 21-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient had ecmo placed along with/prior to the cp. the echo imaging revealed thrombus ingestion. the patient had conversations of tpa to troubleshoot the issue. the patient additionally noted the need for bilateral fasciotomy. the native anatomy and patient condition not appropriate for escalation to the 5.5 pump. support continues with the cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ irreversible issue has been completed since the original report was filed. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30889 | 1220648-2024-06853 | 18722359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722359&pc=OZD | 02/06/2024 | 03/25/2024 |  | 02/16/2024 |  |  | 02/06/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429244 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp discontinued therapy at p2 level with patient hemodynamics at stable condition, but the impella had a clot visualized under echo imaging. the flow rate dropped quickly from 3 l/min down to 1 l/min and suctions alarms. the physicians believe that it was just a cascade of coagulopathy events that lead to clotting. upon removal of the impella no visual clot was seen on the device from the outer parts. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus issues has been completed. the pump was not returned for investigation. data logs show pump flow dropping to 1 liter/min while running at p-level p7 with an active suction alarm. according to the clinical details, the doctor found a small clot on the pump inlet during continuous suction alarm and a drop in pump flow. the cause of the low pump flow issue was most likely biomaterial, based on given clinical details. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ it also states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30890 | 1220648-2024-06854 | 18722406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722406&pc=OZD | 02/04/2024 | 03/26/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408787 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 23-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the pump triggered a placement signal not reliable alarm. upon inspection of the device, it was noted that the catheter had kinked proximal to the red impella plug to the point and that the fiber optic light was visible. the kink was stabilized by applying tape and the staff was advised to verify the positioning on the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the optical signal issue was most likely damaged optical fiber line per clinical and data log review.
30891 | 1220648-2024-06855 | 18722560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722560&pc=OZD | 02/07/2024 | 09/20/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024429815 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 78-year-old male post cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during planned implant of an impella 5.5 pump, the utilized guidewire became stuck within the patient. the anastomosed 10mm graft was cut to remove the coiled wire and the pump was subsequently placed without the guidewire in place. there was no patient harm. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (guidewire damage - peripheral vessels) has been completed since the initial report was submitted. the root cause of the delivery issue was most likely use related resulting in guidewire damage. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06855 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
30892 | 1220648-2024-06856 | 18722577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722577&pc=OZD | 01/30/2024 | 03/05/2024 |  | 02/16/2024 |  |  | 01/30/2024 | 02/16/2024 |  |  | 09/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024405174 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump stopped working during patient support. attempts to restart the pump were unsuccessful and the device was subsequently replaced for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. data logs were reviewed which confirmed pump was unable to be run and received alarm 13 impella stopped motor current high with a motor current spike. however, without the device nor sufficient clinical information for evaluation, the exact root cause of the pump stop could not be determined.
30893 | 1220648-2024-06857 | 18722625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722625&pc=OZD | 02/07/2024 | 04/09/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024412535 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | Material Integrity Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during implant procedure of impella cp for 71-year-old male patient, the physician suspected the easy guide lumen broke but stayed in the outlet cage before the pump was placed. therefore, another pump used and implanted without issues. there was no reported patient harm | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the introducer damage ¿ mechanical issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. according to the clinical, before beginning impella cp insertion the doctor suspected that the red lumen had broken and remained within the outflow cage. product evaluation confirmed a scratch on the epotek of the outflow cage indicating that there was some resistance while removing the red lumen. the most likely cause of the delivery issue was use related because the physician suspected that the red lumen broke, but never confirmed their suspicions.
30894 | 1220648-2024-06858 | 18722745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722745&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024421568 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 55-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the sterile sleeve broke away from the hub while the impella catheter was being cleaned. the plastic was reinserted and there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (sterile sleeve damage) has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the product damage sterile sleeve damage was not determined since product was not returned.
30895 | 1220648-2024-06859 | 18722811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722811&pc=PYX | 02/07/2024 | 03/25/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024435274 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 36-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the impella rp encountered multiple occurrences of suction over multiple days leading to low flows. the patient developed hemolysis and high lactate levels. the pump was ultimately removed, and the patient was put on dialysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the hemolysis and low pump flow issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. the returned complaint device was visually inspected. biomaterial observed in purge gap and around the impeller. the root cause of the hemolysis and the low pump flow issues was biomaterial ingestion.
30896 | 1220648-2024-06860 | 18722817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722817&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432840 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 37-year-old female patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that when the peel away sheath was being removed and the pump repositioning sheath was placed, the team lost left ventricle (lv) position. the pump was unable to be repositioned and the pump was removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related.
30897 | 1220648-2024-06861 | 18722827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722827&pc=OZD | 02/05/2024 | 03/20/2024 |  | 02/16/2024 |  |  | 02/05/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433012 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 81 year-old female noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp case was aborted due to the patient¿s vascular anatomy. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since patient had calcification.
30898 | 1220648-2024-06862 | 18722848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722848&pc=OZD | 02/08/2024 | 03/25/2024 |  | 02/16/2024 |  |  | 02/08/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC, |  |  | US |  | 2024431330 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 52-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. a trans-esophageal echo was performed, and a large clot was noted on the impella cp in the descending aorta. the doctor decided to not remove the cp out of risk of limb loss. patient was taken for immediate ct scan; distal limb flow was decreased, and the physician was unsure if distal aorta was occluded or into iliacs. the impella cp was explanted and patient was escalated to the imeplla 5.5 | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus and the limb ischemia ¿ reversible issues has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus and the limb ischemia was not determined.
30899 | 1220648-2024-06863 | 18722971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18722971&pc=PYX | 02/06/2024 | 09/06/2024 |  | 02/16/2024 |  |  | 02/06/2024 | 02/16/2024 |  |  | 02/28/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024337523 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemolytic Anemia | Lactate Dehydrogenase Increased | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/06/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient was on impella rp support post left ventricle assistance device. the patient's lactate dehydrogenase (ldh) increased to above 900 after the rp was placed. the impella was explanted. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured mechanical hemolytic anemia with a "probable" relationship to the impella device used: due to rp impella bilis elevated, potentially worsening secondary to hemolysis supportive care. the patient recovered with no sequelae. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Section b5, g2 and h6 were updated as per additional information received. | The investigation for the hemolysis/hemolytic anemia has been completed. data logs were reviewed which showed pump ran without issue for entire case. there were some suction alarms in early of the case but were resolved quickly. no positioning alarms were triggered during support. product was not returned for review. the root cause of hemolysis/hemolytic anemia was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use for the impella rp smart assist system with automated impella controller states the follwowing: section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ added- h6 clinical code 1886, h6 component code 925 d4-corrected model number f6/f8 should have been left blank in the initial report -1220648-2024-06863. | The investigation into the hemolysis issue has been completed since the original report was filed. the product was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
30900 | 1220648-2024-06864 | 18723019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18723019&pc=OZD | 02/09/2024 | 03/27/2024 |  | 02/16/2024 |  |  | 02/09/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430171 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | Positioning Problem | [5 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 68-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the pump was in the mitral valve and flows were low. the impella was repositioned. the following day the impella was attempted to be repositioned and placed at appropriate depth on transesophageal echocardiography. the impella was on p2 for repositioning and was put to p6 after repositioning when a motor current high, impella stopped alarm occurred. the level went from p0 to p3 and the impella had normal metrics and flows of 4lpm, so it was decided to leave it on this setting as the impella did not do well on a higher level. shortly after, the patient¿s mean arterial pressure dropped from 75 to 43, then 33. patient experienced syncopal. the patient coded and then was cannulated on va ecmo. during cannulation, the impella was pulled back into axillary artery and placed in surgical mode by perfusion team and later removed. the working theory of code is that inlet tickled the left ventricle wall and caused slow ventricular tachycardia /ventricular fibrillation picture that was masked by the patient¿s automatic implantable cardioverter defibrillator pacer spikes, making the first sign of code, sudden map drop, and syncope. the team does not think the device is at fault but wants impella to be evaluated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the positioning issues, the pump stop, and the ventricular tachycardia/ventricular fibrillation (vt/vf) issue has been completed since the initial report was submitted. the product was returned and investigated. the cause of the positioning issue was use related due to improper technique per clinical review. the cause of the temporary pump stop issue was due to biomaterial per field investigation and clinical review. as the necessary information was not provided, the root cause of the vt/vf was not determined.
30901 | 1220648-2024-06865 | 18723078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18723078&pc=OZD | 02/07/2024 | 04/09/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420870 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 80-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after an hour of impella cp support for 80-year-old female patient, the pump stopped. several attempts were made to restart but were unsuccessful. therefore, the pump was exchanged. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. data logs show that the pump ran for about 33 minutes before having a high motor current pump stop alarm 13. there is a gradual motor current increase prior to the pump stop indicating a gradually increasing resistance on the rotor assembly. upon trying to restart the pump, there are high purge pressure and purge system blocked alarms. the pump was unable to be run in the lab. high purge pressure (hpp) was produced in the lab. the purge line was clear of blood ingress. upon teardown of the pump, buildup was confirmed within the motor. the root cause of the pump stop issue was determined to be a build up of biomaterial in the motor. added- d9 device return date d4-updated model number.
30902 | 1220648-2024-06866 | 18723094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18723094&pc=OZD | 02/12/2024 | 04/26/2024 |  | 02/16/2024 |  |  | 02/13/2024 | 02/16/2024 |  |  | 08/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024391384 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The complainant in japan reported that preliminary ct evaluation showed that the right subclavian artery was small. a left subclavian approach was performed for insertion of impella 5.5, but the catheter could not be advanced to the central side of the artificial blood vessel anastomosis. insertion of 5.5 was not possible so impella cp was placed successfully. no patient harm reported due to the issue. | [5 empty] | The investigation into the delivery issues and the vascular damage - peripheral vessel issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issues was most likely patient condition related due to the patient's small and tortuous vasculature. the root cause of the vascular damage - peripheral vessel was most likely patient condition related due to the patient's small and tortuous vasculature. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30903 | 1220648-2024-06867 | 18723201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18723201&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024422091 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 81-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the echo showed impella cp against the papillary wall and off axis from the apex of the left ventricle. multiple attempts were made to position the pump correctly, finally stopping at 3.87 cm from ao. the physician felt that if the pump was pulled back further there was a risk of crossing over into the aorta. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the positioning issue was patient condition related, specifically the patient's very large papillary muscle.
30904 | 1220648-2024-06868 | 18723266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18723266&pc=OZD | 02/07/2024 | 03/20/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 08/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024386596 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported an 82-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis, evidenced by hematuria, during impella cp support. the pump was initially repositioned, and volume was given to counteract the condition, but the decision was ultimately made to explant the pump the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis. | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed o pump in aorta/ pump in ventricle alarms were triggered. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position as the hemolysis resolved by repositioning.
30905 | 1220648-2024-06869 | 18723312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18723312&pc=OZD | 02/08/2024 | 09/22/2024 |  | 02/16/2024 |  |  | 02/08/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024419049 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient had an attempted impella cp implant due to needing emergent hemodynamic support. the impella cp had difficulty crossing due to tortuosity/ascending aneurysm calcification of aorta and illiacs. the implant was aborted. no patient harm due to the issue. | It was reported that the patient expired as all efforts to hemodynamically support the patient were considered futile. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06869 was provided in sections b2, b5, d6a, d6b, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the delivery issues has been completed since the initial report was submitted. the product was returned and investigated. no abnormalities were observed. the root cause of the delivery issue is determined to be the patient condition since clinical information indicates the patient's calcified anatomy condition.
30906 | 1220648-2024-06870 | 18723317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18723317&pc=OZD | 02/07/2024 | 03/18/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435017 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 81-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support when the pigtail of the catheter was found under the large papillary muscle. the doctor tried to pull the impella back and reposition, but they were unsuccessful. the impella cp was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. data logs showed pump ran without issue for entire case. there was no positioning alarm triggered. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
30907 | 1220648-2024-06871 | 18723372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18723372&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024426996 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 76-year-old female patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the delivery of the impella 5.5 was aborted due to interaction at the anastomosis. impella cp was delivered instead, and no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the clinical details states that the patient had anastamosis in the vessels. the cause of the delivery issue was patient condition related to anastamosis.
30908 | 1220648-2024-06873 | 18724077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18724077&pc=OZD | 02/09/2024 | 03/28/2024 |  | 02/16/2024 |  |  | 02/09/2024 | 02/16/2024 |  |  | 03/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343135 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a (b)(6) female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after over four hours of support, the pump was explanted and the team found the perclose failed to achieve the hemostasis of the impella site. the team also noted some ischemia to be occurring. the patient was sent to the or for cutdown and exploration of the site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed since the initial report was submitted. the product was not returned for investigation. basded on the clinical information provided, the root cause of access site bleeding was determined to be use related as the perclose failed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30909 | 1220648-2024-06872 | 18724100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18724100&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/16/2024 |  |  | 02/07/2024 | 02/16/2024 |  |  | 07/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025442317 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 67-year-old male patient was being escalated from the impella cp to the impella 5.5 due to limb ischemia. the impella cp was removed, and a fasciotomy and repair of groin site was performed due to the patient's pre-existing peripheral arterial disease (pad). the explant was successful with no issues. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
30910 | 1220648-2024-06874 | 18724424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18724424&pc=OZD | 02/10/2024 | 03/28/2024 |  | 02/16/2024 |  |  | 02/10/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423185 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2024 |  |  |  |  | The user facility reported a 38-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was not able to achieve flows greater than 1.0. it was believed that there was possible thrombus ingestion. the cp was explanted, and a new cp placed. the patient is stable. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella pump was returned and visually inspected. a cannula kink near outlet cage was observed. there was no biomaterial nor broken blades observed. data logs were reviewed and impella in aorta alarms with suction were observed during the entire case with low flow. as per clinical, upon starting impella cp, pump moved out into the aorta. it was unable to be re-advanced past the aortic valve. as a result, the cp pump was removed and reinserted. after reinsertion and starting up the pump in auto, spikes in motor current noted on aic and persistent suction alarms began. a new pump was placed and the pump ran without any issues. the cause of positioning issue is use related due to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30911 | 1220648-2024-06875 | 18725328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18725328&pc=OZD | 01/23/2024 | 03/12/2024 |  | 02/16/2024 |  |  | 01/23/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024423693 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Crack | Fluid/Blood Leak | Improper Flow or Infusion |  |  |  |  | Hemoptysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 46-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. multiple complications were encountered during impella 5.5 support. roughly ten days into support, it was documented that the purge cassette was leaking at the connection site. the purge cassette was replaced to resolve the issue and support was maintained. five days later, it was noted that the purge sidearm had a visible crack. tegarderm was utilized to manage the crack and the leak was stopped. seventeen days later, it was reported that the patient began coughing up blood from their lungs. heparin was put on hold as a result of the condition. four days later, a considerable decrease in purge flow was observed coinciding with high purge pressure alarms followed by purge blocked alarms. a tissue plasminogen activator protocol was initiated in an attempt to manage the condition; however, the decision was ultimately made to replace the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the purge leak ¿ cassette, the purge leak ¿ pump, the vascular damage - peripheral vessel, and the high purge pressure issues has been completed since the original report was filed. the device was not returned for investigation. per the provided clinical information, the root cause of the purge cassette leak issue was not determined since product was not returned for investigation. the root cause of the pump- purge leak issue was determined to be a damaged sidearm. in the data log review, it was found that there was a gradual increase in pure pressure suggesting biomaterial buildup. the root cause of the high purge pressure issue was most likely biomaterial. the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information.
30912 | 1220648-2024-06876 | 18725958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18725958&pc=OZD | 02/07/2024 | 04/10/2024 |  | 02/16/2024 |  |  | 02/08/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024413856 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. per the complainant, at the moment of the insertion of the impella 5.5 in the peel away sheet, the 500ml purge bag got rapidly empty without having leakages along the purge cassette. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. additionally, clinical details were not sufficient to determine the root cause. the root cause of the purge leak - pump issue was not determined since product was not returned and data logs were not returned as well.
30913 | 1220648-2024-06877 | 18726505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18726505&pc=OZD | 02/08/2024 | 09/17/2025 |  | 02/16/2024 |  |  | 02/08/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 54-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the impella cp pump stopped during patient support. leading up to the failure, it was noted that both high purge pressure alarms and placement signal not reliable alarms were triggered. the pump flow then became saturated and the pump subsequently stopped. the pump was removed and replaced for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the pump stop issue was not determined since product was not returned and data logs were missing around the time of failure. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. per the data log review, the 5s parameters are similar to seen in a kinked fiber line. with placement signal therefore, the root cause of the optical signal issue was most likely a damaged optical fiber line due to use. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d6a/d6b updated with additional information h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-06877.
30914 | 1220648-2024-06878 | 18726855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18726855&pc=OZD | 02/09/2024 | 03/28/2024 |  | 02/16/2024 |  |  | 02/09/2024 | 02/16/2024 |  |  | 11/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431372 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock (ami/cgs) was on impella 5.5 mechanical circulatory support. on day two of support, the axillary site was seen to be bleeding and was treated with two units of blood. the 5.5 pump remains on for support as the oncology patient in ami/csg awaits hemodynamic recovery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleed was unable to be determined as insufficient clinical information and product was not returned for the investigation.
30915 | 1220648-2024-06879 | 18727911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18727911&pc=OZD | 01/12/2024 | 08/21/2024 |  | 02/16/2024 |  |  | 01/18/2024 | 02/16/2024 |  |  | 04/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330700 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 01/18/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 57 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support via the impella 5.5 device. during insertion of the device, the patient endured "excessive bleeding". as a result,"4 units of packed red blood cells, 2 units of fresh frozen plasma, and 1 unit of platelets" were delivered. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured right axillary site bleeding with a "possible" relationship to the impella device used. blood transfusion was done to address the issue. the patient recovered with no sequelae. in accordance with updated procedures, section d6 has been revised from the initial submission. |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The axillary site bleed is the only adverse event associated with this device. the femoral bleed was captured under manufacturer device report number 1220648-2024-14445. | Added-b2 selected death and added date of patient death, b5 mso review, h6 component code 925, d6a/d6b added in 1220648-2024-06879 -02 but was inadvertently not mentioned in the h11 correction section. f6/f8 should have been left blank in the initial report 1220648-2024-06879. h1 changed type of reportable event to death. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use provides suggestions on management of access site bleeds. it states ¿assess access site for bleeding and hematoma.¿ d4: model number changed to impella 5.5 | B5 has been updated as per additional information received.
30916 | 1220648-2024-06880 | 18727942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18727942&pc=OZD | 10/19/2023 | 03/08/2024 |  | 02/16/2024 |  |  | 01/19/2024 | 02/16/2024 |  |  | 06/30/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024357380 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 64 year old male patient presented with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient was said to have endured an "acute embolic stroke" with a possible relationship to this impella device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva issue has been completed since the original report was filed. the device was not returned for investigation. as no clinical information was provided, the root cause of the stroke/cva could not be determined.
30917 | 1220648-2024-06881 | 18727958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18727958&pc=OZD | 10/12/2023 | 02/29/2024 |  | 02/16/2024 |  |  | 01/18/2024 | 02/16/2024 |  |  | 08/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024384517 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured a cardiac tamponade with a "possible" relationship to the impella device. intervention was required via aspiration. | Additional information provided states that the patient had ventricular tachycardia which was treated with ablation. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the ventricle perforation and ventricular tachycardia issues have been completed. the device was not returned for investigation. data logs do not show that pump was pushed too far in ventricle. the root cause of the ventricle perforation and ventricular tachycardia could not be determined due to limited clinical detail and no evidence of incorrect placement of pump in data logs.
30918 | 1220648-2024-06759 | 18727961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18727961&pc=OZD | 02/01/2024 | 03/11/2024 |  | 02/16/2024 |  |  | 02/01/2024 | 02/16/2024 |  |  | 10/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408597 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a 59-year-old male was on impella 5.5 support for acute myocardial infarction/cardiogenic shock (ami/cgs). the impella 5.5 support was successful for 11 days when saturated flows prompted pump removal and was replaced with an impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported saturated pump flow has been completed. the impella device was returned for evaluation. the product evaluation confirmed biomaterial on the impeller. data log analysis also confirmed positioning issues and a motor current spike before saturated flow. therefore, the cause of the saturated flow was biomaterial.
30919 | 1220648-2024-06758 | 18727962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18727962&pc=OZD | 02/05/2024 | 04/05/2024 |  | 02/16/2024 |  |  | 02/05/2024 | 02/16/2024 |  |  | 08/31/2025 |  | 02/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391590 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 87-year-old patient with acute myocardial infarction cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. after 72 hours on support, the "position in aorta alarm'' went off. an ultrasound examination revealed that the impella cp had pulled out into the aorta. the pump was explanted and not replaced as the hemodynamics were stable. | [5 empty] | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30920 | 1220648-2024-06813 | 18710195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18710195&pc=OZD | 02/03/2024 | 04/09/2024 |  | 02/15/2024 |  |  | 02/05/2024 | 02/15/2024 |  |  |  |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2025437135 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella 5.5 support, a position in aorta alarm was triggered. upon attempting to re-insert the pump, the artificial blood vessel began to protrude through the skin. the ligation between the artificial blood vessel and sheath remained intact with minimal bleeding. the pump was subsequently explanted. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. a small amount of access site bleeding was also reported. the root cause of access site bleeding was most likely use related as the bleeding was resolved after re-suturing the suture pad at access site.
30921 | 1220648-2024-06814 | 18710398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18710398&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/15/2024 |  |  | 02/07/2024 | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024434131 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported an impella 5.5 was attempted to be implanted due to the patient having chronic coronary artery disease. the implant was aborted due to the physician being unable to pass through the patient's anatomy. the implant was aborted and an impella cp was placed. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the initial report was submitted. the root cause of the delivery issue was most likely due to patient condition since pump met resistance at the subclavian artery and was hence aborted due to patient anatomy.
30922 | 1220648-2024-06815 | 18711085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711085&pc=OZD | 01/28/2024 | 03/18/2024 |  | 02/15/2024 |  |  | 01/28/2024 | 02/15/2024 |  |  | 01/31/2024 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181657 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2024 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient's leg became cold to the touch and was a dusky color during impella cp support. due to the noted ischemia, the patient was re-draped and a 6fr reperfusion sheath was placed on the right femoral artery. color and pulses returned to the leg following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. d4-updated model number
30923 | 1220648-2024-06816 | 18711140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711140&pc=OZD | 01/30/2024 | 03/07/2024 |  | 02/15/2024 |  |  | 01/30/2024 | 02/15/2024 |  |  |  |  | 02/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1338086 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) fell off a stretcher and onto the floor during a patient transfer. the patient was not connected to the aic at the time of the event. the console suffered physical damage in multiple locations. there was no patient harm and another aic was utilized for ongoing support. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic ¿ damage has been completed since the original report was filed. the console was returned and tested after arriving in fs. physical examination confirmed the reported complaint. damage to the console¿s rear housing was observed. the cause of the issue was use related.
30924 | 1220648-2024-06817 | 18711141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711141&pc=OZD | 04/19/2022 | 05/19/2025 |  | 02/15/2024 |  |  | 04/19/2022 | 02/15/2024 |  |  | 01/31/2024 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186994 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 62-year-old male patient on impella cp support experienced reduced or absent pulses in the left lower leg and below the knee. amputation was recommended. pulse was never returned for the patient. additionally, the patient developed a clot/thrombus which was detected via ultrasound. heparin was withheld and the impella cp was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30925 | 1220648-2024-06818 | 18711252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711252&pc=OZD | 01/29/2024 | 04/03/2024 |  | 02/15/2024 |  |  | 02/01/2024 | 02/15/2024 |  |  | 10/31/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024425770 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The complainant in japan reported that a patient of unknown age and gender was implanted with an impella cp for mechanical circulatory support. during support, the purge pressure increased, and the purge flow rate became less than 2ml/h. the team was considering replacing the pump, but the first step was to replace the purge set. even after replacing the purge set, the decrease in flow rate did not improve, and the pump was used without being replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the high purge pressure was most likely due to gradual buildup of biomaterial on the pump, per the clinical data provided.
30926 | 1220648-2024-06819 | 18711339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711339&pc=OZD | 02/01/2024 | 03/15/2024 |  | 02/15/2024 |  |  | 02/01/2024 | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435823 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a 64-year-old female undergoing mitra-clip cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the team noted there was a small clot in the left ventricle. the patient was being weaned to remove the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ it also states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30927 | 1220648-2024-06820 | 18711374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711374&pc=OZD | 02/02/2024 | 03/21/2024 |  | 02/15/2024 |  |  | 02/02/2024 | 02/15/2024 |  |  | 07/31/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386914 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | The user facility reported a 76-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient had an echocardiogram which revealed a thrombus in the left ventricle. the decision was made to remove the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury | The investigation into the thrombus issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
30928 | 1220648-2024-06821 | 18711494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711494&pc=OZD | 02/03/2024 | 02/15/2024 |  | 02/15/2024 |  |  | 02/03/2024 | 02/15/2024 |  |  |  |  | 02/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1578302 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/03/2024 |  |  |  |  | The complainant in japan reported that the automated impella controller (aic) had a system self check error alarm. the aic was exchanged. there was no patient involvement. | [5 empty] | The investigation into the self-check error on the automated impella controller (aic) has been completed. data analysis: imc logs confirm that there was loss of communication between the pbm and epc as evidenced by the powermod_1 communication failure error. i.e [(b)(6) 2024 13:58:00 sau_ser10 received duplicate command seq: 43540. send ack eok !]. it was also confirmed from the logs the system self check failure occurred after the multiple powermod_1 communication errors. i.e [(b)(6) 2024 13:58:00 sau_ser10 received duplicate command seq: 43540. send ack eok !]. device analysis: the failure mode was reproduceable upon bootup in the lab during investigation. visual inspection of the pbm reveled corrosion of the copper via in the vicinity of the super capacitors c69 and c70. conclusion: the root cause of the system self check failure was contamination of the pbm board from electrolyte leakage of super capacitors. electrolyte leakage corroded traces of the pbm leading to loss of communication between the epc and the pbm.
30929 | 1220648-2024-06822 | 18711524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711524&pc=PYX | 01/30/2024 | 03/20/2024 |  | 02/15/2024 |  |  | 01/30/2024 | 02/15/2024 |  |  | 01/31/2024 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023179101 |  | 004334 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp and an impella rp device for mechanical circulatory support. after just 11 hours the impella rp was explanted due to optical signal issues. the cp remains on support and ecmo was added with rp explant. the patient has survived the rp signal issues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical signal issue has been completed. the impella device was not returned for evaluation. data logs were reviewed were found that the placement signal was drifting with false drift in the flow rate. there were no motor current spikes observed. therefore, the root cause of the placement signal issue was sensor drift.
30930 | 1220648-2024-06823 | 18711556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711556&pc=OZD | 02/04/2024 | 03/20/2024 |  | 02/15/2024 |  |  | 02/04/2024 | 02/15/2024 |  |  | 08/31/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395760 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2024 |  |  |  |  | The user facility reported a 66-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient developed hemolysis accompanied by a decrease in hemoglobin levels and worsening hematuria. one-unit prbc was transfused, and the pump was repositioned to manage the condition. support remained ongoing with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. data logs were reviewed, and no positioning issue alarm were found. no motor current spike or placement signal trend were observed. also, suction alarm found at the start of the case is unrelated to hemolysis failure event. therefore, the root cause of the hemolysis could not be determined.
30931 | 1220648-2024-06824 | 18711725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711725&pc=OZD | 02/05/2024 | 03/26/2024 |  | 02/15/2024 |  |  | 02/05/2024 | 02/15/2024 |  |  | 09/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024401072 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 60-year-old male patient with was implanted with an impella 5.5 as a bridge to transplant. it was reported that after 2 months of ongoing support the pump stopped. the pump had been used well beyond indication. the pump was explanted and was replaced, the support was continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. data logs were reviewed which confirmed pump was unable to be run and received alarm 13 impella stopped motor current high with a motor current spike. however, without the device nor sufficient clinical information for evaluation, the exact root cause of the pump stop could not be determined.
30932 | 1220648-2024-06825 | 18711853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711853&pc=OZD | 02/06/2024 | 03/07/2024 |  | 02/15/2024 |  |  | 02/06/2024 | 02/15/2024 |  |  |  |  | 02/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1675422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Movement | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported that an automated impella controller (aic) had a loose purge disc holder. the issue was noted immediately prior to an impella 5.5 pump being plugged in. due to the noted issue, the decision was made to swap the aic for the planned support. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer.
30933 | 1220648-2024-06826 | 18711935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18711935&pc=OZD | 02/05/2024 | 04/09/2024 |  | 02/15/2024 |  |  | 02/06/2024 | 02/15/2024 |  |  | 07/31/2025 |  | 02/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024383464 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Computer Operating System Problem | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complaint reported that during impella cp support for 68-year-old female patient, the automated impella controller (aic) displayed "failure purge-system" alarm. the device failed to prime. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) purge issue has been completed. data logs were reviewed, and one event of purge system failure alarm observed. later the device was disconnected. impella cp pump and purge cassette returned for investigation. the pump was visually inspected and no damages or abnormalities found with the purge cassette. purge cassette was connected to the aic and deairing(priming) was initiated. the cassette successfully completed priming, was connected to the pump and the purge parameters were observed to be in normal range (purge pressure - 468 mmhg and purge flows - 11.9 ml/hr) after 10 minutes of purging. the priming issue did not reproduce during field reproduction testing. there was no priming issue found during the case. the purge cassette functioned/operated per specification. therefore, the root cause of the reported issue could not be determined. added-d9 device return date d4-updated model number. h8-changed to reuse.
30934 | 1220648-2024-06827 | 18712013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18712013&pc=OZD | 04/22/2022 | 09/22/2024 |  | 02/15/2024 |  |  | 04/22/2022 | 02/15/2024 |  |  | 12/31/2023 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023170301 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/22/2022 |  |  |  |  | The user facility reported that a 59-year-old male patient was on an impella 5.5 for mechanical circulatory support. it was reported that the impella 5.5 had high purge flow and pressure issues. heparin was added to the purge and the pressure resolved. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported high purge pressure issue has been completed. the pump was not returned for investigation. data logs show a gradual 15-minute rise in purge pressure, which was reduced after heparin was added to the purge. the cause of the high purge pressure issue was most likely biomaterial, based on data log review. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ it also states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06827 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
30935 | 1220648-2024-06828 | 18712025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18712025&pc=OZD | 02/06/2024 | 03/26/2024 |  | 02/15/2024 |  |  | 02/06/2024 | 02/15/2024 |  |  | 06/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375668 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and during support, there was pump flow issue. there was a clot ingestion and pump saturation was noted. the physician made the decision was made to remove the impella and replace it. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported thrombus and saturated flow issues has been completed. the impella device was not returned for evaluation. clinical details states that immediately after starting impella cp support, clot ingestion and saturated flow occurred. the patient's starting activated clotting time (act) was 130 and the team decided to explant the pump and replace it with another. no product was returned for a visual inspection. data log analysis confirmed that the pump was initially unable to start, but after a second attempt a motor current spike occurred followed by saturated flow at p9. the most likely cause of the saturated flow and temporary pump stop was biomaterial. added- b1 selected product problem d4-updated model number. g1-updated manufacturing site name and address.
30936 | 1220648-2024-06829 | 18712137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18712137&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/15/2024 |  |  | 02/07/2024 | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025436771 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed suspected hemolysis during impella cp support. following diagnosis, an attempt was made to wean the pump, but the patient was unable to tolerate the wean. support remained ongoing with the implanted pump despite the noted condition. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. the cause for hemolysis issue cannot be determined since the complaint device not returned for investigation and no sufficient clinical data provided.
30937 | 1220648-2024-06830 | 18712180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18712180&pc=OZD | 02/06/2024 | 03/18/2024 |  | 02/15/2024 |  |  | 02/06/2024 | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024426636 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 68-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support, the purge tubing was leaking. the nurse saw it dripping on to the patient's bed. the staff changed the tubing and the issue resolved. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the purge tubing leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge tubing leak was unable to be determined because the product was not returned. added- h6 impact code 4629.
30938 | 1220648-2024-06831 | 18712231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18712231&pc=OZD | 02/07/2024 | 08/26/2025 |  | 02/15/2024 |  |  | 02/07/2024 | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434111 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that one day into support, an impella 5.5 pump triggered a high purge pressure alarm followed by a purge blocked alarm. flows remained unchanged and tpa was ordered for the purge lines. roughly three hours later, the pump stopped. the device was explanted as a result of the noted failure. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message: could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported high purge pressure and subsequent pump stop has been completed. the impella device was returned for evaluation. visual evaluation confirmed biomaterial on the impeller and blood/biomaterial ingress in the purge line. no kinks were found after opening the red handle. data log analysis confirmed an increasing motor current trend occurred shortly after the high purge pressure and purge system blocked alarms, which is consistent with gradual biomaterial ingestion. the cause of the pump stop was blood ingress. added- d9 device return date, h6 impact code 4644, h6 problem code 1491 d4-updated model number h4-changed to yes. | No change in investigation results mentioned in manufacturer device report 1220648-2024-06831-1 g1 reporting contact email revised on manufacturer device report 1220648-2024-06831 in accordance with updated procedures. h6 investigation conclusions 4307 was erroneously entered on the initial manufacturer device report number 1220648-2024-06831-1.
30939 | 1220648-2024-06832 | 18713033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18713033&pc=OZD | 01/29/2024 | 04/03/2024 |  | 02/15/2024 |  |  | 01/29/2024 | 02/15/2024 |  |  | 10/31/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | GM |  | 2024421034 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The complainant reported during impella cp support for patient of unknown age and gender, the pump stopped. the impella was able to be restarted, but the pump was noted to be in the ventricle. the pump was ultimately explanted and the patient was placed on ecmo. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed since the original report was filed. the pump was returned, tested, and evaluated. biomaterial presence was observed in the purge gap below impeller at motor housing. the cause of the temporary pump stop issue was due to biomaterial interaction per log review and field investigation.
30940 | 1220648-2024-06833 | 18713515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18713515&pc=OZD | 01/31/2024 | 03/15/2024 |  | 02/15/2024 |  |  |  | 02/15/2024 |  |  | 10/31/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024423304 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the repositioning sheath became occlusive following placement. an external fem-fem bypass was subsequently performed to manage the condition. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was filed. the device was not returned. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30941 | 1220648-2024-06834 | 18713780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18713780&pc=OZD | 02/01/2024 | 03/14/2024 |  | 02/15/2024 |  |  |  | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435701 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a 63-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. following impella cp implant, the pump experienced consistent suction alarms. repositioning was attempted, but the issue persisted and it was believed that there was a kink proximal to the motor. due to the noted issue, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. there is a suction alarm just prior to the low flow. upon investigation of the pump a cannula kink was observed. the root cause of the low pump flow was determined to be a kinked cannula.
30942 | 1220648-2024-06835 | 18713834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18713834&pc=OZD | 02/01/2024 | 03/18/2024 |  | 02/15/2024 |  |  | 02/01/2024 | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426680 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the purge system became blocked shortly after the implant. tissue plasminogen activator was added to break up the clot. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure and the thrombus issues has been completed since the original report was filed. data log analysis confirmed gradual biomaterial ingestion as purge pressure rose with no motor current spikes and then gradually decreased as tpa was administered. the most likely cause of the high motor current was biomaterial.
30943 | 1220648-2024-06836 | 18714751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18714751&pc=OZD | 09/13/2022 | 05/23/2025 |  | 02/15/2024 |  |  | 09/13/2022 | 02/15/2024 |  |  | 05/31/2024 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229289 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 42-year-old male patient needed mechanical circulatory support. it was reported that the patient was on impella cp support and developed a hematoma at the right femoral artery (rfa) access site. the impella cp was removed from the rfa site, so the hematoma could be treated. the impella cp was then inserted through the left femoral artery (lfa). | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30944 | 1220648-2024-06837 | 18714771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18714771&pc=OZD | 02/02/2024 | 03/14/2024 |  | 02/15/2024 |  |  |  | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432842 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | The user facility reported a 70-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient developed limb ischemia. the staff attempted to switch the impella access site to the axillary artery. when the incision was made in the axillary, there was a large amount of bleeding. the bleeding was attempted to stop with sutures, and six units of packed red blood cells were given. a femostop was placed and flows were restored to the right lower leg. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ischemia potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. without sufficient clinical details, it was not possible to determine the root cause. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30945 | 1220648-2024-06838 | 18714785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18714785&pc=OZD | 02/04/2024 | 03/27/2024 |  | 02/15/2024 |  |  |  | 02/15/2024 |  |  | 06/30/2024 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023236616 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2024 |  |  |  |  | The user facility reported a 67-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient¿s pulses were unable to be heard on doppler bilaterally. an ultrasound was performed and showed diminished flow to lower extremities, and there was no flow noted on the impella side. a reposition sheath was placed, and some flow was restored, and leg color and temperature returned to within normal limits. the sheath was sutured in place and the patient was weaned off levophed and norepinephrine. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
30946 | 1220648-2024-06839 | 18714814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18714814&pc=OZD | 02/05/2024 | 03/15/2024 |  | 02/15/2024 |  |  | 02/05/2024 | 02/15/2024 |  |  | 04/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351450 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Movement | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella migrated to the aorta during a code blue in icu. all attempts for repositioning were unsuccessful. the patient was brought to the cath lab and the impella cp was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the device was not returned for investigation. based on clinical description and data analysis, the root cause of positioning issue was most likely use related.
30947 | 1220648-2024-06840 | 18714964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18714964&pc=OZD | 02/06/2024 | 03/21/2024 |  | 02/15/2024 |  |  |  | 02/15/2024 |  |  | 10/31/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024423298 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported an (b)(6)female patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had a small left ventricle cavity and short ascending aorta. the catheter had extremely low flows when bent and was positioned near aortic calve annulus. catheter had to be positioned shallow to have any flows but had small kink in the canula right below the motor. the flows were only 2-2.5 l/min. the physician decided to proceed with the case using this catheter. after the hrpci procedure, the impella was removed. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the initial report was submitted. the device was not returned for investigation. the cause of the low pump flow issue was patient condition related since cannula kinked due to patient anatomy.
30948 | 1220648-2024-06841 | 18715030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18715030&pc=OZD | 02/06/2024 | 03/21/2024 |  | 02/15/2024 |  |  |  | 02/15/2024 |  |  | 11/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024436321 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 75-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma superior and lateral to the repositioning sheath during impella cp support. manual pressure was held at the site and suction alarms persisted down to p-3 support level. femostop was then applied to the site and two units packed red blood cells (prbc) were transfused to manage the condition. due to the noted issues, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643 apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the low pump flow and access site bleeding has been completed. the impella device nor the data logs were returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues could not be determined.
30949 | 1220648-2024-06842 | 18715390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18715390&pc=OZD | 02/06/2024 | 04/05/2024 |  | 02/15/2024 |  |  | 02/06/2024 | 02/15/2024 |  |  |  |  | 02/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1752855 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) failed the system self check. the aic was sent back for investigation. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - system self-check error has been completed since the original report was filed. the pump was returned, tested, and evaluated. the reported complaint was confirmed; the console was unable to fully boot due to a system self-check failure. the cause of the aic issue was a defective impellatronics board. due to the defect, adequate voltage was not being delivered to the motor driving circuitry on the impellatronics board.
30950 | 1220648-2024-06843 | 18715582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18715582&pc=OZD | 02/06/2024 | 03/25/2024 |  | 02/15/2024 |  |  | 02/06/2024 | 02/15/2024 |  |  | 08/31/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024385002 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella cp pump immediately began to trigger placement signal unreliable alarms. it was noted that there were no waveforms present and the staff was advised that the optical sensor may have been damaged. due to the noted issue, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the optical signal issue was due to a kinked fiber line based on the changes in 5s parameters observed in the data logs.
30951 | 1220648-2024-06845 | 18715768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18715768&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/15/2024 |  |  |  | 02/15/2024 |  |  | 10/31/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024421299 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support experienced unresolved hemolysis. the impella was exchanged for a 5.5. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the hemolysis issue most likely was improper positioning due to improper technique.
30952 | 1220648-2024-06844 | 18715834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18715834&pc=OZD | 02/07/2024 | 03/20/2024 |  | 02/15/2024 |  |  | 02/07/2024 | 02/15/2024 |  |  |  |  | 02/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1096323 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support and the staff was performing a purge cassette change and the staff was stuck on 'connect purge disc'. staff tried multiple times to reconnect the purge disc without success. the aic (automated impella controller) was replaced however, the purge disc was not detected. a new purge cassette and tubing was used and inserted without issue. all alarms resolved. reportable due to purge issues that resulted in the replacement of the aic. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off/was missing on the console. the root cause of this issue was a missing/broken purge flag.
30953 | 1220648-2024-06846 | 18716197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18716197&pc=OZD | 02/07/2024 | 03/25/2024 |  | 02/15/2024 |  |  | 02/07/2024 | 02/15/2024 |  |  | 06/30/2025 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431102 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 32 hour impella cp support for a patient admitted in ami/cgs (acute myocardial infarction/ cardiogenic shock). the pump's access site had an ooze. the site was treated by pressure and removal of the iib/iiia integrilin. the access site did not need further treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-updated model number.
30954 | 1220648-2024-06847 | 18716226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18716226&pc=OZD | 09/16/2022 | 05/22/2025 |  | 02/15/2024 |  |  | 09/16/2022 | 02/15/2024 |  |  | 02/29/2024 |  | 02/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200216 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | Sepsis | No Clinical Signs, Symptoms or Conditions | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/16/2022 |  |  |  |  | The user facility reported a 65-year-old male patient was on an impella cp device for mechanical circulatory support. it was reported that the impella cp was repositioned under echocardigram. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | Medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
30955 | 1220648-2024-06783 | 18702019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18702019&pc=OZD | 01/27/2024 | 03/18/2024 |  | 02/14/2024 |  |  | 01/27/2024 | 02/14/2024 |  |  | 11/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024423691 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. n impella 5.5 pump was implanted via the axillary to support a patient suffering from heart failure and cardiogenic shock (cgs). the access site had a jackson-pratt (jp) drain and was having a persistent bleed that prompted the team to perform a stitch. the pump remained on for support for a total of just under five days. the patient survived and was bridged to a left ventricular assist device. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleed was most likely due to use as sutures were applied to resolve the bleeding. d4-updated model number. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
30956 | 1220648-2024-06784 | 18702114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18702114&pc=OZD | 02/02/2024 | 03/20/2024 |  | 02/14/2024 |  |  | 02/02/2024 | 02/14/2024 |  |  | 11/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433085 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp for support of the patient admitted with cgs (cardiogenic shock) and cardiomyopathy. the cp accessed limb developed limb ischemia. the ischemia was treated by the placement of a contralateral sheath and perfusion. the team additionally upped the anticoagulation. the pump remained on for support for 4 days and was then explanted. | [5 empty] | The investigation into the limb ischemia ¿ reversible issues has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30957 | 1220648-2024-06785 | 18702214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18702214&pc=OZD | 02/05/2024 | 03/20/2024 |  | 02/14/2024 |  |  | 02/05/2024 | 02/14/2024 |  |  | 11/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423287 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced a retroperitoneal bleed during impella cp support. heparin paused to assist with managing the condition. support remained ongoing with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported retroperitoneal (rp) bleed has been completed. no product was returned for evaluation. based on clinical details, patient did have calcified arteries and clinical team did have some difficulty inserting the pump but was able to have successful insertion after switching to the 25cm x 14fr introducer sheath. it was noted by the rep that "the rp bleed is from the impella cp stick from the outside hospital that originally placed the device", however, the bleed occurred 6 days after insertion and it was not specified when the patient was transferred to another hospital. the retroperitoneal bleed was resolved with pausing of heparin. the root cause of the retroperitoneal bleed cannot be determined as there is insufficient clinical details. d4-updated model number, h6 clinical code updated to 4476.
30958 | 1220648-2024-06786 | 18702241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18702241&pc=OZD | 01/28/2024 | 02/14/2024 |  | 02/14/2024 |  |  | 01/28/2024 | 02/14/2024 |  |  |  |  | 02/14/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1386344 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 33 YR |  | Female |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support for 33-year-old female patient, the automated impella controller (aic) displayed 'controller error' message leading to the console having to be exchanged. hours later when the patient was stable, they transferred to another aic. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30959 | 1220648-2024-06787 | 18702287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18702287&pc=OZD | 01/18/2024 | 03/19/2024 |  | 02/14/2024 |  |  | 02/10/2024 | 02/14/2024 |  |  |  |  | 02/14/2024 |  | AUTOMATIC IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080535 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/10/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support ongoing when the 2nd console utilized (after 1st console failed) had an issue with loading/recognition of the pc. the team replaced the aic (automated impella controller) again and support proceeded. no harm or injury from interruption in the support. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off/was missing on the console. the root cause of the purge flag issue was a missing/broken purge flag. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30960 | 1220648-2024-06788 | 18702415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18702415&pc=OZD | 01/29/2024 | 02/26/2024 |  | 02/14/2024 |  |  | 01/29/2024 | 02/14/2024 |  |  | 02/29/2024 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023183926 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a male patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during support, there was a pump stop leading to the pump being explanted. there was no reported patient harm. | [5 empty] | The investigation into the reported pump stop has been completed since the original report was filed. pump was returned for evaluation. no evidence of failure observed on pump. pump ran in bucket of water for approximately 12 minutes and was unable to reproduce pump stop failure. data logs were reviewed and it was noted that a long bag change, cassette change, and de-air procedure occurred over an hour time period which resulted in an increased purge pressure once completed. a gradual increase in purge was observed after and resulted in high purge pressure alarms which a cassette change did not resolve. no pump stop alarm was noted on logs. the root cause of the pump stop was most likely due to high purge pressure caused by use with a bag change that lasted over 2 minutes allowing blood to enter and occlude the purge gap. | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
30961 | 1220648-2024-06789 | 18702606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18702606&pc=OZD | 01/29/2024 | 05/15/2024 |  | 02/14/2024 |  |  | 01/29/2024 | 02/14/2024 |  |  |  |  | 02/14/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1678477 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Component Incompatible | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The us complainant had an automated impella controller (aic) fail during the setup for the high risk percutaneous coronary intervention (hrpci) with an impella cp for a patient. the aic had multiple personal computer (pc) failures and the purge disc hub ruptured. the team troubleshot by replacing the aic and the impella cp support began on the third pc. there was no harm or injury noted from the delay in support. no patient was involved. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. unable to reproduce the reported failure mode. the root cause for the ruptured purge pressure disc was not determined due to the inability to reproduce. in accordance with updated procedures, sections d2 has been revised from the initial submission. | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
30962 | 1220648-2024-06790 | 18702879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18702879&pc=OZD | 01/29/2024 | 02/14/2024 |  | 02/14/2024 |  |  | 01/30/2024 | 02/14/2024 |  |  | 08/31/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024388738 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported an 81-year-old female patient of unknown race and ethnicity in non-st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that due to coagulopathic issues there was no heparin. the patient was observed to have suffered a stroke. the stroke occurred on day two of the support. the pump was explanted the next day. the patient survived the stroke event. at the time of explant, there was an observation made of thrombosis at the pump inlet. | [5 empty] | The product was returned for investigation, but data logs were not provided. biomaterial was observed on the impeller. the pump ran as expected during a test in a bucket of water after removing the biomaterial from the impeller. the doctor also observed a clot on the pump with the returned product. as per the provided clinical information, heparin was not administered due to coagulopathic issues. during impella support, left-side weakness was noted, and a computed tomography (ct) scan revealed a right-sided posterior cerebral artery (pca) stroke. the pump was explanted the following day. the root cause of the stroke issue was biomaterial observed during the pump support. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury."
30963 | 1220648-2024-06791 | 18703236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18703236&pc=OZD | 01/30/2024 | 03/20/2024 |  | 02/14/2024 |  |  | 01/30/2024 | 02/14/2024 |  |  | 11/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024434343 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a male patient of an unknown age in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma in the right groin and manual pressure was held. one unit of packed red blood cells (prbcs) was transfused as treatment. it was further noted that the pump's positioning remained in the mitral valve despite attempts to maneuver the device. as a result, support levels could not be raised above p-5. the pump was weaned the following day and explanted. | [5 empty] | The investigation for the reported access site bleed and positioning issues has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and positioning issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
30964 | 1220648-2024-06792 | 18703613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18703613&pc=OZD | 01/29/2024 | 03/18/2024 |  | 02/14/2024 |  |  | 01/29/2024 | 02/14/2024 |  |  | 10/31/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418814 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that post coronary artery bypass graft surgery (cabg) the impella 5.5 was placed and coagulopathic issues occurred. the physician stated that cab and impella case times are very similar but only extended due to coagulation issues and bleeding in the chest wall/sternum. the physician after impella insertion cleaned and cleared the open chest wall but the patient continued to ooze, coagulopathic issue. the impella was decreased to p4 and factor 7 was given. the sternum started to look better, and the chest was closed. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30965 | 1220648-2024-06793 | 18703690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18703690&pc=OZD | 01/30/2024 | 08/18/2025 |  | 02/14/2024 |  |  | 01/30/2024 | 02/14/2024 |  |  |  |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1545509 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) failed during preventive maintenance. purge pressure was out of specification. after troubleshooting and replacing the purge assembly the problem was resolved. there was no patient involvement. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d2 common device name revised on manufacturer device report in accordance with updated procedures. e2 health professional revised on manufacturer device report in accordance with updated procedures. e3 occupation revised on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 investigation revised from the initial submission in accordance with updated procedures. h10 additional manufacture narrative revised from the initial submission in accordance with updated procedures.
30966 | 1220648-2024-06794 | 18703723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18703723&pc=PYX | 07/12/2023 | 09/22/2024 |  | 02/14/2024 |  |  | 07/12/2023 | 02/14/2024 |  |  | 07/31/2024 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023232096 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The complainant reported that a patient who had recently been on impella support had encountered a placement signaling not reliable alarm. the nurse attempted to manually re-zero without any success. the customer alleged a positioning issue, however, the investigation found it was actually a case of low pump flow. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06794 was provided in section h6. note: corrected information provided in h6 is an addition to previous coding. | Pump was not returned for investigation. based on the data logs, flow drops at the same time cvp goes negative and suction alarms are triggered. there are no abnormalities with the optical sensor as the 5s parameters are all within specification. the root cause of the low flow is most likely due to the patient¿s condition.
30967 | 1220648-2024-06795 | 18704143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18704143&pc=OZD | 01/30/2024 | 03/20/2024 |  | 02/14/2024 |  |  | 01/30/2024 | 02/14/2024 |  |  | 11/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024435817 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Twisted/Bent | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a (b)(6)female patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp came out of position during the sheath exchange. during the attempted re-delivery of the impella cp it became kinked, and the impella cp was removed. the team made the decision to escalate to impella 5.5 pump for support. no patient harm was reported. | [5 empty] | The investigation into the positioning issues and the product damage (cannula kink) issues has been completed since the original report was filed. the device was not returned for investigation. based on the clinical data, the positioning issue occurred while the peel away sheath was being removed. the cause of the positioning issue was use related due to improper technique. the cause of the cannula kink was most likely use related due to excessive force. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30968 | 1220648-2024-06796 | 18704563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18704563&pc=OZD | 01/15/2024 | 03/15/2024 |  | 02/14/2024 |  |  | 01/15/2024 | 02/14/2024 |  |  | 06/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024378533 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 81 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/15/2024 |  |  |  |  | The user facility reported a 81-year-old female was implanted with an impella device for mechanical circulatory support. complainant in canada reported the patient was on impella cp support when the patient had right groin bleeding. a stemi-dtu study documentation was shared on jan. 15th, 2024 found that the impella procedure was probably related to the access site bleeding. the impella device was found to be unlikely related to the access site bleeding. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30969 | 1220648-2024-06797 | 18705039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705039&pc=OZD | 01/30/2024 | 03/27/2024 |  | 02/14/2024 |  |  | 01/30/2024 | 02/14/2024 |  |  | 06/30/2024 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238344 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 77-year-old female patient was on impella cp support and the patient developed a loss of peripheral pulses in both lower extremities. heparin was withheld and the patient received two ampoules of bicarbonate. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported ischemia has been completed. the device nor sufficient clinical information was available for evaluation. therefore, the root cause of the ischemia could not be determined. this report is being filed as part of a retrospective review of historical records. h6 code for death 1802 was incorrectly selected on the initial report.
30970 | 1220648-2024-06798 | 18705058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705058&pc=OZD | 01/30/2024 | 03/12/2024 |  | 02/14/2024 |  |  | 01/30/2024 | 02/14/2024 |  |  | 03/31/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343007 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 74-year-old female patient was on impella cp support due to having triple vessel disease. while on support, a hematoma was forming lateral to the insertion site. heparin was withheld due to the hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that post procedure, the patient's activated coagulation time (act) was prolonged at 189. therefore, the root cause of the access site bleed was most likely due to improper anticoagulation management.
30971 | 1220648-2024-06799 | 18705122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705122&pc=OZD | 02/02/2024 | 03/21/2024 |  | 02/14/2024 |  |  | 02/02/2024 | 02/14/2024 |  |  | 10/31/2024 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024294493 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | The user facility reported a 38-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the impella was mispositioned in the left ventricle. the pump was removed and reinserted. upon reinserting the impella, the cp kinked. the pump was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues and the delivery issues has been completed since the original report was filed. the device was not returned for investigation. as per clinical, device got mispositioned during procedure and cannula kinked after repositioning attempt. the cause of the positioning issue was most likely use related per clinical review.
30972 | 1220648-2024-06800 | 18705159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705159&pc=OZD | 02/06/2024 | 03/21/2024 |  | 02/14/2024 |  |  | 02/06/2024 | 02/14/2024 |  |  | 10/31/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024423079 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for a patient admitted for tavr (transcatheter aortic valve replacement) workup for as (aortic stenosis). the care was escalated from iabp (intra-aortic balloon pump) to the cp and during the support there was a purge alarm troubleshot by pc troubleshooting. additionally, the team observed an ooze at the access site. the ooze was treated by forward tension suturing. no blood products or other intervention were needed. the pump remains on for support. | [5 empty] | The investigation into the purge cassette failure and the access site bleeding - major issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the "purge system open" alarm was most likely due to a use issue as it was resolved with de-airing the purge line and the pseudo-cassette change. the root cause of the access site bleeding was due to use issue as best practices were not followed and bleeding was resolved with forward-tension suturing. | The investigation for the reported issue has not yet been completed. upon its completion, a follow up report will be submitted. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30973 | 1220648-2024-06801 | 18705171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705171&pc=OZD | 02/06/2024 | 04/05/2024 |  | 02/14/2024 |  |  | 02/06/2024 | 02/14/2024 |  |  | 08/31/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024417101 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that the patient had signs of ischemia in the left lower limb in which the impella was inserted. the patient is being closely monitored. it was further reported that the patient has pale yellow blood urine with good urine outflow, evidence of hemolysis. | [5 empty] | The investigation into the hemolysis and the limb ischemia ¿ reversible issue has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of hemolysis issue cannot be determined since the complaint device, data logs were not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30974 | 1220648-2024-06802 | 18705195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705195&pc=OZD | 02/06/2024 | 03/26/2024 |  | 02/14/2024 |  |  | 02/06/2024 | 02/14/2024 |  |  |  |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383536 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. | It was reported that there was an automated impella controller failure (red alarm). the aic was restarted. remote monitoring was not possible due to the impella connect not connecting after restarting. there was no patient harm as a result of the issue. |  |  |  |  | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console alarm issues has been completed. the console was returned for evaluation. functional testing found that the unexpected shutdown and corresponding alarm was caused by unresponsive software process. data logs show an unexpectedly shut down and then restarted, but the reported alarm was presented after the restart - there were no alarms prior to the shutdown. the root cause of the reported alarm issues was a 3rd party software issue. added- b5 event description, d9 device return date d4-updated model number.
30975 | 1220648-2024-06803 | 18705203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705203&pc=OZD | 02/06/2024 | 09/25/2024 |  | 02/14/2024 |  |  | 02/06/2024 | 02/14/2024 |  |  | 06/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024431798 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Separation | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 14fr sheath placed for support with a cp pump. the patient was admitted with sob (shortness of breath) and confusion and coded on the ccl (cardiac catheterization lab) table. the 14fr sheath had a malfunction during the peel away process. the sheath did not break at the location expected and broke into 3 pieces. the team used hemostats to remove all pieces. | In addition, the patient was placed on cardizem in efforts to reduce systemic vascular resistance and allow impella to provide more flow. the patient¿s blood pressure dropped, became non-pulsatile, and required epinephrine and a dopamine drip to correct. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the introducer damage ¿ mechanical issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visually found no score line deviations. the most likely cause of the mechanical introducer damage was a sheath score line split. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. the impella device was received from the customer on (b)(6) 2024. introducer damage: according to the clinical, while inserting the impella cp device, the team attempted to separate the peel away sheath. however, the sheath did not separate along the score lines and broke into three separate pieces. hemostats and a blade were then used to remove the remaining pieces of the sheath from the catheter. product was returned for a visual inspection and no score line deviations were found. data log analysis was not necessary for this complaint. the most likely cause of the mechanical introducer damage was a sheath score line split. instructions for use for the related event are as follows: vf/vt/af/at: "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined." corrected information for the initial mfr 1220648-2024-06803 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
30976 | 1220648-2024-06804 | 18705213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705213&pc=OZD | 02/06/2024 | 03/22/2024 |  | 02/14/2024 |  |  | 02/06/2024 | 02/14/2024 |  |  | 06/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024372875 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump began to experience suction issues immediately following implant. upon troubleshooting, it was discovered that the cannula on the pump was kinked. due to the noted issue, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure: which triggers the ¿purge pressure high¿ alarm message: could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found a kink on the cannula. the root cause of low flow was a kinked cannula.
30977 | 1220648-2024-06805 | 18705219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705219&pc=PYX | 02/06/2024 | 11/07/2025 |  | 02/14/2024 |  |  | 02/06/2024 | 02/14/2024 |  |  | 12/31/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025442658 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | Tricuspid Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/06/2024 |  |  |  |  | It was further reported that there were concerns that the impella rp had damaged the bio-tricuspid valve. the impella was removed. tee showed wide open tricuspid regurgitation (tr). the patient was placed on ecmo and brought back to the or where the tv was removed, and a new one was placed. the rp seems to be holding the bio valve open and causing wide open tr. | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient has a right atrial (ra) thrombus possibly for the delay in starting heparin. |  |  |  |  | The investigation of the reported failure modes has been completed product was returned and investigated; data logs were not available. small biomaterial was found near the impeller of the pump. there were no visible damages or abnormalities on the pump. case images show the bio-tricupid valve leaflet to be "frozen open". echo images show the valve to be functioning appropriately prior to rp flex placement, with tricuspid regurgitation upon placement of the pump. the cause of the bio-tricuspid leaflet damage was unable to be determined due to lack of clinical information and no abnormalities observed on returned pump. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. instructions for use: rp flex with smart assist system with automated impella controller & rp with smart assist system with automated impella controller: section: contraindications (united states): ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ section: warnings & cautions: cautions: ¿patients with tricuspid or pulmonary valve stenosis or insufficiency, and patients with prosthetic tricuspid or pulmonary valves, may be compromised by the use of the impella rp flex with smartassist system catheter.¿ rp with smart assist system with automated impella controller. section: warnings & cautions: warnings: section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2024-06805. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-06805. d6a and d6b revised from the initial manufacturer device report 1220648-2024-06805 in accordance with updated procedures. d9 device returned to manufacturer date was inadvertently omitted on the initial manufacturer device report number 1220648-2024-06805. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-06805 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-06805 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-06805. h3 was erroneously entered on the initial manufacturer device report number 1220648-2024-06805. h6 type of investigation code 4114 was erroneously entered on the initial manufacturer device report number 1220648-2024-06805. h6 health effect - clinical code 4453, health effect - impact codes 4627 and 4624, and component code was inadvertently omitted on the initial manufacturer device report number 1220648-2024-06805. h10 narrative erroneously stated the device was not returned on the initial manufacturer device report number 1220648-2024-06805. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2024-06805. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus and heart valve damage has been completed. the impella device was returned for evaluation. visual inspection revealed biomaterial near the impeller of the pump. data logs were not returned for review. based upon clinical details, the root cause of the heart valve damage issue was most likely improper positioning of the device. however, the root cause of the thrombus could not be determined without additional clinical details. added-d9 device returned date d4-updated model number.
30978 | 1220648-2024-06806 | 18705253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705253&pc=OZD | 02/06/2024 | 03/22/2024 |  | 02/14/2024 |  |  | 02/06/2024 | 02/14/2024 |  |  | 11/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434925 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 86-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump met resistance during scheduled delivery for patient support. multiple attempts were made to place the pump, but the device would not advance beyond the area of the upper diaphragm. due to the noted resistance, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the clinical details states that the patient had calcified and tortuosity vessels condition. the cause of the delivery issue was patient condition related due to calcified and tortuosity vessels condition.
30979 | 1220648-2024-06807 | 18705958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18705958&pc=OZD | 01/31/2024 | 02/14/2024 |  | 02/14/2024 |  |  | 01/31/2024 | 02/14/2024 |  |  | 09/30/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024406213 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported a (b)(6)male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that pump was found to be leaking at the red plug. the pump was replaced due to the leak. no patient harm was reported. | [5 empty] | The impella cp was returned for evaluation. the pump was visually inspected and the red plug was taped with blood leaking out. a puncture or damage was observed to the catheter close to the 83 cm marking. also, sterile sleeve damage was observed. the root cause of the product damage (hole in the catheter) issue was use related due to catheter being punctured during support per field investigation.
30980 | 1220648-2024-06808 | 18706556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18706556&pc=OZD | 02/01/2024 | 04/04/2024 |  | 02/14/2024 |  |  | 02/01/2024 | 02/14/2024 |  |  | 05/31/2025 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366014 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a male patient of an unknown age in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that approximately two days into support it was documented that the patient had ischemia in their lower extremity. due to the noted condition, the pump was explanted, and an intra-aortic balloon pump (iabp) was placed for ongoing support. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
30981 | 1220648-2024-06809 | 18707228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18707228&pc=OZD | 02/01/2024 | 02/14/2024 |  | 02/14/2024 |  |  | 02/01/2024 | 02/14/2024 |  |  |  |  | 02/14/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1103512 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/01/2024 |  |  |  |  | The user facility biomedical department reported a complainant had a battery failure (red alarm) noted on an automated impella controller (aic). the controller was replaced. there was no patient use nor harm/injury reported. | [5 empty] | The investigation has been completed. the product was returned. the power logs showed that battery one had a cell imbalance on cell two of 2v. the failure mode was reproduced on an engineering bench. batttest confirmed battery one cell two was outputting 2v significantly under the 4v. the rest of the cells were outputting. the battery one pack voltage was measured to be 0 v rather than the expected 16 v, and the battery liquid-crystal display (lcd) indicator was blank (fully discharged) due to the cell imbalance. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure.
30982 | 1220648-2024-06810 | 18707833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18707833&pc=OZD | 09/19/2022 | 05/20/2025 |  | 02/14/2024 |  |  |  | 02/14/2024 |  |  | 03/31/2024 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211263 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2002 |  |  |  |  | The user facility reported a 62-year-old male patient needing mechanical circulatory support. it was reported, prior to being placed on impella cp support, the patient had presented with elevated myocardial infarction. while under impella cp support, the patient¿s labs hemolyzed overnight. the impella cp was decreased from p9 to p7. the urine output was down to 5ml within the last hour and creatinine was up to 3.17. the patient was started on continuous renal replacement therapy (crrt) and was escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30983 | 1220648-2024-06811 | 18707891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18707891&pc=OZD | 04/17/2022 | 05/29/2025 |  | 02/14/2024 |  |  |  | 02/14/2024 |  |  | 01/31/2024 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189340 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2022 |  |  |  |  | The user facility reported a 63-year-old male implanted with an impella cp for mechanical circulatory support experienced a hematoma and excess bleeding at the site of the insertion. it was also noted that heparin was used to stabilize the hematoma. the hematoma was expressed and redressed. after the patient was successfully weaned from the impella cp, the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. this report is being filed as part of a retrospective review of historical records.
30984 | 1220648-2024-06812 | 18707962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18707962&pc=OZD | 04/18/2022 | 03/05/2024 |  | 02/14/2024 |  |  |  | 02/14/2024 |  |  | 12/31/2023 |  | 02/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023180715 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 60 year old male patient for electrophysiology and ablation. it was reported that the impella site was injected with thrombin and ultimately caused vessel closure and acute limb ischemia. this was repaired surgically. when the patient returned at a later date, the decision was made not to access the left groin. | [5 empty] | The investigation into the reported limb ischemia-revesible has been completed since the original report was filed. no device was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
30985 | 1220648-2024-06757 | 18694793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18694793&pc=OZD | 01/29/2024 | 03/21/2024 |  | 02/13/2024 |  |  | 01/29/2024 | 02/13/2024 |  |  | 08/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394055 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella 5.5 fell into the aorta. the medical doctor recommended to reduce the p level to p2. a cardiologist has been called to verify the position of the impella. it was recommended that the impella position to be in the left ventricle. no harm to the patient was reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30986 | 1220648-2023-02282 | 18695513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18695513&pc=OZD | 05/09/2023 | 08/24/2024 |  | 02/13/2024 |  |  | 07/14/2023 | 02/13/2024 |  |  | 12/31/2024 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024326427 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/09/2023 |  |  |  |  | The complainant reported a 70 year old male patient presenting with acute myocardial infarction and cardiogenic shock presenting for impella support. during support, a loss of dopplerable pulses was noted in the impella extremity, at the popliteal artery. the vascular team was made aware and there was discussion on escalating to a 5.5 axillary impella, however, the family ultimately withdrew care due to patient's multiple organ failure and the patient expired. no intervention was provided for the ischemia. | The complainant reported a 70 year old male patient presenting with acute myocardial infarction and cardiogenic shock presenting for impella support. during support, a loss of deplorable pulses was noted in the impella extremity, at the popliteal artery. the vascular team was made aware and there was discussion on escalating to a 5.5 axillary impella, however, the family ultimately withdrew care due to patient's multiple organ failure and the patient expired. no intervention was provided for the ischemia. |  |  |  |  | The investigation into the limb ischemia has been completed since the initial report was submitted. the device was not returned for investigation. based on the clinical information provided, the root cause of the limb ischemia was patient condition related because of the patient's deteriorating condition and noted multiple organ failures. d6a, d6b. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was patient condition related, because patient deteriorating condition and multiple organ failure were noted. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30987 | 1220648-2024-06760 | 18698069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18698069&pc=OZD | 01/26/2024 | 03/29/2024 |  | 02/13/2024 |  |  | 01/26/2024 | 02/13/2024 |  |  | 12/31/2024 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024289241 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Life Threatening | Required Intervention | Required Intervention | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during percutaneous coronary intervention (pci) for (b)(6)patient on impella cp support, the ramus interventricularis anterior (riva) was perforated with the wire, resulting in a pericardial effusion. as a result, the patient had to be resuscitated several times. the doctors decided to explant the impella and implant an ecmo. the patient was then transferred to the icu without the impella. | [5 empty] | The investigation for the reported vessel damage has been completed. the impella product was not returned. the perforation of the ramus coronary artery occurred during the pci intervention process due to the pci wire and is highly unlikely to occur due to impella. the cause of the vascular damage-great vessel issue was use related due to pci wire leading to perforation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30988 | 1220648-2024-06761 | 18698124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18698124&pc=OZD | 01/27/2024 | 03/12/2024 |  | 02/13/2024 |  |  | 01/27/2024 | 02/13/2024 |  |  | 10/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424146 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 66-year-old male patient was on impella cp support due to having acute myocardial infarction. while on support, access site bleeding occurred. the staff was instructed to hold pressure. systematic heparin was withheld, the white of the repo. sheath was exposed, so the sheath was advanced and the physician re-sutured the access site. the bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation the root cause of the access site bleed was determined to be patient movement. it was mentioned that the bleeding was started at the groin during transport.
30989 | 1220648-2024-06762 | 18698785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18698785&pc=OZD | 02/06/2024 | 03/27/2024 |  | 02/13/2024 |  |  | 02/06/2024 | 02/13/2024 |  |  | 10/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024416061 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The user facility reported a 68-year-old male was on impella 5.5 mechanical circulatory support due to cardiogenic shock. there was bleeding observed at the tunnel site. the staff resolved the issue by applying a purse string suture. additionally, a pump stop was observed, however, the impella 5.5 was not explanted and support continued after the pump stop. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the access site bleeding ¿ major and the pump stop issues has been completed since the initial report was submitted. the device was returned and was visually inspected; no biomaterial or catheter, cannula kink was observed. the cause of the pump stop issue was most likely an open electrical line due to likely excessive force during support. the cause of the access site bleeding issue was use related since bleeding resolved with additional suture.
30990 | 1220648-2024-06763 | 18698893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18698893&pc=OZD | 01/29/2024 | 03/20/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 10/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422806 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there was placement signal unreliable alarms. placement was verified and support continued. additionally, the patient had a hematoma and bleeding was noted around the impella site. manual pressure was held and a femostop was placed. | [5 empty] | The investigation for the reported access site bleed has been completed. upon further review, the event of hematoma was found not reportable as the device was not prematurely removed, no blood products were given, and no other intervention was required beyond manual pressure and femostop. the root cause of the access site bleed could not be determined without additional clinical details or a returned device. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30991 | 1220648-2024-06764 | 18699061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699061&pc=OZD | 01/29/2024 | 03/11/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 04/30/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024336979 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 57-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient developed hemolysis during impella 5.5 support. repositioning was attempted, but the staff was unable to torque the pump out of the mitral. due to the noted condition, the decision was made to replace the pump with a new impella 5.5 device. | [5 empty] | The investigation for the reported hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. data logs were reviewed which did not show any positioning alarms being triggered during the case. impella position unknown alarms were found. no other issues were found on the logs. clinical information provided along with the data log review was sufficient to determine the root cause. the root cause of hemolysis was most likely due to improper positioning as per clinical description since it was confirmed pump was stuck at mitral valve. the root cause of positioning issue was most likely due to use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30992 | 1220648-2024-06765 | 18699119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699119&pc=OZD | 01/29/2024 | 03/14/2024 |  | 02/13/2024 |  |  | 01/29/2024 | 02/13/2024 |  |  | 12/31/2024 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024291838 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a 5.5 pump to escalate care from the iabp (intra-aortic balloon pump). the patient had been admitted to nstemi (non-st segment elevation myocardial infarction) and vsd (ventricular septal defect). the 5.5 pump was seen to have damage at the hub. the inner lining at the hub was exposed. the pump remained on for support despite this malfunction. there has been no harm/injury/infection/leak. | [5 empty] | The investigation for the reported product damage has been completed. the impella device was not returned for evaluation. details provided states that seemed like "the cut was rough suggesting that it was a result of excessive force used but the cut seemed to be very close to the bonded area inside the red handle". these facts were inconclusive how the plastic part of catheter was exposed. therefore, the root cause of the product damage issue was not determined since product was not returned for investigation and insufficient clinical information was provided. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30993 | 1220648-2024-06766 | 18699151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699151&pc=OZD | 01/30/2024 | 03/15/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 11/30/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430105 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 52-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient developed access site bleeding, sutures were placed by the surgeon to treat the condition and support with the implanted pump was maintained. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30994 | 1220648-2024-06767 | 18699195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699195&pc=OZD | 01/30/2024 | 03/19/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 10/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024410725 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Difficult to Advance | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 85-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during scheduled impella cp delivery, two used introducers kinked due to the patient's anatomy with one resulting in a crack. as a result of the noted issue, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the difficulty inserting/removing introducer issue has been completed since the original report was filed. the product was not returned for investigation. per the clinical information, the root cause of the introducer damage mechanical issue was an introducer sheath kink caused due to patient's small/ diseased vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30995 | 1220648-2024-06768 | 18699262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699262&pc=OZD | 11/05/2023 | 03/15/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 09/30/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024399851 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 73-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was place in the stemi dtu protocol. the patient was supported with cp for 21 hours before being escalated to the 5.5 pump. months after the completed support the team noted in the case report form (crf) while on impella cp support that there was ventricular tachycardia (vt) that was treated and attributed as a possible relation to the use of impella, and the team performed cardioversion at 200jules to treat. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: nothing beyond study data.
30996 | 1220648-2024-06769 | 18699303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699303&pc=OZD | 01/31/2024 | 03/07/2024 |  | 02/13/2024 |  |  | 01/31/2024 | 02/13/2024 |  |  |  |  | 02/13/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1129223 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported that when starting a new case, the purge cassette was inserted first. then the disk was placed, and an alarm came on the automated impella controller (aic) saying to insert the disk and the alarm would not clear. the alarm would not let the aic advance to the next step. the aic was not sensing the purge disk. there was no patient harm reported. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge flag. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
30997 | 1220648-2024-06770 | 18699332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699332&pc=OZD | 01/22/2024 | 03/28/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 10/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024412518 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Information from a st-segment elevation myocardial infarction (stemi)-door to unloading (dtu) study reported that a patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. during removal of the impella cp, dissection of the external iliac artery occurred, which did not affect flow in the vessel, but resulted in a hematoma and required mechanical compression in addition to the closure system (manta) being applied. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation into the vascular damage - peripheral vessel issue has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the vascular damage issue was not determined due to insufficient clinical information and the relation to impella being unknown.
30998 | 1220648-2024-06771 | 18699338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699338&pc=OZD | 01/26/2024 | 03/12/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 09/30/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395014 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a(b)(6)old male patient was implanted with an impella 5.5 as a bridge to transplant. it was reported the team observed a purge side-arm leak. through close inspection, small crack noted on purge sidearm. the patient continued support and there was no reported patient harm. | [5 empty] | The investigation into purge leak ¿ pump issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the purge leak - pump was reported to be a leak from the sidearm but the exact location of the leak was unknown. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30999 | 1220648-2024-06772 | 18699835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699835&pc=OZD | 01/26/2024 | 05/02/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  |  |  | 02/13/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when removing the purge cassette after impella 5.5 was explanted, the blue ¿wings¿ of the aic where the purge valve is inserted broke. there was no harm to the patient. | [5 empty] | The investigation into the removal issues has been completed since the initial report was submitted. the console was returned and tested, and visually inspected under a digital microscope and was confirmed to be missing the purge retainer and flag. as a result, it was confirmed that the console had aic - purge disc/flag issues the root cause for the missing purge flag was a broken retainer. to conform to updated procedures, section d2 was revised with updated information. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
31000 | 1220648-2024-06773 | 18699878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699878&pc=OZD | 01/28/2024 | 03/12/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 10/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422590 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having st (sinus tachycardia) elevated myocardial infarction. after the pump was implanted, the doctor reported that there were challenges with implanting the impella. the pump was caught under the chordae, and when they attempted to reposition, the impella would jump into the aorta and the doctor was unable to readvance. the impella was eventually removed, flushed, and then replanted. the doctor felt the pump was too deep in the ventricle, however the doctor left the pump in that position due to fear of the pump jumping out of the ventricle. impella support continued. | [5 empty] | The investigation for the reported positioning issue has been completed. no product was returned. as per clinical pump was possibly caught under chordae during implant, and when tried to reposition the pump impella jumped into aorta and doctor unable to readvance. later pump was able to reimplant successfully after flushing the pump. the cause of the positioning issue was use related due to improper technique per clinical review. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
31001 | 1220648-2024-06774 | 18699920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699920&pc=PYX | 01/29/2024 | 03/12/2024 |  | 02/13/2024 |  |  | 01/29/2024 | 02/13/2024 |  |  | 11/30/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024435352 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a bipella support for a patient admitted post syncopal episode and known cardiac disease. the rp flex - cp support was ongoing when after a day of support the rp flex pump stopped. the stopped pump was restarted. the rp had "heparin" running while the cp had bicarb in the purge. the rp and cp remain on for support despite the malfunction. | [5 empty] | The investigation for the reported issue has not yet been completed. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into the pump stop and the removal issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. the pump stop was unable to be reproduced in the lab. data logs and returned product showed no evidence of biomaterial. the clinical states that the distal end of the patient's picc line had been found to be in the inflow cage of the pump upon explant. the root cause of the temporary pump stop was most likely an interaction with the picc line.
31002 | 1220648-2024-06775 | 18699931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699931&pc=OZD | 01/29/2024 | 03/12/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 09/30/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024395913 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Ischemia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to having chronic systolic heart failure, cardiogenic shock and the patient was awaiting a heart transplant. while on support, the patient's right arm was swelling. the staff was concerned with venous compression from the impella. surgery was made aware, but no intervention was planned. additionally, there was suction alarms present throughout the case, however support continued. reportable due to low pump flow and the patient's right arm swelling. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the limb ischemia ¿ reversible and the low pump flow issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the low pump flow and the limb ischemia was unable to be determined as limited clinical details were provided and no product was returned for review.
31003 | 1220648-2024-06776 | 18699951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699951&pc=OZD | 01/30/2024 | 03/14/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 06/30/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370654 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced a bleed of unknown origin during impella cp support. throughout the course of the condition, the patient's hemoglobin fluctuated, and the patient stabilized with blood transfusions. the patient's site looked swollen coinciding with scrotal swelling. due to the bleed, the decision was made to stop systemic heparin. support was maintained following the decision. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-updated model number | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31004 | 1220648-2024-06777 | 18699973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18699973&pc=OZD | 01/30/2024 | 03/13/2024 |  | 02/13/2024 |  |  | 01/30/2024 | 02/13/2024 |  |  | 11/30/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434367 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a male patient of an unknown age in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that pump was pulled back into the aorta and was no longer across the aortic valve. the impella was removed and replaced. | [5 empty] | The investigation into the access site bleeding ¿ minor and the positioning issues has been completed since the original report was filed. the device was not returned for investigation. the cause of the access site bleeding and the positioning issues was most likely patient condition related, per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31005 | 1220648-2024-06778 | 18700001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18700001&pc=OZD | 01/31/2024 | 03/19/2024 |  | 02/13/2024 |  |  | 01/31/2024 | 02/13/2024 |  |  | 10/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419342 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported an 86-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the team had difficulty with traversing the native anatomy. the team attempted all troubleshooting and guidewires. the delivery failed. the cp was aborted, and percutaneous coronary intervention (pci) proceeded without the support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the product was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since the doctor had difficulty delivering the pump due to resistance from aorta.
31006 | 1220648-2024-06779 | 18700006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18700006&pc=OZD | 01/31/2024 | 03/19/2024 |  | 02/13/2024 |  |  |  | 02/13/2024 |  |  | 07/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409233 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported a 67-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced bleeding and developed a hematoma at their impella access site follow implant. a unit of blood was transfused, and pressure was held at the site to manage to condition. roughly four days later, a thrombus was identified in the patient's left ventricle. due to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the thrombus issues has been completed since the original report was filed. the product was not returned for investigation. the root cause of the access site bleeding was most likely due to patient condition since hematoma and bleeding was occurring at multiple sites. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
31007 | 1220648-2024-06780 | 18700016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18700016&pc=OZD | 02/01/2024 | 03/27/2024 |  | 02/13/2024 |  |  | 02/01/2024 | 02/13/2024 |  |  |  |  | 02/13/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1579730 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a 52-year-old male patient of an unknown race with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the automated impella controller (aic) had a controller failure alarm. the aic was exchanged. there was no patient harm reported. | [5 empty] | The investigation into the aic - controller failure (red alarm) issue has been completed since the initial report was submitted. the console was returned and tested. the failure was unable to be reproduced throughout testing. the root cause of the issue could not be determined due to the difficulty in reproducing the failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm with a low probability of reoccurrence. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31008 | 1220648-2024-06781 | 18700028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18700028&pc=OZD | 02/05/2024 | 03/15/2024 |  | 02/13/2024 |  |  | 02/05/2024 | 02/13/2024 |  |  | 07/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390508 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 40 year old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella had suction alarms that would not resolve at p2 level. waveforms were evaluated and there was a very flat motor current and extremely low flows. echo echocardiogram (ultrasound) was performed which revealed a large mobile thrombus on the inlet of the impella device. the patient was taken to the cath lab and the pump was removed over the wire and exchanged for a new cp. when the pump was in the low flow state, presser requirements significantly increased with one drip dose increased 4x and an additional drip was added. the patient¿s blood pressure was borderline throughout the entire exchange, but improved once new cp was implanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported low pump flow has been completed. the impella device was not returned for evaluation. data logs show a spike in motor current and active suction alarm, followed by low pump flow during the start of the case. the cause of the low pump flow issue was most likely biomaterial, based on data log review. d4-updated model number.
31009 | 1220648-2024-06782 | 18700037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18700037&pc=OZD | 02/05/2024 | 03/20/2024 |  | 02/13/2024 |  |  | 02/05/2024 | 02/13/2024 |  |  | 10/31/2025 |  | 02/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434186 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 was not able to be inserted due to the patient¿s anatomical tortuosity. the catheter also kinked near the inlet. the 5.5 insertion was aborted and impella cp was placed without difficulty. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. added- h6 impact code 4627 d4-updated model number.
31010 | 1220648-2024-06725 | 18684258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18684258&pc=OZD | 01/28/2024 | 03/14/2024 |  | 02/12/2024 |  |  | 01/28/2024 | 02/12/2024 |  |  | 05/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024368931 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2024 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the limb became ischemic overnight. the team made the decision to explant the impella cp and the vascular surgeon was consulted to do a cut-down and arterial 'clean out'. the team then placed a new pump via the axillary access. the physician was aware of the life versus limb aspect of the care and had proceeded with the femoral placed impella cp. | [5 empty] | The investigation into the limb ischemia ¿ irreversible issues has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury."
31011 | 1220648-2024-06726 | 18685211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18685211&pc=OZD | 01/31/2024 | 03/15/2024 |  | 02/12/2024 |  |  | 01/31/2024 | 02/12/2024 |  |  | 10/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425004 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported an impella cp was implanted in a 74-year-old female patient needing mechanical circulatory support. it was reported that the patient had a cold foot and was developing ischemia. the impella cp was removed and an impella 5.5 was placed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was filed. the device was not returned. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31012 | 1220648-2024-06727 | 18685382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18685382&pc=OZD | 02/01/2024 | 03/07/2024 |  | 02/12/2024 |  |  | 02/01/2024 | 02/12/2024 |  |  |  |  | 02/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1309660 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | A complainant reported that while setting up an impella cp pump for training, the purge cassette door on the automated impella controller (aic) was opened and the blue purge holder fell out. due to the broken component, a second console was used for training. there was no patient involvement. | [5 empty] | The automated impella controller (aic) has been returned for evaluation but has not yet been evaluated. upon completion of the investigation, a supplemental report will be submitted. | The investigation for the console (aic) purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer.
31013 | 1220648-2024-06728 | 18685713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18685713&pc=OZD | 02/02/2024 | 03/21/2024 |  | 02/12/2024 |  |  | 02/02/2024 | 02/12/2024 |  |  | 10/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406304 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | The user facility reported a 81-year-old male patient needing mechanical circulatory support. it was reported that the team placed the graft and when delivering the impella 5.5, the pump was pushed and perforated the left ventricle (lv) which was known to have "very soft" tissue. the lv was repaired and the impella 5.5 support continued. | [5 empty] | The investigation for the reported ventricle perforation issue has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed negative placement signal metrics (likely ventricular) between two placement signal spikes during insertion. "impella position unknown" alarms triggered upon activating the pump, which resolved following the left ventricle (lv) repair. the root cause of the ventricle perforation was most likely improper positioning of the device during insertion. | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve¿
31014 | 1220648-2024-06729 | 18685912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18685912&pc=OZD | 09/20/2023 | 02/12/2024 |  | 02/12/2024 |  |  | 09/20/2023 | 02/12/2024 |  |  |  |  | 02/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1161393 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported patient of unknown age received hemodynamic support from an impella cp smart assist heart pump. when the automated impella controller (aic) was booted up, there was a ¿error of the purge system¿ displayed prompting another console to be used instead. pump was ultimately explanted the day after. there was no reported patient harm. | [5 empty] | The investigation into the reported aic - purge issue ¿ other has been completed. data analysis showed that low purge pressure and air in purge pressure alarms were most likely due to a kinked purge line on (b)(6). at the same time of those alarms, ic3179 is booted up and immediately had a controller failure for alarm. reviewing the logs, the same alarm, fast toggled on boot up on (b)(6)repeatedly triggering and clearing. the alarm resolved when the purge cassette was inserted. during testing, the failure mode was reproduced on an engineering bench. on boot up, the alarm triggered only resolving when the purge cassette was inserted. swapping out the pud with a known good caused the alarms to not return. while investigating the console, the purge shaft was also found to be difficult to turn due to dried glucose. based on clinical description and data analysis, the root cause of the alarms was caused by a malfunctioning purge unit driver. the blocked purge shaft was most likely caused by a glucose ingress.
31015 | 1220648-2024-06730 | 18685972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18685972&pc=OZD | 01/17/2024 | 03/12/2024 |  | 02/12/2024 |  |  | 01/17/2024 | 02/12/2024 |  |  | 08/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397677 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a male 38-year-old patient of unknown race post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that while on support the patient experienced multiple transient ischemic attacks (tia) and stroke events. it was further reported that the physician explanted the impella and during the left ventricular assist device (lvad) surgery the physician removed a large clot at the apex of the left ventricle. | [5 empty] | The investigation into the thrombus and the stroke/cva has been completed since the original report was filed. the device was returned. as the necessary clinical information was not provided and even though the device was returned, the root cause of the thrombus and the stroke/cva was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury
31016 | 1220648-2024-06731 | 18686042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686042&pc=OZD | 01/23/2024 | 09/20/2024 |  | 02/12/2024 |  |  | 01/23/2024 | 02/12/2024 |  |  | 05/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024353102 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that 74-year-old female patient on impella 5.5 support was in the or for thrombus removal from the left ventricle and left atrium. the thrombus had been suspected to be in the left ventricle at the time of implant. the patient expired on impella support. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ corrected information was provided in sections b2, b5, h1, h6 and h10. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
31017 | 1220648-2024-06732 | 18686115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686115&pc=OZD | 02/04/2024 | 03/25/2024 |  | 02/12/2024 |  |  | 02/04/2024 | 02/12/2024 |  |  | 11/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423449 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2024 |  |  |  |  | The user facility reported an impella cp was implanted in a 51-year-old female needing mechanical circulatory support. it was reported that the impella cp had a spike in motor current, and it appeared that the pump was saturated. the purge pressure high alarm was activated. after the pump stopped, the surgeon explanted the pump, and the patient was stable at the time of explant. there was no patient harm reported. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device nor the data logs were received for evaluation. however, clinical details stated that high purge pressure alarms occurred just prior to the pump stop. therefore, it is likely that the cause of the pump was due to high pressure alarms. however, the root cause of the high pump pressure could not be determined without additional information or the device for evaluation. the instructions for use states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31018 | 1220648-2024-06733 | 18686344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686344&pc=OZD | 01/24/2024 | 03/25/2024 |  | 02/12/2024 |  |  | 01/24/2024 | 02/12/2024 |  |  | 06/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372059 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp implant procedure for 63-year-old female patient, the pump started well but there was a sudden drop in flow to 1l and there were continuous suction alarms. the pump position was checked and even though there were no anomalies, the pump was exchanged due to suspected occlusion of cannula. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation into the low pump flow issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the low pump flow could not be determined since the pump and data logs were not returned for investigation.
31019 | 1220648-2024-06734 | 18686391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686391&pc=OZD | 01/24/2024 | 03/25/2024 |  | 02/12/2024 |  |  | 01/24/2024 | 02/12/2024 |  |  | 06/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374498 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 61-year-old female patient on impella cp support experienced bleeding issues leading to pump being explanted. patient was noted to be stable post-explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31020 | 1220648-2024-06735 | 18686444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686444&pc=OZD | 01/25/2024 | 03/27/2024 |  | 02/12/2024 |  |  | 01/25/2024 | 02/12/2024 |  |  | 10/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024413854 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that patient of unknown age and gender on impella 5.5 experienced high pump flow and pump stop. the automated impella controller (aic) displayed pump error and to change pump. the issue initially resolved following restart of the pump. however, the pump was ultimately explanted and exchanged for another device. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the initial report was submitted. the device was returned, visually inspected, and tested. significant biomaterial blockage observed at outflow cage/wrapped around the impeller. the cause of the temporary pump stop issue was due to biomaterial per log and field investigation.
31021 | 1220648-2024-06736 | 18686497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686497&pc=OZD | 01/24/2024 | 03/12/2024 |  | 02/12/2024 |  |  | 01/24/2024 | 02/12/2024 |  |  | 10/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420857 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient received hemodynamic support from an impella cp smartassist heart pump and venoarterial extracorporeal membrane oxygenation (va-ecmo) due to acute myocardial infarction/cardiogenic shock. on the 12th day after insertion, the impella cp pump suddenly stopped. the intended period of use for an impella cp is five days. the impella cp was removed and replaced with an impella 5.5 smartassist heart pump. the patient suffered no harm from the incident and was successfully weaned off the impella pump support several days later. | [5 empty] | The impella device was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the pump stop could not be determined because no product or logs were returned.
31022 | 1220648-2024-06737 | 18686528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686528&pc=OZD | 01/29/2024 | 03/25/2024 |  | 02/12/2024 |  |  | 01/29/2024 | 02/12/2024 |  |  |  |  | 02/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 10903701 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported that an automated impella controller (aic) failed to move past the blue screen when powered on. there was no patient harm, and the device was removed from service as a result of the failure. | [5 empty] | The investigation into the aic - power on issues/blank screen has been completed since the original report was filed. the console was returned and tested after arriving in fs. the failure mode was not reproduced consistently. upon turning on the aic with the ac power, the console screen was stuck on a white screen and was unable to boot up. the root cause for the intermittent boot-up issue was not determined due to the failure mode no longer reproducing during testing to narrow down a root cause. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31023 | 1220648-2024-06738 | 18686613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686613&pc=OZD | 01/28/2024 | 03/21/2024 |  | 02/12/2024 |  |  | 01/29/2024 | 02/12/2024 |  |  | 07/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374350 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 79-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella fell into the aorta while adjusting the position. the impella was removed. there was no patient harm reported. | [5 empty] | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31024 | 1220648-2024-06740 | 18686691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686691&pc=OZD | 01/31/2024 | 03/19/2024 |  | 02/12/2024 |  |  | 01/31/2024 | 02/12/2024 |  |  | 11/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427026 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Thrombosis/Thrombus | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to having systolic heart failure. during the impella implant, the patient went into ventricular fibrillation (vf) twice and had to defibrillated. additionally, the an ultrasound was performed and a small hematoma was noted, heparin was withheld due to the hematoma. furthermore, the patient developed clots. a washout needed to be performed and a new jp drain was placed. reportable due to patient's vf, hematoma and thrombus. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed, ventricular fibrillation, and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed, ventricular fibrillation, and thrombus could not be determined.
31025 | 1220648-2024-06739 | 18686706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686706&pc=OZD | 01/25/2024 | 03/21/2024 |  | 02/12/2024 |  |  | 01/25/2024 | 02/12/2024 |  |  | 11/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430451 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 74-year-old female patient experienced massive bleeding from the repo sheath between the blue rubber and the hub, leading the patient to requiring a transfusion and device explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleed was most likely use-related since adding additional suture on the repo sheath resolved the bleeding issue.
31026 | 1220648-2024-06741 | 18686722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686722&pc=OZD | 01/29/2024 | 03/15/2024 |  | 02/12/2024 |  |  | 01/29/2024 | 02/12/2024 |  |  | 05/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024351871 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHARMACIST | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 68-year-old patient of unknown ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was oozing from the impella access site. dressing changes were performed, manual pressure was applied, the preclose sutures were tightened and the patient was given one unit of blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643 apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned. the root cause of the access site bleeding issue was use related, as per the clinical description provided, sutures were tightened post-surgery.
31027 | 1220648-2024-06742 | 18686744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686744&pc=OZD | 01/25/2024 | 03/08/2024 |  | 02/12/2024 |  |  | 01/25/2024 | 02/12/2024 |  |  |  |  | 02/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1383042 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a battery failure. the aic was therefore removed from service. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - battery failure (red alarm) issue has been completed since the original report was filed. the console was returned and tested after arriving in fs. the battery failure issue was able to be reproduced. both red alarms indicating battery failure and yellow alarm for battery communication failure were present when booting up the console. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to battery cell imbalance.
31028 | 1220648-2024-06743 | 18686751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686751&pc=OZD | 02/03/2024 | 03/15/2024 |  | 02/12/2024 |  |  | 02/03/2024 | 02/12/2024 |  |  | 09/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410275 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed oozing at the impella access site. the oozing was resolved by pausing heparin and adding extra stitches at the access site. 1 ffp, 2 plts, and 1 prbc were given. it was further reported that the impella was removed due to concerns of hemolysis. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis and access site bleeding has been completed. the impella device was not returned for evaluation. data logs were reviewed and did not reveal any positioning alarms, however, low flows that caused suction alarms were found and remained the rest of case. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. additionally, clinical details provided were sufficient to determine the root cause of the access site bleed. the root cause of access site bleeding was most likely due to anticoagulation management. added-e1-facility name, h6 impact code 4641 d4-updated model number.
31029 | 1220648-2024-06744 | 18686827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686827&pc=OZD | 01/29/2024 | 03/07/2024 |  | 02/12/2024 |  |  | 01/29/2024 | 02/12/2024 |  |  |  |  | 02/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1213645 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported that an automated impella controller (aic) had a broken purge clip. the aic was removed from service as a result of the noted issue. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon examining, reported issue with the broken purge retainer was confirmed. purge disc retainer was completely broken off. the cause of the issue was a broken purge retainer.
31030 | 1220648-2024-06745 | 18686855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686855&pc=OZD | 02/05/2024 | 03/07/2024 |  | 02/12/2024 |  |  | 02/05/2024 | 02/12/2024 |  |  |  |  | 02/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED |  |  | US |  | 1117656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Air/Gas in Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The complainant reported that a purge disc not detected alarm appeared on the console (aic) during impella cp support. multiple attempts of de-airing the tubing and a cassette change was unable to resolve the alarm. the console was swapped and the issue returned. the line was de-aired one last time which resolved the issue. there was no patient harm. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and a hairline fracture was observed on the left side of the retainer. the root cause of this issue was a broken purge retainer. | The console (aic) was returned for evaluation but has not been evaluated. upon completion of the investigation, a supplemental report will be submitted.
31031 | 1220648-2024-06746 | 18686943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686943&pc=OZD | 01/31/2024 | 03/15/2024 |  | 02/12/2024 |  |  | 01/31/2024 | 02/12/2024 |  |  | 11/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024430633 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | The user facility reported a 94-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for an elderly patient with multiple cardiac comorbidities and tavr(transcatheter aortic valve replacement) procedure. the patient had an aortic valvuloplasty. the team observed limb ischemia. the sheath was occlusive to the limb and so the pump was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was filed. the device was not returned. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31032 | 1220648-2024-06747 | 18686957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686957&pc=OZD | 01/26/2024 | 03/12/2024 |  | 02/12/2024 |  |  | 01/26/2024 | 02/12/2024 |  |  | 07/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427250 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported a 64-year-old male patient was on impella cp support due to having a history of cardiac arrest and valvular disease. the patient coded. while on support, the patient had bleeding that was managed by withholding anticoagulants and providing blood products. additionally, the patient partially pulled out the impella, into the aorta. the impella decreased to p2 and then there was a pump stop. the impella was explanted and replaced. | [5 empty] | The investigation for the reported access site bleed and the pump stop has been completed. the impella device was not returned for evaluation. data logs were reviewed finding several "alarm 119" indicating electrical open line related to the three motor driver lines. no other issues were found. as per clinical patient pulled the pump halfway during support leading to pump stop. the cause of the pump stop issue was most likely an open electrical line due to patient movement per data log and clinical review. the cause of the access site bleeding issue was most likely patient condition related per clinical review and unknown relation to impella. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
31033 | 1220648-2024-06748 | 18686967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686967&pc=OZD | 02/05/2024 | 09/21/2024 |  | 02/12/2024 |  |  | 02/06/2024 | 02/12/2024 |  |  |  |  | 02/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1558681 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Computer Operating System Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 02/06/2024 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported an automated impella controller (aic) being used for a 69-year-old male impella 5.5 patient. while in the cath lab for a procedure, an ¿impella controller failure¿ alarm sounded and the device shut down. the device was switched over to another aic and support continued. there was no patient harm reported. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sectionsb1, b2, b5, h1, and h6. | The investigation for the reported aic "red alarm" issue has been completed. the aic was returned for evaluation. the console was tested and functioned as expected with no difficulties with communication. data logs were reviewed which revealed communication errors that persisted beyond a pmd reset. a pump stop alarm occurred and then resolved. no further communication issues were found. there was a pump stop due to pmd communication issues but the cause of the communication issues was not determined. added-d9 device return date. | The controller has not been returned for evaluation. upon completion of the investigation, a supplemental report will be submitted.
31034 | 1220648-2024-06749 | 18686994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18686994&pc=OZD | 02/05/2024 | 03/25/2024 |  | 02/12/2024 |  |  | 02/05/2024 | 02/12/2024 |  |  | 06/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390168 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported an impella cp on a 61-year-old male patient needing mechanical circulatory support. it was reported that the patient was taken to the cath lab, and the physician tried to reposition, but the impella pulled out into the aorta. the physician was unsuccessful at re-crossing the valve. shortly after there was a purge system blocked alarm. after pulling the catheter into the descending aorta, the alarm resolved, but physician was unsure of integrity of the purge system for further use. the pump was exchanged for patient safety. there was no patient harm. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visually inspected and biomaterial observed in the purge gap. the cause of the high purge pressure issue was due to biomaterial blockage in purge gap per field investigation and log review. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. the instructions for use states the following for high purge pressure: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31035 | 1220648-2024-06750 | 18687053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18687053&pc=PYX | 02/05/2024 | 03/12/2024 |  | 02/12/2024 |  |  | 02/05/2024 | 02/12/2024 |  |  | 08/31/2024 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023253355 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported that a 69-year-old female patient had an impella rp pump placed to support an mvr the day prior. an echo was requested for possible clot, prior to rp insertion and was denied. the right ventricle (rv) was found to be akinetic. the rp was placed and supported until suction occurred with multiple pump stops. the team was unable to restart the pump successfully, therefore the pump was wired out and removed. their suggestion was to explant this rp and place a new rp, however the team did not have another un-expired rp pump available to replace it. | [5 empty] | The impella pump has been returned for evaluation but has not yet been evaluated. upon completion of the investigation, a supplemental report will be submitted. | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. upon investigation of the returned pump, a clot was observed around the bottom of the impeller and the purge gap. however, the issue did was not reproduced during testing. the root cause of the pump stop was determined to be biomaterial.
31036 | 1220648-2024-06751 | 18687128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18687128&pc=OZD | 01/27/2024 | 03/01/2024 |  | 02/12/2024 |  |  | 01/27/2024 | 02/12/2024 |  |  |  |  | 02/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1128491 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the clinician was changing the cassette when accidentally slightly bent the disc holder on the automated impella controller (aic). the aic was still able to detect the disc and continued with prompting. the clinician was advised to have another aic in case purge alarms occurred. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge issue ¿ other has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge retainer was broken. the cause of the issue was a broken purge retainer.
31037 | 1220648-2024-06752 | 18687803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18687803&pc=OZD | 01/27/2024 | 03/18/2024 |  | 02/12/2024 |  |  | 01/27/2024 | 02/12/2024 |  |  | 10/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414299 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the impella cp failed delivery despite multiple attempts and troubleshooting measures. the native anatomy of a heavily calcified aortic arch precluded placement. the team performed percutaneous coronary intervention without the support. no harm or injury was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition, patient had heavily calcified aortic arch which caused the difficulty of inserting the pump. d4-updated model number | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31038 | 1220648-2024-06753 | 18688201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18688201&pc=OZD | 01/28/2024 | 03/19/2024 |  | 02/12/2024 |  |  | 01/28/2024 | 02/12/2024 |  |  | 11/30/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024426161 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2024 |  |  |  |  | The user facility reported a 60-year-old patient of unknown race and gender was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the physician was unable to pass the pump through the patient¿s vasculature and the implant was aborted. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the delivery issue could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31039 | 1220648-2024-06754 | 18689638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18689638&pc=OZD | 01/29/2024 | 09/17/2025 |  | 02/12/2024 |  |  | 01/29/2024 | 02/12/2024 |  |  | 07/13/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375078 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant reported the patient was on impella 5.5 support due to having cardiogenic shock and the patient was waiting for a heart transplant. during support, there was impella positions unknown alarms. the impella position was confirmed on echocardiogram (echo). additionally, the sterile sleeve became separated from the white touhy borst. | The customer also reported an optical sensor issue. no further information available. |  |  |  |  | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (sterile sleeve damage) and the optical signal issues has been completed since the original report was filed. the device was not returned. the root cause of the sterile sleeve damage and the optical signal issue was not determined since product was not returned and insufficient clinical information was provided. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. the investigation optical signal issue has been completed. product was not returned for investigation. data logs were investigated. impella position alarms were observed throughout the case. no placement signal low or not reliable alarms were seen. no other issues were observed in the pump. the root cause of the optical signal issue was not determined since product was not returned and insufficient clinical information was provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. g1 updated with correct mdr reporting contact email . d4 model number updated. d6a/d6b updated with additional information. f6 and f8 should have been left blank. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-06754.
31040 | 1220648-2024-06755 | 18690199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18690199&pc=OZD | 01/26/2024 | 10/14/2024 |  | 02/12/2024 |  |  |  | 02/12/2024 |  |  |  |  | 02/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | US |  | 1119399 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Volume Accuracy Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an automated impella controller (aic) began experiencing alarms without any sound. despite multiple alarms being displayed, no sound was audible from the aic. there was no patient harm resulting from the failure. | [5 empty] | The investigation of automated impella controller (aic) - buzzer/ alarm malfunction has been completed. based in data analysis and device analysis, the root cause of the complaint was identified as an unplugged speaker cable. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06755. d2 common device name,d4 model number, f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was received and therefore, an evaluation of the device was is underway. upon investigation closure, a supplemental mdr will be filed.
31041 | 1220648-2024-06756 | 18691621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18691621&pc=OZD | 01/29/2024 | 03/13/2024 |  | 02/12/2024 |  |  |  | 02/12/2024 |  |  | 07/31/2025 |  | 02/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024390713 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support due to stage b cardiogenic shock. while on support, there was bleeding noted from the access site. hemoglobin and platelets were trending down. heparin was stopped and two units of red blood cells were provided to the patient. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding was most likely use issue due to anticoagulation management since hemostasis was achieved by stopping heparin. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿assess access site for bleeding and hematoma.¿
31042 | 1220648-2024-06722 | 18682729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682729&pc=OZD | 09/08/2023 | 03/01/2024 |  | 02/10/2024 |  |  | 01/12/2024 | 02/10/2024 |  |  | 04/30/2025 |  | 02/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024347847 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry, indicating a 43 year old male patient presented with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured signs of hemolysis. as a result, the pump was weaned and removed. the allegation indicates a "definite" relationship to the impella device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported hemolysis has been completed since the original report was filed. data logs showed about 1 hour into the case, pump was in improper position that triggered alarms impella position in ventricle & in aorta. this event occurred for 7 hours & 45 minutes then was resolved & pump ran without issue to the rest of the case. product was not returned for review. based on clinical description & data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position.
31043 | 1220648-2024-06723 | 18682731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682731&pc=OZD | 09/06/2023 | 03/01/2024 |  | 02/10/2024 |  |  | 01/12/2024 | 02/10/2024 |  |  | 06/30/2025 |  | 02/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024379621 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry, indicating this patient to have endured a complete occlusion of the femoral artery with a "probable" relationship to the impella device. as a result, intervention via armada balloon dilation was performed and evidence of a thrombus was identified. treatment via iv of heparin was to be administered. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31044 | 1220648-2024-06724 | 18682735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682735&pc=OZD | 07/29/2023 | 03/05/2024 |  | 02/10/2024 |  |  | 01/12/2024 | 02/10/2024 |  |  | 04/30/2025 |  | 02/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350396 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry, regarding a 78 year old female patient presenting for high risk percutaneous coronary intervention using the impella for hemodynamic support. during support, the patient endured an occlusion of blood flow in the left femoral artery. the allegation was made with a "possible" relationship to the impella device. intervention was discussed, however, the family ultimately chose to withdraw the patient from further care. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added h6 type of investigation:4114 d4-model number changed to impella cp.
31045 | 1220648-2024-06682 | 18675281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18675281&pc=OZD | 09/16/2022 | 02/26/2024 |  | 02/09/2024 |  |  | 09/16/2022 | 02/09/2024 |  |  | 03/31/2024 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023199298 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 60-year-old female patient, the impella was noticed to be in the aorta, and was unable to pass the aortic velocity (aov). the pump was therefore removed. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the positioning issues has been completed. the device was not returned for investigation. according to clinical records, the pump was removed because resistance was encountered while crossing the aortic valve. the cause of the delivery issue was most likely patient condition related.
31046 | 1220648-2024-06683 | 18675354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18675354&pc=OZD | 09/19/2023 | 02/09/2024 |  | 02/09/2024 |  |  | 09/19/2023 | 02/09/2024 |  |  | 07/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024373912 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/19/2023 |  |  |  |  | The user facility reported 52-year-old female patient with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 was not able to be inserted as the patient had severe stenosis in the subclavian artery. the insertion for the 5.5 was aborted and impella cp was placed successfully. no patient harm was reported. | [5 empty] | No product was returned. the most likely root cause of the delivery issue was patient condition related due to severe stenosis in the subclavian artery.
31047 | 1220648-2024-06684 | 18676293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18676293&pc=OZD | 01/25/2024 | 03/06/2024 |  | 02/09/2024 |  |  | 01/25/2024 | 02/09/2024 |  |  | 10/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418772 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported an impella 5.5 was implanted in a 50-year-old female patient for cardiomyopathy. it was reported that a ¿high purge pressure/purge system blocked¿ error message occurred during patient support. the tissue plasminogen activator (tpa) protocol was initiated. the high purge pressure alarms continued and argatroban purge was initiated the following day. support remained ongoing. however, the staff was made aware that the current purge was off label. there was no patient harm. | [5 empty] | The investigation for the reported high purge pressure has been completed. the impella device was not returned for evaluation. however, data logs were reviewed which showed a gradual increase in purge pressure triggering the high purge pressure alarm which was troubleshooted per the tissue plasminogen activator (tpa) protocol. rise in pump flows at p-7 were observed during the high purge pressure event. after about 8 hours the purge pressure gradually decreased with likely tpa addition. this suggests that the block was most likely biomaterial. the root cause of the high purge pressure issue was most likely biomaterial per clinical and data log review. the instructions for use states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted.
31048 | 1220648-2024-06685 | 18676445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18676445&pc=OZD | 01/25/2024 | 03/07/2024 |  | 02/09/2024 |  |  | 01/25/2024 | 02/09/2024 |  |  | 11/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432249 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The complainant reported that the impella cp was implanted in a 37-year-old male patient for mechanical circulatory support. it was reported that the patient experienced suction and the impella pigtail was wrapped in the pap muscle on the transthoracic echocardiography (tte). the impella was removed. the patient is stable. | [5 empty] | Impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was improper positioning due to improper technique since the pigtail was found caught on the pap muscle at the time of mr.
31049 | 1220648-2024-06686 | 18676648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18676648&pc=OZD | 01/26/2024 | 02/23/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 10/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024420112 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had an impella cp placed for support of a st elevation myocardial infarction (stemi) patient. the patient was transferred from the community to a tertiary center for care and monitoring. the pump was explanted after over eight hours due to limb ischemia. the ischemic leg was "dead" and so perfusion was placed and venoarterial extracorporeal membrane oxygenation (va ecmo). the pump was explanted with plan of impella 5.5 care escalation. the limb was treated beyond the explant by cutdown and thrombectomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31050 | 1220648-2024-06687 | 18676885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18676885&pc=OZD | 09/09/2022 | 02/24/2025 |  | 02/09/2024 |  |  | 09/10/2022 | 02/09/2024 |  |  | 05/31/2024 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229019 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2022 |  |  |  |  | The user facility reported a 43-year-old female with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the vessel size was noted to be too small despite using dilators to assess ability to accommodate the impella 5.5. it was decided that an impella cp was more appropriate, and the implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of delivery issues has been completed. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a and d6b revised from the initial submission in accordance with updated procedures on manufacturer device report 1220648-2024-06687. f6/f8-should have been left blank on manufacturer device report 1220648-2024-06687. g1 reporting contact email revised on manufacturer device report 1220648-2024-06687 in accordance with updated procedures. g1 reporting contact fax number missing on manufacturer device report 1220648-2024-06687. h10 instructions for use revised from the initial submission in accordance with updated procedures on manufacturer device report 1220648-2024-06687.
31051 | 1220648-2024-06688 | 18676984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18676984&pc=OZD | 10/15/2023 | 02/09/2024 |  | 02/09/2024 |  |  | 10/15/2023 | 02/09/2024 |  |  | 07/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024380202 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/23/2023 |  |  |  |  | The user facility reported a 43-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant reported the impella pump motor current high alarm x3, unable to restart 3rd time impella removed. | [5 empty] | The impella device was received from the customer on 31-oct-2023 data review: data logs showed pump started & run without issue. about 292 hours into support, pump stopped that triggered alarm impella stopped-mc high. pump could be restarted & run for 2.5 hours more then stopped again. next, pump was stopped manually & disconnected from the console. product analysis: analysis of the pump revealed a large purge gap indicating bearing wear as the root cause of the pump stop. no other issues were found on the rest of the pump. leak test was performed on pc gen2 & no leak was found on the pc. conclusion: the root cause of the pump stop was a wearing bear. this pump stop occurred after 12 days of run time, which is well beyond the design life of 8 days per dtm 0046-9910. the root cause of the pc leak was most likely use related because the failure mode could not be reproduced. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31052 | 1220648-2024-06689 | 18677027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18677027&pc=OZD | 01/26/2024 | 03/14/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 06/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024431214 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 50-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a left ventricle thrombus and lower leg ischemia during impella support. the thrombus was discovered via echocardiogram and the ischemia coincided with symptoms of pain and a cold feeling. due to the noted conditions, the decision was made to explant the pump. an intra-aortic balloon pump (iabp) was subsequently placed for ongoing support. | [5 empty] | The investigation into the limb ischemia ¿ reversible and the thrombus issues has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury"
31053 | 1220648-2024-06690 | 18677057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18677057&pc=OZD | 11/28/2023 | 11/26/2024 |  | 02/09/2024 |  |  | 11/28/2023 | 02/09/2024 |  |  | 06/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374852 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Integrity Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella procedure when inserting the impella long peel away sheath, the sheath kinked and did not allow the device to pass through after removing the dilator. there was a need for the long sheath because of the patient¿s disease, a new introducer kit was opened. a second long sheath was used without issue. there was no patient harm. | [5 empty] | The investigation for the introducer damage has been completed. as per the clinical information provided, the peel away sheath kinked due to calcified vessels. the 2nd sheath was used and was successfully used. therefore, the root cause of the introducer damage issue was patient condition related to patients calcified vessels. added h6 component code 463 corrections to the initial mfr report: d4 model number added d6a/d6b f6/f8 should have been left blank. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31054 | 1220648-2024-06691 | 18677245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18677245&pc=OZD | 08/30/2023 | 09/25/2025 |  | 02/09/2024 |  |  | 08/30/2023 | 02/09/2024 |  |  | 05/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024356159 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Death | Other |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/29/2023 |  |  |  |  | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 45-year-old female patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient experienced a large right side ischemic stroke. the patient had been unresponsive during her stay in the intensive care unit (icu). an electroencephalogram (eeg) and repeat computed axial tomography scan confirmed poor prognosis and recovery. | The complainant also reported that a placement signal not reliable alarm occurred. the staff were walked through disabling the alarm and monitoring the placement. |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation has been completed. the product and data were not returned for review. the root cause of cerebral hemorrhage was most likely due to patient condition and unrelated to the impella. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. product and data logs were not returned for review. the root cause of the optical signal issue was unable to be determined as limited clinical details were provided and no product was returned for review. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. corrections that were made from the initial manufacturer device report number 1220648-2024-06691. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 have erroneous entries on the initial report - should have been left blank. g1 reporting contact email and reporting contact fax number updated.
31055 | 1220648-2024-06692 | 18677468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18677468&pc=OZD | 01/26/2024 | 03/14/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 04/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330385 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported 64-year-old male patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient was taken to the operation room for explant and blood from the insertion wound was noted when opening the pocket and cultures were obtained. impella was successfully removed, and hemostasis achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
31056 | 1220648-2024-06693 | 18677708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18677708&pc=OZD | 10/09/2023 | 02/09/2024 |  | 02/09/2024 |  |  | 10/09/2023 | 02/09/2024 |  |  | 09/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393599 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  |  | Malfunction | No | Yes | Inaccurate Delivery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella 5.5 insertion was aborted due to the 81-year-old female patient's anatomy. the impella 5.5 was unable to be advanced. it was noted that there was a high degree of calcification of all arteries, and it was impossible to use extracorporeal membrane oxygenation (ecmo). there was no reported patient harm or injury. | [5 empty] | The investigation was completed. no product was returned. the root cause of the delivery anatomy issue was most likely patient condition related.
31057 | 1220648-2024-06694 | 18677743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18677743&pc=OZD | 01/28/2024 | 03/12/2024 |  | 02/09/2024 |  |  | 01/28/2024 | 02/09/2024 |  |  | 10/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024424071 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2024 |  |  |  |  | The user facility reported 45-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the physician had to suture the sheath again to improve ooze and removed fem-stop that was replaced overnight. the patient was reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that since the doctor had to res-uture the access site. therefore, the root cause of the access site bleed was most likely use related due to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643 apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31058 | 1220648-2024-06695 | 18677874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18677874&pc=OZD | 01/26/2024 | 03/06/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 10/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411237 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 44-year-old female patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella was pulled back across the valve by accident during heart failure rounds and echocardiogram assessment. an attempt was made to advance back across the valve but was unsuccessful. the impella was turned to p-2 and the patient was monitored for 20 to 30 minutes. the patient remained stable hemodynamically. the impella was explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however clinical information was sufficient to determine the root cause. the cause of the positioning issue was use related due to improper technique because the impella was pulled back across the valve by accident during heart failure rounds and echo assessment.
31059 | 1220648-2024-06696 | 18678003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18678003&pc=OZD | 01/30/2024 | 03/13/2024 |  | 02/09/2024 |  |  | 01/30/2024 | 02/09/2024 |  |  | 11/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428651 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 54-year-old female patient needing mechanical circulatory support. it was reported that resistance was met during attempted implant of an impella 5.5 pump. following right axillary insertion, it was noted that the pump would not advance through the mid subclavian artery. due to the resistance, the decision was made to remove the impella 5.5 implant. an impella cp pump was subsequently inserted for planned support. there was no patient harm. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. | The investigation into the delivery issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue - anatomy was most likely due to patient condition.
31060 | 1220648-2024-06697 | 18678025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18678025&pc=OZD | 01/26/2024 | 03/12/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 10/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024419958 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 35-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the impella limb was ischemic, and perfusion was placed. extracorporeal membrane oxygenation (ecmo) access was suffering from compartment syndrome. the team performed a below knee amputation (bka) on the ecmo limb/right leg. the team considered a left bka as well, but the care was withdrawn after five days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
31061 | 1220648-2024-06698 | 18678097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18678097&pc=OZD | 02/01/2024 | 03/18/2024 |  | 02/09/2024 |  |  | 02/01/2024 | 02/09/2024 |  |  | 08/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397786 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The complainant reported that the 54-year-old male patient using the impella cp lost all pulses in their right leg and a vascular duplex showed no distal flow during impella cp support. due to the noted condition, the decision was made to wean and explant the pump. | [5 empty] | The impella cp device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was filed. the device was not returned. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31062 | 1220648-2024-06699 | 18678127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18678127&pc=OZD | 01/26/2024 | 03/12/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 10/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415499 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 62-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the sterile sleeve was pulled out of the back lock. chlora prep was used to clean the catheter and tegaderm was placed around the sleeve for reinforcement. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue for the sterile sleeve has been completed. the pump was returned for evaluation. the pump was visually inspected and no physical abnormalities or damages observed. the sterile sleeve looked to be good and intact without any ripping/tearing of the sleeve. no other issues were found. the cause of the sterile sleeve damage issue was use related with no issue observed during field investigation. added- d9 device return date, h6 impact code 4641.
31063 | 1220648-2024-06700 | 18678222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18678222&pc=OZD | 02/02/2024 | 03/21/2024 |  | 02/09/2024 |  |  | 02/02/2024 | 02/09/2024 |  |  | 11/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434179 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | The user facility reported an impella cp in a 82-year-old male patient for high risk percutaneous cardiac intervention. the patient experienced bleeding complications from their access site. the patient was also noted to have a recent history of a gastrointestinal bleed since admission. due to the reported bleed, heparin was turned off and the patient was transfused two units of blood (prbcs) and one unit of platelets. as heparin was re-introduced, the gi bleed remained ongoing and a unit of prbc was being given daily. support was maintained throughout the event. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use states that the user should ¿assess access site for bleeding and hematoma.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the root cause of the access site bleeding was unable to be determined as the product was not returned and insufficient clinical details were provided.
31064 | 1220648-2024-06701 | 18678265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18678265&pc=PYX | 02/03/2024 | 09/21/2024 |  | 02/09/2024 |  |  | 02/03/2024 | 02/09/2024 |  |  | 07/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024399899 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | Thrombosis/Thrombus | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/03/2024 |  |  |  |  | Additional information noted care was withdrawn and the patient expired.medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported an impella rp in a 64-year-old male patient needed for high risk surgery. it was reported that the patient developed cardiac tamponade requiring a third pressor to stabilize. large burden of clots were evacuated. additionally, the patient developed severe bleeding and heparin was withheld. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation into the access site bleeding ¿ major and the thrombus issues has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding ¿ major and the thrombus issues was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
31065 | 1220648-2024-06702 | 18678434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18678434&pc=OZD | 02/05/2024 | 03/07/2024 |  | 02/09/2024 |  |  | 02/05/2024 | 02/09/2024 |  |  |  |  | 02/09/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1134173 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Incomplete or Inadequate Connection | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The complainant reported that during inservice while walking through steps for preparation, the automated impella controller (aic) would not recognize the purge cassette disc. another purge cassette was attempted with no success. there was no patient harm reported. | [5 empty] | The aic device has been returned for evaluation but has not yet been evaluated. upon completion of the investigation, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag.
31066 | 1220648-2024-06703 | 18678759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18678759&pc=OZD | 01/26/2024 | 04/02/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 11/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024428660 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Anemia | Hemolysis | Ischemia | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 01/26/2024 |  |  |  |  | It was additionally reported that the patient suffered from post-operative anemia and thrombocytopenia following impella explant. additional information pertaining to the condition and potential intervention was not provided. | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella cp for mechanical circulatory support. the complainant reported that a patient experienced multiple complications during impella cp support. upon initial placement of the repositioning sheath, it was deemed that the sheath was occlusive. an external fem-fem bypass was subsequently placed to provide distal perfusion. an echocardiogram was then performed, and it was found that the device was under the papillary muscle. attempts to reposition the device were unsuccessful resulting in low pump flows. the patient then began to have bloody urine and concerns for hemolysis were noted. due to the reported issues, the decision was made to explant the pump. the patient was successfully weaned. |  |  |  |  | The investigation into the limb ischemia ¿ reversible, the hemolysis and the positioning issues has been completed since the original report was filed. the device was not returned for investigation. based on the clinical information provided, the root cause of hemolysis was most likely due to improper positioning since the pump was caught under pap muscle and was unable to reposition. the root cause of positioning issue was most likely due to use related. and the root cause of the limb ischemia was not determined as the necessary clinical information was not provided. | B5 and h6 have been updated to account for the report of thrombocytopenia and post operative anemia suffered by the patient following impella explant. | The investigation into the limb ischemia ¿ reversible, and the hemolysis, the positioning issues, and the thrombocytopenia issues has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of hemolysis was most likely due to improper positioning since the pump was caught under pap muscle and was unable to reposition. the root cause of positioning issue was most likely due to use related. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the thrombocytopenia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31067 | 1220648-2024-06704 | 18679885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18679885&pc=OZD | 01/26/2024 | 09/05/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 09/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399835 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 68-year-old male was implanted with an impella 5.5 as a bridge to transplant. it was reported upon starting the pump that a placement signal not reliable alarm occurred. there was good pump performance level but no waveforms on the display screen. all connections were checked. the physician made the decision to replace the pump. a new pump was placed using fluoroscopy and transesophageal echocardiography (tee) guidance and was slowly ramped up to p8. there was no patient harm reported. | [5 empty] | Upon further review there is currently no allegation of a malfunction or serious injury associated with the implanted pump. it was confirmed that the issue is with the automated impella controller (aic) only. please see manufacturer device report 1220648-2024-06623 regarding information pertaining to the aic. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31068 | 1220648-2024-06705 | 18680165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18680165&pc=OZD | 01/26/2024 | 03/11/2024 |  | 02/09/2024 |  |  | 01/26/2024 | 02/09/2024 |  |  | 06/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359758 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 47-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that a transthoracic echocardiogram was performed and revealed an apical thrombus. there was no interaction between the thrombus and the impella and support was continued. additionally, the patient had clotted, red urine. the impella was in proper position. continuous renal replacement therapy was started during support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis and the thrombus issues has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the hemolysis and the thrombus was not determined.
31069 | 1220648-2024-06706 | 18680963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18680963&pc=OZD | 01/27/2024 | 03/18/2024 |  | 02/09/2024 |  |  | 01/27/2024 | 02/09/2024 |  |  | 11/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024429943 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp was not able to pass through the introducer. a new impella cp was placed successfully. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. the root cause of the delivery issue was an introducer sheath kink.
31070 | 1220648-2024-06707 | 18681415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681415&pc=OZD | 01/28/2024 | 03/11/2024 |  | 02/09/2024 |  |  | 01/28/2024 | 02/09/2024 |  |  | 08/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024398285 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2024 |  |  |  |  | The user facility reported 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was explanted after 8 hours due to hemolysis with abnormally high plasma free hemoglobin. the patient was escalated to a 5.5 pump. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed pump position was in ventricle that triggered impella position in ventricle & suction alarms. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.
31071 | 1220648-2024-06708 | 18681460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681460&pc=OZD | 01/27/2024 | 03/08/2024 |  | 02/09/2024 |  |  | 01/27/2024 | 02/09/2024 |  |  | 11/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435412 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | Additional information provided states that the impella device was removed and the patient subsequently developed a hematoma and compartment syndrome. a fasciotomy was performed to right lower extremity. | The user facility reported a 36-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the us complainant had placed via the femoral the impella cp for a patient transferred to the tertiary center from the community. the limb became ischemic. there was no pulse and the color was poor. the team discussed a fem-fem bypass but, as the patient was clamped down the team did not place. the life/limb conversation is ongoing and the limb remains ischemic on impella support and there was discussion to possibly escalate to the impella 5.5 pump. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31072 | 1220648-2024-06709 | 18681585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681585&pc=OZD | 01/30/2024 | 09/20/2024 |  | 02/09/2024 |  |  | 01/30/2024 | 02/09/2024 |  |  | 10/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422375 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/30/2024 |  |  |  |  | The user facility reported an impella cp on a 59-year-old male patient with acute myocardial infarction. the patient was taken to the cath lab for emergent impella placement while performing cpr. however, during placement of the pump, a clot ingestion was suspected and the pump had saturated flows and high motor currents. the patient ultimately expired in the procedure area. the cause of death was not attributed to the use of impella. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. however, a supplemental report will be submitted upon completion of the investigation. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06709 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The investigation into the saturated flow issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the saturated flow was most likely due to biomaterial based off clinical details noting no heparin was administered and lv thrombus was observed, along with pump experiencing high motor current.
31073 | 1220648-2024-06710 | 18681608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681608&pc=OZD | 02/01/2024 | 03/18/2024 |  | 02/09/2024 |  |  | 02/01/2024 | 02/09/2024 |  |  | 01/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2024338682 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported an impella cp in a 47-year-old male patient for acute myocardial infarction and cardiogenic shock. it was reported that after placement, the impella was found to be positioned incorrectly against the mitral and the device could not be repositioned and was replaced. the team successfully placed a second impella in the patient which supported the patient until care was withdrawn. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the device was not returned. the cause of the positioning issue was use related due to cpr being done per clinical review. data logs were reviewed and suction alarms, low pump flow observed with placement signal trend. the cause of the low pump flow issue was most likely patient condition related per clinical and data log review. | The impella device was not returned foe evaluation. however, upon completion of the investigation a supplemental report will be submitted.
31074 | 1220648-2024-06711 | 18681623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681623&pc=OZD | 02/02/2024 | 11/20/2025 |  | 02/09/2024 |  |  | 02/02/2024 | 02/09/2024 |  |  |  |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040 |  | No |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | The ventricular assist device (vad) engineers reported that the console (aic) was showing ¿system self-check failure: change console immediately¿ when powered on. the vad engineers alerted team and used a different console for case. there was no patient harm reported. | [5 empty] | The console (aic) was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted upon completion.
31075 | 1220648-2024-06712 | 18681636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681636&pc=OZD | 02/04/2024 | 09/26/2024 |  | 02/09/2024 |  |  | 02/04/2024 | 02/09/2024 |  |  | 11/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2024434915 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/04/2024 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported an impella cp in a 74yo male patient for acute myocardial infarction and cardiogenic shock. it was reported that two days following the implant of the impella cp, persistent suction alarms were being triggered. upon further inspection, it was discovered that a thrombus was present on the inflow of the catheter. due to the noted condition and resulting suction, the decision was made to explant the pump the following day. no harm or injury was reported. |  |  |  |  | The investigation into the thrombus issue has been completed since the initial report was submitted. the device was not returned for investigation. according to the clinical details, the doctor found a thrombus on the pump inlet during continuous suction alarm. the cause of the low pump flow issue was most likely biomaterial, based on provided clinical derails. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06712 was provided in sections b5, d6a, d6b, and h1. | The impella device was not returned for evaluation. however, upon completion of the investigation a supplemental report will be submitted.
31076 | 1220648-2024-06713 | 18681640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681640&pc=OZD | 02/06/2024 | 03/07/2024 |  | 02/09/2024 |  |  | 02/06/2024 | 02/09/2024 |  |  |  |  | 02/09/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024377018 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/06/2024 |  |  |  |  | The complainant reported that the blue retaining clip broke off the controller (aic) during a purge cassette change. the aic was exchanged. there was no patient harm reported. | [5 empty] | The aic has not been returned for evaluation. however, a supplemental report will be submitted upon completion of the investigation. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon examining, reported issue with the broken purge blue retainer was confirmed. the retainer was completely broken off. the cause of the issue was a broken purge retainer.
31077 | 1220648-2024-06714 | 18681790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681790&pc=PYX | 10/21/2023 | 09/22/2025 |  | 02/09/2024 |  |  | 01/10/2024 | 02/09/2024 |  |  | 03/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024394951 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 76 -year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification via abiomed¿s long-term outcome and quality indicator impella registry (loqi) indicating this patient to have endured hemolysis and subsequent acute renal failure that led to intervention via hemodialysis. a "possible" relationship to the impella device was alleged. | The complainant reported a placement signal not reliable alarm during support. the device was not removed and support proceeded with no patient harm. on 2024-01-10, we received a notification via abiomed¿s long-term outcome and quality indicator impella registry (loqi) indicating this patient to have endured hemolysis and subsequent acute renal failure that led to intervention via hemodialysis. a "possible" relationship to the impella device was alleged. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis was most likely biomaterial related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06714. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for investigation. about 41 hours into the case, there is a motor current spike with a placement signal shift and speed deviation independent of p-level change. this indicates a potential biomaterial ingestion. the root cause of the hemolysis was most likely biomaterial related. data logs confirm dp sensor drift 88 hours into the case. placement signal spikes to 3000 mmhg and flow metrics are lost. the root cause of the placement signal issue was determined to be dp sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
31078 | 1220648-2024-06715 | 18681800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681800&pc=OZD | 05/17/2023 | 10/04/2024 |  | 02/09/2024 |  |  | 01/10/2024 | 02/09/2024 |  |  | 12/31/2024 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024305885 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that nonsustained ventricular tachycardia (nsvt) occurred with a possible relationship to the impella 5.5 device used on this patient. | [5 empty] | The investigation for vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06715: d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31079 | 1220648-2024-06716 | 18681815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681815&pc=OZD | 06/09/2023 | 04/23/2024 |  | 02/09/2024 |  |  | 01/10/2024 | 02/09/2024 |  |  | 12/31/2024 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024321746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Thrombocytopenia | Heart Failure/Congestive Heart Failure | [10 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/09/2024 |  |  |  |  | Additional information received that patient experienced significant hypotension and congestive heart failure and was hospitalized to address the issue. | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining a 71-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was noted that worsening thrombocytopenia occurred with a possible relationship to the impella cp device used on this patient. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombocytopenia issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the thrombocytopenia was not determined since there is a lack of clinical details and product was not returned.
31080 | 1220648-2024-06717 | 18681836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681836&pc=PYX | 01/10/2024 | 03/11/2025 |  | 02/09/2024 |  |  | 01/10/2024 | 02/09/2024 |  |  | 12/31/2024 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024289458 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | Abiomed¿s long-term outcome and quality indicator impella registry (loqi) received a notification regarding a "possible" relationship between the impella device and worsening thrombocytopenia. no further details were provided. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The root cause of the thrombocytopenia cannot be determined since the product was not returned and insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
31081 | 1220648-2024-06718 | 18681961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18681961&pc=OZD | 05/04/2023 | 03/11/2025 |  | 02/09/2024 |  |  | 01/10/2024 | 02/09/2024 |  |  | 12/31/2023 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023172850 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry, pertaining to the above case. it was noted that a 57 year old male patient presented with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured hemolysis. in the subsequent 48 hours, the patient developed stage 3 aki, "likely from extensive hemolysis/rhabdomyolysis with decreased uop udall placed. hgb values baseline 17.6 , nadir 9.5 and no resolution" | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | No product returned. data logs show no mc spikes outside p level changes and no relevant alarms. the cause of the hemolysis was not determined as product was not returned, lack of clinical details, and inconclusive data logs. this report is being filed as part of a retrospective review of historical records.
31082 | 1220648-2024-06719 | 18682211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682211&pc=OZD | 07/19/2023 | 03/11/2024 |  | 02/09/2024 |  |  | 01/11/2024 | 02/09/2024 |  |  | 04/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024349739 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry, pertaining to a 72 year old male patient presenting with acute myocardial infarction and cardiogenic shock. during support, a vascular complication that occurred with a probable relationship to the impella cp device used on this patient: ""after removal of the impella through the sheath. subsequent angiogram showed reduced flow in sfa which was followed with 4.0x40mm armada balloon dilation. then angiogram performed of lfa through armada balloon catheter showed improvement in flow. " on (b)(6) 2023, it was noted that, "impella removal and trialysis catheter placement c/b decrease sfa flow s/p lfa balloon angioplasty with restoration of flow." | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31083 | 1220648-2024-06720 | 18682371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682371&pc=OZD | 08/13/2023 | 03/03/2025 |  | 02/09/2024 |  |  | 01/11/2024 | 02/09/2024 |  |  | 03/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024347253 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | Other | Hospitalization | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 01/11/2024 |  |  |  |  | Additional information was received via via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured hemolysis with a "definite" relationship to the impella device. the patient's hospitalization was prolonged, as a result. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 53 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support endured thrombocytopenia with a "probable" relationship to the impella device: "the patient developed thrombocytopenia. negative for heparin induced thrombocytopenia antibodies, and it is noted that the condition is likely related to the impella. baseline platelets on admission prior to impella placement was 138 k/ul, post-impella placement platelet count was 88 k/ul ((b)(6) 2023) and resolved at 192 k/ul ((b)(6) 2023). the lowest platelet count was 63 k/ul on (b)(6) 2023. | Additional information was received from abiomed¿s long-term outcome and quality indicator impella registry (loqi) that the previously reported hemolysis is "probably related" to use of the impella device. |  |  |  | New information added to b(5) regarding hemolysis. the investigation has been reopened. a supplemental mdr will be filed upon completion. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombocytopenia and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical details, the patient experienced thrombocytopenia and hemolysis during pump use. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the thrombocytopenia issue was not determined since product was not returned. d6a, d6b. | The investigation into the thrombocytopenia issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the thrombocytopenia could not be determined because there was insufficient information to make a clinical determination. | Additional information: b6 lab results corrections to the initial mfr report: b5-changed relationship of device to the adverse event hemolysis to "probably related".
31084 | 1220648-2024-06721 | 18682385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682385&pc=OZD | 07/29/2023 | 03/05/2024 |  | 02/09/2024 |  |  | 01/12/2024 | 02/09/2024 |  |  | 03/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024346574 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 55 year old male patient presented with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient reportedly endured hemorrhagic shock with a "definite" relationship to the impella device: "impella complicated by rp bleed: removed impella (b)(6) rpb causing hemorrhagic shock. ctap (b)(6) - moderate to large right retroperitoneal hemorrhage with expansion of right psoas muscle. hemorrhage seen extending into pelvis. hgb 15 on (b)(6), dropped to 7.6 on 7.25 lactic acid 2.8 on arrival to icu hfh, 1.4 on arrival to i5 repeat ctap (b)(6) evening stable rpb repeat ctap (b)(6) show stable rpb that is improving hgb 8.0 today" | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding was not determined since product was not returned.
31085 | 1220648-2024-06561 | 18682388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682388&pc=OZD | 01/22/2024 | 03/04/2024 |  | 02/09/2024 |  |  | 01/22/2024 | 02/09/2024 |  |  | 03/31/2024 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023209571 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 55-year-old male male of unknown race in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported the patient lost pulse in the right arm. inotropes were increased despite being on maximum support from impella. the impella was removed the same day due to the pulse issue and the physician decided to use venoarterial extracorporeal membrane oxygenation (va ecmo). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the access site bleeding ¿ minor and the limb ischemia ¿ reversible issues has been completed. the device was not returned for investigation. the cause of the bleeding issue was most likely patient movement. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31086 | 1220648-2024-06631 | 18682389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682389&pc=OZD | 01/19/2024 | 09/26/2024 |  | 02/09/2024 |  |  | 01/25/2024 | 02/09/2024 |  |  | 10/31/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024414201 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/25/2024 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after 5 days the patient expired from underlying medical causes. there was no pump malfunction or harm/injury/ae attributed to the use of the pump. approximately a week after the explant the team was notified of a cerebral vascular accident (cva) occurring during the support that could be related to the use of impella. the cat scan revealed either a cerebral hemorrhage or infarction had occurred. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06631 was provided in sections b2, g3, and h1, and h6.
31087 | 1220648-2024-06632 | 18682390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18682390&pc=OZD | 01/25/2024 | 03/05/2024 |  | 02/09/2024 |  |  | 01/25/2024 | 02/09/2024 |  |  | 06/30/2025 |  | 02/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024372917 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported an impella cp in a 52-year-old male patient needing mechanical circulatory support. it was reported that during an attempt to reposition the impella cp the cannula bent in several locations after it was pulled back in the aorta. cp was untangled via careful manipulation of the device. the impella cp was removed and replaced with a new device. there was no patient harm. | [5 empty] | The device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. | The investigation for the reported positioning issue has been completed. the impella device was returned for evaluation. visual inspection was performed which found a kink on can about 13mm from outflow cage and can bonding site. no other issues were found on the rest of the pump. the device evaluation and clinical details provided were sufficient to determine the root cause of the reported positioning issue. the root cause of the positioning issue related to improper technique and use of the device outside of product claims, wireless re-access was attempted by the user per clinical details. the impella device is not indicated for wireless re-access.
31088 | 1220648-2024-06628 | 18667372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18667372&pc=OZD | 01/25/2024 | 02/28/2024 |  | 02/08/2024 |  |  | 01/25/2024 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1250419 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The complainant reported that an automated impella controller (aic) powered on with a blinking airplane button that could not be resolved. the power was initially cycled with no change in status. when the device failed to enter maintenance mode, a factory reset was attempted, but also failed to engage. the device was subsequently not used for support. there was no patient harm resulting from the malfunction as no patient was involved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported aic - impella connect - usd card corruption has been completed since the original report was filed. device was returned for investigation. upon boot up, the uboot logs show the rlm had a partswitch entry for partition rp1 ((b)(6)), the console rebooted into a partswitch on rp2 ((b)(6)). "partswitch" indicates that the indicated partition is corrupt and the rlm cannot boot into it which would result in an inability to connect to impella connect as well as not entering maintenance mode. the root cause of inability to enter maintenance mode was due to a corrupt usd card. the root cause of the corrupted usd card was not determined.
31089 | 1220648-2024-06629 | 18668166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18668166&pc=OZD | 01/18/2024 | 03/05/2024 |  | 02/08/2024 |  |  | 01/18/2024 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205398 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | It was reported by a biomedical engineer that the automated impella controller (aic) had a system self-check failure. battery removal was tried to attempt to resolve the issue. the aic was removed from service and a backup will be used for patient support. there was no patient involvement. | [5 empty] | The automated impella controller (aic) was returned for evaluation but has not been evaluated. upon completion of the investigation, a supplemental report will be submitted. | The investigation into the aic: system self-check error has been completed since the original report was filed. the console was returned and tested after arriving in fs. during testing, the issue was not reproduced, and console was able to fully boot up multiple times. the cause of the issue was not determined since issue could not be reproduced.
31090 | 1220648-2024-06630 | 18668395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18668395&pc=PYX | 01/25/2024 | 09/20/2024 |  | 02/08/2024 |  |  | 01/25/2024 | 02/08/2024 |  |  | 08/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024397330 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | Adverse Event Without Identified Device or Use Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/25/2024 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 30-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that complications were encountered while attempting to implant an impella rp flex for patient support. it was noted that the pump was unable to advance due to cannula and swelling of the native heart. after multiple failed attempts, the implant was aborted and the patient succumbed to their condition. there was no allegation of association between the patient's passing and the device. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related as it was reported to placement was difficult due to the anatomy of the patient.
31091 | 1220648-2024-06633 | 18669641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18669641&pc=PYX | 01/25/2024 | 03/08/2024 |  | 02/08/2024 |  |  | 01/25/2024 | 02/08/2024 |  |  | 05/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024354674 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported a female of unknown age, ethnicity, and race was to be implanted with an impella rp device for mechanical circulatory support. it was reported that after multiple attempts the physician aborted the implant due to being unable to pass the right ventricle. | [5 empty] | The investigation for the reported delivery issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the delivery issue could not be determined. added-h6 problem code 2524. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31092 | 1220648-2024-06634 | 18669727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18669727&pc=OZD | 01/19/2024 | 12/03/2025 |  | 02/08/2024 |  |  |  | 02/08/2024 |  |  |  |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132983 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a impella 5.5 pump support ongoing for a patient with multiple cardiac comorbidities and coronary artery bypass grafting (cabg). the automated impella controller (aic) was noted to have alarmed for a controller error. | [5 empty] | The impella device was not received from the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. d9 device return date added. f6 and f8 should have been left blank. updated h3 to device evaluation to anticipated. h6 type of investigation code added.
31093 | 1220648-2024-06635 | 18669770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18669770&pc=OZD | 01/19/2024 | 03/25/2024 |  | 02/08/2024 |  |  | 01/19/2024 | 02/08/2024 |  |  | 11/30/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024426629 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during implant procedure for 74-female-patient, after insertion the 14 fr. x 13cm abiomed (peel away) sheath, and removing the dilator, there was bleeding out of the transition from the elongated part to the sheath head. therefore, the physician switched to a 14 fr. x 25cm abiomed sheath and the procedure continued without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the access site bleed has been completed. the introducer was returned for investigation. a split was noted along one side of the sheath score line between orange introducer hub. no other damage was observed on the introducer. according to the clinical report, bleeding was observed after the insertion of the 14fr x 13cm introducer sheath and upon removal of the dilator. the doctor replaced it with a 14fr x 25cm peel-away sheath without any complications. the cause of the access site bleeding issue is due to a sheath score line split.
31094 | 1220648-2024-06636 | 18669805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18669805&pc=OZD | 03/13/2023 | 02/08/2024 |  | 02/08/2024 |  |  |  | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1684985 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/13/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a controller error alarm. the device was removed from service and returned for evaluation. there was no patient involvement. | [5 empty] | The investigation on the automated impella controller's error alarm has been completed. console logs from the reported day of event could not be analyzed, because they could not be downloaded from the device, due to the fact that the console did not boot up after arriving in danvers. telnet connection to the aic could not be established. messages on the console¿s text screen indicated corruption of compact flash (cf) card(s). after cf cards were temporarily replaced with known-good ones, the issue was resolved and console was able to boot up. logs from ic10077 obtained from the cloud ended before the day of reported event. after replacing cf cards, the console passed full functional check. presumably, the unidentified controller error message observed in the field was related to cf corruption. the cause of the issue was a defective component.
31095 | 1220648-2024-06637 | 18669861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18669861&pc=OZD | 02/20/2023 | 02/08/2024 |  | 02/08/2024 |  |  |  | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 2024286853 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/20/2023 |  |  |  |  | The user facility reported a 45-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the automated impella controller (aic) displayed a controller failure alarm during impella support. the console was replaced because of the failure. there was no patient harm reported. | [5 empty] | Data analysis: automated impella controller (aic) logs did not confirm the reported controller failure alarm. rather, there was an unexpected controller reboot during the clinical procedure. the console was power by ac mains at the time of the reboot which means there were two sources of power (battery and ac power). there were no log entries that could link this reboot to a failed mechanism or process. low tension (lt) power data showed there were normal charger or battery statuses during the reboot which would indicate no apparent logged issue with power. device analysis: console was investigated in the protective multiple earthed(pme) electrical engineering lab. through long-term testing, there were no unexpected reboots reproduced. inspection of the internal printed circuit board (pca)s and associated cables revealed no abnormalities. there were no visibly loose connections that could be impacted by vibration or cable strain. while on ac mains and battery power, epc supply was monitored to look for glitches and no issues were discovered (within normal supply ranges 16-24 v). the root cause of the unexpected controller reboot was not determined due to the inability to reproduce.
31096 | 1220648-2024-06638 | 18669941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18669941&pc=OZD | 03/23/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/23/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1598659 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/23/2023 |  |  |  |  | The biomedical department of the user facility reported that the automated impella controller (aic) had a damaged purge disc retainer. there was no patient involvement. | [5 empty] | Data analysis: after being reviewed, it was determined that the automated impella controller (aic) logs are not relevant to failure mode. device analysis: the console was tested after arriving and the purge disc retainer was confirmed to be broken (broken blue disc retainer). before returning this console back to the customer, was instructed to replace the purge disc retainer, disc detect parts (flag and spring), validate sensor accuracy according to preventative maintenance procedure, and perform a full functional check on the console and optical system. the cause of the issue were broken purge clip and purge disc retainer.
31097 | 1220648-2024-06639 | 18670186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18670186&pc=OZD | 08/29/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 08/29/2023 | 02/08/2024 |  |  | 06/30/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024364438 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2023 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the nurse noticed a small drip on the floor and that the impella 5.5 purge on the device was a little wet. what was on the floor seemed to be sticky so suspect that there is a very small leak from the purge. there was no leak in any of the lines leading to the purge and the yellow to yellow was tight. there have been no alarms for low purge flow or pressure. purge flow and pressures were stable and unchanged. no known patient harm or injury. | [5 empty] | The impella device was received from the customer on (b)(6) 2023 and therefore, an evaluation of the device was performed with the following findings: data review: no purge issues were found on the data. visual inspection found no issues on the pump. no leak was found under leak test. the root cause of the purge issue was use related as no problem found on the pump. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
31098 | 1220648-2024-06641 | 18670682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18670682&pc=OZD | 10/15/2022 | 02/08/2024 |  | 02/08/2024 |  |  |  | 02/08/2024 |  |  | 06/30/2024 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient developed ischemia in their impella leg following impella cp placement. the impella was subsequently removed and the patient was escalated to an impella 5.5 for ongoing support. | [5 empty] | The investigation into the reported ischemia has been completed. no product was returned. the root cause of the limb ischemia issue was patient condition related. the ischemia was observed on the impella which was later extracted as per the clinical description provided. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31099 | 1220648-2024-06640 | 18670794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18670794&pc=OZD | 04/30/2023 | 02/08/2024 |  | 02/08/2024 |  |  |  | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/30/2023 |  |  |  |  | The user facility reported a female patient of unknown age, race, and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that automated impella controller (aic) displayed a blue screen. the aic was replaced, and the support continued. there was no patient harm reported. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show excessive amount of error messages communication, since console boot-up. console was tested, and although the exact issue was not reproduced, there were some anomalies in the system operation. during the visual inspection revealed damage to internal wiring between input and output (i/o) panel and 4-in-1 carrier which - although not directly responsible for portable bitmap (pbm)1 communication (damaged wire was for service dongle connection) - could¿ve resulted in an electrical short, affecting operation of the system. also, a connector for pbm2 serial communication was not fully seated in its socket on the 4-in-1 carrier. it is most likely that operation of the system was related to these issues, but malfunction of 4-in-1 assembly could not be excluded. repair: replaced aic i/o panel and replace the 4-in-1 assembly as a precaution. perform functional check of the console. the root cause of the reported issue was due to a defective cable.
31100 | 1220648-2024-06642 | 18670904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18670904&pc=OZD | 03/13/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/14/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1225719 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Activation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/14/2023 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had an impella 5.5 support with an automated impella controller (aic) with purge disc/pc recognition issues. the team was to follow the instructions for use and replace the aic after other troubleshooting. no lasting harm or injury was noted. | [5 empty] | The investigation has been completed. logs from the complaint date revealed that the console issued the purge disc no detected alarm several times. no issues were observed detecting the pump cassette from the perspective of the logs. the console was tested after arriving. this issue with properly detecting the purge pressure was reproduced and purge flag was confirmed to be broken. the cause of the issue was a broken purge flag.
31101 | 1220648-2024-06643 | 18670909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18670909&pc=OZD | 05/25/2023 | 09/22/2024 |  | 02/08/2024 |  |  | 05/25/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1087349 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Computer Operating System Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/25/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had an aic (automated impella controller) with power/shut down issues. the aic was pulled per ifu recommendations, and a backup console will be used. not reportable. no harm or injury. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Data analysis: aic logs showed that the audio server process had no heartbeat (failed). 3x partial resets did not revive the process, and once a full reset was performed the process was running again. see ic1254 aic logs complaint date. device analysis: the console was investigated on an engineering lab bench. a long-term test with pump and purge was run, but there were no issues with the audio server process reproduced. artificially killing (¿slaying¿) the audioserver process reproduced the partial reset symptom, but process was revived with only partial reset unlike the events in the field. the root cause of the unexpected console reboot was due to an audioserver process failure. the root cause of process failure was not determined due to the inability to reproduce. before sending this console back to the customer, fs was instructed to replace the compact flash (2025-0019) cards as a precaution and perform a full functional check on the console and optical system (0042-7316). | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06643 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
31102 | 1220648-2024-06644 | 18670912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18670912&pc=OZD | 04/17/2023 | 09/30/2025 |  | 02/08/2024 |  |  | 04/17/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087352 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | Black Or African American | Death | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | Other |  |  |  |  | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/17/2023 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 47-year-old male patient of unknown ethnicity with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that automated impella controller (aic) failure alarm that was triggered during patient support. the aic was swapped because of the noted failure. there was no patient harm reported. | The procedure outcome was that patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The complainant also reported a loss of placement signal waveforms followed by a controller error alarm. no previous alarms had been noted. |  |  | This report is being filed as part of a retrospective review of historical records. corrected information for the follow up mdr 1220648-2024-06644-1 was provided in sections b1, b2 and h1. | The investigation into the reported aic purge signal has been completed since the original report was filed. the device was returned for analysis. the error was immediately reproducible in the lab with the same error. this error was resolved by swapping out pud pca. upon observing the pump connector, there is debris around the optical fiber even after cleaning that could cause a small air gap that could lead to the barcoding placement signal but not the oscillating contrast which was not reproduced. the cause of the purge aic issue was firmware corruption on the pud. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06644 was provided in sections b1, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the cause of the oscillating contrast was most likely a defective optical bench specifically a malfunctioning lamp causing out of range contrast calculations but due to a lack of reproduction, issue could not be isolated. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b7 other relevant history, including preexisting medical conditions updated. d4 lot number and udi # corrected. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email and reporting contact fax number updated.
31103 | 1220648-2024-06645 | 18670977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18670977&pc=OZD | 08/04/2023 | 02/08/2024 |  | 02/08/2024 |  |  |  | 02/08/2024 |  |  | 06/30/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024364393 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/04/2023 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 insertion was aborted due to the patient¿s anatomy. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. the root cause of the delivery issue was patient condition related due to difficult patient anatomy.
31104 | 1220648-2024-06646 | 18671043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18671043&pc=OZD | 10/17/2023 | 02/08/2024 |  | 02/08/2024 |  |  |  | 02/08/2024 |  |  | 01/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024334626 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported an unspecified male patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was unable to advance into the descending due to the narrow iliac artery. the case was aborted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. the cause of the delivery issue was patient condition related due to narrow iliac artery impella was unable to advance into descending aorta. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
31105 | 1220648-2024-06647 | 18671044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18671044&pc=OZD | 01/27/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 01/27/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1095648 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Freezes, Becomes Nonfunctional | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 01/27/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) console was unresponsive upon power up. the aic was removed from service as a result of the noted issue. there was no patient involvement. | [5 empty] | The investigation into the power on issues blank screen on the automated impella controller (aic) has been completed. data analysis: no logs were available from the complaint which indicates the might not have been running and generating logs. the cause for this was later confirmed when the device was analyzed. device analysis: the issues powering on were reproduced. the power on signal from the pbm to the epc was lower amplitude than expected (3.3v), which caused the intermittent power on issues. the failure was isolated to the 4-in-1 assembly, which was likely pulling down this signal. replacing the 4-in-1 assembly resolved the issue. the batteries were also found to be depleted which was related to this failure mode. the batteries were likely depleted because the console would partially power on without properly turning on the epc and user did not know it was partially powered on. this would lead to the user leaving the console on in this state and ultimately draining the batteries. the cause of the power on issues was due to a defective 4in1 board low power on signal.
31106 | 1220648-2024-06648 | 18671045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18671045&pc=OZD | 07/13/2023 | 02/08/2024 |  | 02/08/2024 |  |  |  | 02/08/2024 |  |  | 01/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024334531 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Disability | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 62-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient on impella support had the pump stop. the decision was made to extract impella and when it was taken out a thrombus was noticed in the pump. there was no adverse event reported to the patient as a result of this incident. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. pump was not returned for investigation. data logs shows the several motor currant spike at the end of the case followed by pump stop, which is indicative of biomaterial ingestion. the cause of the pump stop issue was ingested biomaterial based on provided clinical information and data logs. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31107 | 1220648-2024-06649 | 18671138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18671138&pc=OZD | 03/15/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/15/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1175897 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/15/2023 |  |  |  |  | The user facility reported a 73-year-old female patient with severe mitral regurgitation was implanted with an impella for mechanical circulatory support. it was reported that the automated impella controller (aic) had purge disc alarm, after placing multiple purge cassettes into an aic for use. the console was replaced because of the failure and the issue resolved. there was no patient harm reported. | [5 empty] | The investigation into the reported aic-purge disc/flag issues has been completed. the device was returned for analysis. logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the alarm was triggering and clearing intermittently. the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the issue was a broken purge flag.
31108 | 1220648-2024-06650 | 18671236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18671236&pc=OZD | 04/01/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 04/01/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1505149 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/01/2023 |  |  |  |  | The user facility reported a (b)(6) year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) experienced a purge failure. the aic was replaced with no harm or injury to the patient. | [5 empty] | The investigation into the automated impella controller (aic) purge issue has been completed. the data logs were not made available and the impella aic was not returned for evaluation. the cause of the issue was not determined since product was not returned.
31109 | 1220648-2024-06651 | 18671748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18671748&pc=OZD | 07/13/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 07/13/2023 | 02/08/2024 |  |  | 04/30/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350824 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 44-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia at the right lower extremity. a 6fr antegrade rfa sheath was placed for fem-fem bypass. some oozing at the groin site which was later controlled. | [5 empty] | The investigation into the patient's access site bleeding and limb ischemia has been completed. the impella products were not returned for evaluation. per the clinical information provided, the root cause of the limb ischemia-reversible was most likely due to the patient's condition related to diseased/ small vessels. the root cause of the access site bleeding - major could not be determined due to insufficient clinical information provided.
31110 | 1220648-2024-06652 | 18671999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18671999&pc=OZD | 03/23/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/23/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1486033 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Activation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/23/2023 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a purge issue against this aic (automated impella controller) console device. discussion was had with the investigating engineer who found, after reviewing available data logs, this was likely a situation where the console failed to recognize a purge disc. a new aic was used. there was no patient impact reported. | [5 empty] | The investigation into the reported aic purge issue other has been completed since the original report was filed. the device was returned for analysis. console logs from the reported day of event are mostly consistent with complaint and show multiple purge related error messages. during troubleshooting of the device it¿s been determined that purge motor shaft was immobilized due to contamination of the motor shaft gasket, caused by prior purge fluid leak. the cause of the aic issue was contamination.
31111 | 1220648-2024-06654 | 18672175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672175&pc=OZD | 05/12/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 05/12/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1314435 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 05/12/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a battery failure alarm. there was no patient involvement. the aic was returned for evaluation. | [5 empty] | The investigation into the automated impella controller (aic) battery cell failure has been completed. logs confirmed the reported failure. there was a battery failure alarm (transport connection blown missing) and a battery 1 communication failure alarm (smbus communication loss or sda short) due to low pack voltage. the ltpowerdata from the complaint date showed cell imbalance with battery 1 which caused battery pack internal electronics to shut down that battery. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure, a known pattern failure.
31112 | 1220648-2024-06653 | 18672204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672204&pc=OZD | 04/07/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 04/11/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1117651 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/11/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported while removing purge cassette and disc, the disc tray detached. the device has been removed from use and a new device will be sent. no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. data analysis: the console was not returned for investigation so the logs had to be pulled from constellation. the last available logs ended mid-procedure on 4/7 and showed the system functioning without any sign of the failure mode. device analysis: device not returned. root cause: the cause of the issue was not determined since product was not returned.
31113 | 1220648-2024-06655 | 18672359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672359&pc=OZD | 03/12/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/12/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1548488 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/12/2023 |  |  |  |  | The user facility reported a 50-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the automated impella controller (aic) had a battery failure. the console was replaced because of the failure. there was no patient harm reported. | [5 empty] | The investigation into the reported aic - battery cell failure has been completed the failure mode was reproduced on an engineering bench. the battery lcd indicator was blank (fully discharged). battery 1 was momentarily replaced to resolve the failure alarm. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure.
31114 | 1220648-2024-06656 | 18672364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672364&pc=OZD | 04/24/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 04/24/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 141816 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/24/2023 |  |  |  |  | The user facility reported a 16-year-old male with heart failure was implanted with an impella 5.5 device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a purge disc failure alarm. the aic was swapped for another aic. there was no patient harm or injury. | [5 empty] | The investigation into the automated impella controller (aic) purge issue has been completed. data analysis: aic logs showed throughout the 39 days running the 5.5 pump there no persistent instances of the purge disc not being detected that could not be resolved. no issues completing the activities performed in the purge wizard. the reported failure mode (purge disc not detected) could not be confirmed based on the device data logs. an example of a purge disc not detected alarm triggering during the during the cassette change procedure was attached. note, the alarm triggered and cleared within six seconds. device analysis: the console was inspected and tested on an engineering lab bench. the associated components to purge disc detection mechanism (purge flag and retainer part of purge pressure sensor assembly) were intact. there were no signs of fluid ingress or substantial build up externally or internally. the aic could detect a known-good purge disc with no issues. the root cause of the reported purge disc detection issues could not be determined due to the inability to confirm via data logs or reproduce. no problem found.
31115 | 1220648-2024-06657 | 18672391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672391&pc=OZD | 03/17/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/17/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1147696 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/17/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) failed to detect a purge cassette. the aic was returned for evaluation. there was no patient involvement. | [5 empty] | The investigation into the purge disc issue on the automated impella controller (aic) has been completed. data review: after reviewing the imc log of this aic, it was found that "purge disc not detected" alarm and "purge pressure low" alarm was triggered several times. device analysis: the console was tested after arriving in field service. it was evident that the purge retainer was cracked.
31116 | 1220648-2024-06658 | 18672434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672434&pc=OZD | 08/12/2023 | 04/16/2025 |  | 02/08/2024 |  |  | 01/09/2024 | 02/08/2024 |  |  | 03/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024342410 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Hospitalization | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 01/09/2024 |  |  |  |  | Abiomed¿s long-term outcome and quality indicator impella registry (loqi) notification received with a "possible" relationship to the impella device. this particular patient presented as a 68 year old male patient presenting with acute myocardial infarction and cardiogenic shock. during support, an "occlusion of external iliac artery with monophasic flow in femoral arteries with minimal tibial artery flow" was noted. the patient's impella inserted leg was determined to be "unsalvageable" and "an above-knee amputation was recommended, however, the patient is too unstable for surgery." the patient was escalated to higher care via an impella 5.5. | An additional report from loqi indicates this patient to have endured hemolysis, in addition to the originally reported ischemia. the patient's hospitalization was prolonged, as a result. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right lower extremity arterial thrombus with a "definite" relationship to the impella device used: ¿the patient has an indwelling arterial sheath of the left external iliac artery for the impella cp access site. the patient was noted to have loss of pulsatility to the left lower extremity. arterial duplex reveals occlusion of external iliac artery with monophasic flow in femoral arteries with minimal tibial artery flow. right lower leg was determined to be unsalvageable. an above-knee amputation was recommended; however, the patient is too unstable for surgery, and extensive anticoagulation is contraindicated due to concern for pericardial and retroperitoneal bleeding. creatine kinase peak 70986 u/l ((b)(6) 2023).¿ the patient not recovered/not resolved. the procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  | The investigation into limb ischemia ¿ irreversible issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation has been reopened to evaluate for hemolysis. a supplemental report will be filed at the conclusion of our second investigation. the initial mdr reported an occurrence date in b3 as (b)(6) 2023, which was erroneous. the correct date for of occurrence is (b)(6) 2023. | The investigation into the limb ischemia ¿ irreversible and thehemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. as all the necessary information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ d6a, d6b. | Medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). b5: additional narrative added based on new information received for thrombus. h1: health effect code added for thrombus, and investigation results for thrombus. the investigation for the reported issue (thrombus) has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1 reporting contact email revised in accordance with updated procedures.
31117 | 1220648-2024-06659 | 18672435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672435&pc=OZD | 04/07/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 04/27/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1356082 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/27/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a console shut down while the patient was in transit to the icu. staff plugged console into wall, restart impella, temporarily recharged battery, and finished transport. patient remained on support for several days on same console without further incident. there was no harm to the patient. | [5 empty] | Data logs from consoles reportedly involved in the complaint were analyzed and confirmed the timeline, however we were unable to identify the events referring to alleged shutdown and/or pump stop. console ic2608 was identified as the one used for transport, and no issues occurred there. console ic6244 did not shut down unexpectedly, neither any of the pumps (419235, 420216) used on it since (b)(6) 2023 stopped unexpectedly. logs from ic6244 revealed isolated instances of temporary loss of placement signal (a known failure mode ¿transient loss of optical data¿, for which there¿s no repair required as issue self-resolves), and few manual pump stops and restarts (presumably as mitigation to suction alarms) towards the end of support with pump 419235. (b)(6) hospital, no pump stops, no shutdowns, no power or battery issues. field troubleshooting was reportedly performed: ¿after hearing of issue, i checked all consoles on site. unplugged them all from the wall (including 6244), turned them on for approximately 15 minutes, and battery charge on all consoles read 100% with green battery in upper right corner.¿ these actions were confirmed, via available cloud log data, to have been conducted on (b)(6) 2023 on consoles belonging to the same account: ic1614, ic6244, and ic4751 (there were no logs available for ic1574 and ic9658). troubleshooting of console ic6244 performed in danvers did not reveal any irregularities, and during long-term testing with a known-good pump the console operated within specification. there were no unexpected pump stops or shutdowns or issues with aic power. it is possible that either the reported issue occurred on yet another console (which we were unable to identify) or the actual issue (temporary loss of placement signal) was misreported. repair: perform functional check of the console.
31118 | 1220648-2024-06660 | 18672441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672441&pc=OZD | 03/29/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/29/2023 | 02/08/2024 |  |  | 09/30/2024 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281678 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Malfunction | No | Yes | Inaccurate Delivery | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant reported upon introduction to left ventricle (lv), the impella cp device prolapsed and the cannula kinked. the device was therefore removed from lv and was pulled down the descending aorta in attempt to catch pigtails on an anatomical structure to straighten device. the device was pulled into left common iliac where it got stuck. all attempts to position pump from this point were unsuccessful and the patient transferred to or for emergent cut down of left femoral artery and removal of device. | [5 empty] | The investigation into the reported delivery issues has been completed. the medical center did not return any impella products for benchtop analysis. based on the clinical information available, the root cause of the delivery issue was determined to be patient condition related.
31119 | 1220648-2024-06661 | 18672460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672460&pc=OZD | 03/28/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/28/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1300416 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/28/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a controller error alarm during setup. there was no patient involvement. | [5 empty] | The investigation into the automated impella controller (aic) controller error alarm has been completed. console logs from the reported day of event are consistent with complaint and show controller error alarm, due to defective optical sensor lamp), that did not resolve after restarting the console. log analysis showed that the same issue occurred 3 days prior, and was preceded by several controller error alarms as early as 2023-03-18. after console was received in danvers, the issue was reproduced, and it¿s been observed that there was no light coming out of aic optical fiber. it¿s been determined that the lamp assembly on the optical bench had failed. console inspection also revealed contamination (that could not be cleaned) of the optical pump connector, cracked purge disk retainer, and purge fluid ingress into the purge pressure sensor area all unrelated to the complaint. the cause of the issue was a defective component.
31120 | 1220648-2024-06662 | 18672468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672468&pc=OZD | 05/25/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 05/25/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1643465 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 05/25/2023 |  |  |  |  | The user facility reported a 54-year-old female undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a high purge pressure alarm. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation into the automated impella controller (aic) high purge pressure issue has been completed. there was a high purge pressure causing purge pressure blocked alarms. the purge cassette was then switched out for another but the alarms did not resolve and purge pressure remained high. the failure continued until the p-level was set to zero for seven seconds which resolved the alarms and dropped the purge pressure. the p-level was ramped back up and purge pressure remained at acceptable levels. the console ran for five more days with no recurrence of errors. the logs support the continued high pressure seen with purge pressure above 1000 mmhg until lowering below 600 mmhg. the high purge pressure was resolved on the same console, not by switching out aics like the clinical review says. the console has functioned without this failure mode reappearing for the past five months in the field. the device was unable to recreate the failure mode and functioned as expected. the cause of the issue was most likely a kinked line or momentary restriction of purge flow to increase purge pressure which is the type of failure that could have been solved without swapping out the console or purge cassette.
31121 | 1220648-2024-06663 | 18672474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672474&pc=OZD | 02/28/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 02/28/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1075037 |  |  |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Calibrate | Difficult or Delayed Activation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/28/2023 |  |  |  |  | No patient involvement: the complainant reported a an aic (automated impella controller) failing to recognize a purge disc . there was no patient impact related. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show multiple occurrences of ¿purge disk not detected¿ alarm (#402) presenting each time after purge cassette is installed. console inspection performed in danvers identified a broken purge disk detect flag, which caused inability to detect purge disk. the inspection also revealed additional deficiencies of the purge system: immobilized purge motor shaft (due to spilled purge fluid residue), broken purge disk retainer, contamination of the purge insert, and evidence of purge fluid ingress compromizing purge pressure sensor (transducer). repair: replace purge pressure sensor (transducer), purge disk retainer, disk detect flag and foil. clean the purge insert and motor shaft gasket (replace gasket if needed). perform functional check of the console.
31122 | 1220648-2024-06664 | 18672477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672477&pc=OZD | 01/26/2024 | 03/12/2024 |  | 02/08/2024 |  |  | 01/26/2024 | 02/08/2024 |  |  | 10/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422934 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported 72-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump had suction and low flows which prompted premature explant and replacement. when explanted the pump was observed to have biomaterial at the cannula. the 2nd pump successfully supports the patient. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the low pump flow issues has been completed since the original report was filed. the console device was returned and tested after arriving in fs. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. the root cause of low flow was biomaterial ingestion.
31123 | 1220648-2024-06665 | 18672517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672517&pc=OZD | 03/30/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/30/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1614917 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/30/2023 |  |  |  |  | The user facility reported a 69-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the blue purge disc holder broke off of the automated impella controller (aic). the aic was swapped because of the noted damage. there was no patient harm reported. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that after one of purge cassette changes, the purge change wizard kept reporting purge fluid not detected, which together with the data showing no purge pressure readings is consistent with broken purge disk retainer. the purge disk retainer was replaced by the field staff and functional check was performed, after which the console was verified to be operating within specification. repair: replace purge disk retainer (and flag, if missing), and perform functional check. the cause of the issue was a defective component.
31124 | 1220648-2024-06666 | 18672521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672521&pc=OZD | 03/15/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 06/26/2023 | 02/08/2024 |  |  | 04/30/2024 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216672 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/26/2023 |  |  |  |  | The user facility reported a 66-year-old female patient post cardiotomy cardiogenic shock low cardiac output syndrome was to be implanted with an impella cp for mechanical circulatory support. it was reported surgical access was gained in the right femoral artery and a 25-centimeter sheath was inserted. upon trying to insert the catheter it was noted that the patient had a large aortic dissection. the patient was reopened, and the dissection repaired. the left ventricle was still akinetic on echo following repair and patient was placed on venous-arteria extra-corporeal membrane oxygenation with little flow noted on cannula (1.5liters/minute). the physician decided, given prognosis of patient, elected not to reinsert impella cp. no known patient harm or injury was reported. | [5 empty] | The pump was not returned for investigation, and data logs were not provided. the cause of the delivery issue was patient condition related to reported dissection in aorta.
31125 | 1220648-2024-06667 | 18672591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672591&pc=OZD | 08/17/2023 | 07/10/2024 |  | 02/08/2024 |  |  | 08/17/2023 | 02/08/2024 |  |  | 05/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368162 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to being hypotensive and began experiencing frequent runs of vt (ventricular tachycardia) . the patient developed a severe gi bleed so, the decision was made to remove the impella. pump was removed after multiple pressors and inotropes were added. reportable due to patient having a gi bleed and the impella was removed. | [5 empty] | The investigation for the reported gi bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the gi bleed could not be determined. the instructions for use for the impella cp with smart assist in section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 clinical code 4476, h6 component code 925 d4-corrected model number added- d6a/d6b. f6/f8 should have been left blank in the initial report. g1 corrected reporting contact email g1-manufacturing site fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
31126 | 1220648-2024-06668 | 18672598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672598&pc=OZD | 08/07/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 08/07/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, UK |  |  | ABIOMED, INC. |  |  | US |  | 1694378 |  | 0042-0040-UK |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Premature Discharge of Battery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 08/07/2023 |  |  |  |  | The complainant reported that the batteries for the automated impella controller (aic) do not charge. the battery icon remains at 50% and the red battery failure alarm is displayed. no patient is involved. | [5 empty] | The investigation into the reported aic battery issue has been completed. the device and logs were returned for evaluation. the failure mode was reproduced on an engineering bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. section e1: the name of the initial reported is unknown, so its left blank.
31127 | 1220648-2024-06669 | 18672631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672631&pc=OZD | 06/15/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 06/15/2023 | 02/08/2024 |  |  | 06/30/2024 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023233630 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2023 |  |  |  |  | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced aortic regurgitation (ar) following implant of the impella pump. the ar was reported to be unavoidable due to the structure of the patient's heart. no medical intervention was required to treat the ar and patient support continued with the implanted pump. there was no patient harm. | [5 empty] | The investigation has been completed. no product or data logs were returned for review. as per clinical review, the pump did not pass through the center of the aortic valve due to left ventricle having the structure of the right ventricle hitting left ventricular wall. the pump was pulled out, reinserted to fix at a position with no suction alarm and certain amount of aortic regurgitation (ar) was unavoidable due to the structure of the heart. the cause of the heart valve damage issue was most likely patient condition related since aortic regurgitation was unavoidable due to structure of the heart during impella insertion and positioning per clinical review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
31128 | 1220648-2024-06670 | 18672734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672734&pc=PYX | 07/06/2023 | 09/19/2024 |  | 02/08/2024 |  |  | 07/06/2023 | 02/08/2024 |  |  | 05/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024346253 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported a 83-year-old male post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. the outflow of an impella rp pump was unable to pass through the tricuspid ring during attempted delivery. after multiple failed attempts, the decision was made to abort the implant. there was no patient harm. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported delivery issue has been completed. no product was returned. as per clinical review, the impella rp was unable to advance further the repaired tricuspid valve. there was resistance when it got to the tricuspid ring with outflow beyond the ring and insertion was aborted. the cause of the delivery issue was patient condition related due to pump unable to advance further the tricuspid ring based on clinical review. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
31129 | 1220648-2024-06671 | 18672780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672780&pc=PYX | 02/28/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 03/22/2023 | 02/08/2024 |  |  | 12/31/2024 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024295458 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2023 |  |  |  |  | The user facility reported a 87-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the right sided impella was placed but was removed due to motor failure. there was no known harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
31130 | 1220648-2024-06672 | 18672803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18672803&pc=OZD | 10/06/2023 | 06/24/2024 |  | 02/08/2024 |  |  | 10/06/2023 | 02/08/2024 |  |  | 06/30/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024359648 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella 5.5 support while awaiting a heart transplant. the pump became positioned under the mitral valve. the staff attempted to reposition, pulled back and the catheter came across the valve into the aorta. attempts to push back in were unsuccessful. the decision was made to bring back to operating room for a new impella 5.5 implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. it was stated that the pump was against the mitral valve and repositioning was attempted but the pump was pulled back and it went into the aorta. the root cause of the positioning issues is most likely due to use. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
31131 | 1220648-2024-06673 | 18673415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673415&pc=OZD | 01/25/2024 | 03/11/2024 |  | 02/08/2024 |  |  | 01/25/2024 | 02/08/2024 |  |  | 02/28/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324147 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported a impella cp that on an 83 year old patient in use for mechanical circulatory support. it was reported that the impella was removed due to poor flow. after removal, clot was seen on the impella. a new impella was placed. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. data logs were reviewed which confirmed low pump flows. additionally, clinical information provided was sufficient to determine the root cause. the cause of the low pump flow issue most likely biomaterial ingested as the details confirmed clot on the removed device. it states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
31132 | 1220648-2024-06674 | 18673520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673520&pc=OZD | 01/25/2024 | 03/05/2024 |  | 02/08/2024 |  |  | 01/25/2024 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER |  |  | ABIOMED |  |  | US |  |  |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | It was reported by biomedical engineer that the console (aic) has a broken purge clip. no patient involvement. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon examining reported issue with the broken purge retainer was confirmed. purge retainer was completely broken off. the cause of the issue was a broken purge retainer. | The controller (aic) has been returned for evaluation but has not been evaluated. upon completion of the investigation a supplement report will be submitted.
31133 | 1220648-2024-06675 | 18673522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673522&pc=OZD | 10/01/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 01/26/2024 | 02/08/2024 |  |  | 05/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024345264 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported 63-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the sterile sleeve had a tear on the distal end. the tear was covered in tegaderm, and the pump will stay in place. no patient harm was reported. | [5 empty] | Due to no returned product and limited clinical details, the root cause product damage to the sterile sleeve could not be determined.
31134 | 1220648-2024-06676 | 18673525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673525&pc=OZD | 01/26/2024 | 03/12/2024 |  | 02/08/2024 |  |  | 01/26/2024 | 02/08/2024 |  |  | 06/30/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024427267 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported 63-year-old male patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump had had frequent suction alarms due to mal positioning. multiple attempts unsuccessfully were made to reposition. the impella was explanted. patient hemodynamically stable. no patient harm was reported. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the device was not returned for investigation. according to the clinical details, multiple suction alarms occurred due to improper positioning. attempts were made to resolve the issue with no success. the most likely cause of the low pump flow was biomaterial. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31135 | 1220648-2024-06677 | 18673527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673527&pc=OZD | 01/26/2024 | 03/15/2024 |  | 02/08/2024 |  |  | 01/26/2024 | 02/08/2024 |  |  | 11/30/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABOIMED, INC. |  |  | US |  | 2024424151 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported 47-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that patient experienced some moderate bleeding at their access site following impella 5.5 implant. a jackson -pratt (jp) drain was placed to manage the condition and prevent a hematoma from forming. the following day, it was noted that the bleeding had slowed, but 1 unit packed red blood cells (prbc) and 1 unit of platelets were transfused to treat the condition. support was maintained following the intervention. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31136 | 1220648-2024-06678 | 18673531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673531&pc=OZD | 11/14/2023 | 02/16/2024 |  | 02/08/2024 |  |  | 11/14/2023 | 02/08/2024 |  |  |  |  | 02/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1104542 |  | 0042-0000-CA |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/14/2023 |  |  |  |  | The user facility reported that during training the blue purge clip that holds the purge disc was broken on one side. another automated impella controller (aic) was used. there was no patient involvement | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into aic - purge disc/flag issues completed since the original report was filed. the console was returned for investigation. console was tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge retainer was broken. the root cause of the issue was a broken purge retainer.
31137 | 1220648-2024-06679 | 18673544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673544&pc=OZD | 11/25/2023 | 10/01/2024 |  | 02/08/2024 |  |  | 11/25/2023 | 02/08/2024 |  |  | 07/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024377078 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Inaccurate Delivery | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp insertion was aborted due to the patient of unknown age and gender's vascular anatomy. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition due to patient vascular anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06679: d6a and d6b revised from the initial submission in accordance with updated procedures. h6 code 2339 was reported incorrectly. new code was added to medical device problem code. h6 component code revised from the initial submission in accordance with updated procedures. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
31138 | 1220648-2024-06680 | 18673546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673546&pc=OZD | 10/20/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 10/20/2023 | 02/08/2024 |  |  | 05/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024359237 |  | 0048-0014 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Ingredient or Reagent Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/20/2023 |  |  |  |  | The user facility reported a 89-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. not possible to implant the pump due to tortuous vascular anatomy, plan to try it again via right axillar next time. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the delivery issue was patient condition related due to tortuous vascular patient anatomy.
31139 | 1220648-2024-06681 | 18673563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18673563&pc=OZD | 08/15/2023 | 02/08/2024 |  | 02/08/2024 |  |  | 08/15/2023 | 02/08/2024 |  |  | 03/31/2025 |  | 02/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024336701 |  | 0048-0014 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2023 |  |  |  |  | The user facility reported an unknown age and gender patient a in acute myocardial infarction cardiogenic shock was implanted with an impella device for mechanical circulatory support. during the impella cp insertion the physician was unable to pass the catheter through the vasculature and the procedure was aborted due to anatomy. no patient harm reported due to the issue. | [5 empty] | The investigation into the reported delivery issue has been completed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the root cause of delivery issue was most likely due to patient anatomy patient condition.
31140 | 1220648-2024-06586 | 18657961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18657961&pc=OZD | 01/27/2024 | 03/11/2024 |  | 02/07/2024 |  |  | 01/27/2024 | 02/07/2024 |  |  | 05/31/2025 |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353908 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 46-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after the impella was inserted suction alarms and low flows occurred. troubleshooting was unsuccessful. the impella was exchanged. the new impella had no issues. there was no visible thrombus on the catheter or visible kinks. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella device was returned for evaluation. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue under bench test. data logs were reviewed and confirmed the low pump flow which remained to the rest of the case. the root cause of low flow was most likely due to patient condition as no problem was found on the pump and the failure mode could not be reproduced.
31141 | 1220648-2024-06618 | 18659897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18659897&pc=OZD | 01/25/2024 | 03/08/2024 |  | 02/07/2024 |  |  | 01/25/2024 | 02/07/2024 |  |  | 06/30/2025 |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024431025 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported a 62-year-old male was on impella cp support and after the impella was removed a thrombectomy was performed. the thrombectomy was required to the left groin and surgical closure was complete. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31142 | 1220648-2024-06619 | 18659926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18659926&pc=OZD | 01/26/2024 | 03/06/2024 |  | 02/07/2024 |  |  | 01/26/2024 | 02/07/2024 |  |  | 06/30/2025 |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024383577 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 78-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the pump was unable to obtain good flow rates. a visual inspection was performed, and it was discovered that a kink was present on the cannula. the pump was removed / replaced to resolve the noted issue. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported low pump flow issue has been completed. the impella device was returned for evaluation. upon visual inspection, a kink was observed near outflow cage. there was no biomaterial or broken blades found. the cause of the low pump flow issue was cannula kink due to improper technique.
31143 | 1220648-2024-06620 | 18659932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18659932&pc=OZD | 01/27/2024 | 03/22/2024 |  | 02/07/2024 |  |  | 01/27/2024 | 02/07/2024 |  |  | 10/31/2025 |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419262 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Pressure Problem | Material Split, Cut or Torn |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump experienced high purge pressure during patient support. as pressures rose above 1008 mmhg, a tpa protocol was initiated. the following day, it was verified that the tpa resolved the high purge pressure and support was able to continue with the implanted pump. there was no patient harm. | Additional information was received pertaining to supplemental allegations associated with this device. one day following intervention for the high purge pressure, it was discovered that the sterile sleeve had been torn during repositioning. the sleeve was taped in an attempt to preserve sterility and support continued. there was no patient harm. | It was also reported that 38 days later while still on support, low pump flow was recorded. after discovering the pump was in the mitral valve apparatus, repositioning was performed. however, this did not resolve the low pump flow issues, so the decision was made to explant the pump and replace it with an impella cp. upon further inspection of the pump no kinks were found, but tissue was observed in the outflow cage. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported high purge pressure and low pump flow has been completed. the impella device was not returned for evaluation. however, clinical details stated that after tissue plasminogen activator (tpa) was added to the purge, the high purge issue resolved. therefore, the most likely root cause of the high purge issue was related to biomaterial. clinical details further states that repositioning of the pump did not resolve the low pump flow. therefore, the pump was subsequently explanted. upon visual inspection at the clinical site, it was reported that tissue had been observed in the outflow cage. the most likely root cause of the low pump flow was also related to biomaterial. added- b5 narrative updated; h6 impact code 4644, 4628, 4627; h6 problem code 2954 d4- updated model number.
31144 | 1220648-2024-06621 | 18660017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18660017&pc=PYX | 01/27/2024 | 02/26/2024 |  | 02/07/2024 |  |  | 01/27/2024 | 02/07/2024 |  |  | 11/30/2025 |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024435373 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 68-year-old male patient of an unknown race in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that upon implantation of the impella, the cannula appeared to be buckling and prolapsing into the inferior vena cava. the slack was removed but upon doing so, bleeding was observed. the rp was removed, and a buddy wire was inserted, however insertion was still difficult. a push pull technique was utilized with the .027 wire and removing the slack in rp flex leading to a successful placement. during support, the flows decreased, and a possible clot was speculated. the flows continued to worsen, and the pump was explanted and replaced. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound | The investigation into the access site bleeding ¿ major, the low pump flow, and the product damage (cannula kink) issues has been completed since the original report was filed. significant biomaterial was observed on the impeller, and the catheter was kinked at the kink protector. low flow was reproduced during the test in a bucket of water. the cause of the access site bleeding issue was use related since the bleeding was observed while the inflow was outside of the hemostasis valve, which allows the blood to come out. the cause of the low pump flow issue was biomaterial found wrapped around the impeller. the cause of the delivery issue was patient condition related since the resistance was encountered at the tricuspid ring, and the doctor was successful with the second attempt.
31145 | 1220648-2024-06622 | 18660453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18660453&pc=OZD | 01/25/2024 | 03/05/2024 |  | 02/07/2024 |  |  | 01/25/2024 | 02/07/2024 |  |  | 02/28/2025 |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321296 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient¿s labs were hemolyzed. echo - echocardiogram (ultrasound)- showed that the pigtail was pointed toward the mitral valve. repositioning will not take place as the patient is scheduled to be escalated to impella 5.5. it was further reported that there was bleeding from the impella access site from taking the peel away out. a second peel away was left in and bleeding subsided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis and access site bleed has been completed. the device was not returned for evaluation. however, data logs were reviewed which revealed that many impella position alarms in ventricle and aorta had occurred. clinical data confirmed impella not repositioned due to escalation to 5.5 decision. the cause of the hemolysis issue was most likely improper position related due to improper technique. the cause of the access site bleeding was lack of vessel recoil as second peel away (25cm impella sheath) ultimately placed to resolve the issue.
31146 | 1220648-2024-06623 | 18660571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18660571&pc=OZD | 01/26/2024 | 11/25/2025 |  | 02/07/2024 |  |  | 01/26/2024 | 02/07/2024 |  |  |  |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1098592 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that placement signal unreliable alarm was active. after troubleshooting the alarm did not resolve. the impella and the aic automated impella controller) were exchanged. the new console and impella were used successfully. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show that during case start with pump 476869, after optical sensor was calibrated, a placement signal not reliable alarm (#1036, ops signal invalid) presented, prior to pump start. pump ran for 3 minutes, and after it was disconnected from the console, a controller error alarm showed (#1039, defective optical sensor lamp), and ¿5s¿ parameters recorded in imc log indicated loss of light in the optical system (light = 0.5v). after console was restarted, the controller error alarm (#1039) repeated, but self-resolved in 5 seconds. case start wizard was launched and new pump (473786) was used. after 7.5 hours, a placement signal not reliable alarm (#1036) presented, and remained unresolved until pump was unplugged 1h 20m later. ltlog data for both pumps used on ic1536 showed that during the psnr condition (#1036) values of contrast parameters fluctuated between -3000 and +3000, which is a calculation artefact resulting from loss of light on the optical system. it is also consistent with controller error alarm (#1039). imc logs did not provide any information relevant to the reported inpella connect malfunction. console was tested in danvers and although the complaint issue (completer loss of light) was not reproduced, a non-conformance of optical bench was discovered, as the amount of light in the system was low, resulting in out-of-spec value of gain parameter. console inspection also revealed some contamination of the optical pump connector. analysis of the rlm revealed corruption of its microsd card, evidenced by ¿partition switching¿. after replacing microsd card, however, the unit was unable to correctly boot up, and further investigation revealed compromised wiring of the aic i/o panel - specifically 5v pin on x19 connector plugged into 4-in-1 carrier, which caused intermittent power to the rlm. placement signals not reliable was due to intermittent failure of optical bench. intermittent issues with impella connect" was due to microsd card corruption. the exact cause for micro sd card corruption was not determined there was no aic pump issue found during the case. the device functioned/operated to specification. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 updated. h3 updated. b7 other relevant history, including preexisting medical conditions updated. d2 common device name updated. d4 model number updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31147 | 1220648-2024-06624 | 18660582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18660582&pc=OZD | 01/26/2024 | 03/14/2024 |  | 02/07/2024 |  |  | 01/26/2024 | 02/07/2024 |  |  | 11/30/2025 |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2024433481 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 53-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed distal limb ischemia on impella side. this was successfully managed with placement of a distal perfusion catheter. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
31148 | 1220648-2024-06625 | 18660622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18660622&pc=OZD | 01/26/2024 | 02/28/2024 |  | 02/07/2024 |  |  | 01/27/2024 | 02/07/2024 |  |  |  |  | 02/07/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1569441 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | A japan user facility reported that when testing on a demo machine, the purge priming did not proceed. the pressure transmitter was not working. automated impella controller (aic) was replaced. there was no patient involvement. | [5 empty] | The investigation into the reported aic - purge disc/flag issues has been completed since the original report was filed. the device and data logs was returned for analysis.imc logs from the complaint date revealed that the console issued the purge disc not detected alarm. the cause for the purge priming not proceeding is a missing purge flag. the root cause for the missing purge flag a defective purge pressure defective part. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31149 | 1220648-2024-06626 | 18660706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18660706&pc=OZD | 01/28/2024 | 03/13/2024 |  | 02/07/2024 |  |  | 01/28/2024 | 02/07/2024 |  |  | 09/30/2025 |  | 02/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410052 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2024 |  |  |  |  | The user facility reported a 76-year-old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the patient developed a hematoma at the insertion site. the hematoma was expressed and a low pressure femstop was applied. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643 apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were not provided.
31150 | 1220648-2024-06593 | 18649325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18649325&pc=OZD | 01/27/2024 | 03/18/2024 |  | 02/06/2024 |  |  | 01/27/2024 | 02/06/2024 |  |  | 11/30/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024430716 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was successfully placed but a hematoma formed, and oozing was uncontrolled. the physician decided to remove the impella and exchanged the impella sheath for a 16fr cook sheath. the impella was reinserted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned. the root cause of the access site bleeding was most likely due to lack of vessel recoil since they needed a larger sheath to stop the bleed.
31151 | 1220648-2024-06594 | 18649466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18649466&pc=OZD | 01/29/2024 | 04/01/2024 |  | 02/06/2024 |  |  | 01/30/2024 | 02/06/2024 |  |  | 09/30/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410562 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 29 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported an impella cp was implanted in a 29-year-old female patient for mechanical circulatory support. the pump was replaced. a blood transfusion was given. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury". | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description provided.
31152 | 1220648-2024-06595 | 18649673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18649673&pc=OZD | 01/26/2024 | 03/25/2025 |  | 02/06/2024 |  |  | 01/26/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153747 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 59-year-old male patient needed mechanical circulatory support. it was reported that the nurse observed the automated impella controller (aic) turn red and display error messages/alarms. no other staff witnessed the screen. the staff members were also unsure if the impella was started dry with the red catheter in place before going inside the patient. there was no patient harm reported. both the aic and the impella were replaced. | [5 empty] | The aic has been returned for evaluation but has not yet been evaluated. upon completion of the investigation, a supplemental report will be submitted. | The investigation of automated impella controller (aic) - pump issue has been completed. the data logs were reviewed and showed an impella stopped alarm had occurred due to the console identifying the motor current as being too high/higher than expected. the real time (rt) logs were reviewed and revealed that the motor current had suddenly risen to 1510 ma for a moment. the console was returned for investigation and a known-good pump was tested with the console and confirmed to be working correctly. no issues were found with the console. the related pump investigation concluded that the reported complaint was most likely caused by a use issue. d2 common device name revised on manufacturer device report in accordance with updated procedures. d4 model number was inadvertently entered incorrectly on the initial manufacturer device report number. f6/f8-should have been left blank on the initial manufacturer device report number. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number.
31153 | 1220648-2024-06596 | 18651494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651494&pc=OZD | 01/18/2024 | 12/19/2024 |  | 02/06/2024 |  |  | 01/18/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1157582 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during prep, the automated impella controller (aic) had 'purge cassette is not removed' alarm leading to console exchange. the new aic was able to be used without issue. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the purge disc/flag issues has been completed. console logs from the reported day of event are mostly consistent with complaint and show that after boot-up and without purge cassette or disk being connected, console presented with controller failure alarm (#418, due to purge pressure sensor defective). it was preceded by error message ¿purge sensor_defect2¿ indicating purge pressure sensor voltage out of range, which did not resolve after console was power-cycled. the console was returned for evaluation. upon functional testing, the reported issue was reproduced. troubleshooting indicated that the issue was caused by malfunction of pud board and defect of purge pressure senor (transducer). added h6 impact code 4629 corrections to the initial mfr report: d2 updated common device name. g1 mdr reporting contact email updated. g1 mfr site name and address updated. h6 med dev prob code 2900 changed to 2994.
31154 | 1220648-2024-06597 | 18651634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651634&pc=OZD | 12/21/2023 | 03/12/2024 |  | 02/06/2024 |  |  | 01/22/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1714346 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) was inspected upon return from the intensive care unit (icu) and the purge cassette insertion was noted to be filled with sticky liquid. the was hard to open and the inner part (drive and pressure sensor) was very dirty, making it impossible to insert the purge system. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported aic "purge issue" has been completed. the console (aic) was returned for evaluation. inspection of the console performed revealed some minor contamination of purge insert area, but also a broken purge disk detect flag, and contamination under the purge disk retainer. console logs recorded during last use of the console on (b)(6) 2023 show multiple attempts to insert purge cassette, each accompanied by purge disk not detected alarm. therefore, the cause of the issue was a defective component.
31155 | 1220648-2024-06598 | 18651702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651702&pc=OZD | 01/24/2024 | 11/20/2025 |  | 02/06/2024 |  |  | 01/24/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1729674 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Complete Loss of Power | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when moving the patient to a transport bed the aic (automated impella controller) completely turned off with no warning or alarms. the screen went black, and the patient was not being supported at that time. the console had a black screen. rn pressed the power button to turn it back on and initiated support and the aic stayed on. the only alarm visible was ¿unexpected controller shutdown¿, no preceding alarms before that. there is no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Data analysis: console imc logs confirm that unexpected controller shutdown occurred after the console abruptly restarted. device analysis: inspection of the internal circuitry of the console revealed no anomaly. the console was able to run a well known pump and purge line after 20 restarts with no anomaly. conclusion: the unexpected controller shutdown is due to a hardware reset. the root cause could not be determined since failure mode could not be reproduced in the lab. d1: corrected common device name: d4: corrected the model number. d9: added the device return information. g1: corrected the manufacturer email address. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation.
31156 | 1220648-2024-06599 | 18651733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651733&pc=OZD | 01/17/2024 | 07/24/2024 |  | 02/06/2024 |  |  | 01/17/2024 | 02/06/2024 |  |  | 08/31/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397802 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 01/25/2024 |  |  |  |  | The user facility reported a 55-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having chest pain. upon the impella implant, when the 0.018" guidewire entered the ventricle, the patient went into ventricular fibrillation. the patient was defibrillated successfully and the impella was implanted successfully. | Additional information was received. abiomed received a st-elevation myocardia infarction study that determined that the patient's limb ischemia was possibly related to the impella procedure. an ultrasound revealed total right common femoral artery occlusion and total left proximal left posterior tibial and peroneal artery occlusions. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt/af/at (vfib) has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. | D.4 revised model number from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. g.3 date was entered incorrectly on the previously submitted report. the date should of been 20-jun-24. h.2. correction was inadvertently omitted from the previously submitted report. h.6 code 4114 was reported incorrectly on the initial report submitted. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative on the initial report submitted stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-06599 and should not have been. h.6 added code 925 since the initial manufacturer device report number 1220648-2024-06599 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the initial submitted manufacturer device report number 1220648-2024-06599. | The investigation into the ventricular fibrillation (vt/vf) issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
31157 | 1220648-2024-06600 | 18651848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651848&pc=OZD | 01/18/2024 | 12/03/2025 |  | 02/06/2024 |  |  | 01/18/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1140349 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the automated impella controller (aic) displayed red alarm with aic failure during prep. the console was exchanged and there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | A.2 age should have been left blank on the initial manufacturer device report number 1220648-2024-06600 as it is unknown. a.3 unknown should have been selected on the initial manufacturer device report number 1220648-2024-06600 as it is unknown. d.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06600 was submitted. d.4 revised model number, catalog number and serial number as they were reported incorrectly on the initial manufacturer device report number 1220648-2024-06600. d.9 added device return date as it was inadvertently omitted from the initial manufacturer device report number g.1 revised reporting contact email in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06600 was submitted. h.3 revised as it was reported incorrectly on initial manufacturer device report number 1220648-2024-06600. h.6 code 2199 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2024-06600. a new code was added.
31158 | 1220648-2024-06601 | 18651872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651872&pc=OZD | 01/21/2024 | 03/25/2024 |  | 02/06/2024 |  |  | 01/21/2024 | 02/06/2024 |  |  | 08/31/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024391000 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump during a high-risk percutaneous coronary intervention (hrpci). on day three of impella support, bleeding at the femoral access site has been reported after moving the patient. the bleeding required transfusion of 4 units of blood. interventions to control the bleeding were: pause of systemic heparin, redressing the wound, optimizing the angle of the repositioning sheath, sandbag compression of the access site. the next day, the patient was successfully weaned from impella support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of access site bleeding was not determined due to a lack of clinical details describing the cause of the bleed, and the product was not returned.
31159 | 1220648-2024-06602 | 18651883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651883&pc=PYX | 01/22/2024 | 03/08/2024 |  | 02/06/2024 |  |  | 01/22/2024 | 02/06/2024 |  |  | 08/31/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024404287 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella rp support, the 63-year-old male patient had oozing from the right femoral artery. the patient reportedly had low platelets and two units of packed red blood cells (prbcs) were given. the impella access site was redressed with surgi-cell and hemostatic discs applied. the patient was noted to be stable after event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
31160 | 1220648-2024-06603 | 18651914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651914&pc=OZD | 01/24/2024 | 03/11/2024 |  | 02/06/2024 |  |  | 01/24/2024 | 02/06/2024 |  |  | 11/30/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024427488 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem | Pressure Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Ischemia | Ventricular Fibrillation | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | [5 empty] | VA ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 66-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient developed oozing and the heparin was withheld and blood products were given to manage the bleeding. additionally, there were suction events noted by staff. intravenous fluids and two units of blood products were given to the patient and the pump flow issues resolved. the patient also experienced a loss of pulsatility, two units of blood products were provided as well. furthermore, the impella had high purge pressure alarms. the purge cassette was replaced, and the staff monitored the patient. one hour later, the patient became hypotensive, arrhythmias, desaturated, and the impella had intermittent suction. the decision was made by the family to withdraw care. the impella was removed after 67.95 hours of support and the patient expired. no allegations were made towards the impella for cause of death. | [5 empty] | The investigation for the reported high purge pressure, low pump flow, limb ischemia, access site bleeding, and ventricular fibrillation has been completed. the product was returned and investigated. upon visual inspection of the pump, heavy biomaterial was found surrounding the impeller and the purge gap. pump was then connected to aic and purge pressure raised to 1051mmhg and purge flow to <1mmhg. after cutting the catheter near the motor housing the purge fluid started passing through confirming that biomaterial blocking the purge gap. no burr or rough/sharp edges were found on the returned product. the root cause of the high purge pressure issue was most likely heavy biomaterial found between the purge gap. as per product investigation and clinical information, the root cause of the low pump issue was most likely biomaterial. the root cause of the access site bleeding issue was not determined due to insufficient clinical information provided. as the necessary clinical information was not provided in conjunction with the returned device, the root cause of the vt/vf was not determined. added-d9- device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
31161 | 1220648-2024-06604 | 18651926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18651926&pc=OZD | 01/25/2024 | 03/12/2024 |  | 02/06/2024 |  |  | 01/25/2024 | 02/06/2024 |  |  | 09/30/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394413 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported a 30-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when the impella 5.5 pump was plugged into the console, an impella defective error message appeared. the cable was unplugged/replugged and the power on the aic was cycled, but the issue persisted. the pump was removed and a new impella 5.5 pump was subsequently placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported pump not detected issue has been completed. the impella device was returned for evaluation. the pump was visually inspected and no physical damages or abnormalities were found. pump was connected to console and no impella defective alarm was observed. data logs were reviewed and an impella defective alarm was observed after pump plug connect. the cause of the pump not detected issue was not determined due to inconclusive evidence based on field investigation and data log review.
31162 | 1220648-2024-06605 | 18652037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652037&pc=OZD | 01/25/2024 | 03/06/2024 |  | 02/06/2024 |  |  | 01/25/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024405807 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported an unknow patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was inserted but was not in the optimal position. the physician decided to exchange it for a new impella. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of positioning issue was most likely use related per clinical review.
31163 | 1220648-2024-06606 | 18652094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652094&pc=OZD | 01/25/2024 | 03/08/2024 |  | 02/06/2024 |  |  | 01/25/2024 | 02/06/2024 |  |  | 11/30/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427011 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported an impella 5.5 was placed as bridge therapy on a 69 year old male patient for mechanical circulatory support. it was reported that the pump was unable to pass through the axillary artery due to the patient¿s anatomy. the impella cp was placed instead. no patient harm was reported. | [5 empty] | The impella device has not been returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. | The investigation into the delivery issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related as per the clinical description provided.
31164 | 1220648-2024-06607 | 18652121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652121&pc=OZD | 01/26/2024 | 03/20/2024 |  | 02/06/2024 |  |  | 01/26/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1708923 |  |  |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | No patient involvement: the complainant reported that an aic (automated impella controller) experienced a system self-check failure during preventative maintenance. there was no patient involvement / harm resulting from the failure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - system self-check error has been completed since the original report was filed. the console was returned and tested after arriving in fs. the root cause of the issue was the purge fw failed to properly install during a system sw installation.
31165 | 1220648-2024-06608 | 18652195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652195&pc=PYX | 01/26/2024 | 03/14/2024 |  | 02/06/2024 |  |  | 01/26/2024 | 02/06/2024 |  |  | 06/30/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024359003 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a (b)(6) year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an impella rp implant due to having a right coronary artery shutdown. upon implanting the rp, several attempts were made to access either pulmonary artery (pa). staff was able to wire in the right pa only with the right ventricle very dilated. the rp would not track past the tricuspid valve area. after multiple attempts, the implant was aborted. reportable due to potential patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine root cause. the root cause of the delivery issue was most likely patient condition related, as after several attempts the rp would not be tracked past tricuspid valve area due to the vasculature as per the clinical description provided.
31166 | 1220648-2024-06609 | 18652290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652290&pc=OZD | 01/25/2024 | 03/07/2024 |  | 02/06/2024 |  |  | 01/25/2024 | 02/06/2024 |  |  | 10/31/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024416001 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed bleeding from the jp drain at the site of the axillary cutdown. compression applied to solve the issue. one unit of prbc was given. the patient is stable. the complainant reported that the patient developed bleeding from the jp (jackson-pratt drain) drain at the site of the axillary cutdown. compression applied to solve the issue. one unit of prbc was given. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since applying forward suturing and manual pressure resolved the bleeding.
31167 | 1220648-2024-06610 | 18652414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652414&pc=OZD | 01/25/2024 | 11/19/2025 |  | 02/06/2024 |  |  | 01/25/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692374 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/25/2024 |  |  |  |  | It was reported that during a biomed check, the team found an automated impella controller (aic) issue when powering on. a backup was used for support. there was no patient involvement. | [5 empty] | The device was received for evaluation. a follow-up report will be submitted upon completion of the investigation.
31168 | 1220648-2024-06611 | 18652428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652428&pc=OZD | 01/22/2024 | 03/27/2024 |  | 02/06/2024 |  |  | 01/24/2024 | 02/06/2024 |  |  | 09/30/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024398093 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 71-year-old male (race unknown) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella cp implant procedure, the impella was inserted via the 0.018¿ wire through the 14fr sheath easily to the aorta. however, while advancing to the left ventricle (lv), there was high friction and it was not possible to insert the impella any further to the correct position in the lv. the impella was therefore withdrawn, flushed, and a second attempt was made. during the second attempt, there was also high friction and kinking when trying to pass the aortic valve. the impella was suspected to be the problem so it was once again withdrawn. a new 0.018¿ wire was opened and inserted into the left ventricle via pigtail catheter. the impella was flushed. a third attempt was made to insert the impella through the new wire to the lv, but when passing the av, there was once again friction and kinking of the cannula. the impella was unable to be withdrawn from the patient through the wire as the wire was blocked in the impella, so both were withdrawn. the patient¿s condition worsened and was therefore put on artificial lung ventilation. a new impella cp was used and was able to be successfully inserted. the patient underwent the high-risk percutaneous coronary intervention and the impella was explanted after the procedure without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues and the product damage (guidewire damage peripheral vessels) issues has been completed since the initial report was submitted has been completed since the original report was filed. the device was returned and visually inspected. the cause of the delivery issue was most likely use related with cannula kinking during advancement in the left ventricle (lv), per clinical review. the cause of the product damage (guidewire) issue was not determined since product was not returned and due to insufficient clinical details.
31169 | 1220648-2024-06612 | 18652465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652465&pc=OZD | 01/26/2024 | 03/14/2024 |  | 02/06/2024 |  |  | 01/26/2024 | 02/06/2024 |  |  | 07/31/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376935 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Alarm | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had attempted impella cp support due to having st elevated myocardial infarction and the patient had severe arm and chest pain followed by ventricular tachycardia and coded. during preparation, the aic (automated impella controller) turned red and displayed an, "impella defective alarm". staff members stated they were unsure if the impella was started dry with the red cheater in place before going inside the patient. both the aic and the impella were replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported "pump not detected" has been completed. the product was returned for investigation, and no physical abnormalities were observed on the returned pump. the pump ran as expected during a test in a bucket of water. the returned pump did not come with the inserted red lumen. data logs show a pump stop alarm with motor current high alarm while outside of the body. according to the provided clinical details, the pump was started with the red cheater in the pump, and the pump was never inserted into the patient's body. the cause of the pump did not start issue was use related since the red lumen was not removed during impella start, as per the provided clinical information. added- d9 device return date d4-updated model number.
31170 | 1220648-2024-06613 | 18652539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652539&pc=OZD | 01/26/2024 | 02/06/2024 |  | 02/06/2024 |  |  | 01/26/2024 | 02/06/2024 |  |  |  |  | 02/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1776556 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 01/26/2024 |  |  |  |  | Us complainant reported the automated impella controller (aic) had an out of box failure. it was reported that there was a substantial amount of screen interference. the aic was sent back for investigation and repair. no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31171 | 1220648-2024-06614 | 18652593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18652593&pc=OZD | 01/18/2024 | 03/19/2024 |  | 02/06/2024 |  |  | 01/18/2024 | 02/06/2024 |  |  | 11/30/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024424523 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 49-year-old patient received hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). after insertion of the impella cp, insufficient impella flow and suction were observed. the intervention was volume administration. the procedure was completed successfully. the patient suffered no harm as a result of the incident. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found no issues on the pump. the low pump flow could not be reproduced under bench test. the root cause of low flow was most likely due to patient condition as the issue could not be reproduced during field investigation.
31172 | 1220648-2024-06615 | 18653001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18653001&pc=OZD | 01/20/2024 | 03/25/2024 |  | 02/06/2024 |  |  | 01/20/2024 | 02/06/2024 |  |  | 02/29/2024 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023186442 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 77-year-old patient received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. the patient was already in advanced cardiogenic shock when the impella was used. blood tests on the first day of impella use indicated evidence of hemolysis, so the correct position of the impella pump in the ventricle was verified using echocardiogram support. two days after insertion of the impella cp, spontaneous bleeding occurred at the femoral access site and the impella cp was removed. blood tests after impella removal no longer showed any signs of hemolysis. the suspicion of hemolysis could not be proven and led to no further intervention. after removal of the impella cp, the norepinephrine dosage was increased. the patient's hemodynamics became progressively worse as the shock progressed, so therapy was discontinued after consultation with relatives. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The investigation for the reported access site bleed and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and hemolysis could not be determined.
31173 | 1220648-2024-06616 | 18653180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18653180&pc=OZD | 01/22/2024 | 03/20/2024 |  | 02/06/2024 |  |  | 01/22/2024 | 02/06/2024 |  |  | 07/31/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024388065 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 63-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had oozing from the right femoral artery and had low platelets. the patient was given two units of packed red blood cells (prbc) and the site was re-dressed with surgi-cell and hemostatic discs were applied. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and (b)(6)apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | He investigation into the access site bleeding ¿ major issues has been completed since the original report was filed. the device was not returned for investigation. the cause of the access site bleeding issue was not determined since product was not returned and not much clinical description is provided.
31174 | 1220648-2024-06617 | 18654376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18654376&pc=OZD | 08/23/2023 | 03/21/2024 |  | 02/06/2024 |  |  | 01/18/2024 | 02/06/2024 |  |  | 02/28/2025 |  | 02/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024323122 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The complainant from a study reported a 55-year-old male patient had a groin hematoma following impella cp explant. the event was determined to be possibly related to the device and the impella procedure. the patient was noted to have recovered without any residual effects. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding - major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue could not be determined because the product was not returned, and the clinical description provided limited information.
31175 | 1220648-2024-06552 | 18639739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18639739&pc=OZD | 01/14/2024 | 03/13/2024 |  | 02/05/2024 |  |  | 01/14/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024416807 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 44-year-old female patient on impella cp support experienced hemolysis. packed blood cells had been given. impella position was noted to be at 4.5cm. therefore, it was retracted 1cm so the impella was mid ventricular again. the patient was stable and was successfully weaned after 14 days on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ | The investigation into the hemolysis issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
31176 | 1220648-2024-06553 | 18640324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18640324&pc=OZD | 01/17/2024 | 03/20/2024 |  | 02/05/2024 |  |  | 01/17/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024412256 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 54-year-old male patient on impella cp support had confirmed hemolysis. therefore, the chief physician made decision to explant impella it could not be ruled out if hemolysis was caused by the pump. the positioning of the pump has been confirmed multiple times via echo. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the hemolysis issue cannot be determined since product was not returned and sufficient clinical information was not provided.
31177 | 1220648-2024-06554 | 18640399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18640399&pc=OZD | 01/18/2024 | 02/29/2024 |  | 02/05/2024 |  |  | 01/18/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024412840 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following high-risk percutaneous coronary intervention for male patient on impella cp support, the impella stopped and the team was unable to restart the pump. the pump stopping occurred at the end of procedure, after revascularization, but before pump weaning process. there was no visible thrombus seen around outlet after explant, and the act > 300 during all procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the pump was returned for investigation, and no physical abnormalities were observed. during a test run in a bucket of water, a pump stop with the motor current high alarm appeared. upon further investigation, biomaterial was observed in the purge gap, and after cutting the catheter near the motor housing, blood was found in the purge line. data logs were reviewed finding saturated pump flows followed by several pump stops. the cause of the pump stop was due to biomaterial. added d9-device return date.
31178 | 1220648-2024-06555 | 18640529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18640529&pc=OZD | 01/17/2024 | 05/20/2024 |  | 02/05/2024 |  |  | 01/17/2024 | 02/05/2024 |  |  | 06/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024378528 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhagic Stroke | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 01/17/2024 |  |  |  |  | A stemi-dtu study reported a 58-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the impella access site while on support. the stemi-dtu study found that the impella procedure was possibly related to the hematoma. there was no known intervention performed for the hematoma. | Additional information was received. additionally, the st-segment elevation myocardial infarction-door-to-unload study found that the patient had a intracranial hemorrhage and had brain surgery. a left craniectomy decompressive for the bleed was performed. the study found that the impella procedure and pci procedure were probably related to the patient's intracranial hemorrhage. the patient expired thereafter. per medical safety officer review, there is no association between impella use and outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer. additional information was received. therefore, the investigation will be reopened and upon investigation re -closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system for use during cardiogenic shock section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h.1 revised as previously entered incorrectly. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4414 was reported incorrectly on the previously submitted report. the correct code has been added to type of investigation. h.10 additional manufacturer narrative instructions for use use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the initial submission stated that the device was not returned but the device was discarded. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. | The investigation for stroke/cerebrovascular accident (cva) was completed. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
31179 | 1220648-2024-06556 | 18640542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18640542&pc=OZD | 01/17/2024 | 08/14/2025 |  | 02/05/2024 |  |  | 01/17/2024 | 02/05/2024 |  |  |  |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1079852 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella support ongoing when the automated impella controller (aic) alarmed for controller failure and aic replacement. the aic was replaced and support proceeded without harm or injury from the interruption. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | D.1 and d.2 revised brand name, common device name and model number in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06556 was submitted. e.1 addded fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06556. g.1 revised mdr reporting contact email in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06556 was submitted. addded contact fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06556. h.6 code 3221 should not have been reported in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06556 was submitted.
31180 | 1220648-2024-06557 | 18640570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18640570&pc=OZD | 01/21/2024 | 02/29/2024 |  | 02/05/2024 |  |  | 01/21/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419328 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2024 |  |  |  |  | The user facility reported an 83-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support that several unsuccessful attempts were made to advance the impella around the subclavian artery. the impella was removed, run in sterile water and pump stop occurred. a new impella was inserted. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency | The investigation into the pump stop issue has been completed. the device was not returned for investigation. based on the clinical information provided, the pump was unable to be delivered around the subclavian despite multiple attempts. the root cause of the delivery issues was most likely patient condition related. data logs show that there was a motor current spike with alarm 13 impella stopped motor current high and the pump was unable to run successfully in the bucket of sterile water. the root cause of the pump did not start issue was not determined since product was not returned.
31181 | 1220648-2024-06558 | 18640752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18640752&pc=OZD | 01/22/2024 | 03/04/2024 |  | 02/05/2024 |  |  | 01/22/2024 | 02/05/2024 |  |  | 03/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342510 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp pump placed for support of a patient admitted in stemi (st elevation myocardial infarction). the pci (percutaneous coronary intervention) was done with the support and the pump was explanted after 16 hours when the patient had hemolysis signs. there was no lasting harm or other intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. the device was not returned for evaluation. data logs were reviewed finding no motor current spikes and few suction alarms were triggered near the end of the case. no other issues were found on the logs. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. d4-model number: impella cp h6-impact code: 4627.
31182 | 1220648-2024-06559 | 18640858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18640858&pc=PYX | 01/20/2024 | 03/05/2024 |  | 02/05/2024 |  |  | 01/20/2024 | 02/05/2024 |  |  | 01/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024328959 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2024 |  |  |  |  | The user facility reported a 68-year-old male patient was on impella rp flex support due to heart failure with reduced ejection fraction. while on support, low pump flows were noted and the pump was explanted. upon explant, a clot was noted in the outflow. | [5 empty] | The impella device has not been returned for evaluation. a supplemental report will be submitted upon completion of the investigation. | The investigation for the reported low pump flow has been completed. the impella rp flex pump was not returned for investigation. clinical details confirmed a clot was observed in the outlet when the pump was removed. data logs were reviewed finding that a motor current spike occurred after the case was initiated and a drop in flows was observed compare to p-levels. therefore, the cause of the low pump flow issue most likely was the biomaterial ingested. the instructions for use provides suggestions on how to prevent thrombus formation. it states "¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ h6-clinical code 4440.
31183 | 1220648-2024-06560 | 18640888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18640888&pc=OZD | 01/21/2024 | 03/08/2024 |  | 02/05/2024 |  |  | 01/22/2024 | 02/05/2024 |  |  | 07/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379821 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 63-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient had a stroke with left sided weakness. neuro suspects cardioembolic origin. the patient was treated with vasopressors to increase cerebral profusion. it was further reported hematoma on the impella insertion site with blood products administered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva and the access site bleeding ¿ major issues has been completed since the original report was filed. the device was not returned for investigation. as no clinical information was provided, the root cause of the stroke/cva could not be determined. the root cause of the access site issue was not determined since product was not returned and not much information is provided in the clinical description.
31184 | 1220648-2024-06562 | 18641203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641203&pc=PYX | 01/22/2024 | 03/05/2024 |  | 02/05/2024 |  |  | 01/22/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024413912 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 31-year-old male with right heart failure was implanted with an impella rp device for mechanical circulatory support post surgery. while the patient was on support, the pump flows were lower than expected. blood volume was given to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. data logs were reviewed which confirmed low pump flow at initiation and it remained the entire case. however, without the device nor sufficient clinical information for evaluation, the exact root cause of the low pump flow could not be determined.
31185 | 1220648-2024-06563 | 18641394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641394&pc=OZD | 01/23/2024 | 02/05/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 11/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426623 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  |  | Initial submission |  |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump would not advance past the aortic arch during attempted delivery due to the patient's vessel size and anatomy. due to the noted resistance, the implant was aborted and the pump was removed from the patient. there was no patient harm. | [5 empty] | The investigation into the delivery issue has been completed. the impella products were not returned for evaluation. based on the clinical review, the root cause of the delivery issue was most likely due to patient condition since the vessel size was too small for the device.
31186 | 1220648-2024-06564 | 18641507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641507&pc=OZD | 01/23/2024 | 03/12/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407863 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump had constant suction despite increasing volume and repositioning. blood products were provided. the patient remained on impella cp support. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. the root cause of low flow was biomaterial ingestion. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31187 | 1220648-2024-06565 | 18641512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641512&pc=OZD | 01/23/2024 | 12/05/2025 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  |  |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320564 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 66-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the pump stopped after a controller error alarm. the automated impella controller (aic) was exchanged, which resolved the issue. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. f6 and f8 should have been left blank. h6 type of investigation code added.
31188 | 1220648-2024-06566 | 18641595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641595&pc=OZD | 01/15/2024 | 03/04/2024 |  | 02/05/2024 |  |  | 01/15/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418202 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old male noted as high-risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient experienced oozing at their access site during impella cp support. the condition was managed with assistance from surgical, as angle matching positioning had not been managed correctly following implant. support continued following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major event has been completed. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved after angle matching.
31189 | 1220648-2024-06567 | 18641622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641622&pc=OZD | 01/23/2024 | 03/08/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 11/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella cp implant that was unsuccessful. the patient had a torturous iliac and multiple attempts were made to balloon the arch, but the staff was unable to pass the impella over the arch. the impella implant was aborted, and pci was finished with anesthesia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition as the patient had heavily calcified aortic arch per clinical details. d4-updated model number.
31190 | 1220648-2024-06568 | 18641645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641645&pc=OZD | 01/23/2024 | 02/05/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419258 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported an 80-year-old male noted as high-risk percutaneous coronary intervention was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 pump was unable to track through the patient's anatomy during attempted delivery. due to the noted resistance, the implant was aborted and an impella cp pump was placed for patient support. there was no patient harm. | [5 empty] | No product was returned. the root cause of the delivery issue was most likely due to patient condition since the vessel size was too small for the device.
31191 | 1220648-2024-06569 | 18641690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641690&pc=OZD | 01/17/2024 | 02/26/2024 |  | 02/05/2024 |  |  | 01/17/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024413574 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 46-year-old male patient on impella cp support experienced leg ischemia. a reperfusion shunt was implanted. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31192 | 1220648-2024-06570 | 18641713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641713&pc=OZD | 01/23/2024 | 03/06/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415372 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that despite a jp (jackson-pratt drain) drain being placed following impella implant, the patient had to be taken to the or for a washout after one day of support. support was maintained during and after the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-updated model number to impella 5.5.
31193 | 1220648-2024-06571 | 18641791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641791&pc=OZD | 01/20/2024 | 09/20/2024 |  | 02/05/2024 |  |  | 01/20/2024 | 02/05/2024 |  |  | 06/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/20/2024 |  |  |  |  | The user facility reported an impella cp in a 64yo female patient for mechanical circulatory support. it was reported oozing at the impella access. the team was able to place sutures to tighten the arteriotomy. additionally, patient experienced a stroke, unknown if it is related to the impella. | It was noted that the patient expired in the operating room. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been returned for evaluation. a follow up will be submitted upon completion of the investigation. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06571 was provided in sections b2, b5, g3/4, h1, and h6. | The investigation for the reported access site bleed and stroke has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause of the access site bleed. the root cause of the access site bleeding was most likely due to use issue since applying additional sutures to tighten up arteriotomy resolved the bleeding. clinical details provided were not sufficient to determine a root cause for the stroke. therefore, the root cause of the stroke could not be determined. the instructions for use provides details on management of access site bleed. it states ¿assess access site for bleeding and hematoma.¿
31194 | 1220648-2024-06572 | 18641806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641806&pc=OZD | 01/24/2024 | 03/12/2024 |  | 02/05/2024 |  |  | 01/24/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411519 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 69-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump which was supporting a patient with a prior asd (atrial septal defect) surgical repair and within this hospital stay for a mvr (mitral valve replacement). the team observed signs of hemolysis on day 5. the patient had lab values rising and the patient was on crrt (continuous renal replacement therapy). the pump remains on for support despite this harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis issue has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed no motor current spikes during the case. the suction alarms were triggered in the first 40 hours then were resolved. logs also showed low flow occurred for entire case and some alarms "impella in aorta" but were resolved immediately. the root cause of hemolysis was not determined since product was not returned and insufficient evidence from data log, clinical review. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ it also states ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31195 | 1220648-2024-06573 | 18641812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641812&pc=OZD | 01/21/2024 | 03/05/2024 |  | 02/05/2024 |  |  | 01/21/2024 | 02/05/2024 |  |  | 03/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347266 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2024 |  |  |  |  | The user facility reported a 74-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that when peel-away sheath was removed, bleeding was noted from the arteriotomy. one of the two percloses was noted to have failed which resulted in the bleed back. due to the noted condition, the patient was weaned, and the pump was explanted. one perclose and manual pressure were utilized to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and "4643" apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since it was determined that failure of one of the two perclose was causing the bleed back from arteriotomy.
31196 | 1220648-2024-06574 | 18641841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641841&pc=OZD | 01/22/2024 | 02/26/2024 |  | 02/05/2024 |  |  | 01/22/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024422064 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was coding during the entire procedure. the first pump was removed due to delivery/kinking issues. this 2nd impella cp was placed but removed shortly thereafter due to vascular damage of the right femoral artery. the damage caused bleeding at the site; manual pressure was applied but the patient expired when cardiopulmonary resuscitation (cpr) failed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta. | The investigation into the vascular damage - peripheral vessel has been completed. the device and data were not returned for investigation. the root cause of the vascular damage - peripheral vessel issue was not determined since product was not returned and there is a lack of clinical details.
31197 | 1220648-2024-06575 | 18641873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641873&pc=OZD | 01/23/2024 | 03/06/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024417623 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported an 83-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had hemolysis and access site bleeding. the access site bleeding was managed by applying sandbags for pressure, a mattress suture was placed and five units of blood products were provided. additionally, the patient had 'cola' colored urine. the impella was repositioned and after repositioning, the patient's urine was tea colored. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the access site bleeding ¿ major, and the hemolysis issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of access site bleeding and the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
31198 | 1220648-2024-06576 | 18641945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18641945&pc=OZD | 01/23/2024 | 03/04/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  |  |  | 02/05/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1671771 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The complainant reported that the aic (automated impella controller) was not able to detect two cassettes after insertion. the aic was removed from service. no patient was involved. | [5 empty] | The impella device was received from the customer and therefore,¿an evaluation of the device was is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. the issue with the aic's inability to detect the purge disc was confirmed to be caused by a broken purge flag.
31199 | 1220648-2024-06577 | 18642007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642007&pc=OZD | 01/24/2024 | 03/12/2024 |  | 02/05/2024 |  |  | 01/24/2024 | 02/05/2024 |  |  | 11/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423599 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Mitral Perforation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having non-st elevated myocardial infarction/cardiogenic shock. while on support, it was found that the patient had a ruptured mitral cord. staff believed that the impella could be a factor to the rupture due to the impella's position being deep in the left ventricle. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the reported heart valve damage and positioning issue has been completed. the impella device was not returned for evaluation. clinical details provided were sufficient to determine the root cause of the positioning issue. the root cause of the positioning issue was most likely patient movement, as the pump was found to be seated deep during echo which was performed after the patient was transferred. however, there was not enough details to determine the root cause of the vascular damage. therefore, the root cause of the heart valve damage issue was not determined. added- b1 selected product problem, h6 problem code 3009, h6 component code 925 h6-updated model number.
31200 | 1220648-2024-06578 | 18642039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642039&pc=OZD | 01/23/2024 | 03/06/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 11/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432397 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Improper Flow or Infusion | Separation Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was high purge pressure error. in addition, the patient had hemolysis and tissue plasminogen activator did not resolve the issue. the patient had cola colored urine so blood products were provided. the decision was made to explant the pump due to high purge pressure and patient's hemolysis. when explanting the pump, the impella came apart at the outlet/cannula junction due to the cross clamp being on the pigtail. the surgeon thought that the impella was pulled far enough, but the clamp was on the pigtail holding the catheter in place. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related hemolysis event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis, high purge pressure, and cannula damage has been completed. the product was not returned for evaluation. data logs were returned, and small motor current spike followed by gradual rise in purge pressure for 10 minutes was observed. saturated flows observed at p-2 during this time. no other issues were found with the logs. as per clinical, hemolysis likely occurred during high purge pressure event. the pump pigtail/cannula got detached at the outflow due to clamping at the pigtail in graft while pulling the catheter. the cause of the hemolysis issue was most likely due to biomaterial per data log and clinical review. the cause of the high purge pressure issue was most likely due to biomaterial per data log and clinical review. the cause of the cannula/pigtail detachment - peripheral vessels issue was most likely use related due to clamping the pump pigtail in the graft and pulling the pump during explant per clinical review. added h6-problem code-1491.
31201 | 1220648-2024-06579 | 18642058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642058&pc=OZD | 01/23/2024 | 03/11/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024417519 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 66-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the pump was removed, an extravasation from the vessel was visualized. a balloon tamponade was performed with addition of iv protamine. the bleed improved, but was still present, so manual pressure was held and a covered stent was placed. stent deployed and patient stabilized with extravasation at site no longer present. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was filed. the device was not returned for investigation. based on the clinical information provided, the most likely cause of the peripheral vessel vascular damage was patient condition.
31202 | 1220648-2024-06580 | 18642072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642072&pc=OZD | 01/23/2024 | 03/04/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423065 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 39-year-old male of an unknown race was noted as high-risk percutaneous coronary intervention was implanted to be with an impella cp device for mechanical circulatory support. it was reported that while on support, the impella migrated into the left ventricle and there was an, ¿impella stopped motor current high¿ alarm and the impella was on p0. the physician repositioned the pump and restarted with success. the staff found that the cannula was bent in the left ventricle. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop and the product damage (cannula kink) issues has been completed since the original report was filed. the device was not returned for investigation. data logs were reviewed which showed about 2 hours into support, the motor current was spiked to 1500 ma and pump position in the ventricle was observed. a high motor current pump stop alarm was found relative to the event of impella migrating into the lv and cannula being bent during repo sheath insertion. pump was able to be restarted later at p-8 and pump ran without any issues during the case. the cause of the temporary pump stop issue was most likely use related due to improper positioning leading to interaction with patient per clinical and log review. the cause of the positioning issue was use related due to improper technique per clinical review.
31203 | 1220648-2024-06581 | 18642220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642220&pc=OZD | 01/24/2024 | 04/05/2024 |  | 02/05/2024 |  |  | 01/24/2024 | 02/05/2024 |  |  |  |  | 02/05/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1088056 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 64-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) displayed a console failure/error alarm. the support continued despite the malfunction. the aic was returned for evaluation once patient was off support. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - controller error (yellow alarm) issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the failure was unable to be reproduced. found a hair crack in the purge retainer which is unrelated to the reported failure. the root cause for the controller error was not determined due to the inability to reproduce.
31204 | 1220648-2024-06582 | 18642334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642334&pc=OZD | 01/24/2024 | 01/06/2025 |  | 02/05/2024 |  |  | 01/24/2024 | 02/05/2024 |  |  |  |  | 02/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263204 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and the automated impella controller (aic) had a controller error (yellow alarm). the console was replaced successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) controller error yellow alarm has been completed. data logs from the console confirm multiple loss of communication pbm1 controller errors during a clinical case. lt logs from the console reveal that there was a single cell failure. the console was returned for evaluation finding a battery failure alarm issued by the console on boot up in the lab. batttest check of the console confirmed that cell 1 of battery 2 had a failure. visual inspection of the pbm revealed corrosion of the copper traces in vicinity of the super capacitors c69 and c70. the root cause for the pbm communication loss errors was due to contamination of the pbm by leakage of electrolyte from supercapacitors causing corrosion of the copper traces in vicinity of the super capacitors. added d9 device return date corrections to the initial mfr report: d2 common device name updated d4 model number updated f6/f8 should have been left blank. g1 mdr reporting contact email updated h6 impact code changed to 4629.
31205 | 1220648-2024-06583 | 18642425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642425&pc=OZD | 01/19/2024 | 03/01/2024 |  | 02/05/2024 |  |  | 01/24/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024414143 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that patient on impella support developed cerebral hemorrhage. the heparin in the purge fluid was changed to sodium bicarbonate, and the patient's progress is being monitored. warfarin was also being administered, and it is unclear whether the anticoagulation had an effect. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva issue has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined.
31206 | 1220648-2024-06584 | 18642536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642536&pc=OZD | 01/24/2024 | 02/26/2024 |  | 02/05/2024 |  |  | 01/24/2024 | 02/05/2024 |  |  | 10/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024424394 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after impella cp placement there was sluggish flow down to the impella accessed limb. the cold limb/ischemic prompted premature explant. the groin site was closed by contralateral leg balloon tamponade and the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31207 | 1220648-2024-06585 | 18642604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18642604&pc=OZD | 01/23/2024 | 03/08/2024 |  | 02/05/2024 |  |  | 01/23/2024 | 02/05/2024 |  |  | 11/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426975 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 64-year-old female noted as high risk percutaneous coronary intervention was to be implanted with an impella 5.5 device for mechanical circulatory support. during the implant, resistance was met causing the motor housing to buckle at anastomoses. the staff exchanged the .018 for a v18. left ventricle access was reobtained and a second attempt was made to pass the impella into the left ventricle. again, the inlet of the impella reached innominate. the motor housing would not pass and buckled at anastomosis. the procedure was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition as the motor housing would not pass and buckle at anastomosis per clinical details.
31208 | 1220648-2024-06587 | 18643536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18643536&pc=OZD | 01/29/2024 | 03/18/2024 |  | 02/05/2024 |  |  | 01/29/2024 | 02/05/2024 |  |  | 09/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401749 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2024 |  |  |  |  | The user facility reported an impella pump in a 54yo male patient for mechanical circulatory support. it was reported the patient exhibited hemolysis symptoms, and the pump was kinked. to provide ongoing care, the team left the patient's pump on. | [5 empty] | The impella pump has not been returned for evaluation. a follow up will be submitted upon completion of the investigation. | The investigation into the product damage (cannula kink) and the hemolysis issues has been completed since the original report was filed. the device was not returned. the cause of the optical signal issue was not determined since product and data logs were not returned. the cause of the hemolysis issue was not determined since product, data logs were not returned and due to insufficient clinical information.
31209 | 1220648-2024-06588 | 18644418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18644418&pc=OZD | 01/24/2024 | 02/07/2024 |  | 02/05/2024 |  |  | 01/24/2024 | 02/05/2024 |  |  | 05/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024352443 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 was unable to pass into the innominate artery. during several attempts at insertion, the catheter began kinking within the part of the graft that was outside the body. ultimately unable to pass the impella due to patient's anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation completion a supplemental will be filed. the root cause of the delivery issue was most likely due to patient condition since it was kinking because it was getting stuck on the anatomy. | The investigation into the delivery issues has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the delivery issue was most likely due to patient condition since it was kinking because it was getting stuck on the anatomy.
31210 | 1220648-2024-06590 | 18644986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18644986&pc=OZD | 01/24/2024 | 03/11/2024 |  | 02/05/2024 |  |  | 01/24/2024 | 02/05/2024 |  |  | 08/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024392998 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to acute heart failure. while on support, there was oozing from the access site and pressure dressing was applied. additionally, there were placement signal unreliable alarms, however, support was continued. furthermore, the impella's position was found to be flipped posterior. the staff attempted to reposition the pump unsuccessfully. | [5 empty] | The investigation into the access site bleeding ¿ minor, the positioning issues, and the optical signal issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the optical signal issue was not determined since product was not returned and based on inconclusive results per data log review. the cause of the positioning issue was most likely use related per clinical review. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31211 | 1220648-2024-06589 | 18645018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18645018&pc=OZD | 01/27/2024 | 03/19/2024 |  | 02/05/2024 |  |  | 01/27/2024 | 02/05/2024 |  |  | 08/31/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380607 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Fluid/Blood Leak | Increase in Pressure | Pumping Stopped | Improper Flow or Infusion |  |  | Renal Failure | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported high purge pressure and purge system blocked alarms. the clinician also stated that there appears to be crystallization within the purge tubing. the clinician believes that purge tubing has not been changed for several weeks. the purge cassette was changed but purge pressures remained high, and alarms continued. it was further reported that the impella had saturated flows. advised was given to replace the impella. there was no harm to the patient. | It was further reported that purge pressures remained high and continued with a subsequent pump stop. patient sustained acute renal failure after prolonged hypotension due to previous pump stop. patient was placed on continuous renal replacement therapy (crrt). no further details were provided. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported high purge pressures, renal failure, pump stop, saturated flow, purge leak has been completed. the impella device nor data logs were returned for evaluation. as per clinical high purge pressure and saturated flows were observed to be chained events. pump stop was followed subsequently which eventually led to acute renal failure due to prolonged hypotension. no purge leak was reported. the cause of the high purge pressure issue was not determined since product and data logs were not returned. the cause of the pump stop issue was not determined since product and data logs were not returned. the cause of the renal failure issue was not determined since product and data logs were not returned. the instructions for use states the following:¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ added- b1 selected adverse event, b2 selected requiring intervention; b5 event narrative; h6 clinical code 2041; impact code 4641, 4629; problem codes 1491, 1503 d4-updated model number. h6-remove 1250 from problem code.
31212 | 1220648-2024-06592 | 18646230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18646230&pc=PYX | 01/28/2024 | 03/15/2024 |  | 02/05/2024 |  |  | 01/28/2024 | 02/05/2024 |  |  | 03/31/2025 |  | 02/05/2024 |  | RP FLEX W SMART ASSIST SET, US |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED INC. |  |  | US |  | 2024390038 |  | 1000323 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2024 |  |  |  |  | The user facility reported an impella rp flex on a 65yo male patient for mechanical circulatory support. it was reported that while on support there was a spike in motor current and device flows dropped from 3.4 l/min on p7 to 2l/min on p7 with patient hemodynamics dropping. physicians elected to replace pump. no additional information was provided. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not returned for evaluation. data logs show a spike in motor current with placement signal shifting, followed by low pump flow during the start of the case. according to clinical records, the pump was replaced after observing the motor current spike and decrease in pump flow. the cause of the low pump flow issue was most likely biomaterial, based on data log review. | The impella rp has not been returned for evaluation. a follow-up report will be submitted upon completion of the investigation.
31213 | 1220648-2024-06591 | 18647365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18647365&pc=OZD | 01/24/2024 | 03/06/2024 |  | 02/05/2024 |  |  |  | 02/05/2024 |  |  | 09/30/2025 |  | 02/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411318 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 62-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that there was bleeding that was noted around the repositioning sheath, and it was coming through the crease of the blue hub. the staff was unable to control the bleeding, so the physician removed the impella, reinserted a new big sheath, and then reinserted the impella. during the exchange, the patient arrested and return of spontaneous circulation was achieved after the impella restarted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the vf/vt/af/at issues has been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
31214 | 1220648-2024-06536 | 18631658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18631658&pc=OZD | 01/22/2024 | 03/12/2024 |  | 02/02/2024 |  |  | 01/22/2024 | 02/02/2024 |  |  | 05/31/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345259 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and upon explanting the pump, there was removal issues. the impella was removed from the left ventricle, but the doctor was unable to pull the impella back into the graft. the doctor made the decision to open the chest to expose the base of the graft at the aorta. graft was un-tunneled and brought into chest cavity and clamped. at this time, patient had bleeding from an unknown source and the patient expired in the operating room after having 98.82 hours of support. the clinical representative was unsure if the impella was related to the cause of death. | [5 empty] | The investigation for the reported issue has not yet been completed. upon its completion, a follow up report will be submitted. | The investigation for the reported removal issues has been completed. the impella device was not returned for evaluation. the root cause of the removal issue was not determined since product was not returned and insufficient clinical information regarding the removal was provided. the instructions for use states "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury."
31215 | 1220648-2024-06537 | 18632758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18632758&pc=OZD | 01/12/2024 | 03/01/2024 |  | 02/02/2024 |  |  | 01/12/2024 | 02/02/2024 |  |  | 11/30/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024425924 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 63-year-old female patient of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while trying to place the pump, the device would no advance due to the patient's groin anatomy. the implant was subsequently aborted because of the noted issue. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related.
31216 | 1220648-2024-06538 | 18632850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18632850&pc=OZD | 01/23/2024 | 05/17/2024 |  | 02/02/2024 |  |  | 01/17/2024 | 02/02/2024 |  |  | 09/30/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399062 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pressure Problem | [6 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 01/17/2024 |  |  |  |  | The user facility reported a 59-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that after the impella placement a high purge pressure and purge blocked alarm occurred. a decision was made to replace the impella with a one. the new impella was inserted successfully. there was no patient harm reported. | Additional informations received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured acute limb ischemia of the left lower extremity (lle) with a "possible" relationship to the impella device used: ¿patient developed acute limb ischemia of the lle after removal of the impella and va ecmo cannulas (left foot cold, no pulse able to be palpated or dopplerable). patient was taken to the or emergently by vascular surgery for left femoral cutdown and embolectomy. doing well post op. |  |  |  |  | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states in section: using the automated impella controller with the impella rp with smartassist system catheter that ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ it also states in section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- d6a/d6b. | The investigation into the high purge pressure issue has been completed since the original report was filed. the impella cp pump was returned for investigation. visually inspected and biomaterial ingested over the purge gap was observed. no purge line kinks, or other abnormalities were found. the cause of the high purge pressure issue was due to biomaterial ingestion over the purge gap per field investigation. | Section b1, b2,b3, b5, and h6 was updated as per additional information. the event is updated from malfunction to serious injury as per new additional information.
31217 | 1220648-2024-06540 | 18632863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18632863&pc=OZD | 01/19/2024 | 07/25/2024 |  | 02/02/2024 |  |  | 01/19/2024 | 02/02/2024 |  |  |  |  | 02/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401769 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 70-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that automated impella controller (aic) had a controller error alarm. the aic was exchanged. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) yellow alarm has been completed. data logs were reviewed which showed that after running for 14 hours 29 minutes, an error ¿processsampleforopmdatapacketcrcerror: calculator placement signal 1045¿, which resulted in ¿controller error (opm comm crc invalid) [optical pump connected] ¿ alarm,¿ event #1045 during the clinical procedure. this confirms the reported failure mode. the event #1045 was resolved by itself after 3 minutes. real time logs confirmed that all signals received from the optical bench (snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the console was returned for evaluation. the console functioned without any issue. the root cause for the controller error and loss of placement signal was an optical bench fw communication protocol anomaly, which resulted in the misalignment of data communication packets received by epc. the aic sw operated as designed. added-d9 device return date d2-corrected common device name f6/f8 should have been left blank in the initial report.
31218 | 1220648-2024-06543 | 18633032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633032&pc=OZD | 01/20/2024 | 03/05/2024 |  | 02/02/2024 |  |  | 01/20/2024 | 02/02/2024 |  |  | 11/30/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427015 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2024 |  |  |  |  | The user facility reported a 71-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to having low ejection fraction. while on support, there was a pump stop alarm. the impella restarted on its own. no intervention was needed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. the impella device was not returned for evaluation. data logs were reviewed and confirmed a pump stop. however, without the actual device and additional clinical details, a root cause could not be determined.
31219 | 1220648-2024-06544 | 18633107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633107&pc=OZD | 01/20/2024 | 03/05/2024 |  | 02/02/2024 |  |  | 01/21/2024 | 02/02/2024 |  |  | 10/31/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420591 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp pump placed for support of a patient admitted in cgs. the pump access site had an ooze that prompted the team to infuse 1 unit of rbcs. the pump had supported for over 95 hours when the family made choice to withdraw care and patient expired. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-model number update to impella cph5-updated to yes
31220 | 1220648-2024-06541 | 18633177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633177&pc=OZD | 01/20/2024 | 10/20/2025 |  | 02/02/2024 |  |  | 01/20/2024 | 02/02/2024 |  |  |  |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1427009 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [17 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/20/2024 |  |  |  |  | The user facility reported a 48-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that automated impella controller (aic) had optical signal issues. the medical team troubleshot and remedied the issue by intervention of replacement of the aic. there was no patient harm reported. | The patient experienced hemolysis, the resolution for this issue is unknown. |  |  |  |  | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed: during preventative maintenance the technician set the pmd hardware revision incorrectly. data analysis imc logs confirmed the reported complaint. there were multiple occurrences of optical sensor not supported alarms. review of the automated impella controller (aic) software design documentation revealed that the cause for the optical sensor not supported alarm is due to the aic determining that its hardware components are not compatible with an impella optical sensor. the aic was running software version 12.2, and in this software version the optical bench (opm) is only compatible if the pmd is of hardware revision 256, 257, or 258. pmd hardware revision can be set manually via a debugging utility. imc logs revealed that the pmd hardware revision of ic7326 had been originally set to 257, but during the most recent preventative maintenance the pmd hardware revision had been incorrectly set to 1. the root cause was determined to be technician error during preventative maintenance. during preventative maintenance the technician set the pmd hardware revision incorrectly. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 updated. f6 and f8 filled out erroneously on initial report. h1, h3, h6, and h10 updated.
31221 | 1220648-2024-06545 | 18633327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633327&pc=OZD | 01/22/2024 | 03/06/2024 |  | 02/02/2024 |  |  | 01/22/2024 | 02/02/2024 |  |  | 11/30/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432765 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during percutaneous cardiac intervention (pci). during pci, there was an ooze from the patient's access site that could not be remedied. the manual compression, angle entry adjustment and figure-eight suture did not remedy the oozing. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
31222 | 1220648-2024-06542 | 18633329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633329&pc=OZD | 01/20/2024 | 03/11/2024 |  | 02/02/2024 |  |  | 01/20/2024 | 02/02/2024 |  |  | 11/30/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024432644 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2024 |  |  |  |  | The user facility reported a 41-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis during impella cp support. the lactate dehydrogenase (ldh) and plasma free hemoglobin labs were verified to be hemolyzed. as levels continued to decline the following day, the decision was made to escalate to impella 5.5 support. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." | The investigation for the hemolysis has been completed. the device was not returned for evaluation. however, data logs were reviewed which revealed that the pump was in improper position and remained the rest of the case. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position.
31223 | 1220648-2024-06546 | 18633497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633497&pc=OZD | 01/22/2024 | 03/06/2024 |  | 02/02/2024 |  |  | 01/22/2024 | 02/02/2024 |  |  | 11/30/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427686 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support post surgery. the team delivered the 5.5 via the graft. the delivery caused a dissection and tear at the graft with bleeding. the team removed the graft and surgical repair was done. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation for the vascular damage has been completed. the impella device was not returned for evaluation. clinical details provided were sufficient to determine root cause. the root cause of the vascular damage - great vessel issue was most likely patient condition related due to the patient's anatomy causing resistance during insertion.
31224 | 1220648-2024-06547 | 18633782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633782&pc=PYX | 01/22/2024 | 03/12/2024 |  | 02/02/2024 |  |  | 01/22/2024 | 02/02/2024 |  |  | 10/31/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024413939 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 31-year-old male post cardiac surgery was implanted with an impella rp device for mechanical circulatory support. while on support there were flows that were lower than expected. blood volume was given. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the device was not returned for investigation. based on clinical description and data review, the root cause of the low flow was most likely biomaterial ingestion.
31225 | 1220648-2024-06539 | 18633901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633901&pc=OZD | 01/18/2024 | 03/01/2024 |  | 02/02/2024 |  |  | 01/18/2024 | 02/02/2024 |  |  | 09/30/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400720 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 71-year-old female noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. the impella was inserted but there was resistance when crossing the aortic arch. the case was aborted due to patient¿s anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related.
31226 | 1220648-2024-06548 | 18633975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633975&pc=OZD | 01/18/2024 | 11/19/2025 |  | 02/02/2024 |  |  | 01/23/2024 | 02/02/2024 |  |  |  |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129496 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 60-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. less than an hour after the implant, the impella had an "air in purge" alarm. the purge cassette was replaced, however following the replacement, flows appeared to be saturated, high purge pressure alarm ensued, and an impella stopped alarm appeared. the impella would not restart. the complainant would like the automated impella controller (aic) investigated to rule out any issues with the console. | [5 empty] | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. d4 model number updated. f6/f8 should have been left blank. h3 has been updated to device eval completed. h6 type of investigation code 10 added. | The investigation has been completed. data analysis: there were frequent purge-related alarms including piston blocked, air in purge and purge pressure high. this was eventually followed by motor current high alarms causing pump stops. these proceeded for the duration of pump 479352 on this console. device analysis: the console was returned and performed as expected. there was no difficulties with either pump performance or the purge assembly. see the -1 for notes on the disposable investigation. root cause: there was no issue found with the console's performance. d9: updated the device returned to status. h6: added codes per the results of the investigation and the clinical code. h10: added the results of the investigation. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31227 | 1220648-2024-06549 | 18633999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18633999&pc=OZD | 01/27/2024 | 03/14/2024 |  | 02/02/2024 |  |  | 01/27/2024 | 02/02/2024 |  |  | 11/30/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434164 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported an impella cp in a 62year old female patient with acute myocardial infarction/cardiogenic shock for mechanical circulatory support. it was reported that there was a sudden drop in flow and continuous suction that was unable to be broken at p2. the pump was replaced since all recommendations failed to resolve issue. there was no patient harm reported. | [5 empty] | The impella device has not been returned for evaluation. a follow-up report will be submitted upon completion of the investigation. | The investigation into the low pump flow has been completed since the original report was filed. the device was returned for investigation. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. the root cause of low flow was biomaterial ingestion.
31228 | 1220648-2024-06550 | 18634011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18634011&pc=OZD | 01/27/2024 | 03/14/2024 |  | 02/02/2024 |  |  | 01/27/2024 | 02/02/2024 |  |  | 10/31/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420791 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported an impella cp in a 72yo male patient for mechanical circulatory support. it was reported that the pump accessed limb, the right, became mottled. pulses were found by doppler. additionally the patient had signs of hemolysis with red/amber urine and hemolyzed labs. the team did repositioning and began continuous renal replacement therapy (crrt). | [5 empty] | The impella device has not been returned for evaluation. a follow up will be submitted upon completion of the investigation. | The investigation into the hemolysis and the limb ischemia ¿ reversible issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of hemolysis and the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned.
31229 | 1220648-2024-06551 | 18634931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18634931&pc=OZD | 01/27/2024 | 03/15/2024 |  | 02/02/2024 |  |  | 01/27/2024 | 02/02/2024 |  |  | 05/31/2025 |  | 02/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369003 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2024 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having heart failure with reduced ejection fraction. while on support, a hematoma formed in the medial groin. a femostop was applies and it was managed with success. angio revealed a low stick just above bifurcation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue due to anticoagulation management since withholding the systemic heparin resulted in resolving of the hematoma.
31230 | 1220648-2024-06497 | 18621770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18621770&pc=OZD | 01/18/2024 | 09/20/2024 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  | 07/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024372380 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 70-year-old white male noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that upon insertion of the impella a perforation was noted while advancing into the left ventricle. the physician repositioned the impella 5.5 and a combination of pledgets and bioglue were used for closure of the perforation. the patient remained on support. | Additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06497 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The investigation into the reported ventricle perforation has been completed since the original report was filed. as per clinical patient was having necrotic tissue condition likely leading to perforation during advancement of pump in lv. the cause of the ventricle perforation issue was most likely patient condition related per clinical and data log review.
31231 | 1220648-2024-06498 | 18621794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18621794&pc=OZD | 01/19/2024 | 03/07/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420066 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 61-year-old white female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient developed a thrombus in their left ventricle during impella cp support. discussions were had about escalating to impella 5.5 support, but due to the thrombus, the escalation was put on hold. support was maintained with the impella cp following the observation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use provides suggestions on how to prevent thrombus formation. it states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
31232 | 1220648-2024-06499 | 18621964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18621964&pc=OZD | 01/19/2024 | 09/19/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 08/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024392982 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 2nd 5.5 placed after the 1st was removed after 17 days of support. the 2nd 5.5 was placed via direct access but with difficulty through valve. the team observed the lv (left ventricle) to have perforated. the lv was repaired. the tissue was fragile and tears occurred in the rv, pa (pulmonary artery), and aorta. patient went into pea (pulseless electrical activity ) and expired. reportable for the harm. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mdr 1220648-2024-07257 was provided in sections b2, b5, d6, h1, h6, h10. | The investigation into the ventricle perforation issue has been completed. the device was not returned for investigation. clinical detailed noted physician had difficulty passing across valve and heart tissue was very fragile and continued to tear easily. the root cause of the ventricle perforation was most likely due to the pump being inserted too deep into the ventricle, with the patient's weak heart tissue possible being a factor.
31233 | 1220648-2024-06500 | 18622957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18622957&pc=OZD | 01/08/2024 | 05/28/2024 |  | 02/01/2024 |  |  | 01/08/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396379 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 77-year-old female (race unknown) with heart failure was to be implanted with an impella 5.5 for mechanical circulatory support. the pump was not able to advance through the left axillary into the left ventricle. the case was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, unable to advance 5.5 into right axillary but bilateral anatomy prohibited impella 5.5 as per the clinical description provided.
31234 | 1220648-2024-06501 | 18623132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623132&pc=OZD | 01/17/2024 | 02/21/2024 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422735 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 69-year-old female undergoing off pump coronary artery bypass was implanted with an impella 5.5 device for mechanical circulatory support. immediately following the planned impella 5.5 explant, a stroke alert was initiated on the patient. an urgent computed tomography (ct) scan verified that an embolic stroke occurred post impella removal. the patient was taken to neurointerventional radiology for a mechanical thrombectomy to treat the condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/cva has been completed since the original report was filed. no product was returned. as this clinical information was not provided, the root cause of the stroke/cva could not be determined.
31235 | 1220648-2024-06502 | 18623319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623319&pc=OZD | 01/18/2024 | 02/23/2024 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIMED, INC |  |  | US |  | 2024418071 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 67-year-old female male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that following the impella cp implant, the patient had weak pedal pulses and the extremity became cool to the touch. concerns for distal perfusion and the patient¿s overall worsening condition were discussed and the decision was made to explant the pump the same day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31236 | 1220648-2024-06503 | 18623481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623481&pc=PYX | 01/18/2024 | 10/07/2025 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024413914 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Activation, Positioning or Separation Problem | Device Sensing Problem |  |  |  |  | Air Embolism | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 46-year-old male with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. it was reported the pump had low flows/suction and had signs of an air embolization were observed, that cleared with a p level adjustment. the pump had migrated to the right ventricle and was unable to redeliver to the pulmonary artery. the pump was explanted and replaced with proteck duo for continued support. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported air embolism, low flow issues, and placement issues has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed no signs related to purge leak issues. clinical details provided were sufficient to determine the root cause of the placement and flow issues. the root cause of the flow issue was most likely use related as the pump was zeroed causing false low flows; motor current remains stable during this time period indicating the pump to be flowing within normal ranges. the root cause of the positioning issue was most likely use related as the pump was repositioned under echo guidance and pulled back so that the pump was shallow in the pa; the patient was not sedated during tee and thus bucked and coughed causing the pump to fall into the rv. air bubbles were seen on the echo while the pump was set to p-6. the root cause of the embolism was not determined since product was not returned and the relation to impella is unknown.
31237 | 1220648-2024-06504 | 18623483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623483&pc=OZD | 12/10/2023 | 03/04/2024 |  | 02/01/2024 |  |  | 01/17/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024393868 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella 5.5 support expired from cerebral hemorrhage. it is unknown if the cerebral hemorrhage was related to the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva issue has been completed. the device was not returned for investigation. as the clinical information was not provided, the root cause of the stroke/cva could not be determined.
31238 | 1220648-2024-06505 | 18623558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623558&pc=OZD | 01/18/2024 | 03/06/2024 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  |  |  | 02/01/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1324937 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) lost the placement signal and waveform suddenly while a patient was on support. the aic was swapped for another aic with no issues. there was no patient harm. patient demographics are not available at this time. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) "loss of waveform" issue has been completed. the aic was returned for evaluation. inspection of the front panel optical connector receptacle revealed significant contamination. the internal circuitry was inspected and revealed that power was being supplied to the optical bench from the pbm but with no light coming out of the lamp. it therefore likely that a component critical to the functioning of the lamp and as a result the optical bench was loose or damaged. data logs were reviewed which confirmed a placement signal not reliable alarm. the root cause of the controller error was a defective optical bench. when the placement signal not reliable alarm is issued, a temporary suspension of the trend line is implemented by the aic as a mitigation.
31239 | 1220648-2024-06506 | 18623609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623609&pc=OZD | 01/18/2024 | 03/01/2024 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  | 11/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425917 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported an 80-year-old female noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that upon the echocardiogram the impella was incorrectly measured. the physician felt the impella was too deep and pulled it back, later the same day the impella came out of the left ventricle. the impella was explanted, and the patient blood pressure was stable. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported positioning issue has been completed since the original report was filed. product & data were not returned for review. the root cause of positioning issue was most likely use related.
31240 | 1220648-2024-06507 | 18623665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623665&pc=OZD | 01/04/2024 | 03/20/2024 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  | 08/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024382468 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a male patient of an unknown age in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was in the left ventricle (lv) when there were concerns of heparin induce thrombocytopenia (hit) and heparin was removed. on day seven of the support, the team observed lv clot to be present. the pump was removed and left ventricular assist device (lvad) was placed. the team did note the thrombus may have developed due to lv not moving well during the large infarct myocardial infarction (mi) that damaged the tissue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
31241 | 1220648-2024-06487 | 18623697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623697&pc=OZD | 01/16/2024 | 03/14/2024 |  | 02/01/2024 |  |  | 01/16/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | GM |  | 2024411434 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The complainant reported 63-year-old male patient on impella cp support had access site bleeding that was potentially caused by coagulation disorder due to septic shock. the patient received one unit of packed red blood cells (prbc). bleeding on the insertion side was ultimately managed | Additional information states that hemolysis was reported on the second day of support. the pump position was confirmed to be in a good position. no further details were provided regarding this issue. |  |  |  |  | The investigation for the reported issue has not yet been completed. upon its completion, a follow up report will be submitted. the instructions for use states the following as it relates to access site bleeding: ¿assess access site for bleeding and hematoma.¿ it also states "¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and hemolysis has been completed. the impella device was not returned for evaluation. clinical details were sufficient to determine the root cause of the access bleed. the root cause of the access site bleeding issue was most likely patient condition related, bleeding was potentially caused by coagulation disorder due to septic shock as per the clinical description provided. however, limited details were provided to be able to determine the root cause of the hemolysis as the pump was in a verified good position.
31242 | 1220648-2024-06488 | 18623699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623699&pc=OZD | 04/25/2023 | 01/31/2024 |  | 02/01/2024 |  |  | 04/25/2023 | 02/01/2024 |  |  | 09/30/2024 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2023280187 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/25/2023 |  |  |  |  | The us medical center had an impella cp inserted for a patient who was denied surgery and sent for high risk percutaneous coronary intervention (hrpci) instead. the patient's condition was observed to be poor with low flows and with imaging the team noted there was a left ventricle (lv) perforation. there was a pericardial drain placed and mass transfusion of blood products (>6 units). the team believed the non-abiomed provided .035 guide wire to have been the cause of the lv perforation, not the abiomed provided .018 wire. the pump was explanted in conjunction with the intervention. the patient has survived the event and intervention. | [5 empty] | The investigation into the lv perforation has concluded. no impella pump or guidewire product was returned for benchtop analysis. only the clinical report was reviewed. the cause of the ventricle perforation was use related. the .035 guidewire is non-abiomed provided nor an approved wire. per the ifu: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." "with the 0.018 inch placement guidewire properly positioned in the left ventricle, remove the diagnostic catheter." "use only abiomed-tested and supplied guidewires with the impella catheter. guidewires are specifically designed with unique characteristics to optimize performance of the impella system. guidewires and catheters should always be used in accordance with abiomed¿s instructions. table b.1 lists the alternative guidewires that have been tested and approved for use with the impella cp with smartassist system. table b.1 alternative guidewires guidewire catalog number boston scientific platinum plus¿ st 0.018 in 46-605, model st/0.018/260 boston scientific v-18 control wire¿ st 0.018 in 46-854, model v18/18/300."
31243 | 1220648-2024-06489 | 18623700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623700&pc=OZD | 01/16/2024 | 03/14/2024 |  | 02/01/2024 |  |  | 01/16/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024402093 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Information from protect iv study: a patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). the day after removal of the impella cp, the patient went into cardiogenic shock. possible sepsis was suspected due to elevated infection levels (il6). the patient's treatment included: defibrillation and administration of epinephrine to treat asystole, volume administration and administration of dobutamine and norepinephrine to treat hypotension, transfusion of two units of packed red blood cells to treat low hemoglobin, insertion of a temporary cardiac pacemaker (icd/crt). , dialysis due to liver and kidney failure, cardiopulmonary resuscitation and intubation due to asystole, administration of a dose of vancomycin and antibiotics with tazobac. the hospital stay was prolonged due to the complications. disseminated intravascular coagulopathy caused occlusion of the visceral arteries and femoral arteries, for which no intervention was planned. the patient died on the same day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the infection/sepsis has been completed since the original report was filed. the device was not returned for investigation. because no evidence has been provided and device was not returned, the root cause of the infection cannot be determined.
31244 | 1220648-2024-06490 | 18623701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623701&pc=OZD | 01/16/2024 | 03/20/2024 |  | 02/01/2024 |  |  | 01/17/2024 | 02/01/2024 |  |  | 08/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024384401 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Break | Pumping Stopped | Detachment of Device or Device Component |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The complainant in japan reported that a patient that has been for several months on impella 5.5 support experienced a decreased in motor current waveform, but the patient's hemodynamics were stable, so weaning was performed. the device could not be removed immediately due to conflicts with the cardiovascular surgery schedule. the device was pulled out of the heart and placed near the axillary and subclavian artery. several days later a surgical operation was performed to remove the product, and when an attempt was made to pull the product out of the body by pulling on the shaft, the shaft broke. the product was torn when the cannula was pulled out to the point where it could be seen, so it was all removed safely without leaving anything behind in the body. there was no harm to the patient. a 30-day mdr report is required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella 5.5 & cp ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ impella rp: ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ pump stop: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency | The investigation for the removal and product damage issues has been completed. the actual device was not returned for evaluation. however, a photograph was provided which revealed clear signs of clamping of the cannula. the catheter was torn from the motor side of the pump during removal due to the clamping which elongated the cannula. per the clinical information provided, the root cause of the product damage - detached inflow/ outflow issue was most likely use related to improper technique of clamping the cannula during removal.
31245 | 1220648-2024-06491 | 18623702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623702&pc=OZD | 01/17/2024 | 03/06/2024 |  | 02/01/2024 |  |  | 01/17/2024 | 02/01/2024 |  |  |  |  | 02/01/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1271201 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a self-check failure and a controller failure alarm upon startup. the aic was swapped to resolve the noted issue. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the controller (aic) red alarm has been completed. the aic was returned for evaluation. during evaluation, the self-check failure error was confirmed and was found to be caused by corrosion on the pbm due to leaking of the supercapacitors, a known pattern failure, which disrupted operation of the battery charging circuitry. data logs were not reviewed as the self check failure prevented the console from booting up. the cause of the aic issue was contamination. there was corrosion of the pbm traces from leakage of electrolyte of the supercapacitors.
31246 | 1220648-2024-06508 | 18623803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623803&pc=OZD | 01/18/2024 | 03/04/2024 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  | 07/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392367 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 64-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the team observed a persistent access site bleed. the bleed was not treated by blood products or vascular repair. the bleed was treated with sutures and manual pressure/dressing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major event has been completed. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved after adding multiple forward pressure sutures.
31247 | 1220648-2024-06509 | 18623824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623824&pc=OZD | 08/27/2023 | 02/01/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 06/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024357119 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 68-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed access site bleeding on day two of support. safeguard placed by cardiologist and the heparin was put on hold. the patient was reported as stable. | [5 empty] | Product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to anticoagulation management. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643, apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31248 | 1220648-2024-06510 | 18623900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623900&pc=OZD | 01/18/2024 | 03/11/2024 |  | 02/01/2024 |  |  | 01/18/2024 | 02/01/2024 |  |  |  |  | 02/01/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1093078 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported that the automated impella controller (aic) displayed a controller failure upon startup. the pump had not yet been connected to the console at the time of the noted failure. the aic was swapped as a result of the failure and support proceeded as intended. there was no patient harm or involvement reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) red alarm has been completed. the aic was returned for evaluation. upon functional inspection, the optical bench was parameters were tested and found to be within specification. data logs were reviewed which confirmed a controller error, which was specifically for a defective optical sensor lamp. however, since the complaint could not be reproduced during testing, the root cause was most likely an intermittent failure of the optical bench.
31249 | 1220648-2024-06511 | 18623952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18623952&pc=OZD | 01/19/2024 | 03/05/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401582 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a (b)(6) patient (unknown gender) noted as high risk percutaneous coronary intervention was to be implanted with an impella 5.5 device for mechanical circulatory support. an impella defective alarm occurred during priming / purge setup of an impella 5.5 pump. multiple attempts were made to restart the process. however the issue recurred and the decision was made to replace both the pump and the automated impella controller (aic). there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported pump not detected issue has been completed. the impella device was not returned for a visual inspection. data log analysis confirmed that an impella defective alarm had been triggered but it could not be confirmed whether this reported issue was due to pump or console issue. therefore, the root cause of the pump not detected could not be determined without the device return and sufficient clinical details.
31250 | 1220648-2024-06512 | 18624020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624020&pc=OZD | 10/22/2023 | 02/07/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 08/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 2024386030 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 62-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported the patient developed a small hematoma at the access site, the pocket was expressed, and the oozing resolved. | [5 empty] | No product was returned. the cause of the access site bleeding issue was not determined due to insufficient clinical information. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643, apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. as per clinical, slight hematoma at access site with high ptts (180 seconds) was observed and heparin was paused. abiomed does not recommend evaluating based on ptt values. the cause of the access site bleeding issue was not determined due to insufficient clinical information.
31251 | 1220648-2024-06513 | 18624108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624108&pc=OZD | 01/19/2024 | 02/01/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420468 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 66-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. upon implanting the impella, multiple attempts were made to insert the impella. however, they were unsuccessful due to the patient's blockage. the impella was removed and angled glide catheter was inserted and angiogram was obtained to view stenosis. the impella implant was aborted. | [5 empty] | The investigation into the delivery issues has been completed. no product was returned. the root cause of the delivery issue was most likely due to patient condition since the patient had stenosis.
31252 | 1220648-2024-06514 | 18624139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624139&pc=OZD | 01/02/2024 | 03/06/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 08/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024392994 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as escalation of care from an impella cp device. while on support, the patient was brought to the operating room for cardiopulmonary bypass (cpb) cannulation. during the cannulation, the right ventricle wall was torn. the impella was left across the valve and turned to p-0. the patient's heart would not arrest so the pump was turned to p-2 and there was a high motor current spike followed by an impella stopped alarm. the impella would not restart and the impella was replaced via direct approach. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the pump stop issue was not determined since product was not returned.
31253 | 1220648-2024-06515 | 18624192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624192&pc=OZD | 01/19/2024 | 03/05/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 11/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024428572 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported an (b)(6) male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that bleeding occurred from the impella access site after the repo sheath was inserted. the bleeding was not able to be controlled with forward tension sutures, manual pressure, and angle matching gauze. the impella cp was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643 apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that post procedure, the patient's activated coagulation time (act) was prolonged at 282. therefore, the root cause of the access site bleed was most likely due to improper anticoagulation management.
31254 | 1220648-2024-06516 | 18624281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624281&pc=OZD | 01/15/2024 | 03/04/2024 |  | 02/01/2024 |  |  | 01/22/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401133 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 47-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. during support, the patient had a deep vein thrombosis (dvt) in the right lower forearm (impella side). the patient had mild pain and swelling in the forearm. the patient was waiting for a heart transplant, so the patient was given medication and a muscle relaxant for discomfort. the patient was also administered a high dose of heparin. the dvt was monitored. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ minor and the thrombus issues has been completed. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
31255 | 1220648-2024-06517 | 18624364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624364&pc=OZD | 01/22/2024 | 03/13/2024 |  | 02/01/2024 |  |  | 01/22/2024 | 02/01/2024 |  |  |  |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1768515 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Self-Activation or Keying | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The complainant reported that the automated impella controller (aic) showed a system self-check error. the aic was exchanged. there was no patient involvement. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - system self-check error has been completed since the original report was filed. the console was returned and tested after arriving in fs and the issue was reproduced. after further investigation, it was determined that the root cause of the aic issue was a firmware corruption.
31256 | 1220648-2024-06518 | 18624454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624454&pc=OZD | 01/19/2024 | 02/01/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  | 11/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024427530 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 66-year-old male patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the pump stopped after 2 days and was unable to be restarted. the pump was removed and replaced by an intra-aortic balloon pump. no patient harm was reported. | [5 empty] | The device was returned and was evaluated. upon investigation of the pump, 3 open electrical lines were found. opening the red handle showed stretching/necking of the motor cables. the ground wire was severed. no catheter or cannula kinks were observed. no biomaterial was found. maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31257 | 1220648-2024-06519 | 18624539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624539&pc=OZD | 01/22/2024 | 09/20/2024 |  | 02/01/2024 |  |  | 01/22/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403455 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Deformation | Adverse Event Without Identified Device or Use Problem | Positioning Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. as the peel-away sheath was removed, the inserted impella cp pump advanced into the left ventricle and folded on itself. the catheter was pulled back into the aorta to straighten out the kink, but the doctor was unable to re-advance the pump across the aortic valve. the pump was explanted and a new pump was placed for support. there was no patient harm. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the device was not returned for investigation. the cause of the positioning issue was use related due to improper technique that led to fold the device on itself. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. note: corrected information provided in h6 is an addition to previous coding.
31258 | 1220648-2024-06520 | 18624547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624547&pc=OZD | 01/19/2024 | 02/20/2024 |  | 02/01/2024 |  |  | 01/19/2024 | 02/01/2024 |  |  |  |  | 02/01/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1259344 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a self-check error. the aic was removed from service and returned for evaluation. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - system self-check error has been completed since the original report was filed. the console was returned and tested after arrival and the failure mode was reproduced so when the console was turned on, it immediately output a system self-check fail. the cause of the aic issue was contamination of a communication line on the pbm.
31259 | 1220648-2024-06522 | 18624824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18624824&pc=OZD | 01/20/2024 | 02/23/2024 |  | 02/01/2024 |  |  | 01/20/2024 | 02/01/2024 |  |  | 11/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429401 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2024 |  |  |  |  | The user facility reported a 79-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the team was evaluating the patient for coronary artery bypass grafting when percutaneous cardiac intervention was done emergently. the pump support was thought to have contributed/caused limb ischemia in the impella accessed leg. the limbs bilaterally lost pulses and the team placed external fem-fem bypass. the pump remains on for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31260 | 1220648-2024-06521 | 18625194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18625194&pc=OZD | 01/22/2024 | 03/05/2024 |  | 02/01/2024 |  |  | 01/22/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423249 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had impella cp support placed for a patient with st elevation myocardial infarction (stemi) and was transferred to a tertiary center for a possible transplant. the access site of the impella cp had an access site bleed/hematoma. the team placed sutures. the support was then explanted after 26 hours and the groin site was closed by contralateral leg balloon tamponade. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since placing forward suture with mattress suture resolved the bleeding and hemostasis was achieved. d4-model number updated to impella cp.
31261 | 1220648-2024-06523 | 18625604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18625604&pc=OZD | 01/09/2024 | 03/05/2024 |  | 02/01/2024 |  |  | 01/17/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024407838 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported the patient was on impella cp smartassist support and during support, the patient had multiple cerebral infractions. the cause of the cerebral infractions is unknown. the patient was on impella support for 465.25 hours before the patient expired. the patient was made do not resuscitate (dnr) due to having the cerebral infractions. the cause of death was the patient's dnr and multiple organ failure. due to the impella being an unknown factor in the cerebral infractions that led to the dnr, there is a possibility that the impella was potentially related to the cause of death. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke/cva has been completed. the impella device was not returned for evaluation. additionally, clinical details were not sufficient to determine a root cause. therefore, the root cause of the reported stroke could not be determined.
31262 | 1220648-2024-06524 | 18626110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18626110&pc=OZD | 01/17/2024 | 03/05/2024 |  | 02/01/2024 |  |  | 01/17/2024 | 02/01/2024 |  |  | 03/31/2024 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204003 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 41-year-old white male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding at their access site coinciding with extreme hematuria. concerns for hemolysis were discussed as ldh levels were over 2000. the pump was subsequently explanted, at which point the bleeding resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis and access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis and access site bleed could not be determined.
31263 | 1220648-2024-06526 | 18626374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18626374&pc=OZD | 01/21/2024 | 02/26/2024 |  | 02/01/2024 |  |  | 01/21/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024407655 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient lost pulse in their impella leg following pump implant. due to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31264 | 1220648-2024-06528 | 18626694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18626694&pc=OZD | 01/22/2024 | 02/01/2024 |  | 02/01/2024 |  |  | 01/22/2024 | 02/01/2024 |  |  | 09/30/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024411379 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that the abdominal blood vessel tortuosity was so severe that the impella could not be inserted, even using a long sheath. the impella cp procedure was aborted. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the delivery issue was most likely due to patient condition since patient had severe abdominal blood vessel tortuosity.
31265 | 1220648-2024-06529 | 18626741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18626741&pc=OZD | 01/21/2024 | 02/07/2024 |  | 02/01/2024 |  |  | 01/22/2024 | 02/01/2024 |  |  | 01/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024314839 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella placement was attempted, but the patient¿s vessel size was too small, and the case was aborted. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. impella placement was attempted, but vessel size was too small and case was aborted. | The investigation into the reported delivery issue has been completed. the medical center did not return the impella device. the root cause of the delivery issue was patient condition related. h.10 investigation was completed at the time the manufacturer device report (mdr) was submitted. the root cause was missing from the additional manufacturer narrative on the original mdr. this information has been inserted into h.10.
31266 | 1220648-2024-06530 | 18627416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18627416&pc=OZD | 12/27/2023 | 03/25/2024 |  | 02/01/2024 |  |  | 01/02/2024 | 02/01/2024 |  |  | 07/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024388797 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 71-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while attempting to place impella 5.5 through right axillary, patient developed thrombus and required fogarty embolectomy of the right axillary/subclavian artery. the support continued for patient for additional 2 days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined.
31267 | 1220648-2024-06533 | 18627610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18627610&pc=OZD | 01/23/2024 | 03/15/2024 |  | 02/01/2024 |  |  | 01/23/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024408705 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 01/23/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had attempted impella 5.5 support due to right coronary artery and left anterior descending artery disease. the distal inlet would not advance though the brachiocephalic artery. multiple attempts were made, the wire was replaced, and the implant was unsuccessful and was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the delivery issue was most likely due to patient condition due to tortuosity and possible calcification of aorta. | G.3 the date on the previous supplemental manufacturer device report was reported incorrectly as 30-jan-2024. the date should of been 01-feb-2024. | The investigation into the delivery issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition due to tortuosity and possible calcification of aorta. | The investigation into the delivery issues has been completed since the original report was filed. the device was returned and tested during investigation. no abnormalities were found. the root cause of the delivery issue was most likely due to patient condition due to tortuosity and possible calcification of aorta.
31268 | 1220648-2024-06532 | 18627625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18627625&pc=OZD | 01/18/2024 | 05/24/2024 |  | 02/01/2024 |  |  | 01/23/2024 | 02/01/2024 |  |  |  |  | 02/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1346265 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | Upon review by abiomed's medical safety officer, the patient expired on support. there is not enough information to associate the use of the device with the patient¿s outcome. | The user facility reported a 42-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that there was a controller error alarm that happened for several days. the aic (automated impella controller) was exchanged, and the alarm issue was resolved. |  |  |  |  | The investigation has been completed since the initial report was submitted. logs confirmed the reported failure mode given there was a controller error alarms triggered during the clinical procedure. this particular error alarm was issued multiple times within the time period of the case complaint. real time (rt) logs confirmed. the product was returned for testing. the failure was unable to be reproduced throughout testing. the cause of the issue was not determined due to the inability to reproduce the error. d.2 revised common device name from the initial submission in accordance with updated procedures. d.4 revised model number from the initial submission in accordance with updated procedures. d.4 revised catalog number as previously entered incorrectly. d.9 device available for evaluation was revised as previously entered incorrectly. e.1 facility name and address line 2 were inadvertently left off the previously submitted report and were added. | The impella device was received from the customer on (b)(6)2024 and therefore,an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
31269 | 1220648-2024-06534 | 18627668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18627668&pc=OZD | 01/23/2024 | 03/12/2024 |  | 02/01/2024 |  |  | 01/23/2024 | 02/01/2024 |  |  | 08/31/2024 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023268460 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was unable to flow over p-3 due to ongoing suction alarms. placement signal is also significantly lower than the patient¿s blood pressure. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump ran without issue under bench test. the root cause of low pump flow was most likely a patient condition as no issues were found on the pump.
31270 | 1220648-2024-06531 | 18627679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18627679&pc=OZD | 01/22/2024 | 03/05/2024 |  | 02/01/2024 |  |  | 01/23/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411501 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to having cardiomyopathy. bleeding was noted in the area of the impella insertion site. echo - echocardiogram (ultrasound)- confirmed pump position and one unit of packed red blood cells were provided. additionally, the patient had ventricular tachycardia (vt) on support and the patient required seven defibrillations and three units of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. additional information: h5-yes for single use device. h6-impact code 4641. d4-model number updated to impella 5.5.
31271 | 1220648-2024-06535 | 18627725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18627725&pc=OZD | 01/23/2024 | 03/06/2024 |  | 02/01/2024 |  |  | 01/23/2024 | 02/01/2024 |  |  | 01/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311099 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | Embolism/Embolus | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2024 |  |  |  |  | The user facility reported a 38-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that bubbles were noted in the left ventricle that appeared to originate in the aorta on echo, following initial implant of impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported embolism has been completed. data logs and product were not returned for review. as per clinical air bubbles were noted in left ventricle which originated in aorta per echo. all other clinical details are unknown. the cause of the embolism issue was unable to be determined as no device or logs were returned and insufficient clinical information was provided. d4-updated model number to impella cp.
31272 | 1220648-2024-06527 | 18627830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18627830&pc=OZD | 01/22/2024 | 04/09/2024 |  | 02/01/2024 |  |  | 01/22/2024 | 02/01/2024 |  |  |  |  | 02/01/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1542729 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic (automated impella controller) displayed a placement signal not reliable alarm during patient support. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation for the reported issue has not yet been completed. upon its completion, a follow up report will be submitted. | The investigation into the optical signal issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the reported failure mode was confirmed. there were multiple placement signal not reliable alarms. the optical parameters were measured, and the signal-to-noise ratio (snr) was below specification. the root cause of the placement signal errors was determined to be the low snr of the original optical bench, and the cause of the low snr was determined to be noise inherent to the charge-coupled device (ccd) array.
31273 | 1220648-2024-06525 | 18627997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18627997&pc=OZD | 01/21/2024 | 09/20/2024 |  | 02/01/2024 |  |  | 01/21/2024 | 02/01/2024 |  |  | 10/31/2025 |  | 02/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024421363 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/21/2024 |  |  |  |  | The user facility reported a 31-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having pulseless electrical activity arrest. while on support, the patient had limb ischemia. the physician noted cold/dead leg and they attempted to place distal perfusion cannula unsuccessfully. vascular was consulted and they determined the leg was too far gone to salvage. the decision was made to withdraw care after 23.45 hours of impella support due to a computed tomography scan showing continued cerebral edema. the patient had become agonal and all support ceased. no allegations were made towards the impella for cause of death. | It was noted that the patient expired in the operating room. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06525 was provided in sections b2, b5, g3/4, h1, and h6. | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31274 | 1220648-2024-06481 | 18613493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18613493&pc=OZD | 04/23/2023 | 02/02/2024 |  | 01/31/2024 |  |  | 04/23/2023 | 01/31/2024 |  |  | 09/30/2024 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273400 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2023 |  |  |  |  | The user facility reported a 60-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. th patient developed hemolysis during impella support. the patient's family was advised to escalate care; however, the family refused escalation and the patient was made dnr/dni. support continued with the implanted pump despite the condition. | [5 empty] | The investigation into the reported, hemolysis has been completed since the original report was filed. product was not returned for review. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not returned. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31275 | 1220648-2024-06482 | 18613548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18613548&pc=OZD | 05/03/2023 | 01/31/2024 |  | 01/31/2024 |  |  | 05/03/2023 | 01/31/2024 |  |  | 11/30/2024 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298858 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2023 |  |  |  |  | The user facility reported a 66-year-old male noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. during attempted delivery, resistance was met in the proximal iliac and the device could not advance due to an existing left iliac lesion. the sheath was exchanged to allow access across the lesion, however the delivery attempt still failed. the implant was subsequently aborted as a result of the existing condition. there was no patient harm. | [5 empty] | The investigation into the failed delivery issue has been completed. product was not returned for investigation. per clinical, both the pump and the 14x25 introducer encountered difficulty when attempting to pass through the iliac lesion. the cause of the delivery issue was patient condition related to iliac lesion, based on provided clinical details. (diseased/small vessels).
31276 | 1220648-2024-06484 | 18613749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18613749&pc=PYX | 06/07/2023 | 01/31/2024 |  | 01/31/2024 |  |  | 06/08/2023 | 01/31/2024 |  |  | 09/30/2023 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023199496 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2023 |  |  |  |  | The user facility reported a 76-year-old male post cardiotomy cardiogenic shock with low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. the following day, a high purge pressure alarm was noted and the medical team was notified. there was no harm to the patient. | [5 empty] | The investigation into the "high purge pressure" alarm has been completed. data analysis: data confirmed the failure mode. about 0.5 hours into the case, the pp increased gradually from 500 mmhg to 1100 mmhg that triggered high pp alarms. high pp continue remaining for about 16 hours then gradually decreased to 600 mmhg & remained to the rest of the case. product was not returned for review. the root cause of the high purge pressure was unable to be determined as limited clinical details were provided and no product was returned for review.
31277 | 1220648-2024-06485 | 18613808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18613808&pc=OZD | 06/12/2023 | 01/31/2024 |  | 01/31/2024 |  |  | 06/12/2023 | 01/31/2024 |  |  | 03/31/2025 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342636 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a pump flow issue occurred most likely due to a clot ingestion of which there was no confirmation of. there was no serious adverse event to the patient. the impella was explanted a day later. | [5 empty] | The investigation into the reported low pump flow has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. data analysis confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. the root cause of the low flow was unable to be determined as no product was returned for review. a risk assessment was performed, and no escalation is required. this report is being filed as part of a retrospective review of historical records.
31278 | 1220648-2024-06486 | 18613832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18613832&pc=PYX | 01/16/2024 | 01/31/2024 |  | 01/31/2024 |  |  | 01/16/2024 | 01/31/2024 |  |  | 10/31/2025 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024428096 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 48-year-old male noted as high-risk percutaneous cardiac insertion was implanted with an impella rp flex device for mechanical circulatory support. the complainant reported that the patient on impella support self-explanted the impella. the patient will continue to be monitored and is stable. | [5 empty] | (h3) the investigation into the reported "positioning issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the positioning issues was determined to be patient movement because the patient pulled the pump fully out of the ij.
31279 | 1220648-2024-06492 | 18617286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18617286&pc=OZD | 01/17/2024 | 03/05/2024 |  | 01/31/2024 |  |  |  | 01/31/2024 |  |  | 11/30/2025 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426164 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 73-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. after initial attempts to deliver the pump via axillary graft were unsuccessful due to anatomical issues and vascular tortuosity, the pump was removed and a graft was placed onto the aorta for a direct approach. the first attempts via direct access was unsuccessful for an unknown indication. on the second attempt via direct insertion, the doctor was unable to view the image via transesophageal echocardiography guidance and the pump perforated the left ventricle. the perforation was repaired surgically without complication and the implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the delivery issues - use and the ventricle perforation issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue and the ventricle perforation issue was use related due to improper technique. proper lv imaging was not available during insertion leading to multiple attempts and perforation in the end.
31280 | 1220648-2024-06493 | 18617393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18617393&pc=OZD | 01/17/2024 | 02/29/2024 |  | 01/31/2024 |  |  | 01/17/2024 | 01/31/2024 |  |  | 10/31/2025 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421265 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Atrial Fibrillation | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the support was weaned and explanted after over 145 hours. the groin had a hematoma that prompted vascular surgeon to do a cut down. vessel loops, sutures, a hemostatic patch, and blood products were given. additionally, there was noted suction/low flows and atrial fibrillation that prompted the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding , low pump flow and vt has been completed since the original report was filed. the device was returned for investigation. the cause of the access site bleeding was patient movement as the patient became confused overnight and began pulling at lines and impella catheter. the cause of the low pump flow was biomaterial ingestion around impella. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined." section d9 updated.
31281 | 1220648-2024-06494 | 18617522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18617522&pc=OZD | 01/17/2024 | 01/31/2024 |  | 01/31/2024 |  |  |  | 01/31/2024 |  |  | 09/30/2025 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396546 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 67-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. resistance was met during attempted impella 5.5 implant. the delivery was ultimately aborted as the pump was unable to pass through the cannulation site due to interference from the aortic ecmo cannula. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed. the product was not returned for evaluation. per the clinical review, the root cause of the delivery issue was most likely due to use issue since the delivery was aborted because there was interference from the aortic ecmo cannula.
31282 | 1220648-2024-06495 | 18617610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18617610&pc=OZD | 01/18/2024 | 03/01/2024 |  | 01/31/2024 |  |  | 01/18/2024 | 01/31/2024 |  |  | 10/31/2025 |  | 01/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415479 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported an 80-year-old female was having an impella 5.5 implant due to needing additional support after coronary artery bypass graft. there were multiple sites of bleeding noted. a left ventricle perforation was observed. the ventricle was repaired using a suture and patch. blood products were given. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve | The investigation into the reported ventricle perforation has been completed since the original report was filed. according to the clinical details, while manipulating the heart they discovered an lv perforation. a sutured patch was used to resolve the issue. no product was returned for a visual inspection. the most likely cause of the ventricle perforation was due to use while manipulating the patients delicate heart.
31283 | 1220648-2024-06496 | 18618075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18618075&pc=OZD | 01/18/2024 | 03/19/2024 |  | 01/31/2024 |  |  | 01/18/2024 | 01/31/2024 |  |  |  |  | 01/31/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1145538 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella 5.5 as a bridge to left ventricular assist device (lvad). on day 42 of support the automated impella controller (aic) displayed repeat red/controller alarms. the aic was successfully replaced and there was no patient harm reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - controller failure (red alarm) issue has been completed since the original report was filed. the console was returned and tested after arriving in fs. the failure modes were not reproduced by running the pump; however, the optical bench failure at the beginning of the failures was able to be reproduced. the cause of the optical bench communication loss was defective software. the cause of the unexpected reboot was the audio processes failing.
31284 | 1220648-2024-06444 | 18605355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18605355&pc=PYX | 06/23/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 06/23/2023 | 01/30/2024 |  |  | 11/30/2023 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022147046 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/23/2023 |  |  |  |  | The user facility reported a 72-year-old male post cardiac surgery was to be implanted with an impella rp device for mechanical circulatory support. the impella rp failed delivery due to thrombus burden and wire interaction. the 5.5 pump and a new rp successfully supported the patient. | [5 empty] | The investigation into the delivery issue of the impella rp has been completed. the device was not returned for evaluation, so no root cause could be positively identified. the cause of the delivery issue was most likely due to an interaction with biomaterial in the patient.
31285 | 1220648-2024-06445 | 18605437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18605437&pc=OZD | 06/27/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 06/27/2023 | 01/30/2024 |  |  |  |  | 01/30/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1205049 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Audible Prompt/Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/27/2023 |  |  |  |  | Unspecified patient: the complainant reported that there were no audible alarms on console. there was no harm to the patient as a result of this incident. | [5 empty] | Data analysis: console logs from the reported day of event were analyzed and showed multiple alarms presenting, that would be associated with audio signals. however, the data does not allow us to infer whether the alarms were audible. therefore, aic logs are not relevant. device analysis: inspection of the console performed at danvers revealed that the speaker cable was not plugged into the 4-in-1 carrier, explaining lack of audio alarms. repair: reinstall speaker cable ((b)(6)), and perform functional check of the console.
31286 | 1220648-2024-06446 | 18605499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18605499&pc=OZD | 07/04/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 07/04/2023 | 01/30/2024 |  |  | 04/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024338718 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Air/Gas in Device | [6 empty] | Embolism/Embolus | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/04/2023 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the device was inserted prior to the purge fluid primed. air was noted on tee (transesophageal echocardiography¿). air aspirated manually through pigtail in lv (left ventricle¿) until it cleared on tee. | [5 empty] | The impella device was not received from the customer. the cause of the embolism issue was use related because the device inserted prior completing priming and air delivered to ventricle. air aspirated manually through pigtail in lv until clear on tee. complaint device not returned for investigation.
31287 | 1220648-2024-06447 | 18605562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18605562&pc=OZD | 01/11/2024 | 05/02/2024 |  | 01/30/2024 |  |  | 01/11/2024 | 01/30/2024 |  |  |  |  | 01/30/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  |  |  | 2024398528 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Failure to Prime | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A 66-year-old patient received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. during preparation, following issues with the first console, the back-up automated impella controller (aic) also skipped two confirmation screens, which would have allowed the operator to start the pump without entering the purge fluid composition and without removing the guidewire. the impella cp heart pump worked without any problems and the patient suffered no harm as a result of the incident. | [5 empty] | The investigation for the reported console (aic) purge issue has been completed. the console was returned for evaluation. upon inspection, the console performed as expected. however, data logs were reviewed which showed that the pump was plugged into a console and recognized as previously initialized. however, when the pump was initially plugged into a different console, it had said uninitialized. the root cause for the aic screen not offering the screen with the purge fluid table, nor the double confirmation of the removed 0,018" wire was due to the pump being previously initialized on the first attempted initialization. this was intended behavior of the console. added- d9 device return date g1-corrected mfr site name and address. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
31288 | 1220648-2024-06448 | 18605572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18605572&pc=OZD | 07/05/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 07/05/2023 | 01/30/2024 |  |  |  |  | 01/30/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1302706 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the purge clip on the automated impella controller (aic) was broken. the device was removed from service as a result of the noted damage. there was no reported patient harm. | [5 empty] | (h3) the investigation into the reported "aic - purge disc/flag issues" has been completed since the original report was filed. product was returned for investigation. during visual inspection of the console, the purge disc retainer was not broken. the console was tested with purge and pump run with no errors. if this was the wrong console reported, the correct console at this site could not be ascertained. during data analysis, the complaint alleges that the purge disc retainer was broken, potentially around or before july 5th. this was not corroborated by any data as the console was able to detect the purge disc the last time the console was run on july 3rd. there was no root cause since there was no failure.
31289 | 1220648-2024-06449 | 18605578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18605578&pc=OZD | 01/12/2024 | 02/28/2024 |  | 01/30/2024 |  |  | 01/12/2024 | 01/30/2024 |  |  | 09/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401819 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had 62-year-old male patient on impella cp support for 5 days with hypothermia protocol for the vfib arrest/ami cgs. the patient suffered from limb ischemia, as the lmpella accessed limb lost pulses. additionally there were signs of hemolysis. the team troubleshot by explanting the pump. | [5 empty] | The investigation into the reported limb ischemia reversible has been completed since the original report was filed. data logs showed pump was in proper position & ran without issue for entire case. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿
31290 | 1220648-2024-06450 | 18605605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18605605&pc=PYX | 03/03/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 07/06/2023 | 01/30/2024 |  |  | 11/30/2023 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022153837 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant reported they were unable to get wire to advance rfv. decision per discussion of lfv insertion attempt. 7fr arrow to pa 1st time success 1st doctor tried .027 then had to switch to .035 wire down to access lpa. successful, during 1st wire pass, 7fr catheter fell out of pa. re-attempt # 1, got catheter deep seated enough into lpa, .027 with j curve to mid lower lobe lpa w/wire tip per best practices. rp inserted zeroed on flouro, torque/rotation push pull technique deployed, nose of rp wouldn¿t make turn with careful attempts by the doctor to push, pull wire and rotate rp. wire fell out of pa. the rp was removed. there were several reattempts that were unsuccessful. decision to explant and put patient on peripheral ecmo. rp explanted, access site used for arterial ecmo. | [5 empty] | The investigation into the reported delivery issues has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the most likely cause of the delivery issue was most likely use related with extreme torque being applied leading to kinking of the product based on the clinical information. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31291 | 1220648-2024-06451 | 18605942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18605942&pc=OZD | 01/11/2024 | 02/28/2024 |  | 01/30/2024 |  |  | 01/11/2024 | 01/30/2024 |  |  | 09/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403918 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp implanted to support a 64-year-old male patient admitted with cgs/ami and undergoing percutaneous coronary intervention (pci) for coronary artery disease (cad). the patient had observed left ventricle (lv) thrombus. the thrombosis cause was unknown and so the pump was explanted. the clot may have been there prior to implant but, explanted when the patient was hemodynamically stable. | [5 empty] | The investigation into the thrombus issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31292 | 1220648-2024-06452 | 18606038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606038&pc=OZD | 01/15/2024 | 02/20/2024 |  | 01/30/2024 |  |  | 01/15/2024 | 01/30/2024 |  |  |  |  | 01/30/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | UNK |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preparation for a high risk percutaneous coronary intervention (hrpci), the automated impella controller (aic) displayed ¿purge system open¿ and ¿purge disc not recognized¿ errors. the issue was resolved by using another cassette and console. there was no reported patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. the purge disc retainer was shown to be broken. the cause of the issue was a broken purge retainer. | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31293 | 1220648-2024-06453 | 18606120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606120&pc=OZD | 01/13/2024 | 09/20/2024 |  | 01/30/2024 |  |  | 01/13/2024 | 01/30/2024 |  |  | 08/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385422 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was noted that the patient expired after explant of the device. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported that the 71-year-old male patient experienced groin site bleeding and bilateral lower extremity ischemia during impella cp support. leading up to support it was noted that the patient required unusually high pressor support due to their existing condition. following implant, bilateral lower extremities (ble) mottling was documented and quickly worsened with symptoms of extensive mottling, blistering, and necrotic toes. discussions were had regarding the need for a bilateral fasciotomy to counteract the condition. heparin administration was then halted seven days into support as a result of the noted oozing/bleeding at the groin site. as the ble ischemia continued to worsen, it was documented that amputation would be necessary if the patient was weaned from support. further decisions related to the patient's course of therapy were left to the patient and their family as a result of their condition. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06453 was provided in sections b2, b5, d6, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31294 | 1220648-2024-06454 | 18606140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606140&pc=OZD | 01/15/2024 | 02/07/2024 |  | 01/30/2024 |  |  | 01/15/2024 | 01/30/2024 |  |  | 10/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420478 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 for support along with va ecmo. the team had escalated from the impella cp to the 5.5. while on the first impella 5.5 support the pump had ps issues. the optical signal issue caused no harm or injury and the pump was explanted and replaced by a new, second impella 5.5 pump. the second pump was noted to be seen to have a left ventricle (lv) perforation. the 5.5 caused perforation was repaired. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported ventricular fibrillation has been completed. the impella pump was returned for investigation, and biomaterial was noted on the impeller and purge gap. purging was started after 10 minutes of soaking in a bucket of water. high purge pressure was not reproduced on the returned pump. the pump ran as expected during testing in a bucket of water, and all the optical 5s parameters were within specified limits. no other physical abnormalities were observed. data logs showed a gradual 10 minute rise in purge pressure while the pump was outside of the body after the initial placement of the pump. as per the clinical details, the doctor found a perforated left ventricle (lv) after live view returned on trans-esophageal echo. the lv perforation was repaired, and the impella started to have high purge pressure alarms. the pump was explanted due to purge related alarms. the cause of the ventricle perforation issue was most likely use related, based on given clinical information and returned product review. added h.6 health effect - impact code: 4627.
31295 | 1220648-2024-06455 | 18606246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606246&pc=OZD | 01/13/2024 | 03/13/2024 |  | 01/30/2024 |  |  | 01/13/2024 | 01/30/2024 |  |  | 09/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432264 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had 70-year-old male patient on impella cp support ongoing when there was bleeding from the swan and impella site. the bleed prompted blood products and pump replacement. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was returned and visually inspected. blood found between blue suture and white hub. the root cause of the access site bleeding was use related issue due to improper technique as the same issue happened on the next pump used. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31296 | 1220648-2024-06456 | 18606261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606261&pc=OZD | 01/16/2024 | 03/07/2024 |  | 01/30/2024 |  |  | 01/16/2024 | 01/30/2024 |  |  | 08/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024392797 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 59-year-old male patient developed a hematoma and severe drop in red blood cells, leading to impella cp explant. the hematoma had to be removed in the operating room (or). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned.
31297 | 1220648-2024-06457 | 18606266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606266&pc=OZD | 01/13/2024 | 02/28/2024 |  | 01/30/2024 |  |  | 01/13/2024 | 01/30/2024 |  |  | 10/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024412065 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device via direct implant for mechanical circulatory support. after the procedure, the team made multiple attempts to pull the impella through the graft. a decision was made to do a re-sternotomy as this is the best option to remove an impella 5.5 device that has been directly implanted. the impella was withdrawn with ease and the graft was trimmed closer to the aorta. the patient is stable. | [5 empty] | The investigation into the reported removal issue has been completed since the original report was filed. no product was returned for investigation. the root cause of the removal issues was not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31298 | 1220648-2024-06458 | 18606338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606338&pc=OZD | 01/14/2024 | 02/29/2024 |  | 01/30/2024 |  |  | 01/14/2024 | 01/30/2024 |  |  | 07/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391355 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 51-year-old male patient became restless in the unit after initially arriving from the cath lab on impella cp support. during the foley and ng tube insertion, the patient moved and fought causing the impella to pull back into the aorta. the patient had a run of svt and then vt, coded once and cpr was done for 5-8 mins. amio was started and the patient went back to the cath lab to reposition the cp. it was further reported that due to some bleeding at the femoral site and the patient¿s high lvedp, it was decided to upgrade him to a 5.5 and explant the cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the vt/vf issues has been completed. the device was not returned for investigation. per the provided clinical details, the root cause of the access site bleeding was most likely due to patient movement as the patient moved and fought causing the impella to pull back likely causing the bleed. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
31299 | 1220648-2024-06459 | 18606343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606343&pc=OZD | 01/16/2024 | 03/14/2024 |  | 01/30/2024 |  |  | 01/16/2024 | 01/30/2024 |  |  | 10/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024412562 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following explant of impella cp, the 60-year-old male patient's puncture site was still bleeding heavily. the access site had to be closed in the or. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31300 | 1220648-2024-06460 | 18606388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606388&pc=OZD | 01/12/2024 | 03/12/2024 |  | 01/30/2024 |  |  | 01/15/2024 | 01/30/2024 |  |  | 06/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024360877 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Vascular Dissection | [11 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 43-year-old male patient on impella 5.5 support was noted to have an aortic dissection and multiple cerebral infarctions. the patient also had a valve replacement and was on extracorporeal life support (ecls) therapy. | [5 empty] | The investigation for the reported stroke and vascular damage has been completed. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the root cause of the stroke and vascular damage could not be determined. the cause of the great vessel vascular damage could not be determined because no product was returned and not enough clinical information was given. the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
31301 | 1220648-2024-06461 | 18606436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606436&pc=OZD | 01/15/2024 | 03/20/2024 |  | 01/30/2024 |  |  | 01/15/2024 | 01/30/2024 |  |  | 10/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024412166 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 55-year-old male patient on cardiogenic septic shock was placed on impella cp support. following successful implant procedure, patient experienced ¿purge flow blocked¿ alarms. there was a reduction of the p-level to p-2. the motor current rose suddenly and the impella stopped working. placement of the impella in the descending aorta did not improve pump performance, leading to the pump being explanted. there was no presence of thrombus noted. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. biomaterial was found in the purge gap. per the product investigation and information provided, the root cause of the pump stop issue was most likely biomaterial.
31302 | 1220648-2024-06462 | 18606574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606574&pc=OZD | 01/15/2024 | 02/21/2024 |  | 01/30/2024 |  |  | 01/15/2024 | 01/30/2024 |  |  |  |  | 01/30/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1310117 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Company representation | Distributor |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the purge disk was not recognized by the automated impella controller (aic). another device was used successfully. there was no reported patient harm. | [5 empty] | The investigation into the aic - purge issue ¿ other has been completed since the original report was filed. the console was returned and tested after arriving in fs. the failure mode was reproduced, and purge flag was confirmed to be missing. the cause of the issue was a missing purge flag. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
31303 | 1220648-2024-06463 | 18606685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18606685&pc=OZD | 01/16/2024 | 03/20/2024 |  | 01/30/2024 |  |  | 01/16/2024 | 01/30/2024 |  |  | 10/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024411274 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported following transfer of 50-year-old female patient to ambulance stretcher, the impella cp pump stopped working. the screen was noted to only show menu items that could be clicked on until the last one. it said ¿please connect the gray connector cables¿. from this menu point onwards it was no longer possible to continue. a second console was used but the same issue happened. the clinic refused to place a new pump. the impella was pulled back into the aorta and catecholamines had to be increased. the patient was transferred, and the pump was explanted. | [5 empty] | The investigation for the reported pump stop has been completed. the pump nor data logs were returned for evaluation. however, without the device nor sufficient clinical information for evaluation, the exact root cause of the pump stop could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31304 | 1220648-2024-06464 | 18607044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18607044&pc=OZD | 01/15/2024 | 03/14/2024 |  | 01/30/2024 |  |  | 01/16/2024 | 01/30/2024 |  |  | 10/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024415157 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 54-year-old female patient underwent coronary intervention due to acute myocardial infarction/cardiogenic shock. during the procedure, the patient's hemodynamic situation deteriorated. the patient was stabilized by inserting an impella cp smartassist heart pump. the introduction of the impella cp occurred during cardiopulmonary resuscitation and without echo support. bleeding occurred at the access site and a hematoma formed. due to the bleeding, the impella cp was removed again. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31305 | 1220648-2024-06465 | 18607282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18607282&pc=OZD | 01/17/2024 | 03/14/2024 |  | 01/30/2024 |  |  | 01/17/2024 | 01/30/2024 |  |  | 08/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393103 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 64-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that patient experienced color change in urine and a decrease in urine production; so, the team switched the purge solution to bicarb and after 33 hours and no improvement the pump was explanted due to hemolysis concerns. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis. | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined.
31306 | 1220648-2024-06466 | 18607634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18607634&pc=OZD | 07/12/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 07/12/2023 | 01/30/2024 |  |  | 04/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024346723 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/12/2023 |  |  |  |  | The user facility reported a patient (unknown age, race, gender) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during pump preparation and prior to insertion, the pump was switched on and immediately a pump stop happened. pump was replaced successfully. | [5 empty] | The investigation into the pump that failed to start during preparation for use has been completed. impella cp was returned for evaluation. visually inspected the pump and no biomaterial at outflow, catheter kink or abnormalities with impeller were not observed. electrical resistance test was done, and all the motor phase values were within normal range. pump cannula was cutdown to inspect the impeller and scratch marks were observed on the impeller blade surface. data logs were reviewed, and motor current high pump stop was observed within 5 minutes of start. the cause of the pump did not start issue was use related with pump started with wire in place and scratch marks observed on the impeller blades per field investigation and clinical review.
31307 | 1220648-2024-06467 | 18607730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18607730&pc=OZD | 04/20/2023 | 10/20/2025 |  | 01/30/2024 |  |  | 04/20/2023 | 01/30/2024 |  |  | 10/31/2024 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024291283 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2023 |  |  |  |  | The user facility reported a patient (unknown demographics) was implanted with an impella cp for circulatory support. suction issues and placement signal unreliable alarms occurred during impella support. it was noted that the patient did not have an anatomically correct cardiac anatomy which led to difficulty with positioning the pump. placement signal unreliable and suction alarms were subsequently ongoing with no resolution after troubleshooting. patient support continued with the implanted pump. there was no patient harm as a result of the noted issue. | It was reported that the patient expired during support. |  |  |  |  | The investigation for the optical signal issue and low pump flow has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the low pump flow issue was most likely due to be improper positioning. positioning the pump properly was difficult due to the patients difficult cardiac anatomy. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. all pertinent information available to abiomed has been submitted at this time.
31308 | 1220648-2024-06469 | 18608251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18608251&pc=OZD | 05/05/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 05/05/2023 | 01/30/2024 |  |  | 12/31/2024 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024322911 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/05/2023 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the impella cp was pulled back into the aorta during patient support. attempts to re-cross the valve were unsuccessful and the cp was subsequently removed and replaced for ongoing support. there was no patient harm. | [5 empty] | The device was not returned for evaluation. the cause of the positioning issue was most likely touhy-borst related since rn found that touhy-borst was unlocked during impella pulled back. this report is being filed as part of a retrospective review of historical records.
31309 | 1220648-2024-06470 | 18608341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18608341&pc=OZD | 05/19/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 05/19/2023 | 01/30/2024 |  |  | 11/30/2024 |  | 01/30/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281975 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported poor positioning due to anatomical structure of the patient's heart. the pump was removed once patient's hemodynamics were deemed stable. | [5 empty] | The investigation into the reported event has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the most likely cause of the issue was use related.
31310 | 1220648-2024-06471 | 18608556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18608556&pc=OZD | 06/08/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 06/08/2023 | 01/30/2024 |  |  | 06/30/2024 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236558 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 77-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient showed a decrease in the hemoglobin/hematocrit values after implanting impella for which the physician suspected it to be due to hemolysis and the patient had to be given one unit of blood. | [5 empty] | The investigation into the reported "hemolysis" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the patient had pulled back during support causing to reposition the device. patient experienced reduction of hemoglobin/hematocrit levels. the data logs were returned and suction, impella position in aorta alarms were observed. the cause of the hemolysis issue was due to improper positioning of the device as a result of patient movement per clinical and data log review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31311 | 1220648-2024-06472 | 18608607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18608607&pc=OZD | 08/01/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 08/01/2023 | 01/30/2024 |  |  | 06/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024360359 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported an 89-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was prepped and inserted but it would not advance past the subclavian artery. an amplatz super stiff successfully advanced past the cannula and into the aorta. this did straighten the artery some, but still the cannula would not advance into the innominate. a decision was made to abort the insertion of impella 5.5 and place the impella cp. the patient is stable. | [5 empty] | The investigation into delivery issues has been completed. product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition. (diseased/small vessels).
31312 | 1220648-2024-06473 | 18608729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18608729&pc=OZD | 08/12/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 08/12/2023 | 01/30/2024 |  |  | 01/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024314482 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/12/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the left femoral artery, impella cp access site, developed a hematoma. staff controlled the issue with manual pressure. | [5 empty] | No product was returned. as per clinical patient was coagulopathic and taken to or during bleeding at access site. blood products were given, and systemic heparin was stopped. the cause of the access site bleeding issue was most likely anticoagulation management related due to patient being coagulopathic during bleeding. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31313 | 1220648-2024-06474 | 18608922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18608922&pc=OZD | 08/24/2023 | 01/30/2024 |  | 01/30/2024 |  |  | 08/24/2023 | 01/30/2024 |  |  |  |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024360329 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support in which there was a minor bleed. the minor bleed did not prompt any intervention. | [5 empty] | Per the given clinical information, the root cause of the cause of the access site bleeding was most likely high patient act values. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31314 | 1220648-2024-06475 | 18609708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18609708&pc=OZD | 01/16/2024 | 02/13/2025 |  | 01/30/2024 |  |  | 01/16/2024 | 01/30/2024 |  |  |  |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1221446 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after the percutaneous coronary intervention (pci) procedure was completed the automated impella controller (aic) alarmed impella stopped and controller failure. the impella was removed from the patient and the aic then stated system self-check failed-change console immediately. there was no harm to the patient. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation automated impella controller (aic) - controller failure (red alarm) and aic - system self-check error has been completed. the logs from confirm the complaint as evidenced by issuance of alarm controller error (loss of comm (cpld)) - alarm issued. later when a pump is plugged into the console, there is an impella stopped. controller failure (motor driver guard error or supply voltage of the motor driver is greater than 22.5 vdc) [detected by cpld] - alarm issued. when the console is rebooted, it is unable to successfully pass system check i.e. [16/01/24 11:13:48 post: first systemcheck failed ! ] due to numerous pmd communication errors. devise analysis: the console was interrogated in the lab. the console passed system self-check initially, but failed after a second reboot. on the initial reboot, system self-check passed, but controller error (loss of communication to and cpld) was issued and cpld buzzer audio was observed. voltage readings confirmed that the pbm was supplying voltage to the impellatronic board within specification. i.e. 20.14v from pbm into pmd impellatronic partition and 4.991v from pbm into 1st channel eep impellatronic board partition. all impellatronic leds were illuminated. after the failed second reboot, the impellatronic board was swapped with a known good one at the same software version and the issue got resolved. the root cause was then confirmed to be isolated to the impellatronics board. the root cause of the aic issue was due to a defective impellatronics board. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-06475 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-06475 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted from manufacturer device report 1220648-2024-06475.
31315 | 1220648-2024-06476 | 18609954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18609954&pc=OZD | 08/29/2023 | 02/01/2024 |  | 01/30/2024 |  |  | 09/13/2023 | 01/30/2024 |  |  | 06/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373681 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported a 47-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was thrashing in bed and the pump was pulled back into the aorta. under echo, pigtail could not be identified so the patient was taken back to the catheterization lab for placement under fluoroscopy. the impella was found to be in descending aorta. attempts were made to re-cross wirelessly but were unsuccessful. the impella was removed, and the patient was placed on dobutamine. patient was sent to icu in stable condition. | [5 empty] | The investigation into the positioning issues has been completed. the product was not returned for review. as per clinical, the root cause of the issue was wireless re-access. the impella device is not indicated for wireless re-access per ifu 0048-9007. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31316 | 1220648-2024-06477 | 18610037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18610037&pc=OZD | 01/16/2024 | 03/07/2024 |  | 01/30/2024 |  |  | 01/16/2024 | 01/30/2024 |  |  | 08/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024394048 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Disability | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after six days male patient on impella 5.5 support had "impella blocked' alarm and pump stop leading to pump explant. following removal, a clot was noted to be in the inflow area. | [5 empty] | The investigation into the pump stop and the thrombus issues has been completed since the original report was filed. the device was returned and inspected. a clot was found in the purge gap and inside the cannula. the root cause of the thrombus cannot be determined due to insufficient clinical information provided. the pump did not pass electrical testing and an open motor cable line was found inside the red handle. the cause of the pump stop was an open motor cable line within the red handle due to excessive force applied to the catheter. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31317 | 1220648-2024-06478 | 18610153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18610153&pc=OZD | 01/16/2024 | 03/20/2024 |  | 01/30/2024 |  |  | 01/16/2024 | 01/30/2024 |  |  | 05/31/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024362014 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp insertion for 56-year-old female patient, there was no flow or motor current. the pump did not start. therefore, another device was used. there was no reported patient harm. | [5 empty] | The investigation into the pump did not start issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. the root cause of the pump did not start issue was determined to be one open electrical line due to excessive adhesive gluing the wire to the bottom of the red handle thus decreasing its slack and mobility. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31318 | 1220648-2024-06479 | 18610174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18610174&pc=OZD | 01/17/2024 | 03/04/2024 |  | 01/30/2024 |  |  | 01/17/2024 | 01/30/2024 |  |  | 09/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409890 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 73-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed a hematoma. the physician made the decision to remove the impella. a manta closure was used followed by manual pressure. the doctor reported that the hematoma was caused by their stick making the impella have an awkward angle. | [5 empty] | The investigation into the reported access site bleeding- major has been completed since the original report was filed. the root cause of the access site bleeding was most likely due to use issue since the per doctor, the hematoma was caused by his stick making the impella take an awkward angle. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31319 | 1220648-2024-06480 | 18610216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18610216&pc=OZD | 01/17/2024 | 09/21/2024 |  | 01/30/2024 |  |  | 01/17/2024 | 01/30/2024 |  |  | 11/30/2025 |  | 01/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426429 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Thrombosis/Thrombus | [11 empty] | Life Threatening | Required Intervention | Disability | Death | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/17/2024 |  |  |  |  | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 81-year-old male (unknown race and ethnicity) undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having coronary disease. after placing the impella, flows were slightly lower, and the waveforms had changed. the physician repositioned the impella, and the ao wave forms immediately changed, and the pressures quickly dropped. an echo was done and revealed an effusion. a pericardiocentesis was performed and blood was removed and auto returned, and cardio-pulmonary resuscitation was performed along with levo and epi given. emergent surgery was performed to remove the patient's clot in the pericardial space and to repair the hole. the physician reported that the left ventricle was very fragile and was deteriorating upon sewing the ventricle for repair. the patient continued with impella support. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation for the reported thrombus and ventricle perforation has been completed. the impella device was not returned for evaluation. without the device or sufficient clinical details, the root cause of the thrombus could not be determined. however, clinical details provided regarding the ventricle perforation were sufficient to determine the root cause. the root cause of the ventricle perforation issue was most likely improper positioning of the device. the device was a little too deep into the left ventricle. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
31320 | 1220648-2024-06366 | 18594803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18594803&pc=OZD | 01/16/2024 | 02/01/2024 |  | 01/29/2024 |  |  | 01/16/2024 | 01/29/2024 |  |  | 09/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024395932 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was having an impella 5.5 implant when the staff was unable to deliver the impella. the impella was not able to cross the aortic valve. the decision was made to proceed with weaning from bypass and abort the impella implant. reportable due to potential patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided are sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition due to patient anatomy. despite numerous attempts under echo guidance and aortic ultrasound, the valve was unable to be crossed with multiple wires and catheters. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon completion of the investigation a supplemental mdr will be filed.
31321 | 1220648-2024-06367 | 18594870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18594870&pc=OZD | 01/17/2024 | 03/07/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 2024426090 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic (automated impella controller) failed to recognize an installed purge cassette during patient support. the cassette was swapped, but the issue remained. the console was subsequently replaced and support was able to continue. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. the reported issue with the purge disc not being able to be detected could not be reproduced. no issues found with the purge flag or retainer after visual inspection. there was no issue found during the case. the device functioned/operated to specification. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31322 | 1220648-2024-06368 | 18594904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18594904&pc=OZD | 01/18/2024 | 02/28/2024 |  | 01/29/2024 |  |  | 01/18/2024 | 01/29/2024 |  |  | 11/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024426001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Integrity Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following a difficult delivery of an impella cp pump, suction alarms were triggered due to a kink near the inflow. support levels were lowered, repositioning was performed, and a unit of volume was given to troubleshoot the suction, but the issue remained. the patient remained stable despite the issue and the pump was subsequently explanted. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. product & data were not returned for review. the root cause of low flow was most likely due to kinked cannula. the root cause of kinked cannula was most likely due to patient condition. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31323 | 1220648-2024-06369 | 18594967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18594967&pc=OZD | 09/18/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/18/2023 | 01/29/2024 |  |  | 01/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024309787 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 42-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient on impella support experienced oozing from the access site and was required to have two stiches to stop it. | [5 empty] | No product was returned. as per clinical review, the patient was oozing at the impella access site. systemic heparin was on hold and two stitches were placed to achieve hemostasis. the cause of the access site bleeding issue was most likely use related due to additional sutures resolving complication per clinical. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31324 | 1220648-2024-06370 | 18594979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18594979&pc=OZD | 07/12/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 07/12/2023 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1193362 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/12/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a red alarm indicating a battery failure. the aic was returned for evaluation. there was no patient involvement. | [5 empty] | The investigation into the automated impella controller's battery (red alarm) has been completed. data analysis: imc logs from the complaint occurrence date confirmed that there were instances of battery failure within the console. we see two types of battery failure alarms with event set #360 (transport connection blown/missing) and event set #350 (low voltage). as these battery failure alarms are occurring, there are instances of event set #23 showing a battery level low alarm (50% or below). device analysis: the battery failure alarm was reproduced on an engineering lab bench. battest showed that the logic fuse (fet) was activated with battery 1 given ¿fu¿ was missing. there no power output with battery 1. battery data was extracted with the ata battery cable, and no safety mechanisms were triggered due to permanent failure. the battery failure alarm was due to a defective battery. the root cause of the defective battery was not determined.
31325 | 1220648-2024-06371 | 18595076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595076&pc=OZD | 01/22/2024 | 03/05/2024 |  | 01/29/2024 |  |  | 01/22/2024 | 01/29/2024 |  |  | 05/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351797 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had attempted impella 5.5 support and upon implanting the impella, there was a left ventricle perforation. the impella 5.5 was advanced over the 0.018 wire and when attempting to torque the cannula away from the mitral valve over the wire, the doctor noticed bleeding and found the patient had a posterior wall perforation. the impella was aborted and two sutures were applied. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation for the reported ventricle perforation has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the most likely root cause of the ventricle perforation was use-related, as there was difficulty torquing the cannula away from the mitral valve per clinical details. d4- model number updated to impella 5.5.
31326 | 1220648-2024-06372 | 18595453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595453&pc=PYX | 07/18/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 07/18/2023 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1125047 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/18/2023 |  |  |  |  | The user facility reported a patient of unknown age, gender, and ethnicity was implanted with an impella rp device for mechanical circulatory support for an unknown indication. it was reported that the purge solution was leaking in and around the purge cassette and the automated impella controller (aic) was wet under the purge door cover was leaking and needed to be changed. no visible damage to the cassette was noted. no patient harm was reported. | [5 empty] | Data analysis: automated impella controller (aic) logs did not have any alarms or issues around the date complaint. device analysis: the console was tested on an engineering lab bench. no issue was found upon running the aic console for an hour with known good pump and purge cassette, and no fluid leakage was found around the purge assembly. root cause: there was no issue found during the case. the device functioned/operated to specification.
31327 | 1220648-2024-06373 | 18595501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595501&pc=OZD | 10/06/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 10/06/2023 | 01/29/2024 |  |  | 04/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024366222 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp met resistance in the iliac on insertion. after multiple attempts to pass the stenosis vessel with no success, the case was aborted. | [5 empty] | No product was returned. the root cause of the delivery issue - anatomy was most likely due to patient condition.
31328 | 1220648-2024-06374 | 18595511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595511&pc=OZD | 08/14/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 08/14/2023 | 01/29/2024 |  |  | 05/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024365111 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/14/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella was primed and inserted but was unable to pass through the iliacs. procedure was aborted. no known patient harm or injury. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. product was not returned for investigation. the cause of the delivery issue was most likely patient condition related.
31329 | 1220648-2024-06375 | 18595631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595631&pc=OZD | 10/06/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 10/06/2023 | 01/29/2024 |  |  | 07/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024376305 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was to be escalated to an impella 5.5 device. the team attempted impella 5.5 insertion into right axillary and was unable to access due to anatomy. several attempts made were unsuccessful. the patient was kept on the impella cp that was already in place and was sent back to intensive care unit. | [5 empty] | The investigation into the delivery issues-anatomy has been completed. the impella product was discarded. the root cause of the delivery issue - anatomy was most likely due to patient condition.
31330 | 1220648-2024-06376 | 18595632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595632&pc=OZD | 09/08/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/08/2023 | 01/29/2024 |  |  | 10/31/2023 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023173318 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2023 |  |  |  |  | The user facility reported an 87-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. the impella 2.5 was unable to be inserted due to the patient¿s anatomy. while the physician was struggling to pass through the curves of the aorta, the aic alarmed ¿purge pressure low¿. all connections were checked, and no fluid was leaking. a new purge cassette was used, and the case started again, and a purge pressure low alarm occurred. the case was aborted. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for evaluation. the most likely cause of the delivery issue was patient condition related due to patient anatomy because patient had severe tortuosity of the aorta.
31331 | 1220648-2024-06377 | 18595683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595683&pc=OZD | 09/14/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/14/2023 | 01/29/2024 |  |  | 05/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024356225 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/14/2023 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 was unable to pass through the right subclavian artery. the physician tried rotating the arm, a buddy wire and a contrast injection of the artery confirmed a very small artery. the procedure was aborted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. delivery anatomy- the root cause of the delivery issue - anatomy was most likely due to patient condition due to patient's small artery size.
31332 | 1220648-2024-06378 | 18595716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595716&pc=OZD | 09/30/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/30/2023 | 01/29/2024 |  |  | 06/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024371888 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient Device Interaction Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/30/2023 |  |  |  |  | The user facility reported a 73-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the peel away sheath was unable to be advanced due the vessel being very calcified and very tortuous. the physician opted to use cook 14 french check flow and had no issues with access or the passage of cp cannula. no patient harm was reported. | [5 empty] | The investigation has been completed. no product was returned for investigation. the root cause of the delivery issue was patient condition related to calcified and tortuous vessels.
31333 | 1220648-2024-06379 | 18595833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595833&pc=OZD | 09/18/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/18/2023 | 01/29/2024 |  |  | 04/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024353338 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/18/2023 |  |  |  |  | The user facility reported an (b)(6) year-old male patient noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported a right guide was inserted to assess right coronary artery which looked fine. the physician then decided to pass a left guide to assess left main coronary artery and was unable to get the left guide to go up and around the aorta. the procedure was aborted. no patient harm was reported. | [5 empty] | No product was returned for investigation. per the clinical information provided, the root cause of the delivery issue was patient condition related to calcified vessels. the root cause of the delivery issue was most likely patient condition related to calcified vessels. the investigation has been completed.
31334 | 1220648-2024-06380 | 18595863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595863&pc=OZD | 09/26/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/26/2023 | 01/29/2024 |  |  | 07/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370527 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2023 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. multiple attempts were made to cross the anastomosis with the impella 5.5. the nose cone continued to get stuck in the anastomosis. attempts were made to lubricate the device, wire, graft, and to align the graft with the artery. torque device was used to assist in crossing. all attempts were unsuccessful, and the case was aborted. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was returned for review. visual inspection found no issues on the pump. based on clinical description and product analysis, the root cause of delivery issue was most likely due to patient condition.
31335 | 1220648-2024-06381 | 18595968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18595968&pc=OZD | 07/25/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/27/2023 | 01/29/2024 |  |  | 05/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344552 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2023 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as escalation of care from an impella cp device. . the patient's native anatomy precluded placement. no reported patient harm or injury. | [5 empty] | The investigation into the delivery issues-anatomy has been completed. the impella products were not returned for evaluation. based on the clinical review, the root cause of the delivery issue - anatomy was most likely due to patient condition of diseased/small vessels.
31336 | 1220648-2024-06382 | 18596115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596115&pc=OZD | 01/14/2024 | 03/01/2024 |  | 01/29/2024 |  |  | 01/14/2024 | 01/29/2024 |  |  | 11/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423055 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2024 |  |  |  |  | The user facility reported a 69-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support the patient experienced access site bleeding. vascular surgeon performed femoral cutdown and purse string suture were placed. the bleeding was resolved and impella support continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding -major has been completed since the original report was filed. the root cause of the access site bleeding was most likely use issue since adding extra suture resolved the bleeding issue.
31337 | 1220648-2024-06383 | 18596127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596127&pc=OZD | 01/11/2024 | 11/12/2025 |  | 01/29/2024 |  |  | 01/11/2024 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1158457 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 67-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a "console error-switch to backup console" error message. the error message resolved itself, but the team switched consoles proactively. the aic was returned for evaluation. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the impella device was returned for evaluation. based on the device and data analysis the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. d4 model number was erroneously entered on the initial manufacturer device report number. | Corrections to the initial mfr report: g1 mdr reporting contact email d2 device name updated f6 and f8 should have been left blank updated h3 to device evaluation anticipated h6 type of investigation code added. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31338 | 1220648-2024-06384 | 18596291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596291&pc=OZD | 01/15/2024 | 09/20/2024 |  | 01/29/2024 |  |  | 01/15/2024 | 01/29/2024 |  |  | 10/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429163 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/15/2024 |  |  |  |  | The user facility reported an 82-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the implanted impella cp pump began to experience suction issues coinciding with a spike in motor current during patient support. the staff was advised that the pump would need to be replaced and the decision was ultimately made to explant the device. there was no patient harm. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported low pump flow has been completed since the original report was filed. data logs confirmed the low flow. logs showed about 50.5 hours into support, mc spiked followed by low flows that triggered auto suction response & suction alarms. pump then was stopped manually & disconnected from the console. visual inspection found biomaterial inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of the low flow was biomaterial ingestion. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
31339 | 1220648-2024-06385 | 18596363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596363&pc=OZD | 01/15/2024 | 03/01/2024 |  | 01/29/2024 |  |  | 01/16/2024 | 01/29/2024 |  |  | 06/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361291 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after one day of support, the pump was explanted as it could not maintain the appropriate position within the patient. the impella cp was replaced with a new impella cp pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported positioning issue has been completed since the original report was filed. the cause of the positioning issue was patient condition related due patient anatomy as he has a wide and short aortic arch that made the impella place properly difficult.
31340 | 1220648-2024-06386 | 18596397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596397&pc=OZD | 01/12/2024 | 02/29/2024 |  | 01/29/2024 |  |  | 01/12/2024 | 01/29/2024 |  |  | 02/28/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338777 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for left ventricle unloading. after the impella implant, it was difficult to stop the access site bleeding. team held pressure for around one hour, tightened perclose sutures, added purse string sutures, maintained angle of insertion, and even attempted to advance blue suture rings into arteriotomy, but all unsuccessful with stopping groin bleeding. the impella was removed and replaced. patient received six units of packed red blood cells for blood loss. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding was most likely due to anticoagulation management since the act levels were high and out the normal range.
31341 | 1220648-2024-06387 | 18596406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596406&pc=OZD | 11/01/2021 | 01/29/2024 |  | 01/29/2024 |  |  | 11/01/2021 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1262855 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2021 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an unspecified impella device for mechanical circulatory support. the purge disk was not recognized and there was no mention of patient harm or involvement requiring medical intervention. | [5 empty] | Logs from the complaint date revealed that the console issued a purge disc not detected alarm several times. the clinical staff attempted to replace and reinsert different purge cassettes and disks with no success. the console was tested. the issue with properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken (missing at blue disc retainer). a brown residue was found on the surface of the purge cassette retainer. it was speculated to be a small amount of dextrose dried up. the dextrose remaining was cleaned right away. the cause of the purge disk not being recognize was the purge flag missing/ broken.
31342 | 1220648-2024-06388 | 18596471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596471&pc=OZD | 01/16/2024 | 03/12/2024 |  | 01/29/2024 |  |  | 01/16/2024 | 01/29/2024 |  |  | 10/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424388 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. bounding distal dorsalis pedis pulse in right leg was noted. hemolysis was also noted, and the p-level was reduced and 500ml albumin bolus was administered. on day three of impella support, impella groin site had no distal pulses present. below knee was cold and mottled. the impella cp was removed, and clot was found at groin site. vascular surgery will make the decision on closure of impella groin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the access site bleeding ¿ minor, the hemolysis, the limb ischemia ¿ reversible and the thrombus has been completed since the original report was filed. the device was not returned for investigation. the cause of the hemolysis issue was most likely patient condition related. as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding ¿ minor, the hemolysis, the limb ischemia ¿ reversible and the thrombus was not determined.
31343 | 1220648-2024-06389 | 18596541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596541&pc=OZD | 01/13/2024 | 08/14/2025 |  | 01/29/2024 |  |  | 01/13/2024 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1160667 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Disconnection | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/13/2024 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, the automated impella controller (aic) plug was found to be disconnected from the terminals. the aic was replaced and there was no patient injury or harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | B.7 added information as it was inadvertently omited from the initial manufacturer device report number 1220648-2024-06389. d.2 and d.4 common device name and moedel number were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06389 was submitted. e.1 added fax number as it was inadvertently omited from the initial manufacturer device report number 1220648-2024-06389. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11956 and should not have been. g.1 revised mdr reporting contact email in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06389 was submitted. h.6 code 3221 should not have been reported on the initial manufacturer device report number 1220648-2024-06389 in accordance with updated procedures. code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2024-06389 and a new code was added.
31344 | 1220648-2024-06390 | 18596623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596623&pc=OZD | 01/16/2024 | 02/14/2024 |  | 01/29/2024 |  |  | 01/16/2024 | 01/29/2024 |  |  | 10/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415458 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 54-year-old female patient female was implanted with an impella 5.5 as a bridge to ventricular assist device (vad). it was reported that distal end of the sterile sleeve became detached from the tuohy-borst. the patient was taken to the operation room and the catheter was cleaned with betadine and sterile sleeve wrapped with tegaderm. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (sterile sleeve damage) has been completed. the device was not returned for investigation. based on the clinical information provided, it was noted by clinical team that anti-contamination sleeve was detached from tuohy-borst lock at distal end. the root cause of the anti-contamination sleeve detachment cannot be determined.
31345 | 1220648-2024-06392 | 18596665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596665&pc=OZD | 03/04/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 03/06/2023 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1390491 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/06/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) purge disc retainer was damaged while the rns were changing out the purge cassette. there are no patient details related to this event. the aic was exchanged. | [5 empty] | The investigation into the broken purge disc retainer has been completed. console logs from the reported day of event do not contain information relevant to this failure mode. visual inspection of the console, performed in (b)(6), confirmed broken purge disk retainer, and missing disk detect flag.
31346 | 1220648-2024-06393 | 18596669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596669&pc=OZD | 06/28/2022 | 01/29/2024 |  | 01/29/2024 |  |  | 06/28/2022 | 01/29/2024 |  |  | 08/31/2023 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2022134805 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. us complainant reported the impella access site was bleeding. a sandbag was placed overnight and controlled, but continued to bleed. additionally, the right leg had low pulses. the staff made the decision to remove the impella cp and implant the impella 5.5. | [5 empty] | No product was returned. the root cause of the limb ischemia was most likely patient condition. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31347 | 1220648-2024-06394 | 18596675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596675&pc=OZD | 08/10/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 08/10/2023 | 01/29/2024 |  |  | 06/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024368874 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/10/2023 |  |  |  |  | The user facility reported a 19-year-old female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. upon delivering the impella 5.5, the axillary artery was noted to be 5-6mm, but placement of 5.5 was still pursued. due to small vessel, 8mm vascutec gel weave utilized and anastomosed onto vessel. several attempts were made to get the 5.5 through. additional, lubricating agents (viper slide and mineral oil) and local vasodilators (papvirne) were utilized to aid in the delivery of the 5.5. ultimately, the vessel was too tight, and the decision was made to remove the 5.5 and implant a impella cp through the right axillary artery. no known patient harm or injury. | [5 empty] | The investigation into the delivery issue-anatomy has been completed. the impella product was not returned for evaluation. per the clinical review, the root cause of the delivery issue - anatomy was most likely due to patient condition of diseased/small vessels.
31348 | 1220648-2024-06395 | 18596748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596748&pc=OZD | 07/03/2022 | 01/29/2024 |  | 01/29/2024 |  |  | 07/03/2022 | 01/29/2024 |  |  | 08/31/2023 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134472 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 51-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a impella cp support ongoing when the team had to explant and replace/escalate to the impella 5.5 pump. this was due to the cp repo sheath being flow occlusive. pulses were lost. | [5 empty] | No product was returned. as per clinical review, the patient had ischemia on right leg with angiogram revealing occluded right femoral artery and the pump was explanted. surgery was planned post explant. the cause of the limb ischemia issue was most likely patient condition related based on clinical review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31349 | 1220648-2024-06391 | 18596854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596854&pc=OZD | 02/28/2023 | 01/31/2024 |  | 01/29/2024 |  |  | 02/28/2023 | 01/29/2024 |  |  | 09/30/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280075 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old asian male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella support and was having hemolyzed labs. there was no central venous pressure (cvp) noted. the staff stated there was no cvp and the position of the impella was correct. the patient eventually expired due to withdrawn care after 71.09 of support. no allegations were made towards the impella for cause of death. | [5 empty] | The investigation into the hemolysis issue has been completed. the product was not returned for investigation. data logs were reviewed, and suction alarms were observed. the cause of the hemolysis issue was not determined since the product was not returned and based on insufficient clinical information. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | (h3) the investigation into the reported "hemolysis" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the data logs were reviewed, and suction alarms were observed. there was no other abnormalities, the cause of the hemolysis issue was not determined since product was not returned and based on insufficient clinical information, data log review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31350 | 1220648-2024-06396 | 18596871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596871&pc=OZD | 08/15/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 08/15/2023 | 01/29/2024 |  |  | 07/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024371646 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2023 |  |  |  |  | The user facility reported a 65-year-old female noted as post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 device for a bailout situation. the team had access site and delivery issues. post placement, the team observed the left ventricle (lv) to be perforated. the 5.5 was removed for surgical repair of the lv. afterwards the team was unable to deliver a new 5.5 pump and the patient was medically managed and expired. mdr 1220648-2023-03113 has been submitted to the fda for perforation. | [5 empty] | No product was returned for investigation. as per clinical the reason the insertion was aborted due to patient anatomy. the root cause of the delivery issue - anatomy was most likely due to patient condition. the investigation has been completed
31351 | 1220648-2024-06397 | 18596930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596930&pc=OZD | 10/30/2022 | 01/29/2024 |  | 01/29/2024 |  |  | 10/30/2022 | 01/29/2024 |  |  | 07/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255205 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support following a stemi. the patient lost pulse in left lower extremity due to distal limb ischemia. the decision was therefore made for emergent escalation to axillary impella 5.5 and removal of left femoral impella. after impella cp was removed, a positive pulse in the left leg was verified. | [5 empty] | The product was not returned for investigation. the cause of the limb ischemia issue was patient condition related since the patient had distal limb ischemia during pump implant and limb ischemia resolved after removal of pump. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31352 | 1220648-2024-06398 | 18596936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18596936&pc=PYX | 03/07/2023 | 01/31/2024 |  | 01/29/2024 |  |  | 03/07/2023 | 01/29/2024 |  |  | 02/29/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023184996 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported 80-year-old white male patient on impella support experienced the pump stop. the device was restarted and did not have any problems since. the medical staff decided to wean off impella and shortly before planned explant, patient needed more support and pressors. family consulted and decision made to withdraw care and patient expired. the patient expiry is not related to the alleged malfunction. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the pump was not returned for investigation. logs were returned and reviewed. during log review, multiple pump-stop alarms were confirmed in the start of the case and a trend in motor current spike was seen throughout the case. the root cause of the pump did not start issue was most likely biomaterial. | (h3) the investigation into the reported "pump stop" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; logs were returned and reviewed, the root cause of the pump did not start issue was most likely biomaterial. during log review, multiple pump-stop alarms were confirmed in the start of the case and a trend in motor current spike was seen throughout the case. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31353 | 1220648-2024-06399 | 18597047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597047&pc=OZD | 08/19/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 08/19/2023 | 01/29/2024 |  |  | 01/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332170 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/19/2023 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the wire and pigtail were placed without issue however, due to what appeared to be calcification in the descending/abdominal aorta the impella was unable to be placed. impella was removed and switched to longer sheath, but was still unsuccessful. angiogram after showed tight calcification. no known patient harm or injury. | [5 empty] | The investigation into the delivery issues-anatomy has been completed. the impella products were not returned for evaluation. based on the clinical review, the root cause of the delivery issue - anatomy was most likely due to patient condition of diseased/small vessels.
31354 | 1220648-2024-06400 | 18597125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597125&pc=OZD | 02/19/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 02/19/2023 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1133152 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/19/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a purge disc not detected alarm. there was no patient involvement. the aic was returned for evaluation. | [5 empty] | The investigation into the automated impella controller (aic) purge disc/flag issues has been completed. logs from the complaint date revealed that the console issued the purge disc not detected alarm several times during one of the first clinical procedures after the console arrived back from annual preventive maintenance. the console was tested after arriving in field service. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). the cause of the issue was a broken purge flag.
31355 | 1220648-2024-06401 | 18597216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597216&pc=OZD | 01/16/2024 | 02/29/2024 |  | 01/29/2024 |  |  | 01/16/2024 | 01/29/2024 |  |  | 09/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024405176 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump had purge pressure/purge flow issues which the team attempted to troubleshoot with purge cassette replacement. this did not resolve the issue. the pump remains on for support aware that the pump may stop. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the purge leak - pump was not determined since product was not returned and there is a lack of clinical details.
31356 | 1220648-2024-06402 | 18597261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597261&pc=OZD | 08/25/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 08/25/2023 | 01/29/2024 |  |  | 03/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347297 |  | 1000080 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2023 |  |  |  |  | The user facility reported a 78 year old female patient in non-st elevated myocardial infarction was to be implanted with an impella cp device for mechanical circulatory support. it was reported a 14fr x 13cm sheath and they were able to place the sheath was placed with some difficulty. the impella cp was attempted several times to be inserted, but would never cross the bifurcation, and unfortunately the procedure was aborted due to access issue. the impella was unable to be inserted into any of the arterial vessels. no patient harm was reported. | [5 empty] | Product was not returned for investigation. the cause of the delivery issue was patient condition related to artery bifurcation, based on provided clinical details.
31357 | 1220648-2024-06403 | 18597267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597267&pc=OZD | 01/17/2024 | 03/04/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108091 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a patient of unknown age, gender, race, and ethnicity implanted with an impella device mechanical circulatory support. it was reported that automated impella controller (aic) failed to recognize an installed purge cassette during patient support. the aic was subsequently swapped out to resolve the noted malfunction. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was tested after arriving in fs. the issue with the aic not being able to detect the purge disc was confirmed. the cause of the issue was a broken purge flag.
31358 | 1220648-2024-06404 | 18597302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597302&pc=OZD | 09/01/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/01/2023 | 01/29/2024 |  |  | 06/09/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372988 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2023 |  |  |  |  | The user facility reported a 69-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support post coronary artery bypass grafting x 4. the team was unable to position the impella in ventricle for optimal position due to anatomy. impella was resting up against wall of ventricle which did not allow for good flows and suction was sustained and unable to break until p1. the procedure was aborted. | [5 empty] | The investigation into the delivery/positioning issues has been completed. the impella products were not returned for evaluation. based on the clinical review, the root cause of the delivery issue - anatomy was most likely due to patient condition because patient's heart in more of a horizontal position. patient had short aorta. impella positioning tried twice. impella inlet resting against ventricle wall.
31359 | 1220648-2024-06405 | 18597441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597441&pc=OZD | 01/17/2024 | 04/01/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1366755 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Malfunction | No | Yes | Use of Device Problem | Optical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic (automated impella controller) displayed a controller error alarm during impella cp start-up. the placement signal was absent at this point in time, but the flow rate remained unaffected. as a result of the noted alarm, the decision was made to swap the aic. following the swap, it was additionally noted that the device was not shut off properly. there was no patient harm associated with the noted events. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm), aic - shutdown or restart issue and the optical signal issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. it was found that the front panel optical connector was heavily contaminated. however, the controller error was not reproduced while running an optical pump simulator. the root cause of the optical signal issue was most likely the air gap caused by the pump connector not being fully connected to the front panel optical connector. the root cause of the controller error was due to an aic software issue.
31360 | 1220648-2024-06406 | 18597562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597562&pc=OZD | 05/24/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 05/24/2023 | 01/29/2024 |  |  | 12/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309627 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported attempted to place two impella cp pumps but both failed. the delivery was complicated by the aortic calcification. the team aborted pump delivery and support. there was no lasting harm or injury. | [5 empty] | (h3) the investigation into the reported delivery issue has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the delivery issue was patient condition related because the patient had so much calcium and a porcelain aorta which prevented advancement of impella. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31361 | 1220648-2024-06407 | 18597657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597657&pc=OZD | 10/16/2023 | 09/19/2024 |  | 01/29/2024 |  |  | 10/16/2023 | 01/29/2024 |  |  | 05/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024363618 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 72-year-old male patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the patient coded upon arrival to cardiac catherization lab and it was reported was cp unable to advance through descending aorta due to severe calcification. the case was aborted and the patient expired. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06407 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | No product was returned for investigation. per the given clinical information, the root cause of the issue was most likely patient condition related.the investigation has been completed.
31362 | 1220648-2024-06408 | 18597665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597665&pc=OZD | 03/08/2023 | 02/01/2024 |  | 01/29/2024 |  |  |  | 01/29/2024 |  |  | 08/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269882 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 66-year-old white male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed hemolysis and an access site bleed during impella support. the hemolysis resolved after the pump was repositioned and the access site bleed was treated with femostop. there was no lasting patient harm. | [5 empty] | The investigation into the reported "hemolysis" has been completed. no product was returned for investigation. the cause of the hemolysis issue was most likely improper positioning of the device since it was resolved by repositioning. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ h.2 device evaluation was selected on the original manufacturer device report (mdr) and should not have been selected as that report was not a follow-up report. h.6 additional codes were added that were inadvertently left off the previously submitted mdr. h.10 the additional manufacturer narrative on the previous mdr was in correct and the correct information is reported in h.10 of this report. | (h3) the investigation into the reported "hemolysis" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the hemolysis issue was most likely improper positioning of the device since it was resolved by repositioning. the cause of the access site bleeding issue was not determined since product was not returned and sufficient information was not returned. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Correction section: h.3 evaluation summary was not attached and was inadvertently checked on the initial manufacturer device report.
31363 | 1220648-2024-06409 | 18597692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597692&pc=OZD | 03/13/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 03/13/2023 | 01/29/2024 |  |  | 10/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024296857 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 78-year-old male patient experienced ventricular fibrillation during pump manipulation ((impella placement was optimized by withdrawing device approximately 1cm) and decreases in the support level. the patient required defibrillation after each ventricular fibrillation arrest episode and support was maintained throughout the event. amiodarone 150mg bolus given post shock. support with the implanted pump continued following the intervention. | [5 empty] | The investigation into the reported ventricular fibrillation (vf) issue has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. as per the clinical review, p-level gradually decreased to p-1 then patient experienced ventricular fibrillation. data logs were reviewed, and the downward trending placement signal was indicative of poor patient condition. any device contacting the heart wall in conjunction with a poor patient condition could result in vf. the vf occurred when impella was pulled back 1 cm to optimize position. the root cause of the ventricular fibrillation issue was determined to be most likely due to patient condition and related to impella based on occurrence of vf events per log and clinical review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ a risk assessment was performed, and no escalation is required at this time. this report is being filed as part of a retrospective review of historical records.
31364 | 1220648-2024-06410 | 18597693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597693&pc=OZD | 10/26/2023 | 11/04/2025 |  | 01/29/2024 |  |  | 10/26/2023 | 01/29/2024 |  |  | 08/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024388737 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Separation | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The complainant also reported that placement signal monitoring was disabled due to issues with the fiber optic cable. | The user facility reported a 32-year-old female patient was implanted with an impella 5.5 as a bridge to transplant. it was reported that while repositioning, they found the back of the anti-contamination sterile sleeve compromised and they placed tape on it, to hold it in place. the clinical representative walked them through connecting it back correctly. no patient harm was reported. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. product damage (sterile sleeve) product was not returned. per the clinical, the back of the anti-contamination sterile sleeve compromised and the hospital staff placed tape on it, to hold it. there was no mention of how this damage happened. thus, the root cause of the product damage (sterile sleeve) is unable to be determined as no product was returned and provided clinical details are insufficient. optical signal issue product was not returned. the data logs were reviewed and a temporary placement signal not reliable (psnr) with intermediate spike in placement signal, gain and contrast was observed. based on the rt log, the first psnr was observed at 02:12 pm on 2023 oct 25 and other 9 psnr alarms were noted in the span of 1 hour and 12 minutes. further analysis of the rt logs showed that the snr value do not vary even though there is a variation in the raw optical sensor reading. furthermore, the observed trend of the 5s parameters - snr, light, gain, lamp and contrast do not correlate with any of the failure modes from the optical failures. thus, the root cause of the optical signal issue is unable to be determined as no product was returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated. b7 other relevant history, including preexisting medical conditions updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31365 | 1220648-2024-06411 | 18597707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597707&pc=OZD | 01/17/2024 | 03/01/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 11/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024426461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported an 81-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that patient developed an access site bleed and a hematoma. the patient was administered red blood cells, applied manual pressure, and a pressure dressing was applied as treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643apply manual pressure above the puncture site. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the access site bleeding ¿ major event has been completed. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical was not provided.
31366 | 1220648-2024-06412 | 18597747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597747&pc=OZD | 03/13/2023 | 09/19/2024 |  | 01/29/2024 |  |  | 03/13/2023 | 01/29/2024 |  |  | 01/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185163 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that the flow rate of an implanted impella cp pump dropped from 2.0 l/min down to 0.0 l/min during support. the drop in flow rate coincided with reports of continuous suction. it was theorized that the patient had an lv thrombus that was acutely ingested by the pump. the patient passed away from their pre-existing condition prior to any discussions of intervention or pump removal. there was no patient injury related to the low pump flow. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06412 was provided. | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the most likely cause of the low pump flow was most likely biomaterial as the physician suspected lv thrombus and the data logs show evidence of ingestion. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
31367 | 1220648-2024-06413 | 18597748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18597748&pc=OZD | 02/20/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 02/20/2023 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1205609 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/20/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) did not sense the purge disc. the event occurred during set-up and there was no patient involvement. the aic was swapped out with a backup aic. | [5 empty] | The investigation into the automated impella controller (aic) purge disc/flag issue has been completed. logs revealed that there were issues priming the impella during the case start. logs showed pressure remained high after forward stroke of the piston, which was likely due to a kinked line or momentary blockage. eventually, there was abnormal fluctuations in the pressure (from 0 to 2000 mm hg) which was likely due to the clinical team troubleshooting the broken purge disc retainer. the console was inspected after arriving in field service the purge disc retainer was confirmed to be broken (broken blue disc retainer) the root cause of the purge issues was due to a broken retainer.
31368 | 1220648-2024-06414 | 18598119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18598119&pc=OZD | 01/17/2024 | 02/29/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 10/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024408800 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 46-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 delivered for support of a patient admitted with cardiomyopathy. the 5.5 was placed and the pump sidearm was damaged and leaking after over a month of use. the sidearm was troubleshot and the pump remains on for support. the patient awaits a heart transplant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿. | The investigation into the leak ¿ pump issue has been completed since the original report was filed. the device was returned to fs for investigation. it was reported that when purged in the lab, the sidearm was found to have been leaking from the filter at the air vent holes. the root cause of the purge leak - pump was determined to be sidearm filter damage allowing purge fluid to leak out of the air vent holes on both sides of the filter.
31369 | 1220648-2024-06415 | 18598201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18598201&pc=OZD | 02/22/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 02/22/2023 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1171907 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/22/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) purge disc retainer was damaged. there was no patient involvement. | [5 empty] | The investigation into the automated impella controller (aic) broken disc retainer has been completed. logs are not relevant to this complaint. the device was received at field service. the cause of the issue was a broken purge retainer.
31370 | 1220648-2024-06416 | 18599269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18599269&pc=OZD | 02/22/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 02/23/2023 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1133145 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/23/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) generated a controller error. there was no patient involvement during this event. | [5 empty] | The investigation into the automated impella controller (aic) purge issue-other has been completed. console logs from the reported day of event are mostly consistent with complaint and showed alarms, opm comm crc invalid and ambient pressure out of range, occurred. there were no purge system failure alarms. there were some purge-related alarms. after the purge cassette was replaced, these alarms did not re-occur. the reported controller error alarms were confirmed in rt logs and identified as a known pattern failure caused by a temporary loss of optical placement signal. abiomed is aware of the issue, and is working on appropriate corrective action. there¿s no need for repairs if the condition self-corrects. this is a low probability event and lasts only a few minutes. if needed, monitor patient hemodynamics and impella position with imaging. purge-related alarms did not re-occur after purge cassette swap. it is reasonable to assume that purge cassette might have been defective. however, the alarms might have been use-related (luers not secured, purge line kinks, etc). the purge cassette has not been returned for analysis. console inspection revealed evidence of contamination (purge fluid spill) in the purge insert area, which might support theory of purge cassette or purge tubing leak. however without the purge to analyze, there is no evidence linking any leaked purge fluid to a specific purge cassette. during testing, the console operated within specification. the cause of the issue was not determined since the issue was not reproduced.
31371 | 1220648-2024-06417 | 18600035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18600035&pc=OZD | 02/26/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 02/26/2023 | 01/29/2024 |  |  | 08/31/2023 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130334 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient began to experience hemolysis coinciding with impella support. the pump was repositioned in an attempt to troubleshoot the issue, but the pump could not be torqued away from the mitral. the support level was subsequently reduced to p6 to counteract the condition. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products. data logs were returned and investigated and no abnormalities were found. per the clinical details provided, the root cause of the hemolysis issue was most likely improper positioning of the device. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31372 | 1220648-2024-06418 | 18600396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18600396&pc=OZD | 01/17/2024 | 01/29/2024 |  | 01/29/2024 |  |  |  | 01/29/2024 |  |  | 03/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024345285 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 72-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that impella cp procedure was aborted as the patient has severe bilateral femoral and iliac arteries calcification. there was no patient harm. | [5 empty] | The medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the delivery issue was most likely due to patient condition since calcification in vessels was observed.
31373 | 1220648-2024-06419 | 18601261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18601261&pc=OZD | 01/17/2024 | 03/05/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 01/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327999 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 33-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to sudden onset chest pain with st (sinus tachycardia) elevation in leads v1-v3. flows were noticed to be decreased with purge flow below 3 and purge pressure increased. the physician repositioned the impella and purge pressure decreased but was still high. there were no kinks or pressure on tubing. the decision was made to replace the cp due to potential clot ingestion. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low flow and high purge pressure issues has been completed. the impella cp was returned for investigation. during visual inspection, biomaterial was observed at the impeller purge gap and outflow cage. no other abnormalities like catheter kink were found. the purge cassette and pump were connected to the console to reproduce high purge pressure. purge pressure increased to 800 mmhg initially and later started to reduce after 15 minutes of continuous purging. purge fluid slowly started to pass at the impeller purge gap. data logs were reviewed, and high purge pressure alarm observed with rise in purge pressure along with motor current spike. purge flows were found to be low during this time. later the high purge pressure event resolved with gradual reduction in purge pressure for next 1.5 hours until explant. therefore, the root cause of the high purge pressure issue was biomaterial ingestion at purge gap. the instructions for use provides details to prevent thrombus formation. it states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ for purge pressure issues, it states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31374 | 1220648-2024-06420 | 18601264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18601264&pc=OZD | 01/18/2024 | 03/01/2024 |  | 01/29/2024 |  |  | 01/18/2024 | 01/29/2024 |  |  | 07/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386921 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 29-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella site experienced oozing during patient support. an additional suture was placed at the site to resolve the condition. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major event has been completed. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved after adding suture.
31375 | 1220648-2024-06422 | 18601382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18601382&pc=OZD | 06/15/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 06/15/2023 | 01/29/2024 |  |  | 12/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298047 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2023 |  |  |  |  | The user facility reported a 78-year-old female was implanted with an impella cp device for mechanical circulatory support. the impella was placed as a bailout peri-arrest. the patient was also cannulated for ecmo.the patient developed access site bleeding around the sheath. to resolve the bleeding, the impella cp was explanted. | [5 empty] | The investigation into patient's access site bleeding has been completed. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to lack of vessel recoil. instructions for use for this event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31376 | 1220648-2024-06421 | 18601406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18601406&pc=OZD | 02/21/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 02/21/2023 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1102538 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Analyze Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a broken purge clip on the automated impella controller (aic). there was no reported patient injury. | [5 empty] | The investigation into the reported events has been completed. console logs did not contain any entries on the reported day of event and log entries from the last use did not contain any relevant data, related to this complaint. visual inspection of the console confirmed broken (missing) purge disk retainer and missing purge disk detect flag, and also showed heavy contamination of the purge assembly. the cause of the issue was a defective component.
31377 | 1220648-2024-06423 | 18601514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18601514&pc=OZD | 01/17/2024 | 03/04/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 11/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426671 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 60-year-old male patient in non-st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that heparin was held for slight right axillary insertion site ooze and hematoma. a hematoma was evacuated at pocket of impella insertion site. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the access site bleeding ¿ major event has been completed. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical was not provided.
31378 | 1220648-2024-06424 | 18601653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18601653&pc=OZD | 01/19/2024 | 05/28/2024 |  | 01/29/2024 |  |  | 01/19/2024 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1086076 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/19/2024 |  |  |  |  | The user facility reported a 61-year-old patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella defective alarm occurred during priming / purge setup of an impella 5.5 pump. multiple attempts were made to restart the process; however the issue recurred and the decision was made to replace both the pump and the aic (automated impella controller). there was no patient harm. | [5 empty] | The investigation into the pump not detected issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the reported failure mode was unable to reproduce. the root cause of the issue was not determined due to the inability to reproduce. d9. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31379 | 1220648-2024-06426 | 18601855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18601855&pc=OZD | 09/08/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/08/2023 | 01/29/2024 |  |  | 01/31/2023 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2023173344 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2023 |  |  |  |  | The user facility reported a 87-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. this was the second 2.5 that was opened and prepped due to low purge pressure alarms on the first one. during the prep, the 13fr 2.5 sheath was exchanged for a long 14fr cp sheath to try and help with vascular issues. physician made multiple attempts to insert this new 2.5 but was unsuccessful. percutaneous coronary intervention (pci) ended with a great result and three stents placed. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, per the given clinical details, the root cause of the delivery issue was patient condition related to small/ diseased vessels.
31380 | 1220648-2024-06427 | 18602100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602100&pc=OZD | 09/13/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/13/2023 | 01/29/2024 |  |  | 07/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375022 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported an 73-year-old male implanted with an impella 5.5 device for mechanical circulatory support. the patient had just undergone an aortic root replacement along with a valve replacement. due to this reason the impella 5.5 would not advance through the innominate. the case was aborted. | [5 empty] | The investigation into the delivery issues (failed delivery) has been completed. no product was returned for evaluation. the clinical review determined the root cause of the delivery issue - anatomy was most likely due to patient condition. the patient just had an aneurysm repair and his aortic root and valve were just replaced.
31381 | 1220648-2024-06425 | 18602175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602175&pc=OZD | 02/28/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 02/28/2023 | 01/29/2024 |  |  | 10/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024296986 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump began to experience flow saturation during patient support. due to the noted issue, the pump was removed and a new impella cp was placed for ongoing support. there was no patient harm. | [5 empty] | The investigation into the saturated flow has been completed. data logs confirmed saturated flow. pump flowed at 3.6 l/min at p-7, which is outside the range of 2.9-3.3 l/min for this p-level. there were also suction and impella position in aorta alarms. upon investigation of the pump, blood was observed within the purge line near the yellow luer. the purge line had yellow/orange discoloration near the 25 cm catheter mark by the motor, which is indicative of denatured blood since the product was returned 200 days following the complaint date. when run in the lab, saturated flow was reproduced and blood was expelled upon starting the pump. no other abnormalities were found. the root cause of the saturated flow was biomaterial as observed by blood ingress in the purge line.
31382 | 1220648-2024-06428 | 18602296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602296&pc=OZD | 06/17/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 06/19/2023 | 01/29/2024 |  |  | 08/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2023260400 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 67-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp that was alarming for purge flow alarms. the low pf did not prompt the team to explant. | [5 empty] | (h3) the investigation into the reported "high purge pressure" has been completed since the original report was filed. product was returned for investigation. the device was visually inspected, and biomaterial ingestion was found to be present in the purge gap. there were no other abnormalities found. the data logs were reviewed, and motor current spike was observed just before the high purge pressure event. the cause of the high purge pressure issue was due to biomaterial ingestion in the purge gap based on field investigation and data log review. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
31383 | 1220648-2024-06430 | 18602405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602405&pc=OZD | 03/21/2023 | 02/08/2024 |  | 01/29/2024 |  |  | 03/21/2023 | 01/29/2024 |  |  | 12/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024305712 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 46-year-old black female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that an impella cp pump was pulled out of position during the course of support. the impella was subsequently kinked, and despite repositioning in the descending aorta to straighten the cannula, the device would re-kink each time it was advanced to the valve. due to the noted difficulty, the decision was made to remove and replace the pump. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned. as per the given clinical details, the root cause of the positioning issue was use related to the improper repositioning technique. h.6 investigation findings code 213 was reported on the original manufacturer device report (mdr) and code 4248 should have been reported. h.10 the additional manufacturer narrative reported that the investigation into the reported limb ischemia and vascular damage has been completed since the original report was filed. this is incorrect as the investigation was completed for the positioning issue and no previous report was filed as it was the original mdr. | (h3) the investigation into the reported limb ischemia and vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the positioning issue was use related to the improper repositioning technique. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | G.3 the date received by manufacturer ((b)(6) 2023) was reported incorrectly on the previously submitted supplemental manufacturer device report (smdr). (b)(6) 2024 should of been reported.
31384 | 1220648-2024-06429 | 18602509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602509&pc=OZD | 01/17/2024 | 03/01/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 10/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419344 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported an 87-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was prematurely explanted due to access site bleed and hematoma. the perclose, manual pressure, and femostop failed. the impella was removed without issue and perclose was deployed. manual pressure was held and femstop was placed. the patient was reported stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major event has been completed. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to patient condition since patient had tortuosity, calcification, and obesity.
31385 | 1220648-2024-06433 | 18602513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602513&pc=OZD | 03/28/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 03/28/2023 | 01/29/2024 |  |  | 10/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292670 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 43-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that an implanted impella cp pump immediately stopped functioning when it was turned on. the pump was able to restart, however, the flow rate was measured to be 3.8 l/min at a support level of p-2. the pump was subsequently replaced to as a result of the reported failure. there was no patient harm. | [5 empty] | (h3) the investigation into the reported "pump stop" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the data logs for the pump show spike in motor current followed by saturated flow and lack of pulsatility in motor current which all points towards biomaterial ingestion. based on the provided clinical information and logs, the root cause of the pump stop issue was most likely biomaterial. the root cause of the saturated flow issue was most likely biomaterial. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31386 | 1220648-2024-06431 | 18602541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602541&pc=OZD | 01/17/2024 | 08/18/2025 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1139181 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The complainant reported that an automated impella controller experienced a battery failure. the console was removed from service as a result of the noted issue. there was no known patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g.1. mdr reporting contact email. d.2. device name has been updated. f.6. and f.8. should have been left blank. updated h.3. to device evaluation anticipated. h.6. type of investigation code added.
31387 | 1220648-2024-06432 | 18602550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602550&pc=OZD | 03/22/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 03/22/2023 | 01/29/2024 |  |  | 08/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270447 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 85-year-old white male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported a kinked cannula/tracking issues. the pump was removed. there was no known harm or injury. upon investigation it was discovered the root cause of the cannula kink issue was use related to excessive force. | [5 empty] | (h3) the investigation into the reported 'product damage (cannula kink)" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the cannula kink issue was use related to the excessive force.
31388 | 1220648-2024-06434 | 18602554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602554&pc=OZD | 04/02/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 04/02/2023 | 01/29/2024 |  |  | 12/31/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309641 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 67-year-old male (unknown race and ethnicity) in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that impella came out of the left ventricle (lv) after it was implanted. the physician was not able to reinsert it in the lv. there was no harm caused to the patient as a result of this incident. | [5 empty] | (h3) the investigation into the reported "positioning issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the issue was wireless re-access. the impella device is not indicated for wireless re-access per ifu. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31389 | 1220648-2024-06435 | 18602648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602648&pc=OZD | 04/04/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 04/04/2023 | 01/29/2024 |  |  | 11/30/2024 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300633 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 45-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella device needed to be repositioned. the doctor was able to reposition the impella successfully to rotate the device away from septum and was sitting freely mid-left ventricle (lv) with elbow just below aov, with catheter at 51cm and red plug impella 5.5 side up. | [5 empty] | (h3) the investigation into the reported "positioning issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the repositioning issue was use related due to improper technique because the pump initially deep at 4.8 cm w/outflow near atrioventricular (av) and attempted to reposition to release the flow. there is no data logs available from one week prior to repositioning date in (4/4). instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31390 | 1220648-2024-06436 | 18602663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602663&pc=OZD | 01/17/2024 | 02/29/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 11/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024421974 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 41-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient had continuous elevated plasma free hemoglobin (pfhgb) since the impella was placed. the impella had to be repositioned continuously and pfhgb kept worsening over the last few days. it was decided to remove the impella. after explant, thrombus throughout the catheter¿s inlet, outlet and cannula were noted. the patient transitioned to balloon pump support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis and the thrombus issues has been completed. the device was not returned for investigation. the root cause of hemolysis and the thrombus was not determined since product was not returned and insufficient evidence from data log, clinical review. d4-model number changed to impella 5.5h5-changed to yes
31391 | 1220648-2024-06437 | 18602678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602678&pc=OZD | 09/28/2023 | 01/29/2024 |  | 01/29/2024 |  |  | 09/28/2023 | 01/29/2024 |  |  | 04/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024352856 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complaint reported that upon the impella implant, after the preclose 6fr sheath was inserted and wire was unable to pass successfully. imaging revealed loss of femoral access and the site closed with perclose. the doctor decided to not impella due access complication and left iliac disease. no known patient harm or injury. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. the root cause of the delivery issue - anatomy was most likely due to patient condition.
31392 | 1220648-2024-06438 | 18602690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602690&pc=OZD | 01/08/2024 | 12/10/2024 |  | 01/29/2024 |  |  | 01/08/2024 | 01/29/2024 |  |  | 10/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422241 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 74-year-old male patient began to experience hemolysis following impella cp implant. due to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ | The investigation for the hemolysis has been completed. product was not returned. data logs were reviewed and do not show any suction alarms or motor current (mc) spikes outside of p level changes. no other abnormalities. the cause of the hemolysis was not determined as product not returned, lack of clinical details, and inconclusive data logs. added h6 component code 925 corrections to the initial mfr report: added d6a/d6b g1 mdr reporting contact email.
31393 | 1220648-2024-06439 | 18602693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602693&pc=OZD | 01/17/2024 | 02/29/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 10/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420064 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 71-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the implant, the impella repositioning sheath continued to have significant bleeding. a femstop was placed, and it was monitored. ultimately, there was still oozing at the site and the impella was explanted. the patient was reported as stable. | [5 empty] | The investigation into the access site bleeding major issue has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the access site bleeding was most likely due to lack of vessel recoil since all the practices have followed correctly with no abnormalities and yet bleeding was observed from repositioning sheath. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31394 | 1220648-2024-06440 | 18602736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602736&pc=OZD | 01/10/2024 | 02/29/2024 |  | 01/29/2024 |  |  | 01/10/2024 | 01/29/2024 |  |  | 11/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423459 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 86-year-old female patient was on impella cp support due to left anterior descending artery, circulatory, and right coronary artery blockages. while on support, the patient had a large hematoma on the impella side and a small one on the femoral artery line. pressure was held and blood products were given. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
31395 | 1220648-2024-06441 | 18602746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602746&pc=OZD | 01/11/2024 | 05/02/2024 |  | 01/29/2024 |  |  | 01/11/2024 | 01/29/2024 |  |  |  |  | 01/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1541270 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Failure to Prime | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A 66-year-old patient received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. during preparation, the automated impella controller (aic) skipped two confirmation screens, which would have allowed the operator to start the pump without entering the purge fluid composition and without removing the guidewire. the aic used initially was swapped for the backup aic. the impella cp heart pump worked without any problems and the patient suffered no harm as a result of the incident. | [5 empty] | The investigation for the reported console (aic) purge issue has been completed. the console was returned for evaluation. upon inspection, the console performed as expected. however, data logs were reviewed which showed that the pump was plugged into a console and recognized as previously initialized. however, when the pump was initially plugged into a different console, it had said uninitialized. the root cause for the aic screen not offering the screen with the purge fluid table, nor the double confirmation of the removed 0,018" wire was due to the pump being previously initialized on the first attempted initialization. this was intended behavior of the console. added- d9 device return date, h6 impact code 4629 g1-corrected mfr site name and address. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
31396 | 1220648-2024-06442 | 18602780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602780&pc=OZD | 01/17/2024 | 02/29/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 11/30/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024428511 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Perforation of Vessels | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported an 84-year-old female patient noted as high-risk percutaneous coronary intervention of the left main artery and right coronary artery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella implant was aborted due to vascular anatomy due to heavy calcification of the iliac and aorta. additionally, a dissection was noted at the distal aorta/iliac junction and a femoral perforation was noted after the sheath was removed. the aorta/iliac dissection received a stent, and the femoral perforation received a balloon tamponade. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation into the delivery issues, the vascular damage - great vessel, and the vascular damage - peripheral vessel issues has been completed. the device was not returned for investigation. based on the clinical information provided, the patient had heavily calcified vessels which created multiple issues while implanting pump. the root cause of the delivery issue, the vascular damage - great vessel, and the vascular damage - peripheral vessel was patient condition related to heavily calcified vessels.
31397 | 1220648-2024-06443 | 18602808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18602808&pc=OZD | 01/17/2024 | 03/01/2024 |  | 01/29/2024 |  |  | 01/17/2024 | 01/29/2024 |  |  | 10/31/2025 |  | 01/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024422168 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 40-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient was off bivalirudin for a few days during impella support and developed a thrombus at the apex of the left ventricle. the patient was placed back on anticoagulants to treat the condition. support was maintained throughout the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus event has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
31398 | 1220648-2024-06323 | 18586229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18586229&pc=OZD | 01/14/2024 | 02/26/2024 |  | 01/26/2024 |  |  | 01/14/2024 | 01/26/2024 |  |  |  |  | 01/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1404002 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/14/2024 |  |  |  |  | The user facility reported a 56-year-old male of an unknown race and ethnicity with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during preparation of the purge cassette, liquid leaked out and the automated impella controller (aic) would not detect the purge cassette. a new purge cassette was prepared, and alarm purge disc not detected was displayed and after two attempts, the aic recognized and the alarm stopped. no patient harm was reported. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge retainer was broken. the root cause of the issue was a broken purge retainer.
31399 | 1220648-2024-06324 | 18586349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18586349&pc=OZD | 01/16/2024 | 03/04/2024 |  | 01/26/2024 |  |  | 01/16/2024 | 01/26/2024 |  |  | 08/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024397720 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following impella explant, the patient complained of right lower extremity pain (rle). unable to doppler dp pulse on right and popliteal pulse present by doppler. the patient was taken back to catheterization lab for rle angiogram and thrombectomy. the patient tolerated it well and had doppler pedal pulses at the completion of the procedure. according to the catheterization lab report, there was a thrombus in the right external iliac and right tibial peroneal trunk, both treated with thrombectomy and right anterior tibial and right peroneal, both treated with balloon inflations. | [5 empty] | The investigation into the reported limb ischemia- reversible and thrombus has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus and limb ischemia was not determined" | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31400 | 1220648-2024-06325 | 18586451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18586451&pc=OZD | 01/17/2024 | 02/02/2024 |  | 01/26/2024 |  |  | 01/17/2024 | 01/26/2024 |  |  | 10/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024415741 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 attempted delivery for the support of a patient but the 5.5 failed delivery despite all measures. the anatomy in the axillary precluded placement. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition since the patient had vascular disease and worsened anatomy.
31401 | 1220648-2024-06326 | 18586501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18586501&pc=OZD | 01/20/2024 | 02/28/2024 |  | 01/26/2024 |  |  | 01/20/2024 | 01/26/2024 |  |  | 11/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024428203 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient experienced oozing at their access site and later lost pulse down their impella leg following impella cp implant. perclose sutures were tightened to manage the oozing and a reperfusion sheath was placed coinciding with an external fem-fem bypass to treat the ischemia. support remained ongoing following the intervention. | [5 empty] | The investigation into the reported access site bleeding and limb ischemia has been completed since the original report was filed. the root cause of the access site bleeding - major cannot be determined as there was a lack of information provided and no product returned. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31402 | 1220648-2024-06327 | 18586509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18586509&pc=OZD | 01/17/2024 | 03/01/2024 |  | 01/26/2024 |  |  | 01/17/2024 | 01/26/2024 |  |  | 11/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024428310 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed some oozing at their impella access site. the stat lock was removed. the angle was matched, and forward tension sutures were applied to manage the condition. support was maintained throughout the event. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved after removing stat lock, angled matching and applying forward tension sutures.
31403 | 1220648-2024-06328 | 18586536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18586536&pc=OZD | 01/20/2024 | 03/07/2024 |  | 01/26/2024 |  |  | 01/20/2024 | 01/26/2024 |  |  | 10/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024424275 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2024 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that echo - echocardiogram (ultrasound) was performed and the impella was noted to be under the pap muscle. the right ventricle assessment was performed and there was normal right ventricle function. the patient was brought to cath lab that same day for impella removal because patient had recovered from support. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the device was not returned for investigation. the cause of the positioning issue cannot be determined since compliant device not returned for investigation. data logs not available to review. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31404 | 1220648-2024-06329 | 18586577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18586577&pc=OZD | 01/22/2024 | 09/05/2024 |  | 01/26/2024 |  |  | 01/22/2024 | 01/26/2024 |  |  |  |  | 01/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1126026 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Pumping Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella rp support due to acute severe right ventricle failure. the impella rp (sn (b)(6)) was inserted and ramped up and the impella motor stopped. staff attempted to restart again and the motor stopped again. the impella was repositioned and restarted again and it was slowly ramped up to p7 then there was a high purge pressure alarm, then alarmed purge blocked. the impella was removed and replaced with an impella rp (sn (b)(6)). when the second impella was ramped up, there was an immediate motor stop. staff replaced the aic successfully and the impella was operational. | [5 empty] | Upon further review the reported issue it was found to be an issue with the pump only, not the automated impella controller (aic). please see manufacturer device report 1220648-2024-06887 for further information regarding the pump. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
31405 | 1220648-2024-06330 | 18587784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18587784&pc=OZD | 11/03/2022 | 02/02/2024 |  | 01/26/2024 |  |  | 11/03/2022 | 01/26/2024 |  |  | 03/31/2024 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support during percutaneous coronary intervention. during the implantation process of the impella cp, the patient became hypertensive as the impella gained access; the patient was weaned to one low dose pressor ¿ levophed. right before implant of the impella, the decision was made to abort the procedure. it was felt that there was an acute on chronic lesion, and ejection fraction (ef) function that was displayed on the echo. pump implantation aborted and percutaneous coronary intervention done without the impella cp. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to pass through the vasculature due to patient anatomy as per the clinical description provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31406 | 1220648-2024-06332 | 18587959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18587959&pc=OZD | 10/10/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 10/10/2023 | 01/26/2024 |  |  | 04/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333540 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 74-year-old female post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. the patient had just undergone multiple cardiac surgical procedures. the impella 5.5 was unable to pass the aortic valve. the tip was getting caught on previous aortic graft repair anastomosis. the impella cp was placed instead. | [5 empty] | The investigation into the delivery issues-anatomy has been completed. the impella products were not returned for evaluation. the root cause of the delivery anatomy issue was most likely patient condition related.
31407 | 1220648-2024-06331 | 18587965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18587965&pc=OZD | 10/06/2023 | 06/20/2024 |  | 01/26/2024 |  |  | 10/06/2023 | 01/26/2024 |  |  | 12/31/2024 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291925 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The patient expired after care was withdrawn. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of device with outcome. | The user facility reported a 63-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported the patient developed an access site bleed during impella 5.5 support. the patient was administered three units packed red blood cells and heparin treatment was halted to manage the condition. a jackson-pratt drain was placed at the axillary site and support with the implanted pump was maintained. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be anticoagulation management because the hemostasis was achieved by temporarily shutting off heparin. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added h6 component code in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31408 | 1220648-2024-06333 | 18588034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588034&pc=OZD | 04/10/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 04/10/2023 | 01/26/2024 |  |  |  |  | 01/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1135089 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/10/2023 |  |  |  |  | The user facility reported a 45-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the automated impella controller (aic) alarmed for purge disk not detected. the aic was swapped out with another aic with no further issues. there was no patient harm. | [5 empty] | The investigation into the automated impella controller (aic) purge disc/flag issues has been completed. data review: after reviewing the imc log of this aic, it was found that "purge disc not detected" alarm (event set #402) was triggered several times. device analysis: the console was tested after arriving in field service. it was evident that the purge retainer was broken (disk detect flag was intact).
31409 | 1220648-2024-06334 | 18588080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588080&pc=OZD | 04/19/2022 | 06/24/2024 |  | 01/26/2024 |  |  | 04/19/2022 | 01/26/2024 |  |  | 12/31/2023 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180709 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was on impella cp support due to having chest pain and the patient coded when they arrived to the emergency room. it was noted that a light fluid leak was noted at the purge filter during impella support. the local team was aware of the leak at the yellow side arm connector. the device was explanted later that day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the purge leak - pump issue has been completed since the original report was submitted. the product was not returned for investigation. the cause of the pump purge leak was a damaged sidearm, but the exact location could not be determined because no product was returned. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
31410 | 1220648-2024-06336 | 18588115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588115&pc=OZD | 10/10/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 10/10/2023 | 01/26/2024 |  |  | 02/28/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024325356 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had an inability to deliver the impella cp to the left ventricle due to native anatomy. there was peripheral arterial disease and calcification. the pump implant was aborted. no pump damage and no lasting harm or injury. | [5 empty] | No product was returned. the root cause of the delivery anatomy issue was most likely patient condition related.
31411 | 1220648-2024-06335 | 18588146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588146&pc=OZD | 05/04/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 05/04/2023 | 01/26/2024 |  |  |  |  | 01/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1123075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | Power Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 05/04/2023 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support conclude. the team then found aic (automated impella controller) issues when disconnecting and powering off. the aic will be investigated, and a backup used for support per ifu recommendations. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that after pump 404456 was manually stopped (the pump was turned off, not the aic as mistakenly reported) and unplugged from the console (as evidenced by optical signals in rtlog data), the console did not detect the pump being disconnected, which prevented the user from shutting down the aic. as means of troubleshooting, the pump was repeatedly reconnected and disconnected few times, but the issue was not resolved. console was eventually powered down by unplugging ac cord and disconnecting batteries via transport switch. also, throughout the case, there were multiple (1700) error messages related to eep/epm communication recorded in imc log: 01/05/23 19:07:55 sau_sens_eep_1 too busy for 2680. 04/05/23 11:05:40 sau_sens_eep_1 too busy for 63467. console ic2081 was tested in danvers by plugging and unplugging a known-good pump and pump simulator, but the issue was not reproduced, and disconnecting of the pump was each time correctly registered. inspection of the console¿s electronic components revealed kinked impellatronic ribbon cable, which might have affected communication between eep circuitry and epc. inspection of the impellatronic board indicated that fcn 98 had been performed, however the workmanship seemed unsatisfactory as evidenced by contamination and flux residue. it is likely that during rework component r330.2 was overheated, at the risk of shorting to the neighbouring component r326.2, which is part of pump motor sensing sub-circuitry, that also includes l6234 (sense1, sense2 - feedback for motor current control for the bridges) and tps767d301 (u101.2). we were unable to establish any relationship between these components and pump detection mechanism. from the three mechanisms used for detecting pump disconnection (optical snr<150, current-sensing resistor, i2c active communication) the snr is related to optical bench, and if any of the other two (located on impellatronic) fails, pump disconnection will not be detected. the issue of pump disconnection not being detected was most likely caused by compromised ribbon cable between impellatronic and 4-in-1 assembly preventing pump detection from being communicated to epc. during testing it¿s been also observed that microsd card on rlm was corrupt, which resulted in loss of wifi configuration every other boot-up (unrelated to this complaint). repair: replace impellatronic communication cable (6000-0141) and impellatronic assembly (0042-1115) as a precaution. replace rlm microsd card, and perform functional check of the console.
31412 | 1220648-2024-06337 | 18588244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588244&pc=OZD | 01/11/2024 | 06/06/2024 |  | 01/26/2024 |  |  | 01/11/2024 | 01/26/2024 |  |  | 10/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418704 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Abdominal Pain | Cardiac Arrest | Hemolysis | Tachycardia | Obstruction/Occlusion | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 01/11/2024 |  |  |  |  | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening ventricular tachycardia with a "possible" relationship to the impella device used: ¿patient had pmh of vt. patient had vt overnight on (b)(6) 2024. limited echo imaging repositioned impella. patient continued on amiodarone drip. it was reported that the patient/event has not recovered/not resolved. | The user facility reported a 70-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, there was unresolved/undetermined cause of suspected hemolysis. the patient had an extremely high plasma free hemoglobin (3800) result. the impella's position was verified. the physician decided to explant the pump and replace it with another 5.5. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have also endured an ischemic bowel with a "probable" relationship to the impella device. this was discovered via a ct scan after the patient had complained of abdominal pain and not having a bowel movement in "2-3" days. the surgical team was consulted and the small bowel resected and a laparotomy was performed. while completing an anastomosis for the bowel resection, the patient endured cardiac arrest. cpr was performed, pressors given until adequate oxygenation and circulation was restored. the device was exchanged "shortly after." lastly, a thrombus was found lodged within the surgical graft being used to place the impella device, during removal. this allegation was made with a "definite" relationship. the thrombus was removed. |  |  |  | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product and data were returned for review. as the necessary information was not provided, the root cause of the thrombus and the vt/vf (afib ablation) was not determined. the root cause of the vascular damage was unable to be determined as limited clinical details were provided and no product & data were returned for review. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | Section b5 and h6 were updated as per additional information received. | The investigation into the reported hemolysis has been completed since the original report was filed. product was not returned for review. the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product & data were returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Additional information was received via abiomed's abiomed¿s long-term outcome and quality indicator impella registry and added to section b(5), with additional coding in h(6). upon investigation closure, a supplemental mdr will be filed.
31413 | 1220648-2024-06340 | 18588282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588282&pc=OZD | 01/01/2024 | 02/07/2024 |  | 01/26/2024 |  |  | 01/12/2024 | 01/26/2024 |  |  |  |  | 01/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1402264 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a broken purge disc clip. this was found during set-up and there was no patient involvement. the aic was replaced. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer.
31414 | 1220648-2024-06339 | 18588284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588284&pc=OZD | 01/10/2024 | 01/26/2024 |  | 01/26/2024 |  |  | 01/10/2024 | 01/26/2024 |  |  | 09/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407088 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the implant, the patient became hypertensive with a systolic pressure of 160. the patient developed a hematoma. manual pressure was held and this did not resolve the issue. a balloon was advanced up and over and the artery was occluded with balloon inflation. the impella was weaned and removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical report was reviewed alone. root cause could not be determined without more clinical details or product. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31415 | 1220648-2024-06338 | 18588291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588291&pc=OZD | 10/24/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 10/24/2023 | 01/26/2024 |  |  | 08/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024392516 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2023 |  |  |  |  | The user facility reported a 41-year-old female post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support as escalation of care from veno-arterial extracorporeal membrane oxygenation (va ecmo) and intra-aortic balloon pump. the pump would not advance through innominate artery into the ascending aorta despite multiple attempts. the procedure was aborted. no known patient harm or injury. | [5 empty] | The investigation into the delivery issues-anatomy has been completed. no product was returned for evaluation. as per clinical review, the impella 5.5 pump would not advance through the innominate artery into the ascending aorta with the patient having stenosis and the case aborted due to vascular complications. the cause of the delivery issue was patient condition related with the patient having stenosis at ascending aorta and the pump was unable to advance due to vascular complications.
31416 | 1220648-2024-06341 | 18588435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588435&pc=OZD | 10/24/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 10/24/2023 | 01/26/2024 |  |  | 08/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024385523 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2023 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the medical doctor attempted to advance the impella 5.5 but was unable to due to vessel tortuosity. the medical doctor stated vessel size was adequate. the device was attempted to be lubricated with viper slide which was unsuccessful. the team also attempted to insert the device over 0.025 wire without success. the decision was made to abort the impella 5.5 and insert axillary impella cp instead. | [5 empty] | No product was returned. the root cause of the delivery anatomy issue was most likely patient condition related.
31417 | 1220648-2024-06342 | 18588622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588622&pc=OZD | 06/01/2022 | 01/26/2024 |  | 01/26/2024 |  |  | 12/21/2023 | 01/26/2024 |  |  | 10/31/2023 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154330 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/01/2023 |  |  |  |  | The user facility reported a (b)(6) female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the medical team had difficulty delivering an impella pump due to the patient's anatomy. two days after implant, the patient was found to have a hematoma at the access site which was managed by the medical team. | [5 empty] | The investigation into the patient's delivery issue and access site bleeding (hematoma) has been completed. no product was returned for evaluation. based on the clinical review, the root cause of the access site bleeding was not determined. there was no clinical description that relates the cause with impella and no other impella related issues as per the case updates and the bleeding was resolved. hence, the root cause is not determined and not enough information was submitted nor the product was returned to determine the root cause. the root cause of the delivery issue - anatomy was most likely due to the patient condition. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31418 | 1220648-2024-06343 | 18588649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588649&pc=OZD | 09/29/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 09/29/2023 | 01/26/2024 |  |  | 07/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024372370 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/29/2023 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella 5.5 was attempted to be placed but due to the anatomy and clot formation the doctors decided to abort the impella and go with extracorporeal membrane oxygenation (ecmo). upon impella insertion, the doctor noticed a large amount of clots. clots were removed and angiogram was performed and revealed a tight lesion just distal to the anastomosis and the artery appeared to be dissected. no known impella harm to patient. | [5 empty] | No product was returned. as per clinical, the impella 5.5 insertion was attempted and aborted due to tight lesion and artery being dissected. the dissection was not related to the impella and unknown what caused the dissection, and this concern was also a cause for the impella 5.5 pump not being inserted past the lesion. the cause of the delivery issue was patient condition related since insertion aborted due to complication per clinical review.
31419 | 1220648-2024-06344 | 18588803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588803&pc=OZD | 06/10/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 06/13/2023 | 01/26/2024 |  |  | 02/28/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024320633 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Deliver | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/13/2023 |  |  |  |  | The user facility reported an 87-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the md had difficulty inserting impella. several attempts were made but they were unsuccessful and the insertion was stopped due to a kink in the shaft. an iabp was inserted and the patient was discharged to the icu without pci. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation into the delivery issues has been completed. product was not returned for evaluation. per the clinical review, the cause of the delivery issue- use was patient condition related because the patient had acute myocardial infarction/cardiogenic shock and tri-vessel lesion. several attempts were made to insert impella, but delivery was difficult and the shaft kinked.
31420 | 1220648-2024-06345 | 18588863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18588863&pc=OZD | 06/22/2023 | 07/29/2024 |  | 01/26/2024 |  |  | 06/22/2023 | 01/26/2024 |  |  | 12/31/2024 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315967 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported a red alarm that wouldn¿t stop and that the team removed the first pump and exchanged to ensure patient was adequately supported. new cp was inserted without complication and patient stabilized on new pump. the impella connect was examined and, there was first a purge disc alarm, then optical sensor alarm, then a low purge pressure alarm. to note, the same cassette was used with the new cp and functioning without issue. no known patient harm or injury. | [5 empty] | The investigation for the pump did not start issue has been completed. according to the clinical, immediately after inserting the impella cp a recurring red alarm was recorded. first a purge disc alarm was recorded, then an optical sensor alarm, and finally a low purge pressure alarm. the decision was then made to remove the pump and replace it with another impella cp device. the new pump was inserted without complication and the patient was stabilized. no product was returned for a visual inspection. data log analysis confirmed the presence of the above alarms. the cause of the pump did not start issue was not determined because no product was returned and not enough clinical information was given. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi in section: cleaning. added-h6 impact code 4629. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
31421 | 1220648-2024-06347 | 18589311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18589311&pc=OZD | 07/06/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 07/06/2023 | 01/26/2024 |  |  |  |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1450701 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) failed at the purge disc. the aic was replaced per the instructions for use recommendation. there was no patient harm or injury. | [5 empty] | The investigation into the purge disk issues has been completed. logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the console was tested after arriving in field service. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause for the purge disc not detected alarm was a missing purge flag. the cause of the issue was a defective component.
31422 | 1220648-2024-06348 | 18589733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18589733&pc=OZD | 06/21/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 07/07/2023 | 01/26/2024 |  |  | 02/28/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024321599 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  |  | White | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 84-year-old female of unknown ethnicity with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the introducer was leaking at the valve. the sheath was removed and replaced to resolve the noted issue. there was no patient harm reported. | [5 empty] | The root cause of the introducer issue could not be determined due to no product returned and a lack of clinical details.
31423 | 1220648-2024-06349 | 18589735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18589735&pc=OZD | 01/13/2024 | 02/15/2024 |  | 01/26/2024 |  |  | 01/13/2024 | 01/26/2024 |  |  | 09/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394401 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2024 |  |  |  |  | The user facility reported a 42-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported the impella 5.5 pump began to experience high purge pressure and low purge flows during patient support. the purge cassette, purge bag, and purge solution were swapped to troubleshoot the issue, but the pressure remained high. after 5 hours, the decision was ultimately made to exchange the pump. there was no patient harm reported. | [5 empty] | The investigation into high purge pressure issue completed since the original report was filed. the device was returned for investigation. according to the clinical, high purge pressure and low purge flow alarms were recorded. product evaluation confirmed that there was biomaterial in the purge gap with no visible purge line kinks. the cause of the high purge pressure is not determined because not enough information was given. | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31424 | 1220648-2024-06350 | 18589876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18589876&pc=OZD | 01/13/2024 | 03/01/2024 |  | 01/26/2024 |  |  | 01/13/2024 | 01/26/2024 |  |  | 08/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024387755 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2024 |  |  |  |  | The user facility reported a 55-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported the impella 5.5 pump began to display a pump in aorta alarm, the physician attempted to reposition unsuccessfully. it was further reported that the patient was taken to the operation room (or) for a hematoma evacuation. while in or a transesophageal echocardiography (tee) showed that catheter was in the aortic valve. the physician made multiple attempts to advance the catheter across the aortic valve without success. a decision was made to explant and not to replace the catheter due to patient being stable without support over the past several hours. multiple blood clots were noted along the catheter when removing the catheter despite the patient being anticoagulated. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported positioning issue thrombus and access site bleeding has been completed since the original report was filed. complaint device not returned for investigation and no data logs to review. the cause of the access site bleeding cannot be determined since the complaint device not returned for investigation and no sufficient clinical data provided. the cause of the positioning issue was wireless re-access as the physician made multiple attempts to advance catheter across av without success. the impella device is not indicated for wireless re-access. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
31425 | 1220648-2024-06351 | 18589945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18589945&pc=OZD | 09/01/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 09/01/2023 | 01/26/2024 |  |  | 04/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024330845 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2023 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the staff was experiencing a great deal of resistance in manipulating the catheter. the catheter was straight, no acute angles. after multiple attempts to slide repo unit, sterile water was applied to catheter shaft. despite this, the doctor still experienced a great deal of resistance in being able to smoothly glide the repo unit. the doctor eventually pulled the graft distally and was able to secure the graft to the blue suture hub. once it was secure, the doctor was able to advance the catheter without issue. no known patient harm or injury. | [5 empty] | No product was returned. as per clinical, the impella 5.5 pump was delivered successfully into the left ventricle and the physician had great resistance while advancing catheter using repositioning unit blue button. after multiple attempts to slide repo unit was unsuccessful and later secured the blue suture hub with yellow pin in place and the physician was able to advance the catheter without any issues. the cause of the positioning issue was not determined since the product was not returned and based on insufficient clinical information.
31426 | 1220648-2024-06352 | 18589963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18589963&pc=OZD | 01/13/2024 | 02/28/2024 |  | 01/26/2024 |  |  | 01/13/2024 | 01/26/2024 |  |  | 09/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024394438 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Unable to Obtain Readings | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to having chest pain and shortness of breath. while on support, the placement signal went out. the staff replaced the automated impella controller and the placement signal and retrograde flow alarms persisted. the impella was removed due to the fear of the pump stopping. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported optical signal issue and low pump flow has been completed since the original report was filed. the product was returned for investigation, and no physical abnormalities were observed. all 5s optical parameters showed a hard failure trend. while the pump passed electrical testing, it failed the optical light testing. a light leak was observed inside the red handle during optical light testing, and the optical line was found to be kinked at the post proximal to the patient cable inside the red handle. data logs indicated that the placement signal spiked to 3000 during the case, and the optical sensor 5s parameters were characteristic of an optical fiber break. as per the clinical details, the doctor reported a 'placement signal not reliable' alarm. the cause of the issue was a damaged optical fiber line during the case. the doctor decided to switch the console, but the controller failure and retrograde flow alarms appeared after plugging in. the cause of the optical signal issue was a kinked optical fiber line resulting from tubing lines within the patient cable sliding out and into the red handle due to an improper manufacturing sequence of tubing line assembly into the patient cable. no component related defect.
31427 | 1220648-2024-06353 | 18590005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18590005&pc=OZD | 12/05/2022 | 01/26/2024 |  | 01/26/2024 |  |  | 12/05/2022 | 01/26/2024 |  |  | 04/30/2024 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217723 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported an impella cp in a 58yo male patient for mechanical circulatory support. it was reported that the introducer sheath could not be inserted due to bilateral iliac occlusions. resistance was met whenever the physician attempted to advance it. the procedure was aborted as a result. there was no patient harm reported. no additional information was provided. | [5 empty] | The investigation for the reported event has not yet been completed. a supplemental reported will be submitted once the investigation has been completed.
31428 | 1220648-2024-06354 | 18590006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18590006&pc=OZD | 01/13/2024 | 02/28/2024 |  | 01/26/2024 |  |  | 01/14/2024 | 01/26/2024 |  |  | 06/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024383615 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2024 |  |  |  |  | The user facility reported a 59-year-old male patient of unknown race noted for st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had a new onset of acute kidney injury (aki) after impella cp placement. the patient had hematuria and oliguria requiring continuous renal replacement therapy (crrt). the impella cp was removed and was replaced with impella 5.5 | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. product was not returned for review. the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product & data were returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31429 | 1220648-2024-06355 | 18590018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18590018&pc=OZD | 01/14/2024 | 02/12/2024 |  | 01/26/2024 |  |  | 01/14/2024 | 01/26/2024 |  |  | 10/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420137 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had attempted to deliver the impella cp to support a patient admitted with st elevation myocardial infarction (stemi). the pump failed to deliver and kinked in the attempts. the patient had known peripheral arterial disease which may have contributed to the issues with iliofemoral delivery. the pump was explanted and replaced by an intra-aortic balloon pump (iabp). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine a root cause. the root cause of the delivery issue due to a patient related condition as it was reported the device buckled due to an aortic occlusion. added impact code: 4642 d4: model number-impella cp.
31430 | 1220648-2024-06356 | 18590090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18590090&pc=OZD | 01/14/2024 | 03/01/2024 |  | 01/26/2024 |  |  | 01/14/2024 | 01/26/2024 |  |  | 11/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024423279 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2024 |  |  |  |  | The user facility reported a 67-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during the impella explant there was a possible small iliac artery dissection that occurred. as intervention, the staff monitored the patient¿s vitals, access site and pulses. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vascular damage - peripheral vessel has been completed. the device was not returned for investigation. clinical details do not specify if excessive force was used during explant, or any patient peripheral vascular conditions. it was noted that vitals, explant site, and pulse were ok. the root cause of the vascular damage could not be determined due to no product returned, limited clinical information on explant procedure, and patient conditions.
31431 | 1220648-2024-06357 | 18590257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18590257&pc=OZD | 09/02/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 09/04/2023 | 01/26/2024 |  |  | 05/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024350565 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/04/2023 |  |  |  |  | The user facility reported an 82-year-old male patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that an attempt was made to insert the impella, but stenosis and meandering were found in the left common iliac artery, and the impella was caught at the base of the pigtail and did not proceed. although the blood vessel was dilated with a balloon, impella did not pass. a stent was also placed and tried again, but it did not pass, and the insertion was abandoned. no patient harm was reported. | [5 empty] | No product was returned for investigation. as per clinical the reason the impella could not be inserted due to the stenosis and tortuousness. physician did everything possible with a balloon and a stent, still impella did not pass, and the insertion was abandoned. the root cause of the delivery issue was most likely patient condition related due to the stenosis and tortuousness. the investigation has been completed.
31432 | 1220648-2024-06358 | 18590428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18590428&pc=OZD | 05/18/2023 | 01/26/2024 |  | 01/26/2024 |  |  | 05/18/2023 | 01/26/2024 |  |  | 11/30/2024 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299958 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2023 |  |  |  |  | The user facility reported a 51-year-old male undergoing coronary artery bypass grafting was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp had optical signal alarms and unreliable flows. the pump was explanted due to clotting concerns but, at the time of explant no thrombosis was noted. the cp was replaced, and support proceeded. no patient harm was reported. this compliant is for pump serial number (b)(6) related to case 01130969 for product malfunction. there is a related case 01130972 that refers to pump serial number (b)(6), which is related to the hematoma. | [5 empty] | Returned complaint cp pump visually inspected. dextrose built up around impella and purge gap. dried blood observed in inlet cage. pump connected to automated impella controller (aic). 5s parameter are within specific range. complaint pump run in water at p-9 after dissolving the dextrose. no low flow reproduced. root cause of the optical signal issue was patient condition related because the optical system was functioning within specification, but the placement signal was low due to poor patient condition. the flow improved when the pump was running on low p-level. the data logs were reviewed and a ventricular placement signal waveform with impella position in the ventricle alarms were observed. the improper positioning resulted in low pump flow and low placement signal.
31433 | 1220648-2024-06359 | 18590570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18590570&pc=OZD | 01/16/2024 | 09/20/2024 |  | 01/26/2024 |  |  | 01/16/2024 | 01/26/2024 |  |  | 05/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024358656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/16/2024 |  |  |  |  | Additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced bleeding and developed a hematoma at their access site following impella cp implant. femostop was placed in an attempt to manage the condition and four units of packed red blood cells were transfused. vascular surgery was consulted to remove the device. however, surgery declined due to the complexity of the case. the patient remained on support for comfort care until they succumbed to their existing condition. |  |  |  |  | Additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. corrected information for the initial mfr 1220648-2024-06359 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31434 | 1220648-2024-06360 | 18590738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18590738&pc=OZD | 01/14/2024 | 03/01/2024 |  | 01/26/2024 |  |  | 01/14/2024 | 01/26/2024 |  |  | 10/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421258 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2024 |  |  |  |  | The user facility reported a 78-year-old male noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump would experience persistent suction issues following implant. to correct the issue repositioning and volume were given, but the suction alarm remained. the pump was then pull back while keeping the pigtail across the valve; however, attempts to readvance the pump into an ideal position were unsuccessful and the decision was made to explant the device. there was no patient harm reported. | [5 empty] | The investigation into the reported positioning issue and low pump flow has been completed since the original report was filed. data logs confirmed the low flow when pump started & remained to the rest of the case that triggered auto suction response & suction alarms. logs also showed about 1 hour into support, pump position was in ventricle then in aortic area corresponded with the clinical narrative that triggered alarms impella position in ventricle & impella position in aorta. pump then was stopped manually & disconnected from the console. the root cause of positioning issue was most likely patient condition related, specifically the patient's anatomy and low sitting aortic valve. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31435 | 1220648-2024-06361 | 18591054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18591054&pc=OZD | 01/16/2024 | 03/05/2024 |  | 01/26/2024 |  |  | 01/16/2024 | 01/26/2024 |  |  |  |  | 01/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1136927 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support and the automated impella controller (aic) had a battery failure (red alarm). the aic was replaced successfully and the affected aic was sent back for investigation. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) "red alarm" has been completed. the aic was returned for evaluation. during evaluation, the console received a persistent battery failure alarm during boot-up. visual inspection found that the battery one lcd screen was blank. the data logs were reviewed confirmed the battery failure occurrence and a cell imbalance with cell 2 of battery 1. the root cause for the battery failure was the cell imbalance with cell 2 of battery 1.
31436 | 1220648-2024-06362 | 18591293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18591293&pc=OZD | 01/15/2024 | 02/07/2024 |  | 01/26/2024 |  |  | 01/15/2024 | 01/26/2024 |  |  | 11/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432272 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/15/2024 |  |  |  |  | The user facility reported a female patient of an unknown age in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump displayed high purge pressure/purge system blocked alarm. attempted to correct the issue by using tissue plasminogen activator (tpa). the pump later stopped and was unable to be restarted. the pump was explanted and there was not patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency. | The investigation has been completed since the original report was filed. the medical center did not return the impella device. data logs showed a gradual 5-hour rise in purge pressure and then a return to normal trends for almost 15 hours after administering tissue plasminogen activator (tpa). again, at the end of the case, a purge pressure rise was reported with several pump stops. as per the clinical details, tpa was given to improve purge flow after a drop, but flows decreased again, leading to a switch to bicarbonate. despite restarting the impella pump, it eventually. efforts to restart it failed, resulting in the decision to explant the pump. based on the returned log trend, the cause of the pump stop issue was most likely biomaterial, which was led to a high purge pressure event.
31437 | 1220648-2024-06363 | 18591394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18591394&pc=OZD | 01/15/2024 | 09/20/2024 |  | 01/26/2024 |  |  | 01/15/2024 | 01/26/2024 |  |  | 10/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024416067 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/15/2024 |  |  |  |  | The user facility reported a 63-year-old male of an unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that purge flow decreased to less than 3 milliliters an hour and increase pressure greater than 1000 millimeters of mercury. the purge system was exchanged without success and the pump remained implanted. no patient harm was reported. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visually inspected and biomaterial ingestion at the purge gap was observed. the cause of the high purge pressure issue was most likely due to biomaterial per field investigation. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06363 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
31438 | 1220648-2024-06364 | 18591443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18591443&pc=OZD | 01/15/2024 | 02/29/2024 |  | 01/26/2024 |  |  | 01/15/2024 | 01/26/2024 |  |  | 08/31/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024392350 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/15/2024 |  |  |  |  | The user facility reported a 55-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having non-st elevated myocardial infarction. upon removal of the abiomed 14fr sheath, there was significant bleeding. percutaneous transluminal angioplasty was performed to the left iliac and a covered stent was placed in the left iliac. one unit of packed red blood cells were provided. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed. the device was not returned for investigation. per the provided clinical details. patient had known history of vessel disease, other comorbidities, and significant calcification. the most likely cause of the peripheral vessel vascular damage was due to the patient's condition. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
31439 | 1220648-2024-06365 | 18592394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18592394&pc=OZD | 05/22/2023 | 12/05/2024 |  | 01/26/2024 |  |  | 12/29/2023 | 01/26/2024 |  |  | 04/30/2024 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023212672 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Tachycardia | Renal Impairment | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 12/29/2023 |  |  |  |  | Additional information received via loqi database indicating the patient to have endured ventricular tachycardia, prompting intervention via an amino-drip: "patient had v tach after impella placement and started on amio drip. transitioned to p.o. amio. heart tx on (b)(6) 2023." | A 40 year old male patient presented with cardiomyopathy for hemodynamic support via the impella 5.5 device. a notification was received via abiomed's long-term outcome and quality indicator impella registry indicating a "possible" relationship to both hemolysis and impella access site hematoma. the patient received unspecified intervention for the alleged hemolysis and an evacuation was performed in response to the access site hematoma. | Additional information received via loqi database indicating the patient to have endured acute kidney injury during support: "patient had shown aki on (b)(6) 2023 labs. baseline cr is 1.3. patient had heart transplant on (b)(6) 2023. aki continued and resolved upon discharge (b)(6) 2024." |  |  |  | New information added to b(5), per notification from loqi database. upon investigation closure, a supplemental mdr will be filed. | The impella device was not received by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into hemolysis, access site bleeding, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. d2 common device name d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact facility name g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. | New information added to b(5), per notification from loqi database. upon investigation closure, a supplemental mdr will be filed. b(3): correct event date has been corrected to (b)(6) 2023.
31440 | 1220648-2024-06113 | 18592416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18592416&pc=OZD | 01/05/2024 | 03/22/2024 |  | 01/26/2024 |  |  | 01/05/2024 | 01/26/2024 |  |  | 09/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403797 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 75-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the access site was oozing and so the team sutured and adjusted the dressing. when lab values ere revealed, hemoglobin and hematocrit drop, so the team also gave one unit of blood. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related, resolved with sheath elevation and a suture. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31441 | 1220648-2024-06279 | 18592417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18592417&pc=OZD | 01/11/2024 | 09/03/2024 |  | 01/26/2024 |  |  | 01/11/2024 | 01/26/2024 |  |  |  |  | 01/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1764189 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Calibrate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and the automated impella controller (aic) had a controller failure (red alarm). the aic was replaced successfully. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Upon further review, the engineer determined the automated impella controller (aic) was not the issue and the purge cassette was the cause. a regulatory report is no longer necessary.
31442 | 1220648-2024-06238 | 18592433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18592433&pc=OZD | 01/10/2024 | 02/02/2024 |  | 01/26/2024 |  |  | 01/10/2024 | 01/26/2024 |  |  | 11/30/2025 |  | 01/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024425029 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 34-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to advance beyond the left subclavian during attempted delivery. after multiple unsuccessful attempts, the implant was aborted and an impella cp was placed for patient support. it was noted that the cannula of the 5.5 had kinked during the attempted delivery. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition since calcification of vessels was observed. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31443 | 1220648-2024-06271 | 18578880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18578880&pc=OZD | 06/20/2023 | 10/14/2024 |  | 01/25/2024 |  |  | 01/25/2024 | 01/25/2024 |  |  | 02/28/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345218 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | The complainant in india had a 58-year old male patient on impella cp support ongoing for 5 days when the pump stopped. the impella cp was due to be explanted later that day, but was instead just explanted and not replaced by another impella cp but rather by an intra-aortic balloon pump (iabp). there was no reported patient harm. | The patient experienced severe bleeding from the impella access site, which was managed through the application of proximal pressure, a reduction in heparin concentration in the purge, 3 units of pack red blood cells, and an adjustment in the angle of the catheter. |  |  |  |  | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the impella stop cannot be determined as the complaint device and the data logs were not returned for investigation. | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the pump stop cannot be determined as the complaint device and the data logs were not returned for investigation. | Instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ h.10 instructions for use were inadvertently omitted from the manufacturer device report. | The investigation into the reported pump stop has been completed. the medical center did not return any impella products. no data logs were returned for review and there was no sufficient clinical data. the cause of the issue could not be determined since the complaint device was not returned. | Additional information for the initial mfr 1220648-2024-06271 was provided in sections b1, b2, b5, b7, d6a, d6b, h1, h6. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
31444 | 1220648-2024-06272 | 18579012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18579012&pc=PYX | 06/29/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 06/29/2023 | 01/25/2024 |  |  | 01/31/2024 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2023175921 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/29/2023 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the staff noticed upon preparation that the sheath tip was damaged and a new tip was used. no known harm or injury. | [5 empty] | No introducer was returned for evaluation. as per the clinical doctor, the 23 fr sheath was damaged and the sheath tip was split. the cause of the introducer damage issue was a sheath tip split most likely due to use related based on clinical information provided.
31445 | 1220648-2024-06273 | 18579164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18579164&pc=OZD | 06/30/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 07/01/2023 | 01/25/2024 |  |  | 04/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349505 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 74-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant had impella cp support that was ongoing with signs of hemolysis. the team adjusted the pump position and support proceeded until care was withdrawn and the patient expired. | [5 empty] | The investigation into the reported hemolysis has been completed. the medical center did not return any impella products. the evaluation revealed that the data confirmed the pump was not in proper position corresponding with the day hemolysis occurred. after two hours, the pump was back in good position and remained for the rest of the case. based on clinical description and data analysis, the root cause of hemolysis was most likely due to the pump was not in the proper position. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31446 | 1220648-2024-06274 | 18579294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18579294&pc=OZD | 08/21/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 08/21/2023 | 01/25/2024 |  |  | 05/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024364930 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported an 80-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support as a bailout effort. first with the shorter peel away sheath, the impella would not pass the diseased area of the descending aorta. longer peel away was placed in attempt to cross. placement of the impella was unsuccessful. the decision was made to terminate efforts and abort the impella insertion. | [5 empty] | The investigation into the delivery issues has been completed. the impella product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition.
31447 | 1220648-2024-06275 | 18579382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18579382&pc=OZD | 01/11/2024 | 02/19/2024 |  | 01/25/2024 |  |  | 01/11/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024418742 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump being delivered for care escalation beyond the prior placed cp. the patient was on support with multiple cardiac comorbidities and brought to the or for the 5.5 placement. when the impella was placed the patient had multiple bleed sites and concerns for thrombocytopenia. the patient was treated with multiple units of prbcs. | [5 empty] | The investigation into the access site bleeding ¿ major and the thrombocytopenia issues has been completed since the original report was filed. the product was returned for investigation. the device was visually inspected the pump and no physical damages or abnormalities were found. the cause of access site bleeding was most likely anticoagulation management related per clinical review. the cause of the thrombocytopenia issue was not determined due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31448 | 1220648-2024-06276 | 18579416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18579416&pc=PYX | 09/12/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 09/12/2023 | 01/25/2024 |  |  | 07/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024374135 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/12/2023 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the second impella rp for the same case was unable to be inserted. the impella was unable to advance beyond the right ventricular outflow tract (rvot) and through the pulmonic valve. the case was aborted. | [5 empty] | The impella device was not received from the customer. the root cause of the delivery issue - anatomy was most likely due to patient condition.
31449 | 1220648-2024-06277 | 18579534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18579534&pc=OZD | 09/26/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 09/26/2023 | 01/25/2024 |  |  | 06/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024360333 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2023 |  |  |  |  | The user facility reported an 82-year-old male with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. due to patient anatomy, the impella 5.5 was not able to be advanced into the left ventricle. it was able to pass through the peel away sheath and just past the anastomosis of the surgical graft. the decision was made to abort placement of 5.5 and place impella cp with axillary access. | [5 empty] | The investigation into the delivery issues has been completed. no product was returned. as per clinical, the pump was able to pass through the introducer sheath past the anastomosis of the surgical graft but did not advance into left ventricle due to patient anatomy. the cause of the delivery issue was patient condition related since the pump was unable to advance into left ventricle due to patient anatomy.
31450 | 1220648-2024-06278 | 18579580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18579580&pc=OZD | 01/11/2024 | 02/12/2024 |  | 01/25/2024 |  |  | 01/11/2024 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | No patient involvement: the complainant reported that an aic (automated impella controller) failed a system check following start up. the aic was rebooted following the controller error and the device was able to start up without issue. there was no patient involvement at the time of the noted issue. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - system self-check error has been completed since the original report was filed. the product was returned for investigation. no console logs were available at the time of the issue occurrence because the self check failure prevented the console from booting up. therefore, console logs were not used in this investigation. the console ic7567 was tested and physically inspected after being returned to field service. a firmware check was performed, which revealed that there was a lack of communication between the 4-in-1 and the battery 1 circuitry on the pbm. the cause of the aic issue was contamination. there was corrosion of the pbm traces from leakage of electrolyte of the supercapacitors.
31451 | 1220648-2024-06280 | 18579930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18579930&pc=OZD | 10/02/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 10/02/2023 | 01/25/2024 |  |  | 07/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024386919 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was undeliverable through the iliac stents. after a second unsuccessful attempt, the case was aborted. | [5 empty] | The impella device was not received from the customer. the root cause of the delivery issue - anatomy was most likely due to patient condition due to iliac stents in the patient.
31452 | 1220648-2024-06281 | 18580234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18580234&pc=OZD | 01/12/2024 | 02/26/2024 |  | 01/25/2024 |  |  | 01/12/2024 | 01/25/2024 |  |  | 11/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024421976 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to a potential heart transplant. while on support, the patient sat in a chair and the red alarm impella in aorta appeared. echo was checked and the impella was pulled into the aorta. the impella was removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported positioning issue has been completed since the original report was filed. the root cause of the positioning issue was patient movement related that led to pull back the impella. this is a use related issue as if properly fixed the impella 5.5 should not get dislodged into the aorta with movement.
31453 | 1220648-2024-06282 | 18580501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18580501&pc=OZD | 01/12/2024 | 02/28/2024 |  | 01/25/2024 |  |  | 01/12/2024 | 01/25/2024 |  |  | 10/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024414242 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for patient admitted with cardiac symptoms and in need of a hrpci. the pump access site was found to be bleeding. the team troubleshot by adjusting dressing, manual holding, and sutures placed. additionally the hematoma was expressed and a femostop placed. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was most likely use issue due to insufficient sutures since extra sutures along with mattress suture had to be placed and then after placing femstop, the hematoma was expressed. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31454 | 1220648-2024-06283 | 18580622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18580622&pc=OZD | 10/13/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 10/13/2023 | 01/25/2024 |  |  | 08/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024390873 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/13/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 insertion was unsuccessful. the patient¿s vessels were occluded. the case was aborted. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer. the cause of the delivery issue was patient condition related because was unable to go direct insertion due to grafts being in the way.
31455 | 1220648-2024-06284 | 18580770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18580770&pc=OZD | 10/23/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 10/23/2023 | 01/25/2024 |  |  | 08/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024384528 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/23/2023 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported that due to aortic heavy calcification, unable to advance impella in aorta and the medical doctor made the decision to remove the pump. no known patient harm or injury. | [5 empty] | No product was returned. the root cause of the delivery issue was most likely patient condition related.
31456 | 1220648-2024-06285 | 18580801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18580801&pc=OZD | 01/11/2024 | 03/18/2024 |  | 01/25/2024 |  |  | 01/11/2024 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, CA |  |  | ABIOMED, INC. |  |  | US |  | 1182681 |  | 0042-0000-CA |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | No patient involvement: the complainant in canada had a console with a yellow alarm as seen by the hospital biomed dept. the console was removed and replaced by a loaner per ifu recommendations while the console is sent for investigation. no patient harm or injury. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm) issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. the issue was initially not reproduced and upon each boot-up no alarms were present. however, visual inspection of the optical pump connector revealed heavy contamination, which could¿ve caused the ¿bad lamp¿ condition. the cause of the aic issue was contamination.
31457 | 1220648-2024-06287 | 18580956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18580956&pc=OZD | 01/12/2024 | 02/29/2024 |  | 01/25/2024 |  |  | 01/12/2024 | 01/25/2024 |  |  | 11/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423460 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 83-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed for a hrpci (high risk percutaneous coronary intervention) who had been turned down by surgeons. the support was placed and was for just under 1 hour when the pump was explanted. the patient had a pericardiocentesis done for an effusion of unknown cause. md was not able to determine if either the wire or pump was the cause of the perforation. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation into the ventricle perforation issue has been completed. the device was not returned for investigation. based on the clinical details provided, the root cause of the ventricle perforation issue was not determined since product was not returned and insufficient clinical information provided to conclude anything.
31458 | 1220648-2024-06286 | 18580969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18580969&pc=OZD | 11/08/2023 | 09/03/2024 |  | 01/25/2024 |  |  | 11/08/2023 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397917 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support ongoing for two months when there was reddening at the access site. the site prompted the team to give the patient antibiotics. the pump remained on despite this infection for a month. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a and d6b added. h6 component code added. instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings and cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06286. b2 required intervention to prevent permanent impairment/damage (devices) inadvertently was not selected. d4 model number. f6/f8-should have been left blank. g1 reporting contact fax number missing.
31459 | 1220648-2024-06288 | 18581038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581038&pc=OZD | 01/12/2024 | 02/26/2024 |  | 01/25/2024 |  |  | 01/12/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403161 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 74-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed hemolysis and was taken back to the cardiac catheterization lab for repositioning, it was suspected the impella was stuck under the papillary muscle. after a few days the labs were still showing hemolysis, but dialysate fluid was clear. one unit of packed red blood cells were given. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis | The investigation into the hemolysis has been completed. the device and data were not returned for investigation. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position.
31460 | 1220648-2024-06289 | 18581316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581316&pc=OZD | 01/09/2024 | 03/06/2024 |  | 01/25/2024 |  |  | 01/09/2024 | 01/25/2024 |  |  | 04/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344072 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The complainant reported immediately after impella cp placement. the female patient experienced suction alarms. the maximum pump flow was noted to be at 1l. x-ray image showed a kinked cannula after removal of the guidewire. the pump ultimately explanted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported cannula kink has been completed. the impella device was returned for evaluation. the pump was visually inspected finding a cannula kink near the outlet. no biomaterial or broken blades found. the cause of the low pump flow was cannula kink due to improper technique. added: d9-device returned date d4-model number-updated to impella cp.
31461 | 1220648-2024-06290 | 18581380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581380&pc=OZD | 01/10/2024 | 01/15/2025 |  | 01/25/2024 |  |  | 01/10/2024 | 01/25/2024 |  |  | 10/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024419831 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Failure to Pump | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient of unknown age and gender was on impella cp support for hemodynamic support. the automated impella controller (aic) displayed console alarm "impella defect". therefore, the pump was explanted. a new pump was placed, and a new aic console was used to resolve the issue. there was no reported patient harm. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. prior to impella support, the patient was cannulated with extracorporeal membrane oxygenation (ecmo) and was intubated. it was reported that after the impella cp with serial number (b)(6) (pump 1) was implanted and connected to the automated impella controller (aic) with serial number (b)(6) (aic 1), an alarm occurred for impella defective. the impella cp was explanted and replaced with a new pump with serial number (b)(6) (pump 2), and the impella defective alarm was not resolved. an additional aic with serial number (b)(6) (aic 2) was connected to the second pump, and the impella defective alarm was not resolved. the aic was replaced again with a third unit, the alarm was resolved, and support was initiated. after 6 days of ecmo and impella support, a cerebral bleed was discovered, and therapy was discontinued. it was noted that the cerebral bleed was not impella related. the patient was given systemic heparin during impella use. it was noted that the patient expired while on impella cp support. there was a delay in needed support when pump 1 would not be recognized by aic 1. additionally, there was a delay of support as the aic 1 and aic 2 would not recognize pump 1 or pump 2. it cannot be said the period of time with loss of hemodynamic support did not contribute to the patient's overall outcome. the patient experienced a cerebral vascular accident while on support with pump 2. while the physician stated it was not impella related, but there is no alternative cause identified or provided as the likely cause of such event. additionally, the patient was given systemic heparin for impella use which may have contributed to the cerebral bleeding. pump 1, aic 1, aic 2, and pump 2 cannot be ruled out as contributing factors in the overall event. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the pump did not start issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the "impella defective" alarms and pump not detected failure cannot be determined as there is no product or data logs to evaluate. | Additional information from initial mfr 1220648-2024-06290 was provided in section b5, d6a, and d6b. this is one of four medical device reports associated with the event. refer to manufacturing report number 1220648-2024-06292 for automated impella controller 1 with serial number (b)(6). refer to manufacturing report number 1220648-2024-06294 for automated impella controller 2 with serial number (b)(6). refer to initial medwatch with date of report 15-jan-2025 for pump 2 impella cp serial number (b)(6). impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information from initial mfr 1220648-2024-06290 was provided in section b1, b2, h1, and h6.
31462 | 1220648-2024-06291 | 18581466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581466&pc=PYX | 01/14/2024 | 02/29/2024 |  | 01/25/2024 |  |  | 01/14/2024 | 01/25/2024 |  |  | 07/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024399893 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2024 |  |  |  |  | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella procedure the impella started at p-2 but immediately after the pump stopped. motor current high alarm followed by impella stopped. the level was reduced to p-0 and started again 4 times with same result when trying to increase the p-level. the decision was made to remove it and a new impella was successfully inserted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned to fs for investigation. upon investigation of the pump, a large biomaterial was found wrapped around the impeller. the root cause of the temporary pump stop was determined to be biomaterial.
31463 | 1220648-2024-06292 | 18581509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581509&pc=OZD | 01/10/2024 | 12/03/2025 |  | 01/25/2024 |  |  | 01/10/2024 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1127918 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Connection Problem | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient of unknown age and gender was on impella cp support for hemodynamic support. the automated impella controller (aic) displayed console alarm "impella defect". therefore, the pump was explanted. a new pump was placed, and a new aic console was used to resolve the issue. there was no reported patient harm. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. prior to impella support, the patient was cannulated with extracorporeal membrane oxygenation (ecmo) and was intubated. it was reported that after the impella cp with serial number 478790 (pump 1) was implanted and connected to the automated impella controller (aic) with serial number 2240 (aic 1), an alarm occurred for impella defective. the impella cp was explanted and replaced with a new impella cp pump with serial number 463963 (pump 2), and the impella defective alarm was not resolved. an additional aic with serial number 5084 (aic 2) was connected to the second pump, and the impella defective alarm was not resolved. the aic was replaced again with a third unit, the alarm was resolved, and support was initiated. after 6 days of ecmo and impella support, a cerebral bleed was discovered, and therapy was discontinued. it was noted that the cerebral bleed was not impella related. the patient was given systemic heparin during impella use. it was noted that the patient expired while on impella cp support. there was a delay in needed support when pump 1 would not be recognized by aic 1. additionally, there was a delay of support as the aic 1 and aic 2 would not recognize pump 1 or pump 2. it cannot be said the period of time with loss of hemodynamic support did not contribute to the patient's overall outcome. the patient experienced a cerebral vascular accident while on support with pump 2. while the physician stated it was not impella related, but there is no alternative cause identified or provided as the likely cause of such event. additionally, the patient was given systemic heparin for impella use which may have contributed to the cerebral bleeding. pump 1, aic 1, aic 2, and pump 2 cannot be ruled out as contributing factors in the overall event. |  |  |  |  | The automated impella controller (aic) device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Additional information from initial mfr 1220648-2024-06292 was provided in section b5. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this is one of four medical device reports associated with the event. refer to manufacturing report number 1220648-2024-06290 for pump 1 impella cp with serial number (b)(6). refer to manufacturing report number 1220648-2024-06294 for automated impella controller 2 with serial number (b)(6). refer to initial medwatch with date of report 15-jan-2025 for pump 2 impella cp serial number (b)(6). corrected information from initial mfr 1220648-2024-06292 was provided in section b1, b2, h1, and h6. | Data analysis: on the complaint date (jan 10, 2024), pump sn (b)(6) was initially used with ic2240, and no "impella defective" alarm or pump detection issue was observed. subsequently, the user decided to switch the pump, likely due to multiple "air in the purge system" alarms (event #404). the rt log indicates that the purge pressure did not increase by 5 mmhg during the piston's first stroke after de-airing each time, resulting in the persistent event set #404. this issue was most likely caused by a piston block (unreported and unrelated to the reported failure). additionally, the snr was low (less than 1000), indicating a secondary, unreported failure mode: a defective optical bench (also most likely unrelated to the reported failure). after the console (ic2240) was rebooted, at step 3 of the case start wizard, the second pump (sn (b)(6)) was connected. an ¿impella defective¿ alarm (event #3) was immediately triggered. the second pump was reconnected twice, but event #3 persisted, confirming the reported failure mode. the user then moved the second pump (sn (b)(6)) from ic2240 to ic5084. however, the ¿impella defective¿ alarm was also displayed on ic5084. device analysis: console ic2240 was returned to fs for further analysis. the aic was thoroughly examined, and no loose cables were observed. voltage measurements for the impellatronic pca were found to be within specification. when tested with a known-good pump and purge cassette, the console did not reproduce the ¿impella defective¿ alarm or pump detection issue. additionally, the purge pressure sensors were inspected and confirmed to be within specification as per the pm procedure. root cause: the root cause of the intermittent "impella defective" alarm could not be determined due to the inability to reproduce it. d2 common device name updated. d10 concomitant medical products and therapy dates updated.
31464 | 1220648-2024-06293 | 18581549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581549&pc=OZD | 01/15/2024 | 09/26/2025 |  | 01/25/2024 |  |  | 01/15/2024 | 01/25/2024 |  |  | 06/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 1000080 |  |  | ABIOMED, INC. |  |  | US |  | 2024383606 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/15/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the cabg (coronary artery bypass grafting) procedure after cross clamping, the impella stopped and restarted three times. a pump stopped alarm occurred. once the impella was placed in surgical mode the cabg procedure finished without a problem. after removing the clamp, the impella was restarted without any problem and had expected flows. it was further reported that as the impella was ramped up an alarm for position in aorta with placement signal in left ventricle occurred. tee (transesophageal echocardiography) showed that the impella was in good position across valve. the alarm was disabled, and pump was left in position for 24hr and then explanted. echo in icu also demonstrated proper positioned and there were no apparent problems with flows. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. no product was returned for investigation. data logs show a motor current spike with a high motor current pump stop. there are suction alarms before and after the pump stop. there are impella position unknown and impella position in aorta alarms upon restart of the pump. the pump's position was confirmed to be in correct positioning. the root cause of the pump stop was most likely use related as the pump was not placed into surgical mode at the start of the cabg procedure; when the pump was put into surgical mode it was able to be restarted and ran with expected flows. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause of the optical signal issue was most likely biomaterial related since there was a motor current spike and the impella in aorta alarm was a false alarm as the pump was confirmed to be in correct position. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b7 updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. h6 updated to include optical signal coding. d4 model number updated. f6 and f8 erroneously filled in initial report. g1 reporting contact email updated.
31465 | 1220648-2024-06294 | 18581582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581582&pc=OZD | 01/10/2024 | 12/03/2025 |  | 01/25/2024 |  |  | 01/10/2024 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1285928 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Connection Problem | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | A patient of unknown age and gender should receive hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. during preparation, the automated impella controller (aic) issued a ¿purge disk not detected¿ alarm. the aic with serial number(b)(6) was exchanged for the aic with serial number(b)(6). this then issued an ¿impella defective¿ alarm. the impella cp smartassist heart pump was replaced with a new impella cp smartassist. the ¿impella defective¿ alarm remained. the aic with serial number (b)(6) was exchanged for a third aic. the impella cp then worked without any further problems. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. prior to impella support, the patient was cannulated with extracorporeal membrane oxygenation (ecmo) and was intubated. it was reported that after the impella cp with serial number (b)(6) (pump 1) was implanted and connected to the automated impella controller (aic) with serial number (b)(6) (aic 1), an alarm occurred for impella defective. the impella cp was explanted and replaced with a new impella cp pump with serial number 463963 (pump 2), and the impella defective alarm was not resolved. an additional aic with serial number (b)(6) (aic 2) was connected to the second pump, and the impella defective alarm was not resolved. the aic was replaced again with a third unit, the alarm was resolved, and support was initiated. after 6 days of ecmo and impella support, a cerebral bleed was discovered, and therapy was discontinued. it was noted that the cerebral bleed was not impella related. the patient was given systemic heparin during impella use. it was noted that the patient expired while on impella cp support. there was a delay in needed support when pump 1 would not be recognized by aic 1. additionally, there was a delay of support as the aic 1 and aic 2 would not recognize pump 1 or pump 2. it cannot be said the period of time with loss of hemodynamic support did not contribute to the patient's overall outcome. the patient experienced a cerebral vascular accident while on support with pump 2. while the physician stated it was not impella related, but there is no alternative cause identified or provided as the likely cause of such event. additionally, the patient was given systemic heparin for impella use which may have contributed to the cerebral bleeding. pump 1, aic 1, aic 2, and pump 2 cannot be ruled out as contributing factors in the overall event. |  |  |  |  | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Additional information from initial mfr 1220648-2024-06294 was provided in section b5. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this is one of four medical device reports associated with the event. refer to manufacturing report number 1220648-2024-06292 for automated impella controller 1 with serial number (b)(6). refer to manufacturing report number 1220648-2024-06290 for pump 1 impella cp with serial number (b)(6) refer to initial medwatch with date of report 15-jan-2025 for pump 2 impella cp serial number (b)(6). corrected information from initial mfr 1220648-2024-06294 was provided in section b1, b2, h1, and h6. | Data analysis: the pump (sn (b)(6)) was moved from ic2240 to ic5084. however, immediately after the pump connection, the ¿impella defective¿ alarm (event #3) was also displayed on ic5084 as it was displayed on the initial console ic2240. note: according to the complaint coordinator, pump 463963 was moved to the third console, ic6874. ic6874 logs confirmed there was no impella detective alarm, and support was provided for five days, indicating the reported failure mode was highly intermittent. device analysis: console ic5084 was returned to fs for further analysis. no hardware issue related to the "impella defective" alarm was found on the aic (ic5084). root cause: the root cause of the intermittent "impella defective" alarm could not be determined due to the inability to reproduce it. d2 common device name updated. g1 reporting contact email and reporting contact us city upated.
31466 | 1220648-2024-06295 | 18581619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581619&pc=OZD | 05/27/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 05/29/2023 | 01/25/2024 |  |  | 01/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024312675 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp support of four days when the pump was explanted. the pump was noted to explant at time of an underlying cerebral hemorrhage. the sub-arachnoid bleed was not known to be due to any pump malfunction. the bleed was of unknown cause, no allegation by the physician, the patient was anticoagulated. | [5 empty] | (h3) the investigation into the reported "stroke/cva" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the stroke was not determined due to insufficient clinical details and unknown relation to impella. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31467 | 1220648-2024-06296 | 18581635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581635&pc=OZD | 01/15/2024 | 01/31/2024 |  | 01/25/2024 |  |  | 01/15/2024 | 01/25/2024 |  |  | 11/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024429986 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/15/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced bleeding at their impella cp access site following patient repositioning. manual pressure was insufficient to manage the condition and the pump was subsequently removed. the perclose failed to deploy and angioseal was also unsuccessful at treating the site. a vascular surgeon was consulted and two units of prbc were transfused. when a second attempt at utilizing perclose and angioseal failed, wire access was lost, and manual pressure had to be held. hemostasis was finally achieved, and no further intervention was required. | [5 empty] | The investigation into the access site bleeding major issue has been completed. the product was not returned for investigation. per the clinical data provided, the root cause of the access site bleeding issue was most likely due to high patient act values. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31468 | 1220648-2024-06297 | 18581704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581704&pc=OZD | 01/18/2024 | 03/01/2024 |  | 01/25/2024 |  |  | 01/18/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024405946 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The user facility reported a 15-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was stuck in vt (ventricular tachycardia) following impella implant. as a result of the noted condition, the decision was made to explant the pump. the vt converted roughly twelve hours after explant. | [5 empty] | The investigation into the ventricular tachycardia (vt) event has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31469 | 1220648-2024-06298 | 18581733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581733&pc=OZD | 01/09/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 01/09/2023 | 01/25/2024 |  |  | 09/30/2024 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275050 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that difficulty was encountered while attempting to place an impella cp pump for patient support. access was initially obtained via the right femoral artery, but resistance was met and access was flipped to the left femoral artery. once again, the pump was unable to advance due to calcification and a tortuous vasculature. the resistance resulted in some minor kinking of the sheath. two more attempts were made utilizing different techniques as well as a stent; however, the device could still not advance pass the iliac. the decision was subsequently made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the reported introducer damage has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the introducer sheath kink was patient condition related to severely tortuous and calcified vessels. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31470 | 1220648-2024-06299 | 18581833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581833&pc=OZD | 09/27/2022 | 03/08/2024 |  | 01/25/2024 |  |  | 01/22/2024 | 01/25/2024 |  |  | 06/30/2024 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023248047 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had enrolled a patient int he protect iv while on support with impella cp. the pump support was successful for 2 hours and allowed for ptca (percutaneous transluminal coronary angioplasty) of the coronary disease. the support was in(b)(6)2022, while later in (b)(6)/2024 the crf - chronic renal failure (kidney failure)- form noted the patient to have suffered from cardiogenic shock 2 days post explant with pea (pulseless electrical activity) and vt (ventricular tachycardia) arresting, all measures were stopped when the family made choice to call the code and the patient expired. the pi alleges the shock had a "probable" relationship to the impella. | [5 empty] | The investigation into the vf/vt/af/at issues has been completed since the original report was filed. the device was not returned for investigation. as no necessary clinical information was provided and the device was not returned, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31471 | 1220648-2024-06300 | 18581871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581871&pc=OZD | 01/08/2024 | 02/13/2025 |  | 01/25/2024 |  |  | 01/08/2024 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1457787 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | No patient involvement: the complainant reported that the aic had a controller error. the aic was removed from service. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) controller failure-red alarm has been completed. data logs from the complaint date confirm the failure mode with the issuance of the controller failure (epm1 sw corruption) alarm. this controller failure happened on boot up and there is no evidence in the logs record of the alarm being issued prior to the complaint date and console was immediately replaced. the console was returned for evaluation. the alarm was reproduced on boot up in the lab. visual inspection of the internal circuitry of the impellatronics board indicated that there was no loose cable connections and that there was power being delivered to the impellatronics board from the pbm. the firmware of the epm was then extracted from the impellatronics board and compared to the product source code revealing a difference. this confirmed that the epm firmware was corrupted. the epm software was then reinstalled onto the impellatronics board which fixed the issue. the root cause of the red controller failure was corruption of the epm software on the impellatronics board. corrections to the initial mfr report: d2 updated common device name f6/f8 should have been let blank.
31472 | 1220648-2024-06301 | 18581894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581894&pc=OZD | 01/12/2024 | 02/02/2024 |  | 01/25/2024 |  |  | 01/12/2024 | 01/25/2024 |  |  | 10/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024413582 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a (b)(6) year-old male patient on impella cp support was bleeding from the left groin. the sheath was repositioned to resolve the issue. the patient was also noted to have been given two units of packed red blood cells (prbcs). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. therefore, the root cause of the access site bleed was most likely use-related as manipulation of repo sheath resolved bleeding.
31473 | 1220648-2024-06302 | 18581936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581936&pc=OZD | 01/09/2024 | 03/06/2024 |  | 01/25/2024 |  |  | 01/09/2024 | 01/25/2024 |  |  | 06/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368576 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu reported a female patient of unknown age on impella cp support, experienced suction alarms immediately after pump placement. maximum pump flow was noted to be at 0.8 liters. pump position was confirmed to be correct using x-rays. act was 286 and the sheath was aspirated and flushed. there was no thrombus noted. pump was ultimately explanted due to reported event. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump ran without issue under bench test. the root cause of low flow was most likely patient condition as the patient has small left ventricle and no issues were found on the return pump.
31474 | 1220648-2024-06303 | 18581968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18581968&pc=OZD | 01/09/2024 | 12/03/2025 |  | 01/25/2024 |  |  | 01/09/2024 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1248778 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant in the eu reported the automated impella controller (aic) was only able to be started up to the point where the purge cassette has to be inserted. therefore, the console was replaced with a new one and the issue resolved. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. h6 type of investigation code added.
31475 | 1220648-2024-06304 | 18582114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582114&pc=OZD | 01/13/2024 | 03/01/2024 |  | 01/25/2024 |  |  | 01/13/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024401329 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a thrombus was suspected in the outflow of an implant impella 5.5 pump. due to the noted condition, the decision was made to replace the pump. following replacement, lactate dehydrogenase¿ (ldh) and creatinine levels were improving and support continued for the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported thrombus has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
31476 | 1220648-2024-06305 | 18582240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582240&pc=OZD | 01/09/2024 | 02/23/2024 |  | 01/25/2024 |  |  | 01/09/2024 | 01/25/2024 |  |  | 10/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420485 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported female patient received impella cp support for percutaneous coronary intervention (pci), but had to be explanted via vascular surgery after 12-hours, due to bleeding. the hemorrhage was treated with a patch. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
31477 | 1220648-2024-06306 | 18582323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582323&pc=OZD | 01/09/2024 | 01/25/2024 |  | 01/25/2024 |  |  | 01/09/2024 | 01/25/2024 |  |  | 01/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312751 |  | 0048-0014 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant from the stemi-dtu study reported patient on impella cp support has bilateral access site bleeding that was moderate in severity. the patient's hemoglobin dropped and patient received blood transfusion. the patient was noted to have recovered. the event was determined by the primary investigator to be probably related to the impella procedure and device. | [5 empty] | The investigation into the reported ¿access site bleeding-major¿ has been completed. the root cause of the access site bleeding was most likely due to patient condition due to bilateral access site bleeding. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ a risk assessment was performed, and no escalation is required at this time.
31478 | 1220648-2024-06307 | 18582375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582375&pc=OZD | 01/13/2024 | 03/06/2024 |  | 01/25/2024 |  |  | 01/14/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024418723 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2024 |  |  |  |  | The user facility reported a 36-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that distal end of the sterile sleeve became detached from the tuohy-borst lock while the patient was ambulating. the 5.5 was in proper 3-point fixation, currently the distal end of the sleeve is covered in a tegaderm. no patient harm was reported. | [5 empty] | The investigation into the product damage (sterile sleeve damage), the thrombus, and the infection/sepsis issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the product damage - sterile sleeve issue was most likely use related. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus and the infection was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31479 | 1220648-2024-06308 | 18582498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582498&pc=PYX | 01/14/2024 | 03/01/2024 |  | 01/25/2024 |  |  | 01/14/2024 | 01/25/2024 |  |  | 07/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024390099 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella rp position is incorrect. the rp was repositioned since it was very shallow, and the physician thinks the pigtail is malpositioned and not across the valve. the impella rp was explanted due to positioning issues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on clinical description & data analysis, the root cause of positioning issue was most likely due to use related.
31480 | 1220648-2024-06309 | 18582584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582584&pc=OZD | 01/16/2024 | 03/01/2024 |  | 01/25/2024 |  |  | 01/16/2024 | 01/25/2024 |  |  | 11/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024425175 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 64-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the physician was concerned about the hematoma at the impella insertion site and the continuing oozing around the sheath. the decision was made to wean and explant the impella. | [5 empty] | The investigation into the reported access site bleeding- major has been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31481 | 1220648-2024-06310 | 18582744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582744&pc=OZD | 01/16/2024 | 02/28/2024 |  | 01/25/2024 |  |  | 01/16/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024412911 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Integrity Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 89-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an inserted 14 fr introducer kinked during impella cp implant. the kink resulted in unplanned resistance during the delivery. there was no patient harm and the pump was able to advance to the intended location despite the complication. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported introducer damage ¿ mechanical has been completed since the original report was filed. no product was returned for evaluation. it was noted in the clinical that the patient has moderate and severe tortuosity femoral arteries. the root cause of the introducer damage was most likely due to the patients tortuous anatomy of the femoral arteries.
31482 | 1220648-2024-06311 | 18582889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582889&pc=OZD | 02/04/2023 | 09/21/2024 |  | 01/25/2024 |  |  | 02/04/2023 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024286029 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Death | Yes | Yes | Device Alarm System | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/04/2023 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported their automated impella controller (aic) displayed a "controller failure" alarm during use. the aic was switched to the back-up aic without any harm to the patient. the patient ultimately expired, however, it was indicated that this event had no bearing on the patient's outcome. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The investigation of the controller failure alarm has been completed. the user facility returned the aic for investigation. the alarm could not be reproduced by the investigation team. the root cause of controller failure alarm was noted to be most likely a defective pbm board. the failure modes will be monitored and trended. console logs from the reported day of event are consistent with complaint and show several controller error alarms (#501, ¿pbm voltage out-of-spec¿) all of which self-resolved, and two controller failure alarms (#800, ¿dsp voltage too low¿). analog voltage measurement subsystems on pbm (powermod_1 and powermod_2) reported voltages to pmd (dsp) and to purge to be below 12v, which triggered critical alarm (#800 - dsp voltage low). according to ltlog and rtlog data for pump 423571, hemodynamic support was not affected. when the console was tested in aachen, the issue was not reproduced. analog voltages measured by pbm were within specification with the console operating on batteries and on ac. this, combined with the intermittent nature of errors and alarms suggests that the corresponding voltages were within specification during the case, but the measurement subsystem (located on pbm) malfunctioned and was reporting incorrect values, triggering erroneous alarms. although inspection of the pbm did not reveal any irregularities (or corrosion), it is recommended to be replaced as a precaution. repair: replace pbm assembly (0042-1125), and perform functional check.
31483 | 1220648-2024-06312 | 18582974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18582974&pc=OZD | 01/16/2024 | 02/06/2024 |  | 01/25/2024 |  |  | 01/16/2024 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1666616 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The complainant reported that the aic (automated impella controller) had a broken purge clip. the aic was removed from service. no patient was involved. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. | The impella device was received from the customer and¿an investigation of the device is ongoing. upon completion of our analysis, a supplemental report will be filed.
31484 | 1220648-2024-06313 | 18583411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18583411&pc=OZD | 03/07/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 03/07/2023 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1471020 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Temperature Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) controller had battery temperature high red alarm. the engineer checked the voltage and noted it was fine. the alarm resolved quickly. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. the evaluation revealed that the most likely cause for the failure mode is most likely due to pbm incorrectly reporting battery temperatures. the failure modes will be monitored and trended. ¿data analysis: in the aic logs for ic8020 (imc ic8020.pdf), the console received a momentary battery temperature high alarm #49 on march 7 which resolved 10 seconds later. additionally, the lt logs show the momentary battery cell voltage spike, which confirms the theory momentary pbm miscommunication (ic8020 battery voltage spike.png). although battery temperature did not spike also in the lt logs, it is believed that the sampling method did not record this transient temperature increase due to miscommunication. ¿device analysis:¿ the device was returned for investigation six months after the complaint date. the failure mode could not be reproduced, and the console had been functioning for the previous six months without the failure mode reoccurring. ¿root cause: the cause for the failure mode is most likely due to pbm momentarily miscommunicating battery temperatures. ¿repair:¿ please replace the pbm (0042-1125), then perform sections 18-20 of pm procedure (0042-7309) and a full functional check (0042-7316).
31485 | 1220648-2024-06314 | 18583629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18583629&pc=OZD | 01/16/2024 | 03/01/2024 |  | 01/25/2024 |  |  | 01/16/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024399101 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced bleeding from their access site during impella cp support. the site was redressed to angle match. the suture hub was advanced, and forward sutures were placed to manage the condition. it was also noted that statseal was applied to the site. support remained ongoing with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major event has been completed. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved after adding extra sutures and adjusting the angle.
31486 | 1220648-2024-06315 | 18583752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18583752&pc=OZD | 04/13/2023 | 10/14/2024 |  | 01/25/2024 |  |  | 04/13/2023 | 01/25/2024 |  |  | 12/31/2024 |  | 01/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300070 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was additionally reported the patient experienced oozing at the impella access site. the site was redressed and the patient received one unit of blood products. | The user facility reported patient who presented for high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the pump was inserted, the repo sheath was removed and by repositioning sheath it was advanced; the pump migrated back in the aorta. multiple attempts were made to re-cross without success. therefore, a side port accessed and pump was exchanged. there was no harm to the patient as a result of this incident. |  |  |  |  | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the positioning issue was most likely due to wireless re-access. the impella device is not indicated for wireless re-access per ifu (b)(4). the failure modes will be monitored and trended. | Corrected information for the initial mfr 1220648-2024-06315 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
31487 | 1220648-2024-06316 | 18583954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18583954&pc=OZD | 05/06/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 05/06/2023 | 01/25/2024 |  |  | 11/30/2024 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312733 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/06/2023 |  |  |  |  | The user facility reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed distal limb ischemia during impella support. a distal perfusion catheter was placed to return flow to the leg and support with the implanted pump continued. it was noted that the patient had known peripheral arterial disease leading up to the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. no product returned for investigation. the root cause of the limb ischemia - reversible issue was most likely patient condition related due to the patient's severe pad. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31488 | 1220648-2024-06317 | 18584202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18584202&pc=OZD | 03/24/2023 | 01/25/2024 |  | 01/25/2024 |  |  | 05/18/2023 | 01/25/2024 |  |  |  |  | 01/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1288838 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the device console malfunctioned during patient support. console displayed 'yellow alarm'. there was no harm to the patient. | [5 empty] | The investigation into the reported product damage has been completed. the evaluation revealed that the cause of the reported failure mode ¿ unexpected console reboot was due to a sw restart to mitigate the sw processes crash. additionally, the cause of the eeprom corruption was not determined since failure mode was unable to be reproduced and pump was not returned. the failure modes will be monitored and trended. data analysis: pump 421368 experienced issues reading eeprom flow characteristic curves 2 on (b)(6) after being unplugged and plugged back in as documented in complaint (b)(4) (imc (b)(6).pdf) . then ic1707 had calculator and audio process crashes followed by the sw watchdog triggering a restart. on the reboot, (b)(6) was unable to read the first eeprom line (calibration data) before triggering an impella defective alarm. the pump was then transferred from (b)(6) to (b)(6). (b)(6) immediately has trouble reading flow characteristic curves 2 which is then followed by the calculator process crashing and a sw watchdog reboot (imc (b)(6). on reboot, the console sees an impella defective alarm with the inability to read the calibration data and no optical data is seen in the rt logs (b)(6) impella defective after reboot.png). the pump is unplugged and then, after being plugged back into (b)(6), performs as expected (b)(4) recovery after unplug pump.png). device analysis: this failure mode was not reproduced on either (b)(6) or (b)(6) and the pump was not returned. root cause: the reported failure mode ¿ unexpected console reboot was due to a sw restart to mitigate the sw processes. the cause of the eeprom corruption was not determined since failure mode was unable to be reproduced and pump was not returned. repair: please perform a full functional check on the console before returning the console to the field.
31489 | 1220648-2024-06318 | 18584824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18584824&pc=OZD | 01/09/2024 | 01/25/2024 |  | 01/25/2024 |  |  | 01/09/2024 | 01/25/2024 |  |  | 11/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426264 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Delivery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported male patient was attempted to be put on impella 5.5 support following aortic valve (av) repair. however, the impella was unable to pass the av, despite several attempts. therefore, the impella implant procedure was aborted and impella and wire were removed. the patient was noted to be stable and there was no reported harm. | [5 empty] | The investigation into the reported ¿delivery issue¿ has been completed. no product was returned for investigation. the root cause of the delivery issue was most likely patient condition related. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31490 | 1220648-2024-06319 | 18584832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18584832&pc=OZD | 01/15/2024 | 09/26/2024 |  | 01/25/2024 |  |  | 01/15/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024402294 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/15/2024 |  |  |  |  | The user facility reported an 83-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during the implant it was difficult to pass the iliac artery. the short sheath was exchanged for the long sheath and upon fluoroscopy it was found that the sheath was kinked in two places. the dilator was placed back into the sheath with the wire and the kinks resolved and the impella was implanted. it was noted that patient¿s leg was mottled and no dopplerable pulses were found, after 4 days the peel-away sheath was removed and replaced with a repositioning sheath and the blood flow improved with pulses being found via doppler. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the reported introducer damage and limb ischemia reversible has been completed since the original report was filed. no product was returned for evaluation. clinical review noted that patient had pad/pvd. the root cause of the introducer damage was most likely due to disease vasculature of the patient. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06319 was provided in sections b5 and h1.
31491 | 1220648-2024-06320 | 18584861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18584861&pc=OZD | 01/16/2024 | 02/23/2024 |  | 01/25/2024 |  |  | 01/16/2024 | 01/25/2024 |  |  | 09/30/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024405259 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 46-year-old female patient of unknown race with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the patient lost pulses distal to the placement site. the impella was removed, and vascular surgery was performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31492 | 1220648-2024-06321 | 18584868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18584868&pc=OZD | 01/16/2024 | 03/12/2024 |  | 01/25/2024 |  |  | 01/16/2024 | 01/25/2024 |  |  | 08/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422945 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 77-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported at the conclusion of the procedure the automated impella controller (aic) alarmed ¿impella stopped and controller failure.¿ the impella was removed, upon removal, the 14x13 impella sheath split. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop, the introducer damage ¿ mechanical, and the access site bleeding - minor issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. the pump was visually inspected; no biomaterial or other physical abnormalities were found. the root cause of the pump stop and the introducer damage ¿ mechanical issues was not determined. however, the root cause of the access site bleeding issue was due to introducer sheath tip split during removal per field investigation and clinical review.
31493 | 1220648-2024-06322 | 18584887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18584887&pc=OZD | 01/16/2024 | 02/29/2024 |  | 01/25/2024 |  |  | 01/16/2024 | 01/25/2024 |  |  | 08/31/2025 |  | 01/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024386035 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hematoma | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility reported a 62-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient developed an access site hematoma which required a pocket revision and evacuation. the pump additionally became dislodged and was unable to be re-advanced. the team was unable to redeliver the pump to correct left ventricle positioning and was replaced. | [5 empty] | The investigation into the access site bleeding major and the positioning issues has been completed. the device was not returned for investigation. per the provided clinical details, the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was due to attempt to re-cross the aortic valve wirelessly. the impella device is not indicated for wireless re-access. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
31494 | 1220648-2024-06239 | 18569560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18569560&pc=OZD | 01/10/2024 | 02/19/2024 |  | 01/24/2024 |  |  | 01/10/2024 | 01/24/2024 |  |  | 09/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024405930 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing when on day 3 of the support the team noted the pump positioning near the mitral to be within the papillary. the pump was removed to reposition. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the product was returned for investigation. the pump was visually inspected and catheter kink was observed. also, biomaterial ingestion over the purge gap at outflow cage was found. the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access. the cause of the access site bleeding issue was patient condition related with patient having calcified vessels and complication resolved without any treatment/intervention.
31495 | 1220648-2024-06240 | 18569616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18569616&pc=OZD | 01/03/2024 | 10/09/2024 |  | 01/24/2024 |  |  | 01/03/2024 | 01/24/2024 |  |  | 12/31/2024 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317942 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to extensive peripheral arterial disease (aorta-bifemoral graft). when the impella was explanted, the physician found and evacuated a clot, then the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added-h6 component code 925 corrections to the initial mfr report: d4-corrected model number added-d6a/d6b f6/f8 should have been left blank. h5 changed to yes.
31496 | 1220648-2024-06241 | 18569626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18569626&pc=OZD | 01/10/2024 | 12/05/2024 |  | 01/24/2024 |  |  | 01/10/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMEND, INC. |  |  | US |  | 2024422862 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [5 empty] | FEM-FEM BYPASS | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that one day following implant, the patient's leg was cold to the touch and a fem-fem bypass was performed to improve the flow. during the dressing change of insertion site, the angle of insertion was not maintained, and a hematoma developed. the patient was provided one unit of fresh frozen plasma (ffp) and two units of packed red blood cells as treatment. the patient remained on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency ischemia: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06241. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email.
31497 | 1220648-2024-06242 | 18569645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18569645&pc=OZD | 01/04/2024 | 03/25/2024 |  | 01/24/2024 |  |  | 01/04/2024 | 01/24/2024 |  |  | 05/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344575 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was explanted after clot was identified in the left ventricle. the patient received sternotomy to remove clot. the patient was switched to iabp - intra aortic balloon (pump) support and is stable. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ it also states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ d4-updated model number | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31498 | 1220648-2024-06245 | 18569710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18569710&pc=OZD | 01/12/2024 | 02/23/2024 |  | 01/24/2024 |  |  | 01/12/2024 | 01/24/2024 |  |  | 09/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398485 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 72-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that pulses were lost in the right lower extremity, the right foot became molted and cool with no pulse present. the physician elected to remove the impella and the patient remained stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31499 | 1220648-2024-06243 | 18569727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18569727&pc=OZD | 12/01/2023 | 02/22/2024 |  | 01/24/2024 |  |  | 01/05/2024 | 01/24/2024 |  |  | 09/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024407352 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Contamination | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a 5.5 support ongoing when the team observed crystallization in the purge side arm. this was observed on day 15 of the support. the pump support remained on for 39 days. this is beyond indication. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the pump was returned for evaluation. visually inspected and the purge sidearm retainer found to have dried dextrose remains. purge fluid was initiated by priming with purge cassette and connecting to the pump for inspecting any leaks at sidearm while purging. there was no issue found during the case. the device functioned/operated to specification per field investigation. there was no issue found during the case. the device functioned/operated to specification per field investigation.
31500 | 1220648-2024-06244 | 18569729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18569729&pc=OZD | 01/07/2024 | 01/24/2024 |  | 01/24/2024 |  |  | 01/07/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMEPLLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024418674 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/07/2024 |  |  |  |  | The user facility reported a (b)(6) year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma on their right femoral artery following impella cp implant. the hematoma was evacuated, and sutures were placed around the site to manage the condition. impella cp support continued following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. based on the provided information, a definitive cause for the noted condition could not be established. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
31501 | 1220648-2024-06246 | 18570136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18570136&pc=OZD | 01/13/2024 | 01/31/2024 |  | 01/24/2024 |  |  | 01/13/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422341 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2024 |  |  |  |  | The user facility reported a 75-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that following the peel away sheath placement the dilator was removed, and the sheath kinked. the wire was not able to pass the kink, a wire a dilator was not able to straighten out the kink either, finally the sheath was backed out over the long impella dilator by advancing the dilator as the peel away was removed, this allowed wire access to the vessel. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the reported difficulty inserting/removing introducer has been completed since the original report was filed. a peel away introducer sheath was returned for evaluation. upon visual inspection, multiple kinks were found along the sheath. the cause of the sheath kinks was most likely due to patient's highly calcified arteries based on visual inspection and patient's peripheral arterial disease/peripheral vascular disease.
31502 | 1220648-2024-06249 | 18570179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18570179&pc=OZD | 01/08/2024 | 05/16/2024 |  | 01/24/2024 |  |  | 01/08/2024 | 01/24/2024 |  |  | 07/31/2024 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023259173 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 62-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the implanted impella cp pump was unable to achieve flows greater than 1 l/min during patient support. after troubleshooting failed to resolve the noted condition, the decision was made to remove and replace the pump. there was no reported patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31503 | 1220648-2024-06247 | 18570192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18570192&pc=OZD | 01/07/2024 | 02/28/2024 |  | 01/24/2024 |  |  | 01/07/2024 | 01/24/2024 |  |  | 11/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024428638 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 68-year-old male patient admitted with a stemi placed on venoarterial (va) extra corporeal membrane oxygenation (ecmo) and impella cp support. during support, the team observed a large hematoma at the groin. the site was held and sutures placed but, the bleed did not resolve. the team treated the condition with 3 untis of red blood cells (rbc) and albumin. the pump was explanted and the team observed vessel damage at the superficial femoral artery (sfa)/profunda that was treated by stenting and transfer for vascular repair. the patient is noted to have survived. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding and vascular damage has been completed since the original report was filed. no damage or abnormalities were found on the returned product. there is no evidence of excessive force used. the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related.
31504 | 1220648-2024-06250 | 18570276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18570276&pc=OZD | 01/08/2024 | 11/04/2024 |  | 01/24/2024 |  |  | 01/08/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024419170 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 69-year-old male noted for high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the impella removal, it was noted there was no initial bleed back from the sheath removal. the left femoral artery (lfa) had been pre-closed prior to impella insertion. post sheath removal, an angiogram of the left femoral artery was performed. it was noted that there was some vessel dissection of lfa and one covered stent was placed successfully without issue. the physician believes the issue was due to poor vessel anatomy of patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. added h6 component code 925 corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank.
31505 | 1220648-2024-06252 | 18570408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18570408&pc=OZD | 01/08/2024 | 05/02/2024 |  | 01/24/2024 |  |  | 01/08/2024 | 01/24/2024 |  |  |  |  | 01/24/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1697069 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The complainant reported that after initial start of automated impella controller (aic) during preventive maintenance, the controller kept rebooting. there was no reported patient harm. | [5 empty] | The investigation for the reported console (aic) shut down or restart issues has been completed. the console was returned for evaluation. upon inspection, the failure mode was reproducible in the lab and evidenced with the console unable to load the boot up sequence. data logs were not relevant for this issue. the root cause for the for the rebooting issue was corrupt compact flash card making the epc unable to read the boot-loader off of the compact flash card. d4-updated model number g1-corrected reporting contact first and last name g1-corrected mfr site name and address. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
31506 | 1220648-2024-06251 | 18570857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18570857&pc=OZD | 01/08/2024 | 02/23/2024 |  | 01/24/2024 |  |  | 01/08/2024 | 01/24/2024 |  |  | 09/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024398913 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a (b)(6) female patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support developed a mottled leg. patient in worsening shock so impella was escalated to a 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31507 | 1220648-2024-06253 | 18572050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18572050&pc=OZD | 12/30/2023 | 02/23/2024 |  | 01/24/2024 |  |  | 12/30/2023 | 01/24/2024 |  |  | 09/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408940 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a 41-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the impella was implanted, the staff was unable to obtain doppler pulses. the right lower extremity had a monophasic doppler pulse to the posterior tibial, but not dorsalis pedis per the heart team. the impella cp was removed and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31508 | 1220648-2024-06254 | 18572164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18572164&pc=OZD | 01/02/2024 | 10/14/2024 |  | 01/24/2024 |  |  | 01/02/2024 | 01/24/2024 |  |  | 11/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424812 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 59-year-old male patient experienced persistent oozing from their access site following impella cp implant. while the condition was originally managed with manual pressure and femostop, but the bleeding eventually worsened and the decision was made to explant the pump. a balloon tamponade was placed above the access site during impella removal and the site was successfully closed. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the pump was retuned for inspection. visual inspection found no issues on the pump. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06254: d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
31509 | 1220648-2024-06255 | 18572396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18572396&pc=OZD | 01/04/2024 | 02/02/2024 |  | 01/24/2024 |  |  | 01/04/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421411 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp placement was aborted due to the patient having severe tortuosity and aneurysmal aorta. the impella was unable to make turn around the aorta. therefore, the device was removed and the patient was placed on veno-arterial extracorporeal membrane oxygenation (va ecmo). there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the delivery issue was most likely due to patient condition since patient had severe tortuosity and aneurysmal aorta which made the placement unsafe to push the impella.
31510 | 1220648-2024-06257 | 18572823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18572823&pc=OZD | 01/08/2024 | 01/30/2024 |  | 01/24/2024 |  |  | 01/08/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415501 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  | White | Malfunction | No | Yes | Inaccurate Delivery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 77-year-old female patient was attempting to have impella 5.5 support due to the patient having ventricular fibrillation and needing cardio-pulmonary resuscitation. upon attempted implant, there was an inability to deliver the impella into the left ventricle. the implant procedure was therefore aborted. there was no reported patient harm. | [5 empty] | The investigation into the delivery issues has been completed. the pump was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31511 | 1220648-2024-06256 | 18572980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18572980&pc=OZD | 01/11/2024 | 02/28/2024 |  | 01/24/2024 |  |  | 01/11/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024423474 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 65-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. at the time of explant, there was an observation made of thrombus. the thrombus was in the right common femoral artery and was treated by ekosonic endovascular system (ekos) with tissue plasminogen activator (tpa). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported thrombus has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. the instructions for use provides information prevention of thrombus. it states "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ added h6 impact code 4644.
31512 | 1220648-2024-06258 | 18573211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18573211&pc=OZD | 01/08/2024 | 02/23/2024 |  | 01/24/2024 |  |  | 01/08/2024 | 01/24/2024 |  |  | 11/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425852 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 77-year-old female patient developed ischemia. the left femoral repo sheath was occlusive so a right 8f sheath with 4f glide catheter was inserted and advanced distal to the left femoral artery reposition sheath to restore flow. the patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31513 | 1220648-2024-06259 | 18573726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18573726&pc=OZD | 01/09/2024 | 02/16/2024 |  | 01/24/2024 |  |  | 01/09/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421340 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 81-year-old female patient was implanted with an impella cp for hemodynamic support. the impella cp pump was unable to achieve flows greater than 1.8 l/min following the implant. troubleshooting failed to resolve the issue and the decision was made to replace the device with a new pump. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. data logs were reviewed and show no signs on clot ingestion and multiple suction alarms. the clinical review stated that repositioning did not resolve the issue and no cannula kinks or biomaterial were observed. the root cause of the low pump flow cannot be determined as there was no product returned and insufficient clinical evidence. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31514 | 1220648-2024-06260 | 18573931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18573931&pc=OZD | 01/09/2024 | 02/06/2024 |  | 01/24/2024 |  |  | 01/09/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415942 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 77-year-old female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. multiple attempts were made for the impella 5.5 to pass through the innominate artery, but it was unsuccessful. the impella cp was inserted instead. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition since patient had small vessels for impella 5.5.
31515 | 1220648-2024-06261 | 18574274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18574274&pc=OZD | 01/09/2024 | 02/29/2024 |  | 01/24/2024 |  |  | 01/09/2024 | 01/24/2024 |  |  | 09/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398963 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Break | Fluid/Blood Leak | Pumping Stopped |  |  |  |  | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the team had to replace the purge cassette on day seven due to a broken luer. the sidearm was broken when the patient was rolled over to change bed sheets. later, the team noted the motor current began to rise and the pump stopped. the patient also experienced a high grade fever and positive blood cultures. antibiotics were given to the patient. the pump was explanted. | [5 empty] | The investigation into the purge leak - pump, the pump stop, and the infection/sepsis issues has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the pump stop was likely a result of a broken sidearm. however, the root cause of the purge leak ¿ pump and the infection could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31516 | 1220648-2024-06262 | 18574327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18574327&pc=OZD | 10/13/2023 | 11/04/2025 |  | 01/24/2024 |  |  | 10/13/2023 | 01/24/2024 |  |  |  |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while on impella 5.5 s2 support, an optical sensor issue was reported. there was an impella sensor failure alarm on the automated impella controller (aic). placement signal monitoring was disabled. the position of the pump was confirmed with echocardiogram. the pump was not removed due to this issue. there was no reported harm to the patient. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
31517 | 1220648-2024-06264 | 18574494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18574494&pc=OZD | 01/10/2024 | 03/25/2024 |  | 01/24/2024 |  |  | 01/10/2024 | 01/24/2024 |  |  | 09/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400482 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a (b)(6)male (unknown race) noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. during the procedure, the patient developed a hematoma. manual pressure was applied, but bleeding got worse when the pump was removed. upon removal, the perclose sutures that were already placed did not adequately stop the bleeding, so more hands were required to get hemostasis. dry closure - pci access site was on the left side so a peripheral balloon was inflated. ultimately, the patient had a significant amount of bleeding to warrant two units of blood. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the introducer damage ¿ mechanical has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. the root cause of the introducer damage was most likely sheath kink due to patient condition.
31518 | 1220648-2024-06263 | 18574499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18574499&pc=OZD | 01/17/2024 | 09/22/2025 |  | 01/24/2024 |  |  | 01/17/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424215 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/17/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having st elevated myocardial infarction. during the impella implant, the patient went into ventricular fibrillation (vf) and the staff aborted the swan placement. while on support, there was placement signal not reliable alarms. the staff confirmed the pump was in good position and impella support continued. additionally, the patient's urine was blood tinged. heparin was increased to aid with the hemolysis. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis issue has been completed. the impella device has not been returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information was provided. also, the relation to impella is unknown. d4-updated model number to impella cp. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product was not returned for investigation. upon investigating data logs, it was found that placement signal went to 3000 for few hours and returned to normal. there was no change in lamp, light or gain suggesting that there was a possible air gap. therefore, per the clinical information, the root cause of the optical signal failure was most likely an air gap from between the console and the patient cable plug or within the middle of the red handle ferrule. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. d6a/d6b updated with additional information. g1 updated with correct mdr reporting contact email. f6 and f8 should have been left blank. h6 medical device problem code revised to include 2917 device sensing problem and component code 925 pump from the initial manufacturer device report number 1220648-2024-06263.
31519 | 1220648-2024-06265 | 18574552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18574552&pc=PYX | 01/17/2024 | 09/20/2024 |  | 01/24/2024 |  |  | 01/17/2024 | 01/24/2024 |  |  | 06/30/2024 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023221308 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/17/2024 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a (b)(6) asian with right sided heart failure was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella rp support due to decreased right ventricle function. while on support, the patient had hemolyzing labs. the decision was made to remove the impella. however, the patient expired when patient went into pulseless electronic activity and no attempts at cardiopulmonary resuscitation were performed. no allegations were made towards the impella for cause of death. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported hemolysis has been completed since the original report was filed. insufficient data was returned for review. data logs on (b)(6) 2024 showed pump ran without issue. last few hours of the logs are not available. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
31520 | 1220648-2024-06266 | 18574586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18574586&pc=OZD | 01/18/2024 | 03/12/2024 |  | 01/24/2024 |  |  | 01/18/2024 | 01/24/2024 |  |  | 10/31/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024417488 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2024 |  |  |  |  | The us complainant reported the patient was on impella cp support due to having shortness of breath. less than an hour after the implant, the impella had an "air in purge" alarm. staff troubleshot, but the alarm kept appearing. the purge cassette was replaced, however following the replacement, flows appeared to be saturated, high purge pressure alarm ensued, and an impella stopped alarm appeared. the impella would not restart. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop and the purge cassette failure has been completed since the original report was filed. the device was returned and tested after arriving in fs. the device was visually inspected and significant biomaterial ingestion was observed at the purge gap and outflow cage, wrapped around the impeller. reproduction testing was done, and purge flow and pressure were observed to be normal after purging. the cause of the pump stop was most likely use related with unresolved air in purge events leading to lack of purge and chained events per clinical and field investigations.
31521 | 1220648-2024-06267 | 18574738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18574738&pc=OZD | 01/11/2024 | 02/02/2024 |  | 01/24/2024 |  |  | 01/11/2024 | 01/24/2024 |  |  | 11/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024423669 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 70-year old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was attempting to have the impella 5.5 support due to the patient coding and return of spontaneous circulation was achieved. upon attempted delivery, a lot of resistance was met at the take off of the subclavian aorta. a buddy wire was inserted to assist and this was unsuccessful. the 5.5 was removed and the impella cp was placed and it passed into the left ventricle. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the delivery issue was most likely due to patient condition since the size of the vessels were small.
31522 | 1220648-2024-06268 | 18574756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18574756&pc=OZD | 01/11/2024 | 11/19/2025 |  | 01/24/2024 |  |  | 01/11/2024 | 01/24/2024 |  |  |  |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1493604 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 70-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) failed to purge the cassette after being inserted. a new purge cassette was attempted with same issue. the team replaced the aic and was able to purge the cassette. | [5 empty] | The investigation has been completed. data analysis: the console (b)(6) displayed the ¿purge system failure (piston blocked) - alarm issued¿, event # 417 twice on 01-11-2024. console imc logs displayed numerous sau_purge communication issues on both (b)(6) 2024. device analysis: the console was tested after arriving in fs. the purge motor shaft was immobilized, and it can¿t be rotated even when the aic was turned off. there is residue around the opening where the purge motor shaft goes through (and under the gasket), and around the mounting screw for the purge cassette ¿horseshoe¿ retainer. console¿s inability to purge was caused by immobilized purge motor shaft, due to dextrose contamination from some prior purge fluid leak. inspection of the console also revealed persistent contamination of the optical pump connector fiber (unrelated to complaint). d10 concomitant medical products and therapy dates updated. | The impella device was received from the customer and an investigation of the device is ongoing. upon completion of our analysis, a supplemental report will be filed. | D.2 common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04722 was submitted. d.4 model number were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04722 was submitted. e.1 fax number was added as it was indavertently omitted on the initial manufacturer device report number 1220648-2023-04722. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04722 and should not have been. g.1 fax number was added in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04722 was submitted. mdr reporting contact email was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04722 was submitted. h.6 code 2199, 3221 and 4144 were reported incorrectly on the initial manufacturer device report number 1220648-2023-04722. a new code was added.
31523 | 1220648-2024-06269 | 18575523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18575523&pc=OZD | 06/21/2023 | 01/24/2024 |  | 01/24/2024 |  |  | 06/21/2023 | 01/24/2024 |  |  | 02/28/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342228 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that 68-year old male was attempted to be implanted with an impella cp device for mechanical circulatory support. after the impella cp was pulled out of position by the physician, attempts to re-advance the pump resulted in the catheter kinking distal to the outlet. the pump was subsequently explanted as a result of the reported issue. there was no patient harm. | [5 empty] | The investigation into the reported positioning issues has been completed. the device was received for analysis. visual inspection found a kink on the cannula about 14 mm from the of cage. there were no other issues found on the rest of the pump. the pump ran without issue under bench test. the root cause of cannula kink was due to attempt to re-access the valve wirelessly. the cause of the issue was wireless re-access. the impella device is not indicated for wireless re-access.
31524 | 1220648-2024-06270 | 18576950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18576950&pc=OZD | 12/17/2023 | 10/29/2024 |  | 01/24/2024 |  |  | 12/26/2023 | 01/24/2024 |  |  | 09/30/2025 |  | 01/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. abiomed¿s long-term outcome and quality indicator impella registry received a notification regarding a patient that had endured an injury with a "definite" relationship to the impella device. as treatment, a self-expanding covered stent was deployed. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the pseudoaneurysm has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the pseudoaneurysm could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added h6 component code 925 corrections to initial mfr report: d4 model number. added-d6a/d6b. f6/f8 should have been left blank.
31525 | 1220648-2024-06209 | 18562334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18562334&pc=OZD | 09/28/2023 | 01/23/2024 |  | 01/23/2024 |  |  | 09/28/2023 | 01/23/2024 |  |  | 08/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024384880 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. attempted 5.5 implant via right axillary artery. first attempt to pass through subclavian resistance was met. md took the 5.5 out and performed angioplasty of right subclavian artery with a 8 x 60 mm mustang balloon. impella 5.5 was attempted again but would not pass. decision was made to abort 5.5 implant and proceed with impella cp through graft. | [5 empty] | The impella device was not received from the customer, therefore no evaluation could be completed. if a device is received a follow up mdr will be submitted. the root cause of the delivery issue - anatomy was most likely due to patient condition related due to calcified vessel condition.
31526 | 1220648-2024-06210 | 18562408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18562408&pc=OZD | 09/28/2023 | 01/23/2024 |  | 01/23/2024 |  |  | 09/28/2023 | 01/23/2024 |  |  | 07/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024371626 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella 5.5 came out of lv (left ventricle) while removing wire. physician was unable to advance back and decision was made to opt for an impella cp device which was successful. no known patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore no evaluation could be completed. if a device is received a follow up mdr will be submitted. review of the clinical notes revealed that the pump was unable to pass through the patient's vasculature as the pump was placed without the guidewire. no product damage was reported. the cause of the issue was wireless insertion. the impella device is not indicated for wireless insertion per the ifu 10003049.
31527 | 1220648-2024-06211 | 18562658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18562658&pc=OZD | 10/22/2023 | 01/23/2024 |  | 01/23/2024 |  |  | 10/22/2023 | 01/23/2024 |  |  | 09/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024395873 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/22/2023 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella was attempted for insertion and after multiple attempts the physician was unable to cross aortic valve with impella. no known patient harm or injury. | [5 empty] | No product was returned for evaluation. as per clinical physician was unable to cross the aortic valve during impella 5.5 insertion after multiple attempts with different catheters and wires. patient had aortic valve stenosis with no movement at aortic valve noted per tee imaging. the cause of the delivery issue was patient condition related with patient having aortic valve stenosis and pump unable to advance aortic valve per clinical review.
31528 | 1220648-2024-06212 | 18563797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18563797&pc=OZD | 10/03/2023 | 02/21/2024 |  | 01/23/2024 |  |  | 01/04/2024 | 01/23/2024 |  |  | 10/31/2024 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2023268821 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | No Visual Prompts/Feedback | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 86-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was connected to the automated impella controller (aic) but the screen would not advance to start. therefore, a second device was used, which worked. there was no known harm or injury and no known patient use. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported pump not detected has been completed since the original report was filed. device was not retuned for inspection. no data logs were created as the device was not able to connect to the console. the cause of this issue can not be determined as the pump was discarded in the field.
31529 | 1220648-2024-06213 | 18563869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18563869&pc=OZD | 01/04/2024 | 02/21/2024 |  | 01/23/2024 |  |  | 01/04/2024 | 01/23/2024 |  |  | 10/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024420754 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during insertion of the impella cp the cannula was kinked just below the motor. several unsuccessful attempts were made to straighten it out with flows being no higher than 1.5. the cp was explanted and a new cp successfully placed. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported product damage has been completed since the original report was filed. according to the clinical, while inserting the impella cp device, the cannula was kinked just below the motor. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the low flow was due to a cannula kink during insertion.
31530 | 1220648-2024-06214 | 18564117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564117&pc=OZD | 01/04/2024 | 01/31/2024 |  | 01/23/2024 |  |  | 01/04/2024 | 01/23/2024 |  |  | 11/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024426393 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was attempting to have an impella cp implant due to high risk percutaneous coronary intervention of left main coronary artery branch. upon attempted implant, the implant was paused due to the patient's lack of airway. patient had a difficult airway, pulmonary hemorrhage. cpr was ongoing and multiple rounds of advanced cardiovascular life support. the patient expired while trying to implant impella for bailout. there were no allegations towards the impella for cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. the medical center did not return the impella device. the root cause of the delivery issue was most likely due to patient condition since the patient had lack of airway and pulmonary hemorrhage.
31531 | 1220648-2024-06215 | 18564122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564122&pc=OZD | 01/04/2024 | 12/09/2024 |  | 01/23/2024 |  |  | 01/04/2024 | 01/23/2024 |  |  | 10/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409550 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to extreme calcification in multiple vessels, especially left main and patient suddenly decompensated. while on support, the patient had hemolysis. the urine was dark red. one unit of blood products were given and fluids. the clincal rep. confirmed the pump was repositioned and the hemolysis certainly appeared unchanged. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ . ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. | The investigation for the hemolysis has been completed. product was not returned. data logs show suction alarms throughout support. there are no motor current (mc) spikes outside of p level changes and no other abnormalities. the cause of the hemolysis was not determined since product was not returned, lack of clinical details, and inconclusive data logs. added h6 impact code 4628, h6 component code 925 corrections to the initial mfr report: d4 model number added d6a/d6b f6/f8 should have been left blank. g1-mdr reporting contact email.
31532 | 1220648-2024-06216 | 18564268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564268&pc=OZD | 01/05/2024 | 03/11/2024 |  | 01/23/2024 |  |  | 01/05/2024 | 01/23/2024 |  |  | 08/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024387399 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp implant procedure for 72-year-old male patient, the pump would not start. the automated impella controller (aic) displayed 'defective impella catheter' prompting new device to be used. the second impella was placed without issue. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported pump stop has been completed. the impella device and data logs were returned for evaluation. upon testing of the pump, an impella defective alarm was reproduced and the pump was unable to be detected to extract the eeprom values. data logs were reviewed which also revealed eeprom problems. the cause of the pump not detected issue was not determined due to inconclusive evidence based on field investigation and log review. added-d9-device return date h6 problem code updated to 1503.
31533 | 1220648-2024-06217 | 18564284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564284&pc=OZD | 01/05/2024 | 02/23/2024 |  | 01/23/2024 |  |  | 01/05/2024 | 01/23/2024 |  |  | 10/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420560 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 73-year-old male patient was on impella cp support due to having non-st elevated myocardial infarction and heart failure. while on support, the patient was having absent pulses in the right lower extremity (impella side). the impella was removed and the patient was escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31534 | 1220648-2024-06218 | 18564498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564498&pc=OZD | 01/05/2024 | 09/09/2024 |  | 01/23/2024 |  |  | 01/05/2024 | 01/23/2024 |  |  |  |  | 01/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1383042 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the patient was on impella support and the automated impella controller (aic) would not purge. the aic was therefore replaced and the affected aic was sent in for investigation. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue has been completed. the logs revealed that a purge disc not detect had occurred, and that priming was unable to commence. during testing, the console was able to accurately detect the presence of the purge disc. the purge flag was confirmed to be operational and non-defective. testing with a pressure calibrator also confirmed that the purge pressure sensor accurately measured purge pressure. however, disassembly of the console¿s purge assembly revealed the dried purge fluid between the surfaces of the purge motor gasket and the interior of the purge receptacle. this dried purge fluid was determined to most likely have been caused by purge fluid ingress. the root cause was due to buildup of purge fluid on the purge stepper motor gasket. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06218. g1 reporting contact fax number missing. h6 4642 was reported incorrectly.
31535 | 1220648-2024-06219 | 18564548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564548&pc=OZD | 12/25/2023 | 03/06/2024 |  | 01/23/2024 |  |  | 01/05/2024 | 01/23/2024 |  |  | 10/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409945 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a patient escalated from cp to the 5.5 pump, which supported the patient for 24 days. the pump may have caused severe ar. the cause of the ar (aortic regurgitation) was not known to the team. the plan was to transfer the patient out to a tertiary center for an avr (aortic valve replacement) to be performed. the causal relationship is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the reported heart valve damage has been completed. the impella device was not returned for evaluation. data logs were reviewed which did not reveal any evidence of improper placement of pump across valve, suction alarms, or incorrect placement alarms. the root cause of the heart valve damage cannot be determined as no evidence suggests that the pump caused any damage.
31536 | 1220648-2024-06221 | 18564637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564637&pc=OZD | 01/08/2024 | 07/16/2024 |  | 01/23/2024 |  |  | 01/09/2024 | 01/23/2024 |  |  | 10/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420781 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing when there was an observation made of consistent (cp was hemolyzing on the unit after being brought from the lab. crrt got initiated as a result. hemolysis persisted despite repositioning. pump kept in place because pt needed flow and not a candidate for anything else.) , and unable to troubleshoot. crrt (continuous renal replacement therapy) was begun and care was withdrawn when the patient was found to not be a candidate for transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. the device, data logs, nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the following sections: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ d4-corrected model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
31537 | 1220648-2024-06220 | 18564650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564650&pc=OZD | 01/06/2024 | 02/07/2024 |  | 01/23/2024 |  |  | 01/06/2024 | 01/23/2024 |  |  | 09/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024422952 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 49-year-old female patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma and a retroperitoneal bleed which was resolved by interventional radiology with coiling. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. based on the clinical information provided, the cause of the vascular damage issue was most likely use related and unknown relation to impella per clinical review.
31538 | 1220648-2024-06222 | 18564660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564660&pc=OZD | 01/09/2024 | 10/10/2024 |  | 01/23/2024 |  |  | 01/09/2024 | 01/23/2024 |  |  | 05/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374841 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 80-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that hrpci (high risk percutaneous coronary intervention) was not performed due to left ventricle (lv) perforation post impella cp insertion. impella loaded onto the 0.18 wire and positioned in the lv without issue. impella started and flows would not increase past p-5. lv waveform extremely negative and suction alarms were present. volume given and issues persisted. echo - echocardiogram (ultrasound) - demonstrated a large pericardial effusion. the patient was prepped for pericardial tap and impella cp weaned down for explant. the impella was explanted without any issues. the patient was intubated in the cath lab and opened sternotomy pericardial window performed. multiple blood products were given to support the patient. during impella insertion the physician felt some tension on the wire as he was crossing into the lv, but everything appeared fine on fluro. suspects this could be how the lv perforation occurred even though best practices were utilized during implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. no product was returned for analysis. the root cause of the ventricle perforation is most likely due to guidewire puncture as the clinical stated that the physician felt resistance on the wire when crossing the lv and thought that is how the puncture may have occurred. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06222. b2 death (patient was in critical condition) was inadvertently selected, changed to required intervention to prevent permanent impairment/damage. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h5 labeled for single use changed to yes. h6 component code revised from the initial submission in accordance with updated procedures. h6 code 2135 was reported incorrectly. new codes were added to health effect - clinical code and health effect - impact code. instructions for use: impella cp with smartassist system use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system use during cardiogenic shock. section: warnings and cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings and cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
31539 | 1220648-2024-06223 | 18564688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564688&pc=OZD | 01/11/2024 | 02/29/2024 |  | 01/23/2024 |  |  | 01/11/2024 | 01/23/2024 |  |  | 04/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024348479 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient experienced bleeding from their right groin during escalation to impella 5.5 support. four units of blood were transfused to counteract the blood loss. support continued with the newly implanted impella 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. per the provided clinical details, the root cause of the access site bleeding issue was most likely due anticoagulation management since reversing the protamine resolved the bleeding as no bleeding was reported in the following days after reversal of protamine.
31540 | 1220648-2024-06224 | 18564793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564793&pc=OZD | 01/12/2024 | 02/20/2024 |  | 01/23/2024 |  |  | 01/12/2024 | 01/23/2024 |  |  | 05/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024358845 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 60-year-old male patient of unknown race and ethnicity with a percutaneous impella cp implantation after cornary artery bypass graft (cabg) surgery for mechanical support. it was reported that the patient developed leg ischemia following iliac vascular injury. the patient was taken to the operating room and the impella was exchanged using surgical axillary approach followed by the repair of the iliac artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: limb ischemia. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. vessel damage: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta. | The investigation into the limb ischemia ¿ reversible and the thrombus issue has been completed since the original report was filed. the product was returned for investigation. the device was visually inspected, and no sharp edges or burrs were found that may promote thrombus growth. as the necessary clinical information was not provided, the root cause of the limb ischemia and the thrombus was not determined.
31541 | 1220648-2024-06225 | 18564809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564809&pc=OZD | 01/12/2024 | 05/15/2024 |  | 01/23/2024 |  |  | 01/12/2024 | 01/23/2024 |  |  |  |  | 01/23/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 2024385956 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 68-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the staff had to de-air the purge line on two separate occasions during patient support due to the console not alerting them of low remaining purge volume. the automated impella controller (aic) was swapped because of the noted issue. there was no patient harm reported. | [5 empty] | The investigation for the reported console (aic) purge issue has been completed. the console was returned for evaluation and functional testing was performed. the console takes the user inputs for the bag size, and based on the remaining reservoir volume, it alarmed either ¿purge volume low¿ or ¿purge volume critically low¿ as intended. the purge pressure sensor accuracy does not meet the criteria. no issue found with the console hardware. unable to reproduce the reported failure mode. data logs were reviewed finding the remaining purge volume was 430.6 ml (the corresponding threshold limit was only 75 ml), the console still displayed ¿air in purge system ¿ alarm,¿ which was most likely due to leakage in the bag or cassette, and the bag might be dry. this aligns with the claim that ¿the purge bag and line has run dry,¿ the root cause of the purge bag and line has ran dry was most likely due to the leakage in the bag or the cassette. though the remaining purge volume does not dropped below the threshold limit, the bag ran dry, most likely making the customer report that the ¿console did not alarm that purge volume was low.¿ in accordance with updated procedures, section d2 has been revised from the initial submission. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31542 | 1220648-2024-06226 | 18564938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564938&pc=OZD | 01/10/2024 | 12/17/2024 |  | 01/23/2024 |  |  | 01/10/2024 | 01/23/2024 |  |  | 10/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415499 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 being placed for support via axillary cut-down for supporting the patient with cgs (cardiogenic shock) and cardiomyopathy. the pump had alarms on the aic (automated impella controller) that prompted the console to be replaced, but that did not resolve the optical signal issue. the pump was replaced and support proceeded. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the optical signal issue has been completed. product and logs were not returned. the root cause of the optical signal issue was not determined since product/logs were not returned and insufficient clinical evidence was provided. corrections to the initial mfr report: d4 model number updated. added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email. h6 impact code added 4629.
31543 | 1220648-2024-06227 | 18564977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564977&pc=OZD | 01/12/2024 | 11/21/2025 |  | 01/23/2024 |  |  | 01/12/2024 | 01/23/2024 |  |  |  |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1590813 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/12/2024 |  |  |  |  | No patient involvement: the complainant reported a battery failure on the aic (automated impella controller). the aic was removed from service. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. a.1 added as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06227. a.2 age should have been left blank on the initial manufacturer device report number 1220648-2024-06227 as no patient was involved. a.3 added as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06227. d.2 common device name was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06227 was submitted. d.4 model number was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06227 was submitted. catalog number was revised since the initial manufacturer device report number 1220648-2024-06227 was submitted. d.9 device return date was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06227. e.1 revised name prefix/title, first and last name as they were reported incorrectly on the initial manufacturer device report number 1220648-2024-06227. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-06227 and should not have been. g.1 revised mdr reporting contact email in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06227 was submitted. added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06227. h.3 updated to device evaluation anticipated as it was reported incorrectly on the initial manufacturer device report number 1220648-2024-06227. h.6 code 2199, 4144 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2024-06227. h.10 additional manufacturer narrative was reported incorrectly on the initial manufacturer device report number 1220648-2024-06227 as the console was returned.
31544 | 1220648-2024-06228 | 18564998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18564998&pc=OZD | 01/10/2024 | 12/24/2024 |  | 01/23/2024 |  |  | 01/10/2024 | 01/23/2024 |  |  | 11/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430024 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella insertion the impella appeared to be caught under the papillary muscle and the patient was experiencing sustained ectopy. the impella attempted to be pulled back from the left ventricle but jumped across aortic valve (av) and could not be re-advanced across av. the physician removed and rewired the left ventricle however the impella would not advance into position past av. the cannula appeared to prolapse and kink. the impella was removed and a new one placed successfully. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation positioning issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the cause of the vt/vf was not determined as the necessary clinical information was not provided. the cause of the positioning issues was most likely use related based on clinical. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06228 b1 adverse event and product problem: d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing. h6 medical device problem code incorrect. new medical device problem code added.
31545 | 1220648-2024-06229 | 18565006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18565006&pc=PYX | 12/25/2023 | 06/11/2024 |  | 01/23/2024 |  |  | 01/12/2024 | 01/23/2024 |  |  | 02/28/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024348531 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Thrombocytopenia | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [5 empty] | IMPELLA 5.5 | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 01/12/2024 |  |  |  |  | The user facility reported a 41-year-old male patient implanted with an impella rp as a bridge to transplant. it was reported that after 42 days it was noticed the pump flows began to decrease. the heparin was increased, however the physician decided to explant the impella rp pump. there was no patient harm reported. | New information received via loqi study indicates an allegation of hemolysis as a result of the reported device malfunction. both device malfunction and hemolysis lead to removal of device. | An additional notification was received via abiomed¿s long-term outcome and quality indicator impella registry, indicating a "possible" relationship between the device and thrombocytopenia. unspecified medical/surgical intervention was reported. no further information was provided. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issues has been completed since the original report was filed. the pump was returned for investigation and biomaterial was found in the inflow cage and wrapped around the impeller. the root cause of the low pump flow was determined to be biomaterial. the root cause of the hemolysis was determined to be biomaterial. the root cause of thrombocytopenia was unable to be determined as limited clinical details were provided. | The investigation into the low pump flow and the hemolysis issues has been completed since the original report was filed. the device was returned and inspected. biomaterial was found in the inflow cage and wrapped around the impeller. the root cause of the low pump flow was determined to be biomaterial. the root cause of the hemolysis was determined to be biomaterial. | B(5): additional information added to report new information of thrombocytopenia. upon completion of our analysis in response to the reported thrombocytopenia, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the product was returned for investigation. data logs show a high motor current spike followed by low flows. there is also a placement signal shift and a speed deviation at the same time as the motor current spike. investigation of the returned product found biomaterial in the inflow cage and wrapped around the impeller. the root cause of the low pump flow was determined to be biomaterial.
31546 | 1220648-2024-06230 | 18565045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18565045&pc=OZD | 01/10/2024 | 01/23/2024 |  | 01/23/2024 |  |  | 01/10/2024 | 01/23/2024 |  |  | 11/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024425277 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Computer Operating System Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Vascular Dissection | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. console alarm "impella defect". new pump and new console solved the issue. | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to left ventricular outflow tract obstruction. the 5.5 was implanted via percutaneous axillary approach. it is suspected that the perclose failed when attempting to tighten around catheter shaft causing dissection of axillary artery. emergency axillary cutdown was required to surgically repair axillary artery followed by ballooning intervention of artery to restore distal limb perfusion. 16 units of blood products were provided. additionally, the patient had loss of distal limb perfusion on the right arm. a penumbra procedure was performed with successful clot removal and restoration of right brachial and right ulnar pulses. signs of rigor mortis to hand were noted. vascular team was consulted. multiple blood products were required. reportable due to vascular damage, limb ischemia and thrombus and necessary intervention. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Section b5, updated to say: "the user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to left ventricular outflow tract obstruction. the 5.5 was implanted via percutaneous axillary approach. it is suspected that the perclose failed when attempting to tighten around catheter shaft causing dissection of axillary artery. emergency axillary cutdown was required to surgically repair axillary artery followed by ballooning intervention of artery to restore distal limb perfusion. 16 units of blood products were provided. additionally, the patient had loss of distal limb perfusion on the right arm. a penumbra procedure was performed with successful clot removal and restoration of right brachial and right ulnar pulses. signs of rigor mortis to hand were noted. vascular team was consulted. multiple blood products were required." section h6, updated to say: 4440 - thromb / 4624 - surgical intervention / 2993 - adverse event without identified device use 1942 - ischemia / 4643 - blood transfusion. 3160 - vascular dissection.
31547 | 1220648-2024-06231 | 18565166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18565166&pc=OZD | 01/10/2024 | 12/03/2025 |  | 01/23/2024 |  |  | 01/10/2024 | 01/23/2024 |  |  |  |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1417342 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/10/2024 |  |  |  |  | The user facility reported that during in lab training the automated impella controller (aic) did not recognize the demonstration-pump. several attempts to reboot were made without success. the console was exchanged to the backup and the issue resolved. there was no reported patient harm or involvement. | [5 empty] | The impella console has been received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g.1. mdr reporting contact email. d.2. device name has been updated. f.6. and f.8. should have been left blank. h.6. type of investigation code added.
31548 | 1220648-2024-06233 | 18565263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18565263&pc=PYX | 01/11/2024 | 02/28/2024 |  | 01/23/2024 |  |  | 01/11/2024 | 01/23/2024 |  |  | 05/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024346896 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 73-year-old male in st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. it was reported that that the impella rp had an immediate stop. the team attempted to restart and troubleshoot the stoppage. the pump was replaced with a new rp pump and was successfully implanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into the reported pump stop has been completed since the original report was filed. no product was returned. data logs were reviewed, and high motor current pump stop was observed immediately within 1 minute of case start. the root cause of the pump did not start issue was not determined since product was not returned and data logs were inconclusive.
31549 | 1220648-2024-06232 | 18565268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18565268&pc=OZD | 01/11/2024 | 03/12/2024 |  | 01/23/2024 |  |  | 01/11/2024 | 01/23/2024 |  |  | 04/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024366068 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | Complainant in japan reported the patient was on impella cp due to acute myocardial infarction/ventricular fibrillation. while on support, the patient had chest compressions and open chest hemostasis was required. the impella moved into the aorta side and the impella was replaced due to the malposition. reportable due to positioning issues and necessary intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
31550 | 1220648-2024-06234 | 18565299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18565299&pc=OZD | 01/12/2024 | 02/29/2024 |  | 01/23/2024 |  |  | 01/12/2024 | 01/23/2024 |  |  | 10/31/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024411458 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | Pressure Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 50-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to being escalated from the impella cp to the 5.5. the impella had a high purge pressure alarm after 15 minutes. impella was inserted and initiated before the cassette could be changed. eventually, there was a purge pressure blocked alarm. there were low flows that were not resolved with repositioning. the impella was replaced successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the reported high purge pressure and low pump flow has been completed since the original report was filed. impella 5.5 pump and secondary purge cassette were returned for evaluation due to high purge pressure alarms and low pump flow. upon visual inspection, a clot was found in the outflow cage of the pump. the pump was placed in a bucket of water and primed with purge cassette. a gradual rise in purge pressure was observed while no purge fluid was seen flowing out of the purge gap. once pump was turned on to p-5, the biomaterial was kicked out of the outflow cage and purge pressure returned to normal values. pump was then turned to p-9 to test for low flow issues and was found to have slightly saturated values that returned to within range as biomaterial on the impeller was kicked off. the pump with returned purge cassette was run a second time with biomaterial cleared out and no failures were observed when running. the data logs were reviewed and showed that the high purge pressure was triggered while outside of the body before insertion. the high purge pressure alarms are likely due to an issue with the primary purge cassette, however, since the first purge cassette was not returned and the alarms could not be reproduced, the root cause cannot be determined. the root cause of the low pump flow was likely due to biomaterial build up at the impeller and purge gap.
31551 | 1220648-2024-06235 | 18565332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18565332&pc=OZD | 01/12/2024 | 09/21/2024 |  | 01/23/2024 |  |  | 01/12/2024 | 01/23/2024 |  |  | 11/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMEPLLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024429236 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 01/12/2024 |  |  |  |  | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 63-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump stopped during patient support. based on the provided information, support continued with the implanted pump following the temporary stop and there was no noted patient harm. | Additional information was obtained providing clarification regarding the event. it was noted that upon arrival to the icu following impella cp implant, purge flows were lower than expected. discussions were initiated on how to counteract the issue; however, the flow rate continued to drop during this time. a tpa protocol was submitted to the pharmacy, but the pump stopped while awaiting delivery. the device was explanted and the patient was reported to be stable following the event. there was no patient harm. |  |  |  | The investigation into the pump stop issue has been completed. the device was not returned for investigation. data log analysis confirms motor current spikes, high purge pressure, and a low purge flow. based on the data logs, the most likely cause of the pump stop was due to high purge pressure. added h6-impact code 4644. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | Sections b5 and h6 updated to coincide with the additional information that was received. the device has not yet been received for evaluation. upon investigation closure, a supplemental report will be submitted.
31552 | 1220648-2024-06236 | 18565415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18565415&pc=OZD | 01/16/2024 | 03/18/2024 |  | 01/23/2024 |  |  | 01/16/2024 | 01/23/2024 |  |  | 11/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024432624 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2024 |  |  |  |  | The user facility in puerto rico reported a 83-year-old undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella procedure a suction alarm occurred. volume and other troubleshooting measures were unsuccessful. the case was aborted. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the device was not returned. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
31553 | 1220648-2024-06237 | 18567802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18567802&pc=OZD | 12/25/2023 | 12/11/2024 |  | 01/23/2024 |  |  | 12/25/2023 | 01/23/2024 |  |  | 09/30/2025 |  | 01/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410752 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 59 YR |  | Male |  |  |  | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2023 |  |  |  |  | The user facility reported a 59-year-old male admitted with angina and vt/vf arrest, in acute myocardial infarction/cardiogenic shock, was implanted with an impella cp device and an impella rp device for mechanical circulatory support. while on support, the patient was maxed out on all pressors when the impella cp exhibited persistent suction alarms and lost aortic and motor current pulsatility. the team began cpr with unsuccessful return of spontaneous circulation and the patient expired. the low flow was not resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the low pump flow has been completed. product not was returned. data logs show few motor current (mc) spikes with drop in flow and suction alarms. overall the ps and flow are trending downwards. clinical details state that the patient was said to be deteriorating and maxed out on all pressors. the cause of the low flow issue was most likely patient condition related. the cause of the ventricular tachycardia (vt)/ventricular fibrillation (vf) could not be determined without additional clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: b1-should have selected product problem in addition to adverse event. b2-should have selected death. d4 model number. added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email.
31554 | 1220648-2024-06168 | 18551532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18551532&pc=OZD | 01/03/2024 | 02/26/2024 |  | 01/22/2024 |  |  | 01/03/2024 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1424423 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Erratic or Intermittent Display | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It has been reported that the automated impella controller showed a flickering display. the controller has been returned to the manufacturer for investigation. | [5 empty] | The automated impella controller (aic) device was received from the customer but has not yet been evaluated. upon completion of the evaluation , a supplemental mdr will be filed. | The investigation into the reported aic ¿ shutdown or restart issue has been completed since the original report was filed. the device failure was reproduced during testing (ic7392 flickering.mp4). the lcd screen was replaced with a known good but there were still remnants of flickering. the 4 in 1 was also replaced and the flickering went away. the cause of the issue was defective 4 in 1 and lcd screen.
31555 | 1220648-2024-06169 | 18551572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18551572&pc=OZD | 01/04/2024 | 02/07/2024 |  | 01/22/2024 |  |  | 01/04/2024 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1571192 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 75-year-old male patient was implanted with impella cp for mechanical circulatory support. however, automated impella controller (aic) was switched on, 'problem to the battery' red alarm and 'low charge battery' alarms appeared. the device was switched for another console and the procedure was completed successfully. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console battery issues has been completed. the console was returned for evaluation. upon testing, a single cell battery failure occurred on battery 1 cell 3 that was reproduced. logs were reviewed which showed that battery 1 had a cell imbalance on cell 3 of 4000 mv. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure.
31556 | 1220648-2024-06170 | 18551649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18551649&pc=OZD | 12/30/2023 | 03/11/2024 |  | 01/22/2024 |  |  | 01/05/2024 | 01/22/2024 |  |  | 05/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024351569 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported patient of unknown age and gender on impella cp support, experienced 'purge pressure high' alarm 15-20 minutes after it started running. purge system was exchanged but issue was not resolved. alarm sounded for 64 hours and decision was made to wean patient off pump. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported high purge pressure has been completed. the impella device was returned for evaluation. data logs were not returned. product evaluation confirmed no evidence of biomaterial, kinks, or other abnormalities. the pump was then run in the lab and the high purge pressure alarm could not be reproduced. the cause of the high purge pressure was not determined because no logs were returned for analysis and not enough clinical information was given. the instructions for use states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ added: d9-device return date h6-impact code 4644 h6-problem code updated to 1491.
31557 | 1220648-2024-06171 | 18551833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18551833&pc=OZD | 01/05/2024 | 02/08/2024 |  | 01/22/2024 |  |  | 01/05/2024 | 01/22/2024 |  |  | 10/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024414797 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported patient on impella cp support experienced 'low purge pressure' alarm. the purge cassette was changed but issue did not resolve. a leak was noted on the purge side arm luer lock. the problem resolved after purge bypass. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak ¿ pump issue has been completed. the device was not returned for investigation. a purge leak was found at the side arm yellow luer lock. it did not resolve with purge cassette change and a purge bypass was performed. low purge pressure alarm was resolved. no data logs were returned for evaluation. the root cause of the purge leak was likely due to yellow luer damage and it was resolved with bypass. the cause of the yellow luer damage is unknown as the product was not returned.
31558 | 1220648-2024-06172 | 18551932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18551932&pc=OZD | 01/06/2024 | 09/20/2024 |  | 01/22/2024 |  |  | 01/06/2024 | 01/22/2024 |  |  | 10/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419323 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 56-year-old female patient on impella cp support, developed hemolysis and had mottling of the left foot indictive of ischemia. unable to get left distal pulses. complications were not successfully managed so the plan is to explant the impella. however, patient expired on support 3 days post implant procedure. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06172 was provided in sections b2, b5, d6a, d6b, h1, and h6.
31559 | 1220648-2024-06173 | 18552181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552181&pc=OZD | 01/08/2024 | 11/10/2025 |  | 01/22/2024 |  |  | 01/08/2024 | 01/22/2024 |  |  | 02/28/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342194 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a ps issue observed at the time of wean and explant. the lv tracing for edp and ps/flow were not correlating. there was no harm or injury or intervention needed. | [5 empty] | Data review: data showed auto suction response when pump start but flow was in specification. no alarm related to flow issue were found. the rt logs showed more detail of flow with fluctuation indicated that the patient's volume in ventricle was unstable. all 5s parameter was in specification. product analysis: visual inspection found no issues on the pump. pump ran without issue under bench test. the root cause of flow issue was patient condition
31560 | 1220648-2024-06174 | 18552453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552453&pc=OZD | 05/22/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 05/22/2023 | 01/22/2024 |  |  | 12/31/2024 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325649 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 56-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support.. the us complainant had a cp support ongoing when there was high purge pressure alarms and purge system blocked. the team replaced the pc and support proceeded. the team later removed the pump. the cp is not a life sustaining/full support pump. there was no reported harm or injury. | [5 empty] | The investigation into the reported product damage has been completed. returned pump was evaluated for the high purge pressure issue. during the evaluation, biomaterial was discovered in the purge gap and on the impeller. several alarms were triggered, indicating issues with the purge system being blocked and experiencing low/high volume. furthermore, during the observation, a kink was noticed in the catheter, and as a response, the catheter was cut open at the site of the kink. however, even after this intervention, the catheter kink failed to reproduce a high purge pressure. per data log review, suction and a motor current spike were observed prior to the purge pressure spike indicating that biomaterial ingestion occurred prior to the catheter kink and was the root cause of the high purge pressure. the evaluation revealed that the root cause of the issue was biomaterial found in purge gap. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31561 | 1220648-2024-06175 | 18552473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552473&pc=OZD | 05/28/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 05/28/2023 | 01/22/2024 |  |  | 03/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328766 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/28/2023 |  |  |  |  | The user facility reported a 62-year-old white male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the impella cp support was delayed as the guidewire would not deliver with the pump. the delay was when the cp was replaced and new cp placed. no harm or injury to the patient. | [5 empty] | The investigation in the delivery issues malfunction has been complete. product was not returned for review. the root cause of delivery issue was unable to be determined as limited clinical details were provided and no product was returned for review.
31562 | 1220648-2024-06176 | 18552490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552490&pc=OZD | 06/07/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 06/07/2023 | 01/22/2024 |  |  | 03/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024329821 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/07/2023 |  |  |  |  | The user facility reported a patient (unknown age, gender and race) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp optical had loss of the signal but, the team left the support ongoing for just under one day when then explanted. there was no harm or injury. | [5 empty] | The investigation into the optical signal issue has been completed. the root cause of the optical signal issue was not determined since no data logs were returned and the returned product was investigated but no abnormalities relative to optical signal issue were found during evaluation.
31563 | 1220648-2024-06177 | 18552530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552530&pc=OZD | 01/02/2024 | 10/27/2025 |  | 01/22/2024 |  |  | 01/02/2024 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1319386 |  | IC5633 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | Device Sensing Problem | Electrical Power Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported "controller failure" leading to console being exchanged. there was no reported patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: the last case performed at the customer site was from on (b)(6) 2023. imc log shows that the first occurrence of controller error (opm comm crc invalid) [optical pump connected] event set #1045 ¿ alarm on (b)(6) 2023, which resolved in 1 minute 26 seconds. the second occurrence of #1045 was on (b)(6) 2024, which matches with the reported complaint date and confirms the reported failure mode (it seems like there was a typo in the description, assuming ¿control failure¿ as ¿controller failure¿), following the second occurrence of controller error, the event set #1045 cleared in 5 minutes 46 seconds and reoccurred for the third time within a second and revolved in 1 minute 41 seconds. in addition to that, the console received the below alarms, controller error (optical bench ambient pressure out of range) [optical pump connected] ¿ alarm, event set #1040 placement signal not reliable (optical signal-to-noise ratio) ¿ alarm, event set #407 real-time (rt) logs show that all signals received from the optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. while the psnr alarm was on, the snr fluctuated between 0 and 5. see the attachments section for ic5633 imc and rt logs. device analysis: the console was tested on an engineering lab bench. the reported failure mode(fm) controller error and the additional fm psnr were not reproduced by running the console with the optical pump. the fringe pattern looks good, no debris was found in the front panel optical connector, and there was no issue with the console hardware. the root cause for the controller error was not determined due to the inability to reproduce. the root cause for the placement signal not reliable was not determined due to the inability to reproduce. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: updated to include placement signal coding. b7 other relevant history, including preexisting medical conditions updated. d2: common device name updated. d4: model number and serial number updated. d10: concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller failure (red alarm) has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure was not reproduced. the root aic - controller failure (red alarm) was not determined due to the inability to reproduce.
31564 | 1220648-2024-06179 | 18552588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552588&pc=OZD | 01/03/2024 | 10/30/2025 |  | 01/22/2024 |  |  | 01/03/2024 | 01/22/2024 |  |  | 03/31/2024 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023188227 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported male patient had digestive bleeding that was managed with low heparin. impella 5.0 was placed for support but experienced high purge pressure episodes that occurred within the first 48 hours post-insertion. the high purge pressure was managed, however, the pump was ultimately explanted. a new pump was used. there was no reported patient harm. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the high purge pressure cannot be determined as no product or data logs were returned, and clinical troubleshooting did not resolve the issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation of high purge pressure has been completed. the purge cassette and the pump were returned for investigation. there were no cannula or catheter kinks, no biomaterial or foreign material deposits around the impeller cage or inside the cannula. the high purge pressure was not reproduced when the pump was purged in the lab, with purge pressure stabilizing around the 750 mmhg mark. an attempt was made to run the pump but high motor current alarms were triggered and the pump did not start. upon further inspection, it was found that the impeller blades were not freely rotating as they normally would, the cause of which was revealed to be biomaterial stuck in the purge gap, on the impeller shaft, and between the impeller shaft and motor wall. a small cut was made in the catheter near the motor to see if there was any blood ingress in the purge line, but no blood was found. no data logs were returned for review. clinical details note that heparin was used in purge solution, lysis protocol was used, and p-levels being dropped and ramped up again, and none of these measures were able to fully resolve the high purge pressure alarms. high purge pressure: the root cause of the high purge pressure was biomaterial in the purge gap and the motor based on the product evaluation. it was not able to be determined whether it was biomaterial buildup or ingestion due to lack of data logs for analysis. a thrombus failure mode was also investigated in case of biomaterial ingestion. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added d6a and d6b revised from the initial manufacturer device report 1220648-2024-06179 in accordance with updated procedures g1 reporting contact email revised on manufacturer device report 1220648-2024-06179 in accordance with updated procedures.
31565 | 1220648-2024-06178 | 18552744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552744&pc=OZD | 01/08/2024 | 02/21/2024 |  | 01/22/2024 |  |  | 01/08/2024 | 01/22/2024 |  |  | 06/30/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024383018 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Stretched | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported after pushing the impella cp in the left ventricle, the guidewire was pulled. at the end, about 80% was pulled back. due to the pressure to pull/raise, the guidewire has "unwound" style and appeared to be getting "longer?". there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the product damage (guidewire damage - great vessel) issue has been completed since the original report was filed. the product was returned for investigation. the guidewire tip was found uncoiled/starched and partially damaged. additionally, a kink was found on the returned introducer at sheath tip. the cause of the guidewire damage issue was not determined since sufficient clinical information was not provided and due to inconclusive evidence from product investigation.
31566 | 1220648-2024-06180 | 18552769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552769&pc=OZD | 06/10/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 06/10/2023 | 01/22/2024 |  |  | 02/28/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341868 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/10/2023 |  |  |  |  | The user facility reported a 38-year-old male post cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the impella cp had high purge pressure alarms within five minutes of initial impella insertion. purge tubing and consoles changes were made with no resolution. new impella catheter implanted. five minutes after insertion of new device, the automated impella controller displayed aic purge pressure failure and then the motor stopped. pump successfully restarted after multiple attempts. after successful pump restart no issues or alarms noted. md aware and decision to leave catheter in on support. | [5 empty] | The investigation into the pump stop has been completed. based on the clinical information provided, the root cause of the pump stop issue was not determined since product was not returned and data logs were not returned as well. instructions for use for this event: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31567 | 1220648-2024-06182 | 18552779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552779&pc=OZD | 12/13/2023 | 04/01/2024 |  | 01/22/2024 |  |  | 01/10/2024 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 8679 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported an 80-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the during support the automated impella controller (aic) displayed yellow alarm (low placement signal and suction), the p level would drop, and the impella flow would go below 2 liters a minute. proper placement was confirmed using echocardiogram and fluoroscopy, the volume was normal, no obstruction was noted at the inflow or outflow, and no clots were noted in cannula after removal. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - controller error (yellow alarm) has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the frequent suction alarms across cases were not reproduced during testing. these alarms can be intermittently influenced by patient's condition. the cause of the positioning discrepancies was not determined.
31568 | 1220648-2024-06181 | 18552781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18552781&pc=OZD | 06/16/2023 | 05/16/2025 |  | 01/22/2024 |  |  | 06/16/2023 | 01/22/2024 |  |  | 02/28/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024342285 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/16/2023 |  |  |  |  | The user facility reported an unspecified male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella pump would not cross iliac stent. folded several times and got caught on stent. doctor was unable to pass through vasculature and the procedure was aborted. no known patient harm or injury. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided. | Product was not returned for investigation. the cause of the delivery issue was patient condition related to diseased vessel condition, which resulted in product damage. this damage, a kink in the cannula, was noted during insertion based on provided clinical information.
31569 | 1220648-2024-06183 | 18553086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553086&pc=OZD | 01/10/2024 | 12/19/2024 |  | 01/22/2024 |  |  | 01/10/2024 | 01/22/2024 |  |  | 07/31/2024 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023249992 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a male of unknown age and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the physician believes the impella was behind the papillary muscle and stated that during placement the pigtail of the catheter was wrapped behind papillary muscle. the decision was made by the physician not to reposition the impella cp. while on support the patient arrested which required cardio-pulmonary resuscitation; after 20 minutes return of spontaneous circulation was achieved. it was noted that the left lower extremity had no pulse and the patient¿s toes were blue. the decision was made to explant the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the limb ischemia and positioning issues has been completed. no product returned. data logs confirm few suction alarms and positioning alarms. no motor current (mc) spikes outside of p level changes. the cause of the low pump flow was most likely improper positioning as stated by the clinical and the cause of the improper positioning was not determined due to lack of clinical details. the root cause of the limb ischemia could not be determined without additional information. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: added d6a/d6b as it was missing. g1 mdr reporting contact email h6 impact code 2199 was incorrectly coded.
31570 | 1220648-2024-06185 | 18553169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553169&pc=OZD | 06/21/2023 | 09/22/2024 |  | 01/22/2024 |  |  | 06/21/2023 | 01/22/2024 |  |  | 10/31/2023 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022143083 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a fluid leak from the purge cassette was identified during patient support. the purge cassette was replaced to resolve the issue. roughly nine days later, it was documented that a gradual decrease in purge flow coinciding with an increase in purge pressure was observed. heparin was increased, however the the purge flow eventually decreased to 1ml/h. this caused the purge pressure to increase to 900mmhg or more and an alarm was triggered. although the function of the pump was unaffected, the decision was made to replace the pump for continued patient support. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported purge leak/high purge pressure has been completed. review of data logs confirmed the reported failure mode. the logs show a gradual increase in purge pressure over the course of the 10th day of support from 600 mmhg to 1100 mmhg within 20 hours which is indicative of biomaterial buildup. the pump then was stopped manually and disconnected from the console. during functional testing at the bench level, a leak from the piston ¿ cylinder of the purge cassette was revealed. during analysis of the pump, the pump was primed and run to reproduce the failure mode. high pp occurred during bench test due to glucose residue inside pss. when glucose residue was removed, the purge pressure was back to normal range. visual inspection found biomaterial inside the purge gap and a kinked purge line inside the red pump plug. there is no kink on the purge line inside the catheter. mh then was torn down to observe if there was biomaterial in the purge gap. blood was also found on the stator and in the purge gap indicating that the biomaterial built up in the purge gap is a cause of the high purge pressure during the case. a kink was found on the purge line inside the pump plug but not a cause of high pp as the failure mode could not be reproduced with it. the root cause of the reported leak was determined to be a leak from piston - cylinder. the root cause of the high purge pressure was most likely was due to biomaterial built up in the purge gap. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ a risk assessment was performed, and no escalation is required. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06185 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
31571 | 1220648-2024-06184 | 18553184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553184&pc=OZD | 01/12/2024 | 12/03/2025 |  | 01/22/2024 |  |  | 01/12/2024 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1629362 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/12/2024 |  |  |  |  | The user facility reported that during routine maintenance, the power fail alarm of the automated impella controller (aic) was not working. there was no patient harm or involvement reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. a.1 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06184. a.2 age should have been left blank on the initial manufacturer device report number 1220648-2024-06184 as no patient was involved. a.3 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06184. d.2 common device name was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04722 was submitted. d.9 device return date was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06184. e.1 revised name prefix/title as it was reported incorrectly on the initial manufacturer device report number 1220648-2024-06184. e.2 and e.3 revised as they were reported incorrectly on the initial manufacturer device report number 1220648-2024-06184. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-06184 and should not have been. g.1 revised mdr reporting contact address and manufacturing site name and address as they were reported incorrectly on the initial manufacturer device report number 1220648-2024-06184. h.3 updated to reflect device evaluation anticipated as it was reported incorrectly on the initial manufacturer device report number 1220648-2024-06184. h.6 code 4144 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2024-06184. h.10 additonal manufacturer narrative was reported incorrectly on the initial manufacturer device report number 1220648-2024-06184 as the device was returned.
31572 | 1220648-2024-06186 | 18553265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553265&pc=OZD | 06/27/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 06/27/2023 | 01/22/2024 |  |  | 06/30/2023 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1597336 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 13fr peel away cracked upon insertion of dilator during prep. the sheath was discarded and replaced with a 14fr x 13 cm cp sheath. there was no harm to the patient as a result of this incident. | [5 empty] | (h3) the investigation into the reported "introducer damage - mechanical" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the introducer damage issue was not determined since product was not returned and with insufficient clinical information.
31573 | 1220648-2024-06187 | 18553353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553353&pc=OZD | 06/23/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 06/23/2023 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1108094 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/23/2023 |  |  |  |  | The user facility's biomedical engineer reported that their automated impella controller (aic) displayed a battery failure alarm. there was no patient involvement. the device was removed from service. | [5 empty] | The investigation into the automated impella controller (aic) battery failure has been completed. data analysis: no aic logs are available from the complaint date. reproduced the failure mode to obtain the aic logs and confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ) performed batttest using mobaxterm and found that the 4th cell of battery 2 is not functioning as intended; this confirms the cell imbalance with battery 2, which caused the battery pack's internal electronics to shut down that battery. single cell failure. device analysis: the console was tested on an engineering lab bench. the failure mode was reproduced while running the aic without the ac power. also confirmed the battery 2 failure by visual inspection of the battery lcd screen. the battery failure alarm was due to a defective battery 2. the root cause of the defective battery was due to cell imbalance, a known pattern failure.
31574 | 1220648-2024-06188 | 18553397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553397&pc=OZD | 07/03/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 07/03/2023 | 01/22/2024 |  |  | 03/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345308 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a (b)(6) white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a bipella cp-rp support initiated for post-op bentall patient. the cp kinked within the new valve and was explanted and replaced. a new cp was placed, and support proceeded until care withdrawn and patient expired. | [5 empty] | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the issue was a kinked cannula. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31575 | 1220648-2024-06189 | 18553575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553575&pc=OZD | 06/23/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 06/23/2023 | 01/22/2024 |  |  | 02/28/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338588 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient developed a hematoma around the impella access site during support. two mattress sutures were placed around the repositioning sheath, however, the suture needle punctured the catheter and blood started to come out of the red sidearm. the pump was subsequently removed as a result of the noted damage. there was no lasting patient harm. | [5 empty] | The investigation into the reported events has been completed. the device was not returned for evaluation, therefore no root cause can be determined. the evaluation revealed that the most likely cause of the product damage (hole in catheter) issue was most likely use related with the catheter being punctured during the process of placing sutures around repositioning sheath. additionally, the cause of the access site bleeding issue was not determined since product was not returned and there is insufficient clinical information. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
31576 | 1220648-2024-06190 | 18553702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553702&pc=OZD | 07/18/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 07/18/2023 | 01/22/2024 |  |  | 03/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024329132 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/18/2023 |  |  |  |  | The user facility reported a 70-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had successfully delivered the .018 wire. the team, however, was unable to deliver the 5.5 pump. the acute angulation of the anatomy precluded placement. the team was able instead to deliver the cp. no harm or injury. | [5 empty] | The impella device was not received from the customer. the cause of the delivery issue was patient condition related due to anatomy issue that unable to pass the 5.5 into the aorta due to the acute angle from left side and motor of the impella pump would not make the turn due to calcified vessels.
31577 | 1220648-2024-06191 | 18553770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553770&pc=OZD | 12/30/2023 | 02/02/2024 |  | 01/22/2024 |  |  | 12/30/2023 | 01/22/2024 |  |  | 06/30/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024373652 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a (b)(6) asian female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was attempting to have an impella cp implant and the staff had delivery issues. staff pulled the easy lumen guide out of the pump, but they were unable to wire the pump. a new impella was opened and placed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for investigation. according to clinical details, the root cause of the delivery issue was most likely due to use issue since the staff were unable to wire the pump due to premature removal of the easy guide lumen.
31578 | 1220648-2024-06192 | 18553871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553871&pc=OZD | 06/28/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 06/28/2023 | 01/22/2024 |  |  | 12/31/2024 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024300128 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2023 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the motor housing on an impella cp pump could not be advanced into the left ventricle during attempted delivery. an iabp (intra-aortic balloon pump) was subsequently placed to provide some support for the patient. there was no patient harm. | [5 empty] | There was no device returned for evaluation, therefore no root cause could be determined. information from the clinical reports pointed to the patient's anatomy as a likely root cause.
31579 | 1220648-2024-06193 | 18553895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553895&pc=OZD | 01/03/2024 | 03/22/2024 |  | 01/22/2024 |  |  | 01/03/2024 | 01/22/2024 |  |  | 11/30/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422101 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed a hematoma extending from the access site into the pectoral area during impella 5.5 support. as a result of the noted condition, heparin administration was temporarily halted. support was maintained with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-updated model number. g1-reporting contact email updated.
31580 | 1220648-2024-06194 | 18553911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18553911&pc=OZD | 01/04/2024 | 03/15/2024 |  | 01/22/2024 |  |  | 01/04/2024 | 01/22/2024 |  |  | 02/28/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 JP |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024323843 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a support with an impella 5.5 for approximately one month when the pump was explanted due to low flow. the team questioned whether or not there was thrombus causing the low purge flows. the 5.5 was explanted and replaced by an impella cp pump. no lasting harm or injury. | [5 empty] | The investigation for the reported low pump flow has been completed. product was returned for investigation. there was clear sign of biomaterial in the purge gap. no kink was observed. per the data log investigation, there was a gradual increase in purge pressure suggesting a bio material buildup in the purge gap. no motor current spikes were observed. therefore, the root cause of the high purge pressure issue was most likely biomaterial. added- d9 device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31581 | 1220648-2024-06195 | 18554054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18554054&pc=OZD | 01/04/2024 | 01/22/2024 |  | 01/22/2024 |  |  | 01/04/2024 | 01/22/2024 |  |  | 10/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024416512 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported after successful weaning, the impella 5.5 was removed from the patient of unknown age and gender. however, a thrombus was found at the outlet after explant. there was no effect on the patient's outcome. the anticoagulation was ptt controlled. | [5 empty] | (h3) the investigation into the reported "thrombus" has been completed since the original report was filed. product was returned for investigation. according to the clinical review, upon removal a thrombus was observed on the outflow cage. the returned product was visually inspected and the thrombus was not found. it is possible the thrombus was dislodged from the pump in the time between explant and arrival for evaluation. the most likely cause of the thrombus is patient condition. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31582 | 1220648-2024-06196 | 18554373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18554373&pc=OZD | 07/07/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 07/07/2023 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1218786 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/07/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a battery failure. the instructions for use were followed and the aic was replaced. the aic will be returned for evaluation. there was no patient involvement during this event. | [5 empty] | The investigation into the battery failure on the user facility's automated impella controller has been completed. data analysis: aic logs confirmed the reported failure. there was a battery failure alarm and a battery 1 comm. failure alarm due to loss of communication to the battery. though there is no ltpowerdata from the complaint date, batttest confirms the cell imbalance with cell 2 of battery 1, which caused the battery pack internal electronics to shut down the battery 1. device analysis: the failure mode was reproduced on an engineering lab bench. confirmed the battery failure and the communication issues by performing batttest in the terminal window and by visual inspection of the battery lcd screen. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to cell imbalance, a known pattern failure.
31583 | 1220648-2024-06197 | 18554434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18554434&pc=OZD | 01/04/2024 | 03/19/2024 |  | 01/22/2024 |  |  | 01/04/2024 | 01/22/2024 |  |  | 11/30/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024424051 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 69-year-old white female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported a 69-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as escalation of therapy from an impella cp device. while on support, the patient had access site bleeding. a jackson pratt (jp) drain was placed and it had significant volume. the physician was able to fix the bleeding by going through the jp drain tunnel and suturing the pectoral muscle. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleed was most likely due to user technique, the issue was resolved after suturing the pec muscle. h6: impact code: 4641.
31584 | 1220648-2024-06198 | 18554470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18554470&pc=OZD | 01/05/2024 | 03/06/2024 |  | 01/22/2024 |  |  | 01/05/2024 | 01/22/2024 |  |  | 10/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024413829 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aneurysm | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced a pseudoaneurysm at their impella cp access site. when the pump was explanted, the bleeding was initially stopped via compression. however, afterwards, a pseudoaneurysm formed at the area of compression. manual pressure was once again held to achieve hemostasis. although surgical intervention was considered, it was not deemed necessary as the manual pressure managed the condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported vascular damage has been completed. the impella device was not returned for evaluation.. insufficient clinical information was provided and the relation to impella is unknown for the vascular damage. the root cause of the vascular damage issue could not be determined since product was not returned and insufficient clinical information. the instructions for use states ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ it also states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31585 | 1220648-2024-06199 | 18554705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18554705&pc=OZD | 08/10/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 08/10/2023 | 01/22/2024 |  |  | 03/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024346539 |  | 100080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/10/2023 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the physician attempted to advance impella over aortic arch due to calcium and snaring on old axillary stent. the impella implant was aborted and staff proceeded to fix lm disease without support. 14f sheath removed and site closed with perclose x2 at conclusion of case. no known patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore no root cause can be positively identified. from information in the clinical report, the patient's was most likely due to patient condition.
31586 | 1220648-2024-06200 | 18556000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18556000&pc=OZD | 09/05/2023 | 01/22/2024 |  | 01/22/2024 |  |  | 09/05/2023 | 01/22/2024 |  |  | 03/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024343814 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella peel away was removed. the physician was unable to advance the repo sheath into the artery. staff attempted to place 0.35 wire through re-access port but was unsuccessful. physician decided to remove impella and hold manual pressure. after the impella was explanted the patient expired in the procedural area. no allegations were made towards the impella for the cause of death. | [5 empty] | The impella device was not received from the customer, therefore no root cause could be positively identified. the probable cause of the delivery issue was patient condition related as patient had pad/pvd. if a device is returned a follow mdr will be submitted.
31587 | 1220648-2024-06201 | 18556848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18556848&pc=OZD | 01/07/2024 | 04/01/2024 |  | 01/22/2024 |  |  | 01/07/2024 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1226867 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu reported the automated impella controller (aic) showed 'controller failure' alarms prompting unit to be exchanged for another console. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for reported console (aic) "controller failure" alarm has been completed. the aic was returned for evaluation. the controller alarm could not be reproduced while running an optical pump. there was no issue with the console. data logs were reviewed which confirmed the reported failure. therefore, the root cause of the controller error was due to an aic software issue.
31588 | 1220648-2024-06202 | 18556887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18556887&pc=OZD | 01/08/2024 | 11/19/2025 |  | 01/22/2024 |  |  | 01/08/2024 | 01/22/2024 |  |  |  |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | 0042-0000-UK |  |  | ABIOMED, INC. |  |  | US |  | 129286211 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Difficult to Setup or Prepare | Failure to Read Input Signal | Device Handling Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic was unable to detect an installed purge disc during patient support. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. d4 model number has been updated. | The investigation has been completed. data review: after reviewing the associated logs for ic7847, the observed purge disc detection issue was confirmed. numerous ¿purge disc not detected¿ alarms (event set #402) were observed in the imc logs on the reported occurrence date. however, imc logs from the following day (2023.11.24) contains evidence of the console being able to detect the purge disc (indicated by ¿imcgui: purger stat: sensor=1¿ in the imc logs). see "ic7847_imclog.pdf" in attachments for logs. device analysis: console serviced by field service on 2024.05.23. console passed pm and functional check. conclusion: suspect that the observed purge disc detection issues were use related. there was evidence of purge disc not detected alarms in the imc logs from the reported occurrence date but logs from the following day contained evidence of the console being able to detect the purge disc once it was inserted. d4: corrected the model number, serial number, lot number, and udi. h3: added information regarding the investigation status. h4: corrected the device manufacturer date. h6: added codes per the results of the investigation and corrected the medical device problem codes. h10: added the results of the investigation. f6, f8: omitted these dates, this section was filled out in error. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31589 | 1220648-2024-06204 | 18557261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18557261&pc=OZD | 01/05/2024 | 03/29/2024 |  | 01/22/2024 |  |  | 01/05/2024 | 01/22/2024 |  |  | 10/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422309 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 64-year-old female patient was on impella cp support for ventricular fibrillation and the patient had cardio-pulmonary resuscitation. the patient had runs of ventricular tachycardia (vt) while on impella support. additionally, there were purge pressure high/ purge blocked alarms. the y connector was removed and tpa purge was started. shortly after, the motor current was trending high with flow saturation. decision was made to leave the impella in and monitor the patient. eventually, the impella stopped. there were three attempts to restart with no success. the impella was therefore explanted. | [5 empty] | The investigation into the pump stop issue has been completed since the initial report was submitted. the product was not returned for investigation. according to clinical reports, the cause of the pump stop was high purge pressure. the cause of the high purge pressure was unknown. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31590 | 1220648-2024-06205 | 18557489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18557489&pc=OZD | 01/02/2024 | 11/19/2025 |  | 01/22/2024 |  |  |  | 01/22/2024 |  |  |  |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1668279 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Life Threatening | [23 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump stopped working during patient support. after three unsuccessful attempts to restart the pump, the patient suffered a sudden cardiac arrest and cpr was initiated. the impella was replaced, and the aic (automated impella controller) was swapped to assist with stabilizing the patient. it was noted that the patient had been agitated and mobile prior to the pump stop. | [5 empty] | The impella device was returned from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | Corrections to the initial mfr report: g1 mdr reporting contact email. b1 updated to malfunction only. d2 device name updated. d4 model number updated. f6 and f8 should have been left blank. h6 type of investigation code added.
31591 | 1220648-2024-06206 | 18557820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18557820&pc=OZD | 01/03/2024 | 11/13/2025 |  | 01/22/2024 |  |  | 01/03/2024 | 01/22/2024 |  |  | 10/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423022 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Increase in Pressure | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/03/2024 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump experienced a high purge pressure / purge system blocked alarm and subsequently stopped working. the pump was successfully restarted; however, it stopped a second time and the decision was made to explant the device. leading up to the event, it was noted that the purge sidearm was damaged. the staff had been advised that the pump would need to be replaced at the time of the observed purge issue. there was no noted patient harm as a result of the failure. | [5 empty] | The investigation has been completed with the returned product. the pump was returned for investigation, however, only a cut catheter attached to the pump was returned, devoid of the electrical side of the wire connecting to the console and the red handle containing the circuitry. therefore, the pump was unable to run in the lab, be checked for an electrical issue, or check for reported sidearm damage. no abnormalities were seen in the returned part. data logs were returned for analysis. data logs confirm high purge pressure and purge system blocked alarms at the time of purge pressure spikes. purge pressure low alarms were triggered at the end of the case, with an air in purge system alarm. a motor current spike was seen at the end of the case, with motor current high alarms triggering, which caused the pump stop. no other motor current spikes or trends outside of p-level changes were seen for the remainder of the case. no placement signal spikes were seen throughout the case. all 5s parameters were stable and exhibited no unexpected trends. pump stop: the root cause of the pump stop was not determined due to partial product returned, inconclusive evidence from data analysis, and insufficient clinical details. high purge pressure: the root cause of the high purge pressure was not determined due to partial product returned, inconclusive evidence from data analysis, and insufficient clinical details. purge leak - pump: the root cause of the purge leak - pump was damaged sidearm based on clinical details. however, the cause of the sidearm damage was unknown due to partial product returned and insufficient clinical details. d9 updated. h3 updated. b7 other relevant history, including preexisting medical conditions updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported high purge pressure, purge sidearm damage, and pump stop has been completed. the impella pump was not returned for evaluation. data logs confirm high purge pressure and purge system blocked alarms at the time of purge pressure spikes. the purge pressure and flow are otherwise stable throughout the case and do not indicate a leak or broken line, most likely because that part of the logs are missing. the root cause of the high purge pressure was not determined since product and sufficient data logs were not returned. the root cause of the pump stop was not determined since product and sufficient data logs were not returned. the cause of the purge sidearm damage was not determined since product was not returned.
31592 | 1220648-2024-06207 | 18558270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18558270&pc=OZD | 01/06/2024 | 12/10/2024 |  | 01/22/2024 |  |  |  | 01/22/2024 |  |  | 09/30/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400721 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2024 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed ischemia in their impella leg and had hemolyzing labs during support. as a result of the noted conditions, the decision was made to escalate to impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis and limb ischemia has been completed. product was not returned. data logs were reviewed which show no suction alarms, no motor current (mc) spikes outside of p level changes, and no other abnormalities. the cause of the hemolysis was not determined since product was not returned, lack of clinical details, and inconclusive data logs. the root cause of the limb ischemia could not be determine without additional information. added-h6 component code 925 corrections to the initial mfr report: d4 model number added d6a/d6b, g1 mdr reporting contact email. f6/f8 should have been left blank.
31593 | 1220648-2024-06208 | 18558476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18558476&pc=OZD | 01/03/2024 | 02/21/2024 |  | 01/22/2024 |  |  |  | 01/22/2024 |  |  | 10/31/2025 |  | 01/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC |  |  | US |  | 2024412584 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported an unspecified male patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following insertion of the impella cp peel away sheath, the hemostatic membrane was noted to be very leaky. therefore, another device was used. there was no reported patient harm. | [5 empty] | The investigation into the reported product damage has been completed since the original report was filed. no product or sufficient clinical imaging was returned. the cause of the introducer damage issue was not determined since product was not returned and due to insufficient clinical imaging. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31594 | 1220648-2024-06134 | 18542175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18542175&pc=OZD | 01/05/2024 | 03/25/2024 |  | 01/19/2024 |  |  | 01/05/2024 | 01/19/2024 |  |  | 09/30/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412898 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported an 80-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was not able to be positioned properly causing arrhythmias. the impella cp was explanted and the issues resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrhythmia could not be determined.
31595 | 1220648-2024-06135 | 18542201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18542201&pc=OZD | 01/02/2024 | 02/12/2024 |  | 01/19/2024 |  |  | 01/02/2024 | 01/19/2024 |  |  | 10/31/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024416324 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during the implant procedure for 71-year-old female patient, the physician tried to thread the 0.018" wire through the easy guide lumen, but was reported to have been extremely difficult. the doctor had to use so much pressure that a hole was created in the pigtail of the impella cp. the physician therefore used another impella device. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (hole in catheter) has been completed since the original report was filed. the device was returned for investigation. the returned product had the red lumen in place. a hole was noted on the pigtail. the red lumen was removed, and the only abnormality observed was a hole aligned with red lumen and pigtail damage. as per the clinical details, the physician encountered extreme difficulty threading the 0.018" wire through the easy guide lumen, necessitating excessive pressure, and resulting in a hole in the pigtail of the impella. based on the returned product investigation and the provided clinical information, the cause of the delivery issue was use related since the damage was noted after applied excessive force.
31596 | 1220648-2024-06136 | 18542263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18542263&pc=OZD | 01/09/2024 | 06/06/2024 |  | 01/19/2024 |  |  | 01/09/2024 | 01/19/2024 |  |  | 10/31/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419341 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had mottling on the right lower extremity. angio revealed a thrombus and occlusiveness distal to the repo. sheath. patient was weaned and the impella was explanted. there was a notable thrombus exiting the artery at the same time as the impella explant. a 5 fr fogarty catheter-balloon was utilized obtaining additional residual thrombus. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia and thrombus has been completed. the device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia and thrombus was not determined. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added codes 925 and 4114 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
31597 | 1220648-2024-06137 | 18542631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18542631&pc=OZD | 01/09/2024 | 02/21/2024 |  | 01/19/2024 |  |  | 01/09/2024 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024385422 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Loose or Intermittent Connection | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the purge disc holder on the automated impella controller (aic) was loose which caused purge issues. the aic was swapped to resolve the noted issue. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported aic -purge disc issue has been completed since the original report was filed. device analysis: console was tested after arriving in fs. after analysis, purge retainer was found to be broken (see attached image). portion of the purge retainer that is supposed to be fastened to the purge assembly was no longer intact. section d9 was updated.
31598 | 1220648-2024-06138 | 18543174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18543174&pc=OZD | 01/02/2024 | 03/11/2025 |  | 01/19/2024 |  |  | 01/02/2024 | 01/19/2024 |  |  | 04/30/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337513 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during surgical access to the axillary artery of the (b)(6)male patient, the subclavian was lacerated. initial attempts to manage the laceration were unsuccessful and a covered stent was subsequently placed. the impella 5.5 pump was then placed via direct access and support was successfully initiated. four days into support, the patient developed hemolysis with a downward trending hemoglobin. the patient was given two units of packed red blood cells and 1 unit albumin to stabilize the condition. support remained ongoing following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Product not returned. data logs show brief suction event that resolves. no mc spikes outside of p level changes. no other abnormalities. the cause of the hemolysis was not determined due to no product returned, lack of clinical details, inconclusive data logs. the cause of the vascular damage peripheral vessel was not determined since product was not returned and lack of clinical details.
31599 | 1220648-2024-06139 | 18543277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18543277&pc=OZD | 01/11/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 01/11/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1116341 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 01/11/2023 |  |  |  |  | The user facility turned on the automated impella controller (aic) for educational purposes. upon startup, the aic alarmed with a "battery failure" message. the power was cycled in an attempt to troubleshoot the issue, however, the failure remained. there was no patient involvement. | [5 empty] | The investigation into the battery failure alarm has been completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure, a known pattern failure. before sending this console back to the customer. field service replaced the battery kit (0042-3016), validated charging, and performed a full functional check on the console and optical system (0042-7316).
31600 | 1220648-2024-06140 | 18543433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18543433&pc=OZD | 12/27/2023 | 03/04/2024 |  | 01/19/2024 |  |  | 12/27/2023 | 01/19/2024 |  |  | 10/31/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408822 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 63-year-old male patient on impella 5.5 support, post coronary artery bypass graft (cabg) surgery, suffered altered mental status with right sided weakness. a ct of the head was performed and showed acute ischemic infarction of the right basal ganglia. the pump was explanted the following day and patient was noted to continue to show signs of improvement. the cause of stroke is undetermined. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva issue has been completed. the device was not returned for investigation. as the clinical information was not provided, the root cause of the stroke/cva could not be determined.
31601 | 1220648-2024-06141 | 18544106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544106&pc=OZD | 10/07/2022 | 03/11/2025 |  | 01/19/2024 |  |  | 10/07/2022 | 01/19/2024 |  |  | 01/31/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023182859 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 77-year-old white male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a patient developing a hematoma during impella support that prompted vascular intervention to repair. during closure, a complication occurred resulting in a second vascular injury that prompted additional vascular intervention to repair. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
31602 | 1220648-2024-06143 | 18544145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544145&pc=OZD | 02/08/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 02/08/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 60-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a controller error that occurred during support. investigation a breakdown in communication of multiple components. a partial reset was attempted, but a full reset was required to resolve the issue that spurred the alarm. the device continued to support the patient with no harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. the evaluation revealed that the most likely root cause was a defective component.
31603 | 1220648-2024-06142 | 18544155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544155&pc=OZD | 01/23/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 01/23/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1570900 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 01/23/2023 |  |  |  |  | The user facility reported a 78-year-old white male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the user facility's automated impella controller (aic) unexpectedly shutdown during impella support. the console rebooted and patient support continued without further issue. there was no harm to the patient this was day four of support and support continued for 4 more days after. | [5 empty] | The investigation into the aic shutdown/restart issue has been completed. the user facility did not return the impella console for evaluation. the clinical report was used to conclude the cause of the aic issue was a third party hardware/software issue.
31604 | 1220648-2024-06144 | 18544157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544157&pc=OZD | 03/15/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 03/15/2023 | 01/19/2024 |  |  | 05/31/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227660 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/15/2023 |  |  |  |  | The user facility reported a 57-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the cp pump had a lower than expected flow during support. the device was removed and replaced with another impella cp pump. there was no harm to the patient. | [5 empty] | The investigation into the low flow on the cp pump has been completed. the user facility did not return impella products. based on the clinical review the root cause of the low pump flow is most likely due to biomaterial that was ingested. the physician stated the pump was initially placed right into a clot.
31605 | 1220648-2024-06145 | 18544177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544177&pc=OZD | 03/01/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 03/01/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1156148 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 55-year-old white male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had an impella 5.5 support ongoing when the purge was leaking from a purge disc failure. the ifu was followed and backup controller used. no harm or injury from the replacement. | [5 empty] | (h3) the investigation into the reported low pump flow has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found that it¿s been confirmed that purge disk retainer was broken. there was also evidence of purge fluid leak, as well as fluid ingress inside the purge assembly, compromising purge motor shaft gasket and purge pressure transducer. the exact cause of the leak was not determined, however it might have been related to the broken purge retainer not holding the disk firmly against pressure sensor, resulting in pressure being under-reported, which lead to over-pressurizing the disk, and eventually it rupturing. during data analysis, the console logs did not contain any entries on the reported day of event (2023-03-01), however records on (b)(6)2023 and on (b)(6) 2023 are consistent with complaint and show intermittent detection of purge disk, after purge cassette change with that purge cassette, there is also evidence supporting claim of purge leak: purge flow increased and purge pressure low alarms in imc log as well as purge flow and purge pressure values recorded in ltlog and rtlog. this report is being filed as part of a retrospective review of historical records.
31606 | 1220648-2024-06146 | 18544240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544240&pc=OZD | 03/18/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 03/18/2023 | 01/19/2024 |  |  | 09/30/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281817 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 52-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that upon insertion of the guidewire for impella implant, a type b dissection was noted in the descending aorta. medical intervention was not required, however, the decision was made to abort the scheduled implant. there was no lasting patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the investigation into the reported vascular damage has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the vascular damage as the abiomed wire, an alternate wire, and the impella all faced resistance at the descending aorta.
31607 | 1220648-2024-06147 | 18544591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544591&pc=OZD | 04/15/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 04/15/2023 | 01/19/2024 |  |  | 11/30/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308474 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 35-year-old white female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp optical pump stopped during patient support. the pump was explanted as a result of the failure and a new impella cp pump was implanted for ongoing support. there was no patient harm. | [5 empty] | The impella was returned for evaluation. visual inspection of the returned pump revealed a large impeller purge gap. additionally, the catheter was severed near the repo unit, with the motor cable cut and signs of excessive force as both ends of the catheter were torn. during functional evaluation, the device ran for at least 5 minutes after activation and resulted in a stoppage. the root ause of the pump stop was bearing wear after 16 days of use, which is beyond 8-day design life. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31608 | 1220648-2024-06149 | 18544641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544641&pc=OZD | 05/16/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 05/16/2023 | 01/19/2024 |  |  | 11/30/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312169 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump experienced persistent suction issues as a result of a thrombus that had entered the inlet of the pump. due to the noted issue, the pump was explanted. | [5 empty] | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the low pump flow issue was most likely biomaterial (thrombus). the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31609 | 1220648-2024-06148 | 18544651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544651&pc=OZD | 04/13/2023 | 09/19/2024 |  | 01/19/2024 |  |  | 04/13/2023 | 01/19/2024 |  |  | 11/30/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310469 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed hemolysis coinciding with impella support. it is unknown if intervention was required for the condition; however, support with the pump continued until the decision was made to withdraw care due to other aspects of the patient's health. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | (h3) the investigation into the reported "hemolysis" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the hemolysis issue was not determined since product was not returned, data logs were not returned, and insufficient information provided. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06148 was provided.
31610 | 1220648-2024-06150 | 18544666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544666&pc=OZD | 05/21/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 05/21/2023 | 01/19/2024 |  |  | 11/30/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303537 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/21/2023 |  |  |  |  | The user facility reported a (b)(6) black male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the pump shifted out of the patient's ventricle during support and was unable to be re-advanced past the aortic valve. a new pump was used, with no adverse impact to the patient. | [5 empty] | The investigation into the "positioning problem" has been completed. the cause of the issue was not determined since product was not returned and sufficient information was not provided.
31611 | 1220648-2024-06151 | 18544764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544764&pc=OZD | 05/27/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 05/27/2023 | 01/19/2024 |  |  | 02/28/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322105 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/27/2023 |  |  |  |  | The user facility reported an 80-year-old white male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the pump was delivered with a 14fr long sheath placed through the femoral artery. after the impella cp was advanced through, the sheath kinked. the kinked sheath was removed and replaced the a shorter sheath. there was no patient harm. the pump supported the patient for three days until it was weaned and explanted. | [5 empty] | The investigation into the introducer issue of kinked sheath has been completed. product was not returned for review. the root cause of the introducer kink was unable to be determined as limited clinical details were provided and no product was returned for review.
31612 | 1220648-2024-06152 | 18544786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544786&pc=OZD | 03/22/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 06/05/2023 | 01/19/2024 |  |  | 12/31/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298075 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Failure to Deliver | Material Deformation |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported a 65-year-old white female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during insertion of the guidewire, the wire began to buckle. this was due to the excessive force being used. the physician utilized a new device and discarded the initial device. there was no patient harm related to this event. | [5 empty] | The investigation into the impella cp delivery issue has been completed. product was not returned for investigation. the cause of the delivery issue was most likely kinked guidewire since doctor had to apply excessive force inserting guidewire into red lumen, based on provided clinical information. the failure modes will be monitored and trended.
31613 | 1220648-2024-06153 | 18544979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18544979&pc=OZD | 06/08/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 06/09/2023 | 01/19/2024 |  |  | 01/31/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024322653 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 1 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/09/2023 |  |  |  |  | The user facility reported a 1-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support.the user facility reported male (age unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump stopped after approximately 30 days of impella support. the pump was removed with no adverse impact to the patient. | [5 empty] | The investigation into the pump stop has been completed. product & data were not returned for review. the pump stop occurred on day 30 which is 22 days over the 8 day design life of the impella cp optical. the root cause of the pump stop was unable to be determined as limited clinical details were provided and no product was returned for review.
31614 | 1220648-2024-06154 | 18545047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18545047&pc=OZD | 06/15/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 06/15/2023 | 01/19/2024 |  |  | 02/28/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341985 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 73-year-old male patient required to have impella support. the physician accidentally removed the pump from the left ventricle (lv) at the end of the case and then was unable to get it back across the valve without the wire. decision was made to remove impella without replacement since patient was stable and on minimal ionotropic support. the patent did not encounter any serious adverse events as a result of this incident. | [5 empty] | (h3) the investigation into the reported "positioning issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the physician was unable to cross aortic valve without guidewire at the end of case after impella was implanted. the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access per ifu. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31615 | 1220648-2024-06155 | 18545054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18545054&pc=OZD | 12/25/2022 | 01/19/2024 |  | 01/19/2024 |  |  | 12/25/2022 | 01/19/2024 |  |  | 09/30/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274906 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 63-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following insertion of the repositioning sheath, a large hematoma formed at the impella access site. due to the noted condition, the decision was made to remove the impella and place a large edwards sheath (14f x 30 cm ) in order to achieve hemostasis. | [5 empty] | The investigation into the reported hematoma has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that a lack of vessel recoil onto the sheath was the most likely cause of the hematoma. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31616 | 1220648-2024-06156 | 18545738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18545738&pc=OZD | 05/19/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 05/19/2023 | 01/19/2024 |  |  | 03/31/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327521 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had impella 5.5 support ongoing. the 5.5 was seen to have a purge leak from the pc. the leak was seen on day 10 of the support. the pc was replaced and the 5.5 continued to support the patient without harm or injury. additionally, there was high purge pressure observed. the purge system was blocked with alarm and bicarb was running. | [5 empty] | The investigation into the reported event has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. data log data was reviewed and there were many purge pressure alarms, purge system blocked alarm, purge flow alarms that occurred. the cause of the issue was not determined since product was not returned for analysis. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31617 | 1220648-2024-06157 | 18546128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18546128&pc=OZD | 12/27/2023 | 12/05/2024 |  | 01/19/2024 |  |  | 12/27/2023 | 01/19/2024 |  |  | 10/31/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 68-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon insertion, the impella was backloaded and attempted to be inserted, but met significant resistance. impella was removed, access was regained and was attempted a second time with same issue. impella was removed again and access was regained with a supercore wire instead of 0.18 and the impella placed successfully in left ventricle. additionally, there was significant groin bleeding around the repo sheath. patient was given two units of blood products and a femstop was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into access site bleeding and delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. the root cause of the delivery issue was most likely patient condition related due to resistance met upon insertion. instructions for use as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings and cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings and cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system with automated impella controller. section: warnings and cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06157. b1 both adverse event and product problem. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use missing.
31618 | 1220648-2024-06158 | 18546157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18546157&pc=OZD | 06/01/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 06/01/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1133146 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the screen on their automated impella controller (aic) was flickering. the aic was swapped as a result of the noted issue. there was no patient harm. | [5 empty] | (h3) the investigation into the reported "aic - display issue" has been completed since the original report was filed. product was returned for investigation. the device was inspected and found the failure mode was not able to be reproduced. cables¿ connectors were seated properly. isolated testing over three hours showed no signs of flickering screen. visual inspection revealed residue build up in the purge unit area which was unrelated to reported failure mode. the automated impella controller (aic) logs are not relevant to this failure mode.
31619 | 1220648-2024-06159 | 18546527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18546527&pc=OZD | 06/08/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 06/08/2023 | 01/19/2024 |  |  | 02/28/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341963 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that prior to the venoarterial extra corporeal membrane oxygenation (ecmo) cannulation purge alarms "high purge pressure/purge system blocked" started to occur. troubleshooting was attempted but without success. the decision was made by the doctor to proceed and insert a new impella cp. procedure completed successfully. | [5 empty] | The investigation into the reported high purge pressure has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. data analysis confirmed the failure mode. data showed that about 10 hours into the case, the pump was in the aorta area. fifteen minutes later, pp started to increase gradually from 400 mmhg to 1100 mmhg that triggered alarms high pp and purge system blocked. the high pp then remained to the rest of the case. based on clinical description and data analysis, the root cause of the high purge pressure was most likely a kinked catheter due to a tight aortic arch patient condition. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ a risk assessment was performed, and no escalation is required. this report is being filed as part of a retrospective review of historical records.
31620 | 1220648-2024-06160 | 18547419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18547419&pc=OZD | 01/26/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 01/26/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1215718 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 01/26/2023 |  |  |  |  | The user facility biomedical engineer reported the automated impella controller (aic) had a "controller failure" alarm at the conclusion of a case. there was no patient harm. | [5 empty] | The investigation into the aic controller failure alarm has been completed. the aic was returned for evaluation. the cause of the aic issue was a firmware corruption.
31621 | 1220648-2024-06161 | 18547422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18547422&pc=OZD | 06/14/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 06/14/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1178000 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/14/2023 |  |  |  |  | The user facility reported an 86-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) blue purge disc manifold became loose and the pressure was unreliable. the aic was swapped out for another aic. there was no patient harm related to this event. | [5 empty] | The investigation into the automated impella controller (aic) purge disc/flag issues has been completed. data analysis: the aic log from the complaint date shows that the console issued a purge pressure low ¿ alarm issue which is due to the misalignment of the purge retainer and the dislodged flag, which confirms the reported failure that the blue purge disc manifold loose and pressure unreliable. device analysis: the console was tested after arriving in field service. upon visual inspection, it was confirmed that the disk detect flag was dislodged, and the retainer was not flushed with the aic. no dextrose residue was found. the root cause for the purge pressure low was because of the dislodged disk detect flag and the misaligned blue retainer.
31622 | 1220648-2024-06162 | 18547431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18547431&pc=OZD | 02/16/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 02/16/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1679798 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had an aic with a battery failure, red alarm, when plugged in by staff. the ifu will be followed and the aic returned and a backup controller will be used. there was no reported harm or injury. | [5 empty] | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis. console logs from the reported day of event are consistent with one of described scenarios, and show persistent battery failure alarm while console is plugged into ac. the scenario where ¿console when turned on and unplugged alerts staff to plug in¿, is a normal behavior when operated on batteries only. there¿s no evidence of the aic being switched on and operated without ac either with or without battery alarm (as there were no alarms found within aic logs) on the day of reported event. starting with aic boot-up, the battery failure alarm was no longer present. battery data from (b)(6) 2023 shows no evidence of battery failure or battery cell imbalance. all available evidence is consistent with battery transport switch being disengaged on the day of reported event, and most likely being result of the staff being unfamiliar with operation of the device. the evaluation revealed that the cause of the issue was use related (operator did not disengage transport switch prior to using the console).
31623 | 1220648-2024-06163 | 18547518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18547518&pc=OZD | 03/03/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 03/03/2023 | 01/19/2024 |  |  |  |  | 01/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1629292 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/03/2023 |  |  |  |  | The user facility reported that after storing their automated impella controller (aic) in a box for over one year, there was a battery failure alarm when powered up for first time. there was no patient involvement. | [5 empty] | The investigation into the automated impella controller (aic) battery failure (red alarm) has been completed. console logs from the reported day of event are consistent with complaint and show battery failure alarm (and battery 1 and battery 2 communication failure alarms upon boot-up. prior boot up was 14 months before that, and there were no battery alarms. after console was received in danvers, the issue was reproduced. inspection of battery packs revealed both batteries fully depleted (led indicators were blank), and batttest diagnostics showed that batteries could no longer be changed due to internal lock-up caused by cell voltages below safety thresholds. during testing, batteries were temporarily replaced with known-good ones, and the investigator was able to charge them. per the ifu, the console should be stored with its power cord plugged into wall ac outlet, which was not the case for ic9380. the cause of the aic issue was a depleted battery due to user facility technique/use error.
31624 | 1220648-2024-06164 | 18547662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18547662&pc=OZD | 04/12/2023 | 10/08/2025 |  | 01/19/2024 |  |  | 04/12/2023 | 01/19/2024 |  |  | 01/31/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024315304 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/12/2023 |  |  |  |  | The user facility reported a 54-year-old black male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the pump was inadvertently pulled into the aorta during patient support. attempts were made to reposition the device, however the pump would not re-cross the valve. the pump was subsequently explanted as a result of the noted issue. there was no patient harm. | It was also reported the patient experienced hemolysis while on support. due to hemolyzed labs, the device decreased from p level 9 to level 7 which assisted in managing the complication; along with medication treatment. | The complainant reported that the pump was inadvertently pulled into the aorta during patient support. attempts were made to reposition the device, however the pump would not re-cross the valve. the pump was subsequently explanted as a result of the noted issue. it was also noted that heparin was on hold, as patient bleeding from face and dialysis catheter site. also, labs showed evidence of hemolysis. p level dropped from 9 to 7. a 30-day mdr is required for the report of positioning issues, access site bleeding, hemolysis. |  |  |  | Per the provided clinical information, the root cause of the positioning issue was use related to wireless reaccess. the impella device is not indicated for wireless re-access per ifu 0048-9007. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information for the initial mfr 1220648-2024-06164 was provided in sections b5, d6a, d6b, h1, and h6. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the failure modes will be monitored and trended. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
31625 | 1220648-2024-06165 | 18548041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18548041&pc=OZD | 04/20/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 04/20/2023 | 01/19/2024 |  |  | 04/30/2024 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210732 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that difficulty was met during delivery of an impella 5.5 pump which led to the perforation of the left ventricle. after crossing the aortic valve, it was documented that the doctor was unable to turn the pump away from the mitral valve. poor imaging made repositioning difficult and it was eventually noted that the device had perforated the left ventricle. the pump was pulled back and the lv was surgically repaired. following the intervention, the pump was re-advanced and the support level was increased to p-7 to help stabilize the patient. support then proceeded with the implanted pump as planned. | [5 empty] | The investigation into the reported ventricle perforation has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the data log shows no abnormalities other than the recorded impella position in the aorta at the start of the case. the evaluation revealed that the most likely cause of the ventricle perforation issue was most likely improper positioning and related to excessive force applied when placing impella with difficulty getting impella into the correct position. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
31626 | 1220648-2024-06166 | 18548078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18548078&pc=OZD | 05/18/2023 | 01/19/2024 |  | 01/19/2024 |  |  | 05/18/2023 | 01/19/2024 |  |  | 01/31/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328127 |  | 1000080 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Insufficient Flow or Under Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a cp support initiated for support of a ami/cgs patient. the cp was found to have low flows and was replaced. at the time of explant the cp was observed to have clot/thrombosis. the new 2nd pump was placed without harm or injury. the pump data logs were analyzed and found to have a pump stop. | [5 empty] | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. data showed the low flow occurred when pump started that triggered auto suction response and suction alarms. shortly after activation, a high motor current pump stop occurred that triggered alarm 13 impella stopped - mc high. an attempt to restart the pump failed. the pump was then stopped manually. the evaluation revealed that the most likely cause of the issue was biomaterial (thrombus). instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31627 | 1220648-2024-06167 | 18548310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18548310&pc=OZD | 12/22/2023 | 01/29/2024 |  | 01/19/2024 |  |  |  | 01/19/2024 |  |  | 09/30/2025 |  | 01/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024405048 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient developed a hematoma at their access site during impella 5.5 support. the patient was taken to the operating room for a washout / evacuation to treat the condition. the patient was transfused one unit of blood and heparin administration became sub-therapeutic to reduce the chance of recurrence. support remained ongoing with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, the root cause of the access site bleeding issue was not determined since product was not returned.
31628 | 1220648-2024-06112 | 18533884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18533884&pc=OZD | 12/05/2023 | 03/04/2024 |  | 01/18/2024 |  |  | 01/05/2024 | 01/18/2024 |  |  | 09/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024407346 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Inadequate Filtration Process | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 42-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on day 14 of impella support, white crystals were observed inside the cover of the purge side arm. there was no problem with purge related operations or alarm generation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak ¿ pump issue has been completed. the device was not returned for investigation. per the clinical information provided, the most likely cause of the pump purge leak was a damaged sidearm. the specific location of the leak could not be determined since product was not returned.
31629 | 1220648-2024-06114 | 18534137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18534137&pc=OZD | 01/05/2024 | 02/12/2024 |  | 01/18/2024 |  |  | 01/05/2024 | 01/18/2024 |  |  | 09/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024410754 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 80-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced recurrent bleeding issues during impella cp support. the bleeding resulted in repeated drops in hemoglobin levels and blood transfusions were administered at intervals to counteract the condition. the bleeding was managed via dressing changes and additional sutures at the access site. support was maintained with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. as per clinical review, the root cause of the access site bleeding was most likely use related as the bleeding was resolved by adjusting the angle of insertion and adding an additional suture.
31630 | 1220648-2024-06115 | 18534662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18534662&pc=OZD | 12/20/2023 | 06/12/2024 |  | 01/18/2024 |  |  | 12/20/2023 | 01/18/2024 |  |  | 09/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409912 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 74-year-old female patient had impella cp support due to having non-st elevated myocardial infarction. the facility alleged that the companion sheath, "appeared to not lock correctly". the staff therefore used a different sheath successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the introducer damage - mechanical issue has been completed since the original report was filed. the pump was returned for investigation. no leaks were observed from the luer-lock connect on the returned device when pressurized to 400 mmhg for 30 seconds. the loose clear collar luer-lock is not abnormal for a luer-lock connection and has no negative affect on the seal of the luer lock. the companion sheath luer-lock functioned/operated to specification.
31631 | 1220648-2024-06116 | 18534899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18534899&pc=OZD | 01/07/2024 | 02/15/2024 |  | 01/18/2024 |  |  | 01/08/2024 | 01/18/2024 |  |  | 11/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024424846 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was inserted with no issue but when the impella started an impella stopped, and pump failure alarmed occurred. the team tried to restart the impella with no success. a new impella cp was inserted and started successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue completed since the original report was filed. the device was returned for investigation. kink found on catheter but determined to not be related to failure. the root cause of the pump stop, and pump failure alarm cannot be determined due to inability to reproduce failure and no data logs from field to analyze.
31632 | 1220648-2024-06117 | 18534914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18534914&pc=OZD | 01/07/2024 | 02/23/2024 |  | 01/18/2024 |  |  | 01/07/2024 | 01/18/2024 |  |  | 06/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024371823 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/07/2024 |  |  |  |  | The user facility reported a 50-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following implant of an impella cp pump, the patient developed ischemia in their limb. as a result of the noted condition, the decision was made to explant the pump prior to the planned coronary artery bypass grafting (cabg). | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31633 | 1220648-2024-06118 | 18535113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18535113&pc=OZD | 01/08/2024 | 03/25/2024 |  | 01/18/2024 |  |  | 01/08/2024 | 01/18/2024 |  |  | 11/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024421972 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to traverse a patient's anatomy during attempted implant. the delivery was subsequently aborted and and impella cp pump was implanted for planned support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
31634 | 1220648-2024-06119 | 18535159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18535159&pc=OZD | 01/08/2024 | 09/17/2024 |  | 01/18/2024 |  |  | 01/08/2024 | 01/18/2024 |  |  | 10/31/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024422304 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after the patient on impella support was sent to the intensive care unit, lower leg ischemia was noted and confirmed by ultrasound. options were discussed but the family wanted impella to be removed. after the impella was removed, the patient hemodynamically decompensated with two rounds of cardiopulmonary resuscitation (cpr) completed. the family requested no further medical treatment and the patient passed in the catheterization lab. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06119 was provided.
31635 | 1220648-2024-06120 | 18535423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18535423&pc=OZD | 01/08/2024 | 02/09/2024 |  | 01/18/2024 |  |  | 01/08/2024 | 01/18/2024 |  |  | 11/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024425903 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Improper Flow or Infusion | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced hemolysis and access site bleeding during impella cp support. the patient was given one unit of prbc to treat the conditions; however, the patient then showed signs of renal failure. continuous renal replacement therapy was initiated, and the pump was repositioned for better flow. left ventricle distention began to worsen and support levels were increased. despite the indicated flow and support level, the condition remained and the patient was escalated to impella 5.5 support. following successful placement, it was noted that the flows on the 5.5 were considerably lower indicating that the registered flows during cp support were not accurate. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the reported hemolysis, access site bleeding and saturated flow has been completed since the original report was filed. the pump was returned for investigation, and biomaterial was observed on the impeller. no other physical abnormalities were found on the returned pump. data logs indicated saturated flow throughout the case with a reported spike in motor current before stopping the pump at the end of the case. the cause of the hemolysis issue was determined to be biomaterial, based on the returned product and data log review. the cause of the saturated pump flow issue was biomaterial found wrapped around the impeller. the root cause of access site bleeding is not determined.
31636 | 1220648-2024-06121 | 18535427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18535427&pc=OZD | 01/09/2024 | 12/09/2024 |  | 01/18/2024 |  |  | 01/09/2024 | 01/18/2024 |  |  | 09/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407662 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella insertion after exchange from 14fr to repo sheath, the patient developed significant hematoma and bleeding from the site. the physician performed a cutdown to repair the artery and femostop was applied. later that same day, the patient was taken to the operation room for another femoral vascular repair at the impella site. bleeding was controlled. several blood products were given. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into access site bleeding and vascular damage - peripheral vessel has been completed. pump was not returned for investigation. as per clinical information provided, the bleeding was related to improper angle matching and therefore vascular repair was required to control bleeding. therefore, the root cause of the vascular damage issue was most likely use related to improper technique. access site bleeding was a chain event to vascular damage therefore the failure access site bleeding has ben deleted. instructions for use as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings and cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06121 d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email g1 reporting contact fax number missing. h6 health effect - impact code 4624 was missing h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use missing.
31637 | 1220648-2024-06122 | 18535621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18535621&pc=OZD | 01/09/2024 | 12/11/2024 |  | 01/18/2024 |  |  | 01/09/2024 | 01/18/2024 |  |  | 09/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404605 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed hemolysis during impella cp support. as a result of the noted condition, the patient was escalated to impella 5.5 support and the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. d4 model number d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
31638 | 1220648-2024-06123 | 18535623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18535623&pc=OZD | 01/09/2024 | 02/12/2024 |  | 01/18/2024 |  |  | 01/09/2024 | 01/18/2024 |  |  |  |  | 01/18/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1129855 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Noise, Audible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support and the automated impella controller (aic) had a loud noise in the motor of the purge cassette. the aic was replaced. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic ¿ other has been completed since the original report was filed. the console was returned for investigation. the data logs from the complaint date show numerous user interactions with the purge system of the console via purge wizards. there were no persistent issues observed with the completion of the purge-related actions. the right data logs show many purge flow ticks during the purge cassette change. the console was visually inspected. there was residue build up inside and behind the purge unit assembly inside the aic, but this did not impact the performance of the purge system. there was a crack on the small purge retainer visible under a digital microscope, but the purge pressure sensor operated without any issues. the console was run with a known good pump and purge for about an hour with no noticeable issues. there was no issue found during data and device analysis. the user likely heard routine de-airing procedure and thought there was a problem. updated d9: device return date.
31639 | 1220648-2024-06124 | 18535692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18535692&pc=OZD | 01/09/2024 | 12/30/2024 |  | 01/18/2024 |  |  | 01/09/2024 | 01/18/2024 |  |  | 10/31/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | Device Handling Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had an impella cp placed for support but after delivery to the left ventricle (lv), there was kinking occurring due to calcified leaflets. the pump kinked further with repositioning and was removed/explanted. the pump supported for approximately 30 minutes. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The failure mode was changed from "product damage (cannula kink) was changed to "low pump flow" because low pump flow was reported. no product was returned for analysis. the clinical details mention calcified aov leaflets caused cannula kink after which flows were reduced. the data logs confirm reduced pump flow with no motor current spikes. there was no mention of excessive force or improper technique. the root cause of the low pump flow was most likely patient condition related the following have been reported incorrectly or missing from manufacturer device report. d4 model number d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing.
31640 | 1220648-2024-06125 | 18535866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18535866&pc=OZD | 12/30/2023 | 03/25/2025 |  | 01/18/2024 |  |  | 01/09/2024 | 01/18/2024 |  |  | 10/31/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418794 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Ischemia | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 01/09/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to weaning off bypass onto the impella. while on support, there was suction alarms and losses in pulsatility noted. the patient was taken to the operating room for a washout, tamponade draining and an additional chest tube placement. the patient had a surgical clot in the chest wall causing tamponade that required evacuation. this resolved the suction alarms. | The complaints team received new information about an impact study regarding acute kidney injury (aki) stage i. the study found the impella 5.5 and impella procedure to be possibly related. creatinine elevation could be secondary to procedure to place impella or to impella functioning. impella possible source of microemboli. no further details were provided for the treatment of the renal issue. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible and the low pump flow issues has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the low pump flow was due to patient condition as the left ventricle wall was likely being pushed up against the pump due to tamponade. however, the root cause of the limb ischemia was not determined. | B.5 additional information was received. h.6 health effect - clinical code was added due to new informatin received. d.4 revised model number since the initial manufacturer device report number 1220648-2024-06125 was submitted in accordance with updated procedures. d.6a and d.6b added implant and explant date since the initial manufacturer device report number 1220648-2024-06125 was submitted in accordance with updated procedures. e.1 added fx number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-06125. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-06125 and should not have been. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-06125 was submitted in accordance with updated procedures. added reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-06125 was submitted. h.6 codes 2954 and 403 were reported incorrectly on the initial manufacturer device report number 1220648-2024-06125. | G.2 added study selection as it was inadvertently omitted from follow-up 2 manufacturer device report number 1220648-2024-06125.
31641 | 1220648-2024-06126 | 18536120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18536120&pc=OZD | 01/11/2024 | 02/26/2024 |  | 01/18/2024 |  |  | 01/11/2024 | 01/18/2024 |  |  | 11/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024423694 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support ongoing for a coronary artery bypass grafting (cabg¿) patient with cardiomyopathy. the pump had saturated flows. there was no harm or injury from the malfunction. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the saturated flow issue has been completed. the device was not returned for investigation. data logs show that a few seconds after starting the pump there is a motor current spike. the root cause of the saturated flow was most likely biomaterial.
31642 | 1220648-2024-06127 | 18536167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18536167&pc=OZD | 01/12/2024 | 02/26/2024 |  | 01/18/2024 |  |  | 01/12/2024 | 01/18/2024 |  |  | 08/31/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024380549 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to being escalated from the impella cp. while on support there was high purge pressure noted. eventually, the pump was stopped. the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop and high pump pressure has been completed since the original report was filed. no product was returned for evaluation. data logs returned for this pump end on (b)(6) 2024, however, pump did not fail until (b)(6) 2024. no clinical details to indicate why high purge pressure alarm and pump stop occurred. the root cause of the high purge pressure cannot be determined as no pump was returned, no data logs at the time of failure, and no clinical details at time of alarm. the root cause of the pump stop cannot be determined as no pump was returned, no data logs at the time of failure, and no clinical details at time of alarm.
31643 | 1220648-2024-06128 | 18537290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18537290&pc=OZD | 12/21/2023 | 12/24/2024 |  | 01/18/2024 |  |  | 12/21/2023 | 01/18/2024 |  |  | 09/30/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407781 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 73-year-old female patient on impella cp support had a ct scan that showed cerebral ischemia. the patient had also been experiencing a gastrointestinal (gi) bleed requiring blood products. it was further reported that the right hand is blue due to the ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation stroke/cva, limb ischemia, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06128. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
31644 | 1220648-2024-06129 | 18537990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18537990&pc=OZD | 01/07/2024 | 09/20/2024 |  | 01/18/2024 |  |  | 01/07/2024 | 01/18/2024 |  |  | 10/31/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024409518 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | Positioning Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/07/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella support experienced suction events. echo - echocardiogram (ultrasound) confirmed that the impella catheter appeared to be behind the papillary muscle. repositioning at bedside was attempted without success. the patient was brought to the cardiac cath lab for impella repositioning. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type corrected information for the initial mfr 1220648-2024-06129 was provided. note: corrected information provided in h6 is an addition to previous coding.
31645 | 1220648-2024-06130 | 18538147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18538147&pc=OZD | 01/09/2024 | 02/21/2024 |  | 01/18/2024 |  |  | 01/09/2024 | 01/18/2024 |  |  |  |  | 01/18/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1076289 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Calibration Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | No patient involvement: the us hospital found a syscal alarm on the aic (automated impella controller) during a preventative maintenance. there was no patient use and no harm or injury. | [5 empty] | The impella device was received from the customer and an evaluation of the device is not yet completed. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported aic - system self-check error has been completed since the original report was filed. due to the nature of the failure mode, no aic logs were able to be recovered for review that were related to the failure mode. device analysis: the failure mode seen in the field was reproduced and the device outputted system self check when booted up. then even with a service dongle inserted, the console could still not get through boot up. since boot up is run by the compact flash cards ((b)(6)), the cfcs were replaced by known-goods but the failure mode remained. the cause of the system self check failure was a defective 4-in-1 and corrupted compact flash cards.
31646 | 1220648-2024-06131 | 18538234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18538234&pc=OZD | 01/09/2024 | 01/29/2024 |  | 01/18/2024 |  |  | 01/09/2024 | 01/18/2024 |  |  | 10/31/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410749 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 48-year-old male of an unknown race in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed oozing and a hematoma at the impella access site following implant. due to the noted condition, the decision was made to discontinue the administration of tirofiban. support was maintained throughout the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
31647 | 1220648-2024-06132 | 18539477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18539477&pc=OZD | 12/28/2023 | 03/27/2024 |  | 01/18/2024 |  |  | 12/28/2023 | 01/18/2024 |  |  | 10/31/2025 |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024412232 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a 43-year-old female was implanted with an impella cp as a bridge to transplant. it was reported that the pump exhibited low purge pressure alarms with a high flow rate. it was discovered that the purge sidearm to be snapped in half between fluid reservoir and 0.2-micron filter. blood was seen leaking from the purge sidearm, and a purge bypass was performed, staff reconnected the purge fluid and tissue plasminogen activator (tpa) was ran. the pump ran for around 12 more hours, but ultimately the flows were saturated, the pump stopped and was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency. pump stop: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency. purge leak: do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir. clean the connector cable with 70% isopropyl alcohol. | The investigation into the pump stop, the access site bleeding ¿ major and the purge leak ¿ pump issues has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the pump stop issue was most likely leak from the broken sidearm due to excessive force, based on given clinical details. purge leak allowed the blood to enter the purge line. the cause of the access site bleeding issue was most likely use related, as the sidearm was damaged, leading to blood leaking from it.
31648 | 1220648-2024-06133 | 18540172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18540172&pc=OZD | 10/20/2023 | 10/14/2024 |  | 01/18/2024 |  |  | 12/19/2023 | 01/18/2024 |  |  |  |  | 01/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369439 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. abiomed¿s long-term outcome and quality indicator impella registry received a notification of a patient enduring hemolysis with a "possible" allegation to the impella device: unspecified intervention was required to prevent serious injury/death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs from the time of the event (initial 3 days) were not returned. therefore, the root cause of the hemolysis issue was not determined since product was not returned and data logs at the time of the event were not returned.
31649 | 1220648-2024-06077 | 18525190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18525190&pc=OZD | 01/06/2024 | 02/23/2024 |  | 01/17/2024 |  |  | 01/06/2024 | 01/17/2024 |  |  | 10/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423602 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2024 |  |  |  |  | The user facility reported a post-partum, and post-surgical 29-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp was placed in the femoral and the limb became ischemic it was, cool to touch, pulseless, and had mottled skin. the impella cp explanted, and patient was escalated to impella 5.5 pump via the axillary for support. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
31650 | 1220648-2024-06078 | 18525287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18525287&pc=OZD | 01/07/2024 | 01/23/2024 |  | 01/17/2024 |  |  | 01/07/2024 | 01/17/2024 |  |  |  |  | 01/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1415626 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/07/2024 |  |  |  |  | The user facility in japan reported a patient of unknown age and gender in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. while setting up the impella cp for support, the purge set up failed, replaced with a separate purge set, with no improvement. the automated impella controller (aic) needed to be replaced, but there was no backup controller available. the impella insertion was aborted and there was no patient harm reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation and was tested upon return. data logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the console was tested after arriving. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. the cause of the issue was a missing purge flag.
31651 | 1220648-2024-06079 | 18525588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18525588&pc=OZD | 01/07/2024 | 02/06/2024 |  | 01/17/2024 |  |  | 01/07/2024 | 01/17/2024 |  |  |  |  | 01/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1121971 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/07/2024 |  |  |  |  | The user facility reported a 51-year-old male patient of unknown ethnicity was implanted with an impella 5.5 as a bridge to transplant. on day 4 of the support, the purge cassette was changed per usual and when the new purge cassette was placed the automated impella controller (aic) failed to recognize. the aic was swapped out with a backup console. no patient harm was reported. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs could not confirm that the reported complaint of the purge cassette not being detected but confirmed that purge disc could not be detected through multiple issuances of the alarm. the console was was visually inspected to reveal a missing purge flag. the root cause for the purge disc not detected alarm was a missing purge flag. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31652 | 1220648-2024-06080 | 18525722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18525722&pc=OZD | 01/08/2024 | 09/16/2024 |  | 01/17/2024 |  |  | 01/08/2024 | 01/17/2024 |  |  | 09/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024396419 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 64-year-old male of unknown ethnicity was implanted with an impella 5.5 device for mechanical circulatory support due to being hemodynamically unstable when coming off bypass. there was a left ventricle perforation after the impella placement. the origin of perforation is unknown however, the impella is considered a possible factor. the perforation required covered stent and/or balloon tamponade. surgeons noted severe difficulty with lack of clotting after heparin was discontinued and protamine was given. the patient expired in the procedural area after 4.20 hours of support. no allegations were made towards the impella for cause of death. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury . dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06080 was provided.
31653 | 1220648-2024-06081 | 18526062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18526062&pc=PYX | 01/06/2024 | 07/15/2024 |  | 01/17/2024 |  |  | 01/06/2024 | 01/17/2024 |  |  | 08/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024400552 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2024 |  |  |  |  | The user facility reported a 61-year-old male was on impella bipella due to right and left sided heart failure for mechanical circulatory support. it was reported there was pump saturation on the impella rp motor current with no flow. the implla rp was replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported saturated flow has been completed. according to the clinical details, on the second day of impella rp support no flow was recorded with stable motor current. the decision was made to explant the pump and replace it with another rp. no product was returned for a visual inspection. data log analysis confirms sensor drift as raw dp sensor and flow were gradually decreasing. the most likely cause of the placement signal issue was sensor drift. the most likely cause of the placement signal issue was sensor drift. added- d6a/d6b f6/f8 should have been left blank in the initial report.
31654 | 1220648-2024-06082 | 18526254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18526254&pc=OZD | 01/09/2024 | 02/02/2024 |  | 01/17/2024 |  |  | 01/09/2024 | 01/17/2024 |  |  | 09/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409932 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2024 |  |  |  |  | The user facility reported an 81-year-old male undergoing an aortic valvuloplasty was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was loaded and placed into the left ventricle outflow tract (lvot). after several attempts the impella would not advance further. the physician removed the wire and started the impella, stating that the impella pulled itself in when turned on. the impella started and flows were normal. however, the motor current was barely pulsatile. the physician did not like this placement and after much manipulation the impella was removed. the impella was flushed while a pigtail and .018 wire was placed back into the ventricle. the impella was re-prepped and loaded onto the wire. the impella once again got caught in the lvot and would not pass any further. the impella insertion was aborted due to the patient¿s anatomy. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the delivery issue was use related as the location of the resistance was the aortic valve as the patient had transcatheter aortic valve replacement (tavr).
31655 | 1220648-2024-06083 | 18527288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18527288&pc=OZD | 08/24/2023 | 01/17/2024 |  | 01/17/2024 |  |  | 12/21/2023 | 01/17/2024 |  |  | 06/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373667 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a (b)(6) female/male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. due to patient body habitus, patient experienced hematoma in icu. manual pressure held and hematoma got softer. patient also experiencing skin tears due to nature of patient skin. due to concern about access site, patient trial weaned at p2 and decision made to remove pump. pump removed in ccl deploying the 2 percloses. additional manual pressure held. hemostasis obtained. | [5 empty] | The investigation into the patient's hematoma has been completed. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to patient condition. the failure modes will be monitored and trended.
31656 | 1220648-2024-06084 | 18527549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18527549&pc=OZD | 12/25/2023 | 02/07/2024 |  | 01/17/2024 |  |  | 12/25/2023 | 01/17/2024 |  |  |  |  | 01/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1386957 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported automated impella controller (aic) was defective. no other information is available. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic ¿ other has been completed since the original report was filed. the console was returned for investigation. the console logs from the reported day of event are consistent with complaint and show pbm1 communication failure which started approx. 2 days after initiating support with the pump. inspection of the console performed in aachen fs revealed corrosion on pbm, affecting its communication with the epc. this is a known pattern failure. this pbm¿s sn falls outside of previously identified range. the cause of the aic issue was contamination.
31657 | 1220648-2024-06085 | 18527569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18527569&pc=OZD | 12/27/2023 | 03/05/2024 |  | 01/17/2024 |  |  | 12/27/2023 | 01/17/2024 |  |  |  |  | 01/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1094052 |  | IC1399 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the flap in front of the purge cassette on the automated impella controller (aic) was broken off. a different aic was able to be successfully used. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the console (aic) damage has been completed. the aic was returned for evaluation. the aic was evaluated finding that the purge pressure disk retainer was missing. data logs were not reviewed after the device evaluation as they were irrelevant to the finding. therefore the root cause of the reported damage was due to damaged component.
31658 | 1220648-2024-06086 | 18527771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18527771&pc=OZD | 12/29/2023 | 02/26/2024 |  | 01/17/2024 |  |  | 12/29/2023 | 01/17/2024 |  |  | 02/28/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325304 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported impella cp moved in aorta of patient of unknown age and gender, during cpr and transport to hospital. the pump was unable to be re-positioned and had to be explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the delivery issues has been completed. the device was not returned for investigation. the root cause of the positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review.
31659 | 1220648-2024-06088 | 18527927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18527927&pc=OZD | 01/08/2024 | 09/20/2024 |  | 01/17/2024 |  |  | 01/09/2024 | 01/17/2024 |  |  | 06/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024368904 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/09/2024 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient developed signs of hemolysis after the implant. there was renal failure after 12 hours. after 76.50 hours of support, care was withdrawn and the patient expired. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06088 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The investigation for the reported hemolysis issue has been completed. the impella device was returned for evaluation. visual inspection found the catheter was cut and a big clot was found inside cannula, the inflow cage, and the pump housing. no other issues were found on the rest of the pump. data log analysis also confirmed that suction alarms were recorded when the patient's urine began to change color. therefore, the cause of the hemolysis was due to biomaterial.
31660 | 1220648-2024-06089 | 18527966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18527966&pc=OZD | 12/27/2023 | 03/14/2024 |  | 01/17/2024 |  |  | 12/27/2023 | 01/17/2024 |  |  | 08/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024395108 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was explanted but the patient of unknown age and gender had to be transferred to vascular surgery due to issues at the impella groin side. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. correction : the clinical code and impact code in section h6, was submitted incorrect inadvertently in the initial report which has been corrected to this report. section b1 and b2 was submitted incorrect inadvertently in the initial report which has been corrected to this report.
31661 | 1220648-2024-06092 | 18528122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528122&pc=OZD | 12/28/2023 | 02/23/2024 |  | 01/17/2024 |  |  | 12/29/2023 | 01/17/2024 |  |  | 10/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420281 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user reports a patient in the stemi-dtu study who was supported by an impella cp smartassist heart pump experienced bleeding at the access site during the course of support. because manual compression failed to stop the bleeding and a decrease in hemoglobin and a hematoma at the access site were diagnosed, the pump was ultimately removed. | [5 empty] | The impella device was not received from the customer as it was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding ¿ major been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical was not provide to determine.
31662 | 1220648-2024-06091 | 18528165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528165&pc=OZD | 01/12/2024 | 02/29/2024 |  | 01/17/2024 |  |  | 01/12/2024 | 01/17/2024 |  |  | 05/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024365994 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2024 |  |  |  |  | The user facility reported a 21-year-old male of unknown race with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the impella purge filter on the side arm was cracked. the impella was replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the access site bleeding ¿ minor and the purge leak ¿ pump issues has been completed. the device was not returned for investigation. per the provided clinical details, the root cause of the purge leak was most likely sidearm filter damage due to excessive force. however, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided.
31663 | 1220648-2024-06093 | 18528262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528262&pc=OZD | 01/05/2024 | 01/23/2024 |  | 01/17/2024 |  |  | 01/05/2024 | 01/17/2024 |  |  | 10/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424278 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was access site bleeding. a total of eight units of blood products were given to the patient. pressure dressing was placed and the access site improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The medical center did not return the impella device. no benchtop analysis was possible. the root cause of the access site bleeding could not be determined since the product was not returned and enough clinical info was not provided.
31664 | 1220648-2024-06094 | 18528285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528285&pc=OZD | 01/06/2024 | 05/02/2024 |  | 01/17/2024 |  |  | 01/06/2024 | 01/17/2024 |  |  |  |  | 01/17/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721589 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem: Dose Calculation Error | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/06/2024 |  |  |  |  | The user facility reported a (b)(6) female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) prompted a purge volume critically low alarm despite a full bag being attached with the correct volume entered. a review of the infusion history screen identified that the aic was incorrectly calculating the volume being delivered. the purge flow and purge pressure remained unchanged despite the noted issue. as a result of the observed failure, the staff was advised to swap out the aic to the back up console. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) purge issues has been completed. the console was returned for evaluation. the reported complaint could not be reproduced. the purge assembly was tested and confirmed to be working correctly. data logs revealed a miscommunication between the pud and the 4-in-1 that resulted in an incorrect number of purge flow ticks being sent by the pud to the 4-in-1. the number of purge flow ticks was far outside of the normal range, and this resulted in an incorrect calculation of the remaining purge fluid. the miscommunication only happened a single time in a 6 month period, and the issue could not be reproduced. the cause of the aic issue was a defective purge assembly causing an incorrect number of ticks to be measured/communicated. to conform to updated procedures, section d2 was revised with updated information.
31665 | 1220648-2024-06096 | 18528294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528294&pc=OZD | 12/28/2023 | 02/29/2024 |  | 01/17/2024 |  |  | 12/28/2023 | 01/17/2024 |  |  | 11/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426255 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 22 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 22-year-old female patient on impella 5.5 support experienced partly high flow values (6.2l/min). the left ventricle (lv) received adjustments which made it somewhat better, but the lv was noted to be narrow. however, the patient was well supported and there was no reported harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported high purge pressure has been completed since the original report was filed. no product was returned. data logs were reviewed, and motor current spike followed by saturated flows observed at higher p-levels. pump flow of 6.2 l/min at p-9 was found and then later low flows seen when reduced to p-4. at the end of the case pump was again observed to have saturated flows at p-2. the cause of the saturated flow issue was most likely biomaterial per data log review.
31666 | 1220648-2024-06095 | 18528325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528325&pc=OZD | 01/02/2024 | 03/20/2024 |  | 01/17/2024 |  |  | 01/02/2024 | 01/17/2024 |  |  | 04/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332363 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient developed an infection. the 5.5 was removed due to being a likely source. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen. | The investigation into the infection/sepsis issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the infection/sepsis was not determined.
31667 | 1220648-2024-06097 | 18528394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528394&pc=OZD | 01/02/2024 | 12/09/2024 |  | 01/17/2024 |  |  | 01/02/2024 | 01/17/2024 |  |  | 09/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403165 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed hemolysis during impella cp support. the patient was given one unit of blood to manage the condition and the p-level was decreased while plans were made to escalate to impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. no product was returned. data logs were returned which showed a few suction alarms that resolve throughout the case. no motor current (mc) spikes outside of p level changes and no other abnormalities. the cause of the hemolysis was not determined since product was not returned, lack of clinical details, and inconclusive data logs. added h6 component code 925 corrections to the initial mfr report: d4 model number added d6a/d6b g1 mdr reporting contact email. f6/f8 should have been left blank.
31668 | 1220648-2024-06098 | 18528409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528409&pc=OZD | 01/06/2024 | 12/09/2024 |  | 01/17/2024 |  |  | 01/06/2024 | 01/17/2024 |  |  | 07/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391373 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2024 |  |  |  |  | The user facility reported 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient on impella cp support developed hemolysis. the impella was repositioned, and the level was lowered in hopes of decreasing hemolysis. impella cp was exchanged for an impella 5.5. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient condition including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06098.
31669 | 1220648-2024-06099 | 18528431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528431&pc=OZD | 12/31/2023 | 09/25/2024 |  | 01/17/2024 |  |  | 12/31/2023 | 01/17/2024 |  |  | 10/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410903 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Increase in Pressure | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Sepsis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user reported a patient with st-elevation myocardial infarction was on impella cp smartassist for hemodynamic support. the pump experienced high purge pressures and low flows during support that could not be attributed to visible kinks in the purge path. the anticoagulation was also within the intended framework. no harm to the patient was reported. | It was also reported the patient had a fever and was septic. it was noted that the patient expired while on support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06099 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The investigation into the reported high purge pressure and low pump flow has been completed since the original report was filed. no product was returned for investigation. data logs show a motor current spike prior to purge system blocked and high purge pressure alarms. placement signal begins to drift slightly upwards. purge flow also drops immediately and remains low for 1 hour before beginning to increase but still remaining in low ranges; this is most likely from the addition of tpa. the root cause of the high purge pressure was most likely biomaterial.
31670 | 1220648-2024-06100 | 18528436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528436&pc=OZD | 01/07/2024 | 02/23/2024 |  | 01/17/2024 |  |  | 01/07/2024 | 01/17/2024 |  |  | 04/30/2024 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219812 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/07/2024 |  |  |  |  | The user facility reported a 45-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp was placed and limb ischemia was observed bilaterally. the pulses were lost. the team explanted the impella cp and had to perform surgical repair of the femoral artery. hemostasis achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ irreversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31671 | 1220648-2024-06101 | 18528463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528463&pc=OZD | 01/02/2024 | 01/24/2024 |  | 01/17/2024 |  |  | 01/02/2024 | 01/17/2024 |  |  | 09/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMEPLLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410572 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 77-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. attempted to advance impella cp but due to the tortuous and calcified left iliac the imella cp implant was aborted. decision was made to place intra-aortic balloon pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related.
31672 | 1220648-2024-06103 | 18528615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528615&pc=OZD | 01/04/2024 | 08/18/2025 |  | 01/17/2024 |  |  | 01/04/2024 | 01/17/2024 |  |  | 10/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024418705 |  | 1000100 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that several hours after the purge tubing was changed, a purge system failure alarm occurred which led to a controller failure alarm. after the controller was switched out, all alarms were resolved. there was no patient harm. | [5 empty] | The impella device was received from the customer and an investigation regarding the returned device has been completed. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g.1. mdr reporting contact email. d.1. brand name updated. d.4. model number updated. d.9. updated device availability for evaluation to yes. f.6. and f.8. should have been left blank. h.6. type of investigation code added.
31673 | 1220648-2024-06104 | 18528793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18528793&pc=OZD | 01/04/2024 | 12/09/2024 |  | 01/17/2024 |  |  | 01/04/2024 | 01/17/2024 |  |  | 08/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404414 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was explanted prematurely due to bleeding from the access site. there was an observed iliac artery dissection that needed to be treated by balloon tamponade and manta closure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the vascular damage and access site bleed has been completed. pump was not returned for investigation. the clinical information provided is very little and insufficient to investigated the failure mode. therefore, the root cause of the vascular damage issue and the access site bleed was not determined since product was not returned and insufficient clinical information was provided. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 corrections to the initial mfr report: added d6a/d6b. g1 mdr reporting contact email. f6/f8 should have been left blank.
31674 | 1220648-2024-06105 | 18529214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18529214&pc=OZD | 01/04/2024 | 07/17/2024 |  | 01/17/2024 |  |  | 01/04/2024 | 01/17/2024 |  |  | 09/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395644 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a 5.5 pump for support of a patient with multiple cardiac and systemic comorbidities. the 5.5 was placed but the access site bleeding prompted the team to adjust the anticoagulation and remove the heparin. the pump was later explanted and implanted with venoarterial extracorporeal membrane oxygenation (va-ecmo). | It was reported that the patient developed a small hematoma at the impella cutdown site so heparin was paused temporarily. the impella was to remained in place until heart transplant or durable lvad placement. the hematoma at the impella sit resolved after heparin was paused and the team attempted to restart but the hematoma came hack. the heparin was still paused with the plan to start slow. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and enough clinical information was not provided. b.5 revised as this information was inadvertently omitted from manufacturer report number 1220648-2024-06105. d.4 revised model number from manufacturer report number 1220648-2024-06105 in accordance with updated procedures. d.6a and d.6b revised from manufacturer report number 1220648-2024-06105 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer report number 1220648-2024-06105 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer report number 1220648-2024-06105 was submitted. h.6 added codes 1884 and 4641 as they were inadvertently left off manufacturer report number 1220648-2024-06105. code 925 was added in accordance with updated procedures since manufacturer report number 1220648-2024-06105 was submitted. h.10 additional manufacturer narrative instructions for use (ifus) were reported on manufacturer report number 1220648-2024-06105 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
31675 | 1220648-2024-06109 | 18530120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18530120&pc=OZD | 01/05/2024 | 01/26/2024 |  | 01/17/2024 |  |  | 01/05/2024 | 01/17/2024 |  |  | 11/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024422020 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to traverse a patient's anatomy during attempted implant. the delivery was subsequently aborted and impella cp pump was implanted for planned support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, the root cause of the delivery issue was most likely due to patient condition.
31676 | 1220648-2024-06090 | 18530959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18530959&pc=OZD | 12/26/2023 | 03/07/2024 |  | 01/17/2024 |  |  | 12/26/2023 | 01/17/2024 |  |  |  |  | 01/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1175801 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the purge disk was not detected during prep. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the pump did not start issue has been completed since the original report was filed. the console was returned and tested after arriving in fs. broken purge disk detect flag was confirmed during console inspection. the cause of the issue was a broken purge disc detect flag.
31677 | 1220648-2024-06108 | 18531273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531273&pc=OZD | 01/05/2024 | 01/31/2024 |  | 01/17/2024 |  |  | 01/05/2024 | 01/17/2024 |  |  | 10/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024420675 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 51-year-old female undergoing cardiac surgery was to be implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 was unable to pass through the artery and impella cp was placed instead via axillary insertion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the delivery issue was most likely due to patient condition due to small vessel.
31678 | 1220648-2024-06106 | 18531319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531319&pc=OZD | 12/31/2023 | 09/20/2024 |  | 01/17/2024 |  |  |  | 01/17/2024 |  |  |  |  | 01/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1669290 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Crack | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 41-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) had purge disc holder damage that occurred when the purge cassette was changed. the purge disc holder got damaged by the replaced purge cassette. the aic was replaced. there was no patient harm. the aic was returned for repair and evaluation, however has not been logged in as received as of the date of this report. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06106 was provided. note: corrected information provided in h6 is an addition to previous coding.
31679 | 1220648-2024-06110 | 18531544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531544&pc=PYX | 01/05/2024 | 03/11/2025 |  | 01/17/2024 |  |  |  | 01/17/2024 |  |  | 02/28/2024 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2023204347 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 83-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella placement, there was oozing from the impella access site. an additional purse string suture was added. heparin was not started until hemostasis was achieved. it was further reported that the impella migrated into the right ventricle and the patient developed hemolysis. patient care was withdrawn. death was not related to the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | Pump was not returned for investigation. as per limited clinical information provided, the slight oozing from the access site was controlled with sutures but the reason of bleeding is unknown. therefore, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information provided. clinical mentions that the pump was pulled in rv. therefore, the root cause of the positioning issue was most likely use related. clinical mentions that the pump migrated in rv and hemolysis was observed. the patient was withdrawn care and therefore its unknown how hemolysis resolved. therefore, the root cause of the hemolysis issue was most likely improper positioning of the device.
31680 | 1220648-2024-06111 | 18531815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531815&pc=OZD | 12/09/2023 | 07/18/2024 |  | 01/17/2024 |  |  | 12/18/2023 | 01/17/2024 |  |  | 09/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409941 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | Abiomed¿s long-term outcome and quality indicator impella registry received a notification regarding a 67 year old male patient presenting for high risk percutaneous coronary intervention. the impella cp device was used for hemodynamic support. during support, the patient developed an event of "right femoral artery access site bleeding" with a probably relationship to the impella device. as a result, multiple blood products were transfused. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-corrected model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
31681 | 1220648-2024-05996 | 18531823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531823&pc=OZD | 12/24/2023 | 02/14/2025 |  | 01/17/2024 |  |  | 12/24/2023 | 01/17/2024 |  |  | 10/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024413012 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user reported a 68-year-old patient receiving hemodynamic support from an impella cp heart pump experienced alarms. the pump position was checked and slightly corrected, but no improvement was observed. the pump was described as lying freely in the left ventricle but not pumping blood. an echocardiogram examination could not detect any ventricular thrombus. when the user attempted to reposition it again, the pump dislocated completely into the aorta. the pump was ultimately successfully removed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow and positioning issues has been completed. the root cause of the positioning issue was use related to improper technique. the root cause of the low flow issue was a kinked cannula the cause of the cannula kink was not determined as there are no clinical details indicating how the kink occurred. d6a and d6b revised on manufacturer device report 1220648-2024-05996 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-05996 in accordance with updated procedures.
31682 | 1220648-2024-05998 | 18531824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531824&pc=OZD | 12/22/2023 | 10/01/2024 |  | 01/17/2024 |  |  | 12/22/2023 | 01/17/2024 |  |  | 02/28/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319628 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported, that an impella cp implant attempt was unexpectedly aborted. when the patient of unknown age and gender, went into ventricular tachycardia (vt), after the guidewire was inserted. as the patient had difficulty coming out of the vt, the decision was made to abort the implant. and proceed with the planned percutaneous coronary intervention (pci) without impella support. | [5 empty] | The impella device was not received from the customer. and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05998: d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
31683 | 1220648-2024-05994 | 18531825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531825&pc=OZD | 12/22/2023 | 09/19/2024 |  | 01/17/2024 |  |  | 12/22/2023 | 01/17/2024 |  |  | 06/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374873 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the team found previous coronary artery bypass grafting had failed. the patient on impella cp support had red urine with confirmed spun renal panel showing pink/red serum. pump repositioning was unsuccessful as the hemolysis did not resolve. also, the right ventricle was failing with an increasing central venous pressure (cvp) and right coronary artery (rca) circulation disruption caused initial event. the patient was made do-not-resuscitate (dnr) and expired shortly after the dnr decision. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. corrected information for the initial mfr 1220648-2022-01256 was provided in sections b2, b5, d6a, d6b, h1, and h6.
31684 | 1220648-2024-05995 | 18531826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531826&pc=OZD | 12/22/2023 | 01/23/2024 |  | 01/17/2024 |  |  | 12/22/2023 | 01/17/2024 |  |  | 08/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394708 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 87 -year-old male (unknown race and ethnicity) undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported the impella cp had delivery issues due to the 87-year-old male patient's calcified vessels. the user tried two different long sheaths on both sides but kinked badly. dilators were also used but did not work. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. a risk assessment was performed, and no escalation is required at this time. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related. h.6 investigation findings code 174 was reported on the original manufacturer device report (mdr) in error as the investigation was not complete yet. codes 3233 and 3221 should of been reported on the original mdr. investigation conclusions code 4307 was reported on the original manufacturer device report (mdr) in error as the investigation was not complete yet. code 11 should of been reported on the original mdr. the codes that reflect the investigation have been completed are reflected in h.6 of this supplemental manufacturer device report.
31685 | 1220648-2024-05999 | 18531827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531827&pc=OZD | 12/30/2023 | 12/04/2024 |  | 01/17/2024 |  |  | 12/30/2023 | 01/17/2024 |  |  | 10/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024415320 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was pushed in too deep to the point the pigtail bounced off the apex and moved up towards the aortic valve when the repo sheath was being placed. the physician attempted to reposition the pump and the entire cannula went past the aortic valve and into the aorta. the physician attempted to insert the long peel away and rewire the pump, but then the white connector cable became disconnected and the purge fluid stopped. the physician decided to remove the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the pump positioning issue and the purge leak pump issue has been completed. pump was not returned. data logs show the pump running only about 3 minutes before turned off but continued to purge. there are also pump disconnected while running alarms. the pump was connected and reconnected to the console a few times. the root cause of the positioning issue was most likely a use issue based on clinical description. the cause of the purge issue was most likely a use issue as the pump was unplugged from the aic and thus was no longer running or purging. added-h6 impact code 4628 corrections to the initial mfr report: added d6a/d6b.
31686 | 1220648-2024-05997 | 18531828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531828&pc=PYX | 12/22/2023 | 02/22/2024 |  | 01/17/2024 |  |  | 12/22/2023 | 01/17/2024 |  |  | 07/31/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024399894 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that the patient was on impella rp support and pedal pulses were not present even with doppler. they attempted to place a dialysis catheter with no success due to stenosis. on arrival to icu, pedal pulses were monitored again with no results. the color of the plantar surface of the foot was turning blue to purple quickly. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, seere valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31687 | 1220648-2024-05832 | 18531829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18531829&pc=OZD | 12/10/2023 | 09/25/2024 |  | 01/17/2024 |  |  | 12/10/2023 | 01/17/2024 |  |  | 06/30/2025 |  | 01/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported a 37yo male was implanted with an impella cp for mechanical circulatory support. it was reported that the patient experienced ventricular tachycardia (vt) and ectopy during impella insertion. support was able to proceed as planned once insertion was successful. however, during planned escalation to impella 5.5 support on a later date, a thrombectomy was required on the femoral artery from the cp insertion site. following the intervention, the site was closed with sutures and the skin was stapled. the patient was subsequently escalated to impella 5.5 support. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. a supplemental report will be submitted when the investigation has been completed. the instructions for use states the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported ventricular fibrillation (vf) and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vf and thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
31688 | 1220648-2024-06044 | 18517233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18517233&pc=OZD | 01/01/2024 | 02/22/2024 |  | 01/16/2024 |  |  | 01/01/2024 | 01/16/2024 |  |  | 09/30/2024 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET (CHINA, INDIA) |  |  | ABIOMED, INC. |  |  | US |  | 2024312057 |  | 004413 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Fluid/Blood Leak | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2024 |  |  |  |  | The user facility reported a n unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in india reported a patient attempting to be on impella 2.5 support had a purge issue. during set-up the purge solution was connected to the iv set-up by mistake. after this, the iv set (other end) was connected to yellow luer on the impella pump. the white connector cable was connected to automated impella controller (aic) and impella pump respectively. after this, there was a constant red alarm mentioning the low purge pressure. the purge bag was changed a couple of times however, the alarms stayed. the iv set-up was disconnected and purge spike was connected to purge bag. however, the alarms persisted. restarting the aic did not work. the impella was procedure was abandoned. the physician believes there was an issue with the pump. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed. the device was not returned for investigation. the cause of the purge leak issue was not determined since product was not returned and due to insufficient clinical information. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
31689 | 1220648-2024-06045 | 18517309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18517309&pc=OZD | 12/27/2023 | 02/22/2024 |  | 01/16/2024 |  |  | 12/27/2023 | 01/16/2024 |  |  | 08/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024385013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to multiple organ failure and their ejection fraction was 20%. while on support the patient lost pulses in their right foot. the pump was explanted. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
31690 | 1220648-2024-06046 | 18517421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18517421&pc=OZD | 12/19/2023 | 02/16/2024 |  | 01/16/2024 |  |  | 12/29/2023 | 01/16/2024 |  |  | 05/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024352120 |  | 10000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Syncope/Fainting | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had enrolled a patient in the protect iv protocol and placed the impella cp for support. the patient had a vaso-vagal event at the time of the insertion which the pi of the study alleged a "probable" relationship to the use of impella. the team had to treat the patient for the hypotension with ns iv bolus, 2 amps of atropine and neosynephrine. dopamine was started and patient was given zofran. the case proceeded and the cp allowed for the patient's hrpci (high risk percutaneous coronary intervention) and was explanted after 5 hours. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for investigation. according to clinical information, the root cause of the delivery issue - anatomy was determined to be patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31691 | 1220648-2024-06047 | 18517634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18517634&pc=OZD | 01/02/2024 | 02/09/2024 |  | 01/16/2024 |  |  | 01/02/2024 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024418654 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Pumping Stopped | [6 empty] | Ventricular Fibrillation | [12 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump stopped working during patient support. after three unsuccessful attempts to restart the pump, the patient suffered a sudden cardiac arrest and cpr was initiated. the impella was replaced and the patient was stabilized. it was noted that the patient had been agitated and mobile prior to the pump stop. heparin had also been discontinued since the previous day due to oral bleeding. | [5 empty] | The investigation into the pump stop, the ventricular fibrillation, and the vascular damage - peripheral vessel issues has been completed since the original report was filed. the product was returned for investigation. based on the device investigation provided, the root cause of the pump stop was determined to be three open electrical lines most likely caused by patient movement. the root cause of the ventricular fibrillation was most likely patient condition related since it occurred after the pump stopped. any device contacting the heart wall in conjunction with a poor patient condition could result in vf/vt but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the vf is unable to be determined. | The impella device was received from the customer on 01/05/2024 ,an evaluation of the device is under way. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ . ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ . ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ . ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
31692 | 1220648-2024-06048 | 18517882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18517882&pc=OZD | 12/26/2023 | 10/08/2024 |  | 01/16/2024 |  |  | 12/26/2023 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419188 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported a 45-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed mitral regurgitation during impella 5.5 support. the device was repositioned multiple times, but the condition persisted. as a result, the pump was removed and replaced with a new device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported heart valve damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the heart valve damage could not be determined. the instructions for use for the impella 5.5with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 impact code 4628, component code 925 corrections to the initial mfr report: added-d6a/d6b f6/f8 should have been left blank.
31693 | 1220648-2024-06049 | 18518103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18518103&pc=OZD | 12/27/2023 | 10/11/2024 |  | 01/16/2024 |  |  | 12/27/2023 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419233 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a (b)(6) female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support as escalation from an impella cp device. the 5.5 delivery failed despite all troubleshooting measures. the pump was removed and a graft thrombosis was observed. the decision was made to consult with vascular surgery for clot removal. it is unknown at this time if surgery was required. the team instead continued with the impella cp support with no issues. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of the presence of clot presented restricting the insertion of impella 5.5. the case was aborted. the root cause of the thrombus could not be determined without additional clinical details. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. added-b1-selected product problem corrections to the initial mfr report: added-d6a/d6b f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31694 | 1220648-2024-06050 | 18518333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18518333&pc=OZD | 12/28/2023 | 01/16/2024 |  | 01/16/2024 |  |  | 12/28/2023 | 01/16/2024 |  |  | 09/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024403704 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and a left ventricle gram was performed. a left ventricle perforation was noted and was caused by the impella pigtail. the impella was removed. emergent pericardiocentesis. mass transfusion. cts was in house. patient was placed on ecmo and chest was opened. attempted closure with pledgets after impella removal. surgeon stated lv was ¿mush¿. multiple efforts attempted. cardiac massage. defibrillation. unable to maintain pressure. despite best efforts, patient did not survive. | [5 empty] | The pump was returned for investigation. no physical abnormalities, such as cannula kinks or damages to the inlet cage or pigtail, were observed. as per clinical details, the ventricle was noted to be diseased prior to impella use. the cause of the ventricle perforation issue was patient condition related since no physical abnormality were reported on returned pump and also doctor reported diseased lv.
31695 | 1220648-2024-06051 | 18518439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18518439&pc=OZD | 12/25/2023 | 04/19/2024 |  | 01/16/2024 |  |  | 12/25/2023 | 01/16/2024 |  |  | 06/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376511 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2023 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was taken back to the operation room to fix insertion site bleeding. the physician cut down to the femoral artery and placed three purse-string sutures around the impella. the bleeding stopped. two units of red blood cells were given. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding is determined to be use issue because the bleeding was resolved upon cut down to fem art and placing 3 purse string sutures around the impella. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31696 | 1220648-2024-06052 | 18518463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18518463&pc=OZD | 12/29/2023 | 12/19/2024 |  | 01/16/2024 |  |  | 12/29/2023 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420698 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for a patient admitted with a stemi (st elevated myocardial infarction) and in need of bipella support. the patient support was successful for over 8 hours when the patient became agitated and pulled the pump. the team had to explant and deliver a new cp. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on the clinical details, the root cause of the positioning issue is due to use as the patient pulled the impella out. added h6 impact code 4629 corrections to the initial mfr report: d4-updated model number. added d6a/d6b which was missing. f6/f8 should have been left blank. g1 mdr reporting contact email
31697 | 1220648-2024-06053 | 18518608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18518608&pc=OZD | 12/27/2023 | 01/16/2024 |  | 01/16/2024 |  |  | 12/27/2023 | 01/16/2024 |  |  | 06/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374874 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on support when there was low pump flow. the staff decided to replace the impella as a clot injection was suspected. | [5 empty] | The investigation into the low pump flow has been completed. the impella cp pump was returned for investigation. it was visually inspected and biomaterial ingestion was observed over the inflow cage. no biomaterial was found to present in the outflow cage or impeller. no other abnormalities were observed. the impella was running at auto with low flows of 2.0 l/min during field testing. the cause of the low pump flow issue was most likely biomaterial ingestion over inflow cage per field investigation.
31698 | 1220648-2024-06054 | 18518747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18518747&pc=OZD | 12/28/2023 | 05/29/2024 |  | 01/16/2024 |  |  | 12/28/2023 | 01/16/2024 |  |  | 07/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024380931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp being placed for support during a percutaneous coronary intervention (pci). the cp was placed via the improperly placed 23fr sheath. the 23fr sheath was placed in error, rather than the 14fr sheath. the access site was bleeding. the team explanted the pump and contralateral leg access was closed with a balloon. the stent was covered and manual hold was done. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding is a use issue, as the physician used a 23fr rp sheath in place of the 14fr. there is no evidence to indicate the packaging contained incorrect components. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31699 | 1220648-2024-06057 | 18518924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18518924&pc=OZD | 12/30/2023 | 01/16/2024 |  | 01/16/2024 |  |  | 12/30/2023 | 01/16/2024 |  |  | 06/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024356792 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | Pumping Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an impella cp implant when there were no pump flows due to a kinked cannula. the aic (automated impella controller) showed the impella never started and the pump was removed and replaced. additionally, the staff could not locate the 0.018" guidewire upon implanting due to packaging issues. a different guidewire was used instead. | [5 empty] | The pump was returned for investigation without the guidewire. a cannula kink was observed on the returned pump, while no other physical abnormalities were noted. low flow issue was there request log. the cause of the packaging issue was not determined since product was not returned and sufficient clinical information was not provided.
31700 | 1220648-2024-06058 | 18518976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18518976&pc=OZD | 12/29/2023 | 11/19/2025 |  | 01/16/2024 |  |  | 12/29/2023 | 01/16/2024 |  |  | 09/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024413084 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/29/2023 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a utilized introducer split upon removal of the dilator during preparation for impella implant. as a result, the introducer was replaced and the subsequent impella cp implant was successful. there was no patient harm. | [5 empty] | Corrections to the initial mfr report: g1 mdr reporting contact email f6 and f8 should have been left blank component code has been added to h6. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation was completed. ntroducer damage - mechanical. per the clinical, the 14fr x 13 cm sheath was split upon removing the dilator. the returned product was evaluated. return product looked to be split below the introducer hub where it is joined to the sheath. the damage looks like it occurred possibly during insertion. there was also minor resistance during insertion but was advanced with further torquing even after split. a bulge is also observed on the sheath just below the hub which possibly resulted due to improper insertion of the sheath during the impella procedure. at the area of the bulge, the outer diameter of the sheath measured around 6.003 mm and 6.06 mm on front and back sides respectively which is clearly out of spec. and, the sheath outer diameter measured approximately 5.97 mm which is in spec (standard outer dimeter is 5.97 mm+/- 0.0254 mm). therefore, the root cause of the introducer damage - mechanical issue was most likely use issue due to improper technique(s) followed during insertion. h6: added codes per the results of the investigation and the clinical code. h10: added the results of the investigation.
31701 | 1220648-2024-06059 | 18519115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18519115&pc=OZD | 01/01/2024 | 01/26/2024 |  | 01/16/2024 |  |  | 01/01/2024 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024418835 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2024 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to advance through the descending aorta due to existing clot burden that wrapped around the catheter. as a result of the patient's anatomy, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, the root cause of the delivery issue was most likely due to patient condition as the case was aborted due to extensive clots while attempting the device insertion. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31702 | 1220648-2024-06061 | 18519459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18519459&pc=PYX | 01/01/2024 | 12/09/2024 |  | 01/16/2024 |  |  | 01/01/2024 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024413645 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2024 |  |  |  |  | The user facility reported a 37-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella rp flex support experienced significant hemolysis. the pump was repositioned which helps to improve the hemolysis, but ldh (lactate dehydrogenase) remained elevated. one unit of prbc's was provided as treatment. the patient is stable. | [5 empty] | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. the cause of the hemolysis was most likely patient condition related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06061. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31703 | 1220648-2024-06055 | 18519733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18519733&pc=OZD | 12/30/2023 | 02/16/2024 |  | 01/16/2024 |  |  | 12/30/2023 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024421428 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a (b)(6) female of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was delivered through the 14fr sheath, but the pump was unable to be delivered via the sheath due to anatomy. the team aborted the impella cp and performed the percutaneous coronary intervention (pci) procedure via the 14fr sheath. no patient harm was reported. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for investigation. based on clinical description, the root cause of the delivery issue was most likely due to the patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31704 | 1220648-2024-06062 | 18519773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18519773&pc=OZD | 01/02/2024 | 11/30/2025 |  | 01/16/2024 |  |  | 01/02/2024 | 01/16/2024 |  |  |  |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1130666 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 01/02/2024 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported the automated impella controller (aic) showed failure controller alarm prior to prep for potential implant. therefore, a new aic was used. there was no reported patient harm. | [5 empty] | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name updated. f6 and f8 should have been left blank. updated h3 to device evaluation anticipated. h6 type of investigation code added. | The impella device was received from the customer on (b)(6)2024 and therefore, an evaluation of the is under way. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. console logs from the reported day of event are mostly consistent with complaint and show controller error alarm (#1039, defective optical sensor lamp), preceded by optical bench resets in response to failing ¿5s¿ parameters: 02/01/24 10:23:51 ossiopsens state ossi_bad_lamp_reset_bench_state: lamp: 100.00%, light: 0.52v, cont.: -272.54%, gain: 1.02, snr: 0.5, mano: 747.1. console log data from prior case with pump (b)(6) indicated sudden loss of light in the optical system, manifested by erratic values of contrast between -3000 and +3000. there were no controller failure alarms, either on aic boot-up or during prior case. the issue was not reproduced during testing in aachen, however, his type of event (manifested by negative and oscillating contrast of the optical signal) is known to be caused by intermittent light source, either due to lamp assembly failure or opm electronics failure. there were no controller failure alarms. the cause of the aic issue was a defective optical bench. d4 model number updated.
31705 | 1220648-2024-06056 | 18519806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18519806&pc=OZD | 12/28/2023 | 11/30/2024 |  | 01/16/2024 |  |  | 12/28/2023 | 01/16/2024 |  |  | 08/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024392139 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | Unspecified Heart Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a 83-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and experienced pericardial effusion. the cause is unknown for the pericardial effusion. the impella was explanted. | [5 empty] | The investigation into ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06056: d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. h6 code 4454 was reported incorrectly. new code was added to health effect - clinical code. h10 instructions for use missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31706 | 1220648-2024-06060 | 18519817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18519817&pc=OZD | 12/29/2023 | 05/01/2024 |  | 01/16/2024 |  |  | 12/29/2023 | 01/16/2024 |  |  |  |  | 01/16/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146768 |  | 0042-000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | No patient interaction: us complainant reported the automated impella controller (aic) came back from preventative maintenance service and was found that the rotary knob was not working correctly. the aic was unable to change selections. the aic was replaced. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) rotary knob issues has been completed. the console was returned for evaluation. physical inspection of the console confirmed the reported complaint. the rotary knob did not consistently respond to rotations, and it responded inconsistently to button presses. this issue made it impossible to properly navigate through the user interface. the data logs were reviewed finding no errors or malfunctions were identified in the logs. the root cause of the issue was identified as the ribbon cable connecting the impellatronic board to the 4-in-1 being incorrectly inserted in the cable connector on the impellatronic board. d4-updated model number.
31707 | 1220648-2024-06064 | 18519833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18519833&pc=OZD | 01/03/2024 | 09/20/2024 |  | 01/16/2024 |  |  | 01/03/2024 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024410632 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to st elevated myocardial infarction. the patient experienced access site bleeding post implant. femstop was placed and oozing slowed. no blood products were given. additionally, hematuria was noted in the patient after the implant. the impella was found too deep and the pump was repositioned. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06064 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31708 | 1220648-2024-06065 | 18519841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18519841&pc=OZD | 01/02/2024 | 12/19/2024 |  | 01/16/2024 |  |  | 01/02/2024 | 01/16/2024 |  |  | 09/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410680 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a impella cp support ongoing when the pump came out of position. the pump was being adjusted by the md in the icu. the pump was explanted and replaced. at the time of explant the team placed a dry-seal sheath to alleviate bleed and placed a antegrade sheath for limb perfusion. the pulses were hard to locate in both limbs. | [5 empty] | The investigation for the reported positioning issue and limb ischemia has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue was most likely use related. the root cause of the limb ischemia could not be determined without additional clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: d4 updated model number added d6a/d6b was missing in initial. f6/f8 should have been left blank. g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ assess access site for bleeding and hematoma
31709 | 1220648-2024-06067 | 18520113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18520113&pc=OZD | 01/03/2024 | 02/21/2024 |  | 01/16/2024 |  |  | 01/03/2024 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024418807 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support for heart failure/left ventricular assist device work up. while on support, suction was occurring and they were concerned that there was a clot or fibrin formation. the impella was explanted and replaced. after the explant, a clot formation was found on the outlet cage and motor housing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the reported low pump flow has been completed since the original report was filed. pump was returned for evaluation. biomaterial was found on the outlet cage and impeller of the pump. no other issues with the pump were observed. no data logs were available for day complaint occurred. the root cause of the low pump flow was most likely due to biomaterial on the outflow cage in impeller.
31710 | 1220648-2024-06068 | 18520359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18520359&pc=OZD | 12/29/2023 | 09/23/2025 |  | 01/16/2024 |  |  | 12/29/2023 | 01/16/2024 |  |  | 06/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390571 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 35-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that one day following impella cp implant, the patient developed some bleeding at the cutdown site. one unit of blood was transfused to treat the noted condition and support was maintained. it was further noted that the placement signal was unreliable on multiple occasions and the alarm was ultimately ignored by the staff. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. impella cp was returned for evaluation. visually inspected and no catheter kink, physical abnormalities or biomaterial was observed. no other issues were found during product investigation. imc logs were verified, and no abnormalities were found. the cause of the access site bleeding issue was not determined due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical bench 5s parameters were verified and all the values observed to be in a normal range. light continuity test was done, and light was successfully passing through the sensor without any issues. no other issues were found during product investigation. data logs were reviewed, and placement signal unreliable alarms were observed with some suction alarms. the 5s parameters were observed to be varying with no particular failure trend. the cause of the optical signal issue was not determined since issue did not reproduce during field investigation and no failure trend observed per log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-06068. d4 model number updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should not have had entries in the initial report. g1 reporting contact email and reporting contact fax number updated.
31711 | 1220648-2024-06070 | 18520437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18520437&pc=OZD | 12/30/2023 | 02/22/2024 |  | 01/16/2024 |  |  | 12/30/2023 | 01/16/2024 |  |  | 09/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2024406979 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility in japan reported a 68-year-old male with arrhythmogenic heart failure was implanted with an impella cp for mechanical circulatory support. it was reported that the patient developed hemolysis. the pump flow was lowered, and blood products were provided. the decision was made to explant the impella cp pump due to only slight improvement was noted after intervention. | [5 empty] | The investigation into the hemolysis issue has been completed. the device was not returned for investigation. the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31712 | 1220648-2024-06069 | 18520543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18520543&pc=OZD | 01/04/2024 | 07/26/2024 |  | 01/16/2024 |  |  | 01/04/2024 | 01/16/2024 |  |  | 11/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429120 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Suction Problem | Structural Problem | Deformation Due to Compressive Stress |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp delivered to support a patient s/p avr (aortic valve replacement) and mvr (mitral valve replacement) from day prior. the pump was delivered to the lv (left ventricle). the cp failed at the setup of the cassette. the cp was replaced. once the cp was delivered the pump kinked and had suction. the team had to replaced the pump when team unable to troubleshoot the suction. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage/kink and low pump flow has been completed. data logs were returned and investigated. per the clinical information, the cannula was kinked during insertion and patient had small lv which resulted in suction events and low flow. the root cause of the low pump flow issue was patient condition related to patient's anatomy, when the cannula kinked upon insertion. added- h6 med dev prob code 1248 d4-corrected model number added- d6a/d6b f6/f8 should have been left blank in the initial report. h6 med dev prob code 2506 changed to 2889.
31713 | 1220648-2024-06071 | 18520737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18520737&pc=PYX | 01/04/2024 | 01/26/2024 |  | 01/16/2024 |  |  | 01/04/2024 | 01/16/2024 |  |  | 01/31/2024 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023167658 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 58-year-old female was to be implanted with an impella rp due to a pulmonary embolism. the physician had delivery issues. the wire was placed without difficulty, but despite multiple attempts, they were unable to advance the impella rp into the pulmonary artery. physician also tried 0.25 platinum plus wire without success. other doctors attempted the implant without success. the impella implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, the root cause of the delivery issue was most likely due to patient condition as the case was aborted; patient had very dilated rv and is only 160cm.
31714 | 1220648-2024-06073 | 18521403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18521403&pc=OZD | 01/06/2024 | 09/19/2024 |  | 01/16/2024 |  |  | 01/06/2024 | 01/16/2024 |  |  | 07/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024391314 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2024 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to anterior st elevated myocardial infarction. while on support, the patient was hemolyzing and anuric. the impella was repositioned, and the patient received two units of blood products and followed each infusion with 40 of lasix. the staff was monitoring the kidney. it is unknown if positioning issues helped or resolved the hemolysis. the patient expired after 29.44 hours of support. no allegations were made towards the impella for cause of death. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06073 was provided. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
31715 | 1220648-2024-06074 | 18521689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18521689&pc=PYX | 01/10/2024 | 05/15/2024 |  | 01/16/2024 |  |  | 01/10/2024 | 01/16/2024 |  |  | 10/31/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024431799 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 50-year-old male of unknown race undergoing coronary artery bypass grafting was implanted with an impella rp device for mechanical circulatory support. it was reported that while in the catheterization lab using the same internal jugular (ij) access site, an 8 french sheath and 15 french repositioning unit were placed when the oozing started. the physician attempted to gain vessel control using modified purse string suture, but it broke. several more mattress sutures were place around the site until the oozing was under control. once the patient started to wake up, they started to cough continually. this resulted in the bleeding to resume from the rp access site. two units of blood were given, and manual pressure was applied multiple times. a pressure dressing was applied, and the patient was sent on support to the unit. the patient was reported as stable. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the pump was returned for investigation and there were no damages noted. the root cause of the access site bleeding was most likely due to use issue. high purge pressure was noted in the case. the product was inspected and there was biomaterial seen on the impeller and in the purge gap. the root cause of the high purge pressure is due to biomaterial. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ "be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿
31716 | 1220648-2024-06075 | 18522760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18522760&pc=OZD | 01/04/2024 | 02/07/2024 |  | 01/16/2024 |  |  | 01/04/2024 | 01/16/2024 |  |  | 03/31/2024 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2023206890 |  | 0048-0014 EU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 48-year-old female (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the purge reservoir broke. the team replaced the impella cp with another cp. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported purge reservoir break has been completed. the impella cp was returned with purge system bypassed with needle. the purge sidearm was returned with purge line cut at the filter end to the red plug. the purge sidearm was visually inspected and the sidearm joint to the reservoir was observed to be cracked and damaged due to characteristic of excessive force. no other abnormalities were found. the root cause of the purge reservoir break was user-related as excessive force caused damage to the yellow luer to the reservoir joint. change h6 impact code 2199 to 4629.
31717 | 1220648-2024-06076 | 18522766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18522766&pc=OZD | 01/05/2024 | 01/29/2025 |  | 01/16/2024 |  |  | 01/05/2024 | 01/16/2024 |  |  | 09/30/2025 |  | 01/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024398958 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 71-year-old male of unknown race noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced pulsatile bleeding at the impella access site during the transition between the peel-away sheath and repositioning sheath. no bleeding or hematoma was noted following advancement of the repositioning sheath. approximately, four days later, the pump was noted to be mispositioned in the aorta. as a result of the poor positioning the decision was made to explant the device. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a and d6b revised from the initial submission in accordance with updated procedures on manufacturer device report 1220648-2024-06076 f6/f8-should have been left blank on manufacturer device report 1220648-2024-06076 g1 reporting contact email revised in accordance with updated procedures on manufacturer device report 1220648-2024-06076. g1 reporting contact fax number missing on manufacturer device report 1220648-2024-06076.
31718 | 1220648-2024-06031 | 18510441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18510441&pc=OZD | 12/20/2023 | 04/18/2024 |  | 01/15/2024 |  |  | 12/20/2023 | 01/15/2024 |  |  | 08/31/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024384980 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported a 45-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was noted that the patient was bleeding at the insertion site and a decision was made to place a purse suture to stop the bleeding, which was successful, and the bleeding stopped. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related. purse suture was placed successfully to stop the bleeding as per the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31719 | 1220648-2024-06032 | 18510773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18510773&pc=OZD | 08/14/2023 | 01/15/2024 |  | 01/15/2024 |  |  | 12/18/2023 | 01/15/2024 |  |  | 01/31/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310720 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the impella had bad purge cassette which resulted into high purge pressure/purge system blocked. issue resolved by changing purge cassette. no patient harm reported. | [5 empty] | The investigation into the reported high purge pressure has been completed. as per clinical high purge pressure and purge system blocked alarm were observed. impella cp pump was not returned, but the purge cassette was returned. the purge cassette was connected to console and prepped for priming. the green fluid started to purge with flow of 30 ml/hr but later after 10 minutes of continuous purging with sample pump it normalized with normal purge flow of 15 to 16 ml/hr and purge pressure was between 400 to 450 mmhg. no leaks or other abnormalities were found. data logs were reviewed and there was immediate purge pressure rise within a minute, there was a stepwise purge flow decrease and purge flow ticks were stalling during this high purge pressure event. the cause of the high purge pressure issue was most likely kinked purge line in the catheter per data log review. the failure modes will be monitored and trended.
31720 | 1220648-2024-06033 | 18513057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18513057&pc=OZD | 01/05/2024 | 02/20/2024 |  | 01/15/2024 |  |  | 01/05/2024 | 01/15/2024 |  |  | 09/30/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024398330 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Insufficient Flow or Under Infusion | Improper Flow or Infusion |  |  |  |  | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported an 80-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp pump was unable to achieve flows above 2 liters a minute during patient support. despite attempts to reposition the pump, the issue persisted, and the decision was made to replace the device. upon removal of the original pump, a clot was noted at the inlet of the device. planned support continued with the newly implanted pump. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the product was returned for investigation. the pump was visually inspected. no cannula kink/ no broken blades observed. biomaterial was observed around the impeller and the inflow cage. the cause of the low pump flow issue was biomaterial ingestion. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31721 | 1220648-2024-06034 | 18513357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18513357&pc=OZD | 12/26/2023 | 02/23/2024 |  | 01/15/2024 |  |  | 12/26/2023 | 01/15/2024 |  |  | 09/30/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403887 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported a 73-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient developed hemolysis during impella 5.5 support. the patient was given one unit of platelets to counteract the condition and support continued with the implanted pump. | [5 empty] | The investigation into the reported hemolysis been completed since the original report was filed. according to the clinical, three days into impella 5.5 support hemolysis was suspected. hemolysis was resolved two days later, but no information was given as to how this issue was resolved. no product was returned for a visual inspection. data log analysis confirms that there were no sustained positioning issues. the cause of the hemolysis was not determined because no product was returned and not enough information was given. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31722 | 1220648-2024-06035 | 18514532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18514532&pc=OZD | 01/03/2024 | 02/23/2024 |  | 01/15/2024 |  |  | 01/03/2024 | 01/15/2024 |  |  | 10/31/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420130 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support via femoral artery access. the pump was placed however after just over 10 hours, the pump was explanted due to concerns of limb ischemia. the foot was cold. upon explant of the pump the team found no flow beyond the iliac artery. vascular surgery was consulted. it is unknown if surgery was required at this time. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ irreversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31723 | 1220648-2024-06036 | 18514712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18514712&pc=OZD | 12/29/2023 | 02/23/2024 |  | 01/15/2024 |  |  | 12/29/2023 | 01/15/2024 |  |  | 10/31/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMEPLLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024420077 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | FEM-FEM BYPASS | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 74-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that two days after the impella cp implant, it was difficult to detect a distal pulse in the left leg. due to the noted condition, a contralateral fem-fem bypass was placed. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31724 | 1220648-2024-06037 | 18514871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18514871&pc=OZD | 12/31/2023 | 10/31/2024 |  | 01/15/2024 |  |  | 12/31/2023 | 01/15/2024 |  |  | 10/31/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420111 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Retroperitoneal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/31/2023 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a retroperitoneal bleed during impella cp support. the source for the bleed was unknown. as a result of the condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported retroperitoneal bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the retroperitoneal bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections for the initial mfr report: added d6a/d6b. f6/f8 should have been left blank.
31725 | 1220648-2024-06038 | 18514907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18514907&pc=OZD | 12/31/2023 | 02/20/2024 |  | 01/15/2024 |  |  | 12/31/2023 | 01/15/2024 |  |  |  |  | 01/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1547850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/31/2023 |  |  |  |  | The user facility reported a 31-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the automated impella controller (aic) was measuring high purge flow and low purge pressure during patient support. upon inspection of the unit, it was discovered that the purge disc housing was broken. the aic was swapped as a result of the noted damage. there was no patient harm reported. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. purge retainer was confirmed to be broken. the cause of the issue was a broken purge retainer. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31726 | 1220648-2024-06039 | 18514916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18514916&pc=OZD | 01/02/2024 | 03/05/2024 |  | 01/15/2024 |  |  | 01/02/2024 | 01/15/2024 |  |  |  |  | 01/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1276240 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) purge disk transducer broke during a cassette change. the aic was replaced. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon visual inspection of the console, it was found that the purge retainer was broken. the cause of the issue was a broken retainer likely occurred throughout normal use of the device. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31727 | 1220648-2024-06040 | 18514962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18514962&pc=OZD | 01/03/2024 | 02/26/2024 |  | 01/15/2024 |  |  | 01/03/2024 | 01/15/2024 |  |  |  |  | 01/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1109488 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) kept shutting down and restarting on its own. the aic was being set-up for use and there was no patient involvement at the time of the malfunction. the aic was replaced without issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the reported pump stop has been completed since the original report was filed. this console was tested after arriving in fs. unable to reproduce the reported failure mode. no automatic reboot occurred, the console logged the data, and no issue occurred upon running the aic with the pump for 3 hours. no issue was found with the flash cards. found a hair crack on the purge retainer (unrelated to the reported failure). the root cause of intermittent unexpected controller shutdown was not determined due to the inability to reproduce.
31728 | 1220648-2024-06041 | 18514992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18514992&pc=OZD | 01/03/2024 | 01/15/2024 |  | 01/15/2024 |  |  | 01/03/2024 | 01/15/2024 |  |  |  |  | 01/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1161034 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 01/03/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a battery failure alarm on start-up. there was no patient involvement at the time of the noted issue. after troubleshooting was unsuccessful in resolving the alarm, the decision was made to remove the device from service. | [5 empty] | Data review: after reviewing the imc logs, the issues with the battery failure alarms on this console were confirmed. there were multiple instances of battery failure alarms (transport connection blown/missing) (event set #360) and single instance of battery failure alarm (low voltage) [v < 12v] (event set #350) found from the reported occurrence date. results of running batttest on telnet revealed that ¿fu¿ was missing on battery 2 and that cell 1 in battery 2 had a voltage of 3871mv which is about ~200mv less than the rest of the cells in that battery. ltlog power data supports the finding that there is a cell imbalance in battery 2 found in cell1 (the line for cell1 shown in light blue, varies significantly from the rest of the cells in the battery). see attached images for batttest results and ltlog power data. device analysis: console was tested after arriving in fs. battery failure alarm issue was able to be reproduced. lcd charge level indicator on battery 2 was malfunctioning (see attached image). issue was determined to be caused by internal battery cell imbalance (found in cell1). before sending this console back to the customer, fs was instructed to replace battery set [3900-0032 aic - battery (6120mah)], verify battery functionality via sections 18 through 20 of pm procedure [0042-7309], and perform a full functional check on the console and optical system [0042-7316].
31729 | 1220648-2024-06042 | 18515055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18515055&pc=OZD | 01/03/2024 | 12/10/2024 |  | 01/15/2024 |  |  | 01/03/2024 | 01/15/2024 |  |  | 10/31/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424149 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 65-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support due to non-st elevated myocardial infarction. upon impella implant, there was access site bleeding. one of the perclose sutures failed or dislodged. staff placed purse-string suture at site to help with oozing. additionally, the patient lost pulses however, they came back after after warming and starting dobutamine. furthermore, the patient had a red tinged in their urine. the impella was turned down and labs were sent. the impella was removed the following day. a clot burden was found in the left iliac post closure of the impella. it is unknown if the clot is related to the impella, however, the patient was not on heparin. | [5 empty] | The investigation into access site bleeding, hemolysis, and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06042: d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email. h6 component code revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31730 | 1220648-2024-06043 | 18515067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18515067&pc=OZD | 01/03/2024 | 03/25/2024 |  | 01/15/2024 |  |  |  | 01/15/2024 |  |  | 10/31/2025 |  | 01/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024419223 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 47-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported the patient developed a small hematoma with oozing at the insertion site. the heparin was stopped and the oozing resolved. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of access site bleeding was not determined due to a lack of clinical information detailing the cause of the bleed, and the product was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury
31731 | 1220648-2024-06020 | 18502382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18502382&pc=OZD | 12/15/2023 | 09/20/2024 |  | 01/12/2024 |  |  | 12/15/2023 | 01/12/2024 |  |  | 01/31/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333651 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant was on impella cp support due to the 66-year-old male patient having chest pain. the patient coded and return of spontaneous circulation was able to be achieved. however, the impella was pulled out, due to anatomy of the left ventricle. the cp was retrograded out of the left ventricle and was not able to cross aortic valve. a second device was inserted successfully. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-06020 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
31732 | 1220648-2024-06021 | 18503481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18503481&pc=OZD | 12/21/2023 | 10/08/2024 |  | 01/12/2024 |  |  | 12/21/2023 | 01/12/2024 |  |  | 12/31/2024 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024307478 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed ischemia in their impella limb during pump support. it was noted that no pulse was present in the right lower extremity and the decision was made to escalate to impella 5.5 support. upon cp explant, a femoral endarterectomy and thrombectomy were performed prior to site closure via femoral cutdown. the patient was stable following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system states the following section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 impact code 4629, 4625 corrections to the initial mfr report: d4-corrected model number added-d6a/d6b f6/f8 should have been left blank h5-labeled for single use changed to yes.
31733 | 1220648-2024-06022 | 18503505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18503505&pc=OZD | 12/28/2023 | 05/06/2024 |  | 01/12/2024 |  |  | 12/28/2023 | 01/12/2024 |  |  | 10/31/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420590 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, some oozing was noted at the groin sight of the 58-year-old female patient. heparin was stopped as a result of the condition. support was maintained with the implanted pump. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31734 | 1220648-2024-06023 | 18503536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18503536&pc=OZD | 12/22/2023 | 01/12/2024 |  | 01/12/2024 |  |  | 12/25/2023 | 01/12/2024 |  |  | 06/30/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024381564 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/25/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump was unable to advance beyond the descending aorta during attempted delivery. it was noted that the patient's medical history was significant for an aortic arch artificial blood vessel replacement and descending aorta stent graft. due to the encountered resistance, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer. no product was returned. the root cause of the delivery issue was most likely patient condition related.
31735 | 1220648-2024-06024 | 18503575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18503575&pc=OZD | 12/25/2023 | 12/09/2024 |  | 01/12/2024 |  |  | 12/25/2023 | 01/12/2024 |  |  | 02/28/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321759 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2023 |  |  |  |  | The user facility reported a 34-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed hemolysis following impella cp implant. although repositioning was initially discussed, the decision was ultimately made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. product was not returned. data logs showed no motor current (mc) spikes outside of p level changes. there are brief suction alarms on day 1 that resolve. the cause of the hemolysis was not determined due to no product returned, lack of clinical information, and inconclusive data logs. added h6 impact code 4628, component code 925 corrections to the initial mfr report: d4 model number. added d6a/d6b. g1 mdr reporting contact email. f6/f8 should have been left blank.
31736 | 1220648-2024-06025 | 18503622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18503622&pc=OZD | 12/26/2023 | 01/31/2024 |  | 01/12/2024 |  |  | 12/27/2023 | 01/12/2024 |  |  |  |  | 01/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1140917 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was on support when the aic had a purge disc not detected alarm. the staff pushed on the purge disc and the alarm resolved. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation. after reviewing the imc logs, the issues with the automated impella controller (aic) not being able to detect the purge disc was confirmed. there were multiple instances of purge disc not detected alarms on the day after the reported occurrence date. the console was tested after arriving. the purge disc not detected issue was able to be reproduced and the purge flag was found to be broken. the root cause of this issue was a broken purge flag. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31737 | 1220648-2024-06026 | 18505198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18505198&pc=OZD | 12/27/2023 | 10/02/2024 |  | 01/12/2024 |  |  | 12/27/2023 | 01/12/2024 |  |  | 10/31/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414379 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 73-year-old male patient was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support when they experienced a hematoma at the access site. blood products were given, the pump was removed and the patient had vascular surgery repair. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleed leading to the vascular damage issue was most likely use related, as surgeon noted small branch at insertion site which may have been sheared with peel away removal as per the clinical description. added- h6 component code 925 corrections for the initial mfr report: d4-corrected model number added- d6a/d6b f6/f8 should have been left blank. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31738 | 1220648-2024-06027 | 18505324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18505324&pc=OZD | 12/28/2023 | 04/11/2024 |  | 01/12/2024 |  |  | 12/28/2023 | 01/12/2024 |  |  | 09/30/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408555 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | Air/Gas in Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a (b)(6) female patient was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and the pump would not prime with purge. multiple de-airing attempts were made to de-air. staff made the decision to replace pump and purge cassette. new device was successful. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported pump issue has been completed. neither the impella device nor data logs were returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues could not be determined. d4-updated model number.
31739 | 1220648-2024-06028 | 18505856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18505856&pc=OZD | 12/28/2023 | 02/19/2024 |  | 01/12/2024 |  |  | 12/28/2023 | 01/12/2024 |  |  | 10/31/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024413258 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was scheduled for impella implant for mechanical circulatory support. the us complainant was attempting to deliver an impella 5.5 to support a patient with cardiomyopathy and cgs (cardiogenic shock). the patient's anatomy precluded placement. the 5.5 could not traverse and deliver due to small artery and anatomic limitations. the team instead placed a cp. no harm or injury. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for investigation. based on clinical description, the root cause of the delivery issue was most likely due to the patient condition. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31740 | 1220648-2024-06029 | 18509036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18509036&pc=OZD | 12/18/2023 | 11/08/2024 |  | 01/12/2024 |  |  | 12/18/2023 | 01/12/2024 |  |  | 10/31/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420122 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a (b)(6). female (race unknown) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the team found the native anatomy to preclude placement and so vascular consulted and ballooned the native anatomy. the cp was removed as it failed delivery and kinked. the team placed a new second pump with difficulty. this report is for the second pump that had delivery issues. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage and delivery issues has been completed. pump was not returned for investigation. as per clinical information provided, the patient had tortuous arteries making it difficult to pass pump through the vasculature. therefore, the root cause of the delivery issue was patient condition related to patients tortuous vessels. the root cause of the product damage issue was a catheter kink due to patients tortuous vessels. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ corrections to the initial mfr report: b1-selected adverse event. b2-selected death and added death. b5-mso review. added d6a/d6b. f6/f8 should have been left blank. h1-type of reportable event changed to death. h6 impact code 4629.
31741 | 1220648-2024-06030 | 18509316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18509316&pc=OZD | 12/16/2023 | 12/31/2024 |  | 01/12/2024 |  |  | 12/16/2023 | 01/12/2024 |  |  | 10/31/2025 |  | 01/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418150 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2023 |  |  |  |  | The user facility reported a (b)(6) female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was on impella support for 0.05 hours before being removed and a new impella was inserted. it was reported while taking out the peel away sheath the device was advanced all the way into the ventricle and the cannula kinked on itself. the device was removed from the heart and turned off. it was then pulled back, unkinked in the iliac, and removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage-kink has been completed. the pump was not returned for investigation. without the product the failure cannot be further investigated. therefore, the root cause of the product damage -cannula kink issue was not determined since product was not returned and insufficient clinical information was provided. corrections to the initial mfr report: d4 udi number updated. added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email.
31742 | 1220648-2024-06012 | 18491154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18491154&pc=OZD | 12/31/2023 | 12/12/2024 |  | 01/11/2024 |  |  | 01/01/2024 | 01/11/2024 |  |  | 09/30/2025 |  | 01/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410023 |  | 0042-0040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2024 |  |  |  |  | The user facility reported a 31-year-old black male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that one day after impella cp implant, the patient developed bleeding at their access site. the patient was given three units of blood to counteract the blood loss and both stitching and femostop were placed to manage the condition. support continued following the intervention, but positioning issues were encountered the following day. the patient was noted to be sensitive to movement and the pump repeatedly moved out of position, resulting in alarms and suction. as a result of the ongoing complication, the decision was made to escalate to impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the access site bleed and the positioning issue has been completed. pump was not returned for investigation. clinical mentions that on day 3 "impella in aorta" alarms were observed but the reason of the pump being pulled into aorta is unknown. therefore, the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. on day 2, access site was observed to be bleeding but the reason of bleeding is unknown. the bleeding was stopped using stitching and femostop. therefore, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added h6 impact code 4629 corrections to the initial mfr report: b1-should have selected product problem. d4 model number. added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email.
31743 | 1220648-2024-06015 | 18491777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18491777&pc=OZD | 12/22/2023 | 12/04/2024 |  | 01/11/2024 |  |  | 12/22/2023 | 01/11/2024 |  |  | 09/30/2025 |  | 01/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397376 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 34-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there was bleeding from the impella access site. a femstop was placed and the bleeding was stable. additionally, the doctor made the choice to wean and remove the impella due to bleeding and hemolysis. the hemolysis was going to be confirmed by blood test of plasma free hemoglobin. results of that test are unknown. upon removal, one of the perclose suture broke when tying down. femostop placed for extra support. | [5 empty] | The investigation into hemolysis and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-06015. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
31744 | 1220648-2024-06013 | 18491813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18491813&pc=OZD | 01/04/2024 | 02/23/2024 |  | 01/11/2024 |  |  | 01/04/2024 | 01/11/2024 |  |  | 09/30/2025 |  | 01/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024407246 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a (b)(6)female of unknown race with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp pump stopped working roughly twenty-one days into support. although the device successfully restarted, the motor current remained high, and the decision was made to escalate to impella 5.5 support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. according to the clinical, on the 21st day of impella cp support a pump stop occurred. the pump was explanted and replaced with a 5.5. no product was returned for a visual inspection. data log analysis confirms high motor current trends and low/high flow around the time of the pump stop. the cause of the pump stop is not determined because not enough product or clinical information was returned.
31745 | 1220648-2024-06017 | 18491873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18491873&pc=OZD | 12/26/2023 | 04/16/2024 |  | 01/11/2024 |  |  | 12/26/2023 | 01/11/2024 |  |  | 08/31/2025 |  | 01/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385386 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to left main disease and for high risk percutaneous coronary intervention. the patient had access site bleeding and the patient received 2 units of blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that post procedure, the patient's activated coagulation time (act) was prolonged at 280. therefore, the root cause of the access site bleed was most likely due to improper anticoagulation management. d4-updated model number
31746 | 1220648-2024-06016 | 18494838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18494838&pc=OZD | 12/22/2023 | 12/10/2024 |  | 01/11/2024 |  |  | 12/22/2023 | 01/11/2024 |  |  | 07/31/2025 |  | 01/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376892 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient had an increase of plasma free hemoglobin (pfhgb) and was also showing signs of distal limb ischemia. the patient was brought to the operation room and the impella cp was removed and impella 5.5 was inserted. the patient is stable. | [5 empty] | The investigation the hemolysis and limb ischemia has been completed. product was not returned. data logs were reviewed which showed no relevant alarms and no motor current (mc) spikes outside of p level changes. no other abnormalities. the cause of the hemolysis was not determined since product was not returned, lack of clinical details, and inconclusive data logs. the root cause of the limb ischemia could not be determined without additional clinical details. corrections to the initial mfr report: d4 model number. added d6a/d6b. g1 mdr reporting contact email. f6/f8 should have been left blank. h6 impact code 4627 should be 4629. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ . ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31747 | 1220648-2024-06018 | 18497704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18497704&pc=OZD | 12/27/2023 | 11/17/2025 |  | 01/11/2024 |  |  | 12/27/2023 | 01/11/2024 |  |  |  |  | 01/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1154074 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 12/27/2023 |  |  |  |  | The user facility reported a patient was to have an impella device for mechanical circulatory support. the complainant reported the the automated impella controller (aic) experienced a red alarm, battery failure. the staff attempted to troubleshooted and tested, but was unsuccessful. the aic has been sent in for investigation. | [5 empty] | The automated impella controller was returned on 01/03/2024, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
31748 | 1220648-2024-06019 | 18497916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18497916&pc=OZD | 12/27/2023 | 02/21/2024 |  | 01/11/2024 |  |  |  | 01/11/2024 |  |  |  |  | 01/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1126030 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported via the hospital biomedical department there was a console error. us complainant reported the automated impella controller (aic) did not recognize the purge disc. the aic was replaced. | [5 empty] | Automated impella controller was returned to abiomed on 01/04/2024, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. aic's inability to detect the purge disc was confirmed. the cause of the issue was a broken purge flag.
31749 | 1220648-2024-05976 | 18483171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483171&pc=OZD | 12/27/2023 | 10/08/2024 |  | 01/10/2024 |  |  | 12/27/2023 | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414379 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when inserting the impella cp an uncontrolled hematoma formed in the groin leading to impella removal and vascular repair. six units of blood were administered. the patient was escalated to impella 5.5 insertion. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05976 d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. h6 health effect - impact code 4641 and 4627 were missing.
31750 | 1220648-2024-05977 | 18483260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483260&pc=OZD | 12/26/2023 | 11/19/2025 |  | 01/10/2024 |  |  | 12/28/2023 | 01/10/2024 |  |  |  |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 2024376658 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a 46-year-old black male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that their aic (automated impella controller) started to alarm during patient support and the screen froze. when the touch screen could not be manipulated, the decision was made to replace the aic. the pump flow remained unchanged throughout the event. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data analysis: on dec 27th, 2023, since the console displayed ¿purge volume low ¿ alarm,¿ the user performed the ¿cassette and bag change procedure.¿ during the fast prime, rt logs indicate that the purge pressure amplitude did not exceed 100 mmhg (complaint_date_purgepressure_during_fast_prime.png). typically, the purge pressure amplitude reaches above 500 mmhg. this indicates that there may have been an open or kink in the purge line, or air may have been introduced into the purge line due to an inverted bag. the console remained in ¿fast prime mode¿ for 32 seconds and in ¿slow prime mode¿ for 28 seconds. during this period, the user touched the soft key (most likely attempting to end the procedure). on the priming screen, the console displayed the message ¿complete the purge wizard ¿ alarm,¿ event set #420 (imc_logs_ic1827.pdf). however, per the design of the priming screen, the aic does not provide an option to interrupt the procedure using the soft key or rotary push button. the 'complete the purge wizard' alarm on the priming screen and the priming screen not progressing were misreported as a frozen screen. device analysis: this console was tested upon arrival at fs, but the issue could not be reproduced despite following the same sequence of actions that occurred at the customer site. after triggering the ¿purge volume critically low ¿ alarm,¿ event #412, the ¿cassette and bag change procedure¿ was completed successfully, and event #412 was resolved. the rt log shows that the purge pressure periodically reached 500 mmhg during the fast prime mode. no issues were found with the soft keys or rotary push button. no issue was found on the screen upon wiggling the groups of wires that connect the display to the 4-in-1 connectors x26 and x51. no loose contact was found. (note: any loose connection on these two connectors/wires could freeze or turn the screen black.) x26- lvds connector (for high-speed differential signal for lcd) x51- backlight connector (to control the power for the lcd backlight unit) a mild fluid ingress was observed between the motor spline gasket, and the rear of the purge cassette insert. however, no issues were found with the purge flow ticks on the complaint date. found a hair crack on the purge retainer (unrelated to the reported failure mode). root cause: the 'complete the purge wizard' alarm on the priming screen and the priming screen not progressing were misreported as a frozen screen. the non-progressing priming screen could result from an open or kinked purge line, or air introduced into the purge line due to the inverted bag. therefore, the root cause is most likely a use issue. d9: updated the device returned to status g1: corrected the manufacturer email address h6: added codes per the results of the investigation and the clinical code h10: added the results of the investigation
31751 | 1220648-2024-05978 | 18483296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483296&pc=OZD | 12/29/2023 | 05/14/2024 |  | 01/10/2024 |  |  | 12/29/2023 | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419057 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 92-year-old white female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support pre-percutaneous coronary intervention. during support, the impella had saturated pump flows. a clot had migrated into the left ventricle and into the pump causing pump motor saturation. the impella was replaced with no patient harm. | [5 empty] | The investigation for the reported saturated flow has been completed. the pump was not returned for investigation. data logs show the pump reaching saturated flow after reported motor current spike. the doctor explanted the pump due to elevated motor current and saturated pump flow. the cause of the saturated pump flow issue was most likely biomaterial, based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31752 | 1220648-2024-05979 | 18483306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483306&pc=OZD | 12/29/2023 | 01/24/2024 |  | 01/10/2024 |  |  | 12/29/2023 | 01/10/2024 |  |  | 11/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422095 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a (b)(6) male (race unknown) in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the physician attempted to insert the impella 5.5 via the right axillary site but was unable to pass the innominate artery due to the patient's anatomy. the vessel was tortuous. a left axillary approach was attempted afterwards, but unsuccessful also due to the patient¿s anatomy. the impella cp was placed successfully via left axillary. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31753 | 1220648-2024-05980 | 18483370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483370&pc=OZD | 12/24/2023 | 09/05/2024 |  | 01/10/2024 |  |  | 12/25/2023 | 01/10/2024 |  |  |  |  | 01/10/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1502288 |  | 0042-0040-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/25/2023 |  |  |  |  | The japan user facility reported a 62-year-old male patient was on impella cp support and there was pump stop. the team would like the automated impella controller (aic) to be inspected because the cause is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency. | An investigation of the automated impella controller (aic) was requested by the customer because the cause of the event was unknown. the engineer evaluated the aic and determined there was no issue with the console. a regulatory report is no longer necessary.
31754 | 1220648-2024-05981 | 18483393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483393&pc=OZD | 12/29/2023 | 02/22/2024 |  | 01/10/2024 |  |  | 12/29/2023 | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024419218 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 62-year-old black female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. two hours into impella 5.5 support, the pump ingested a clot and stopped. the pump was explanted and replaced with a new impella 5.5 device for ongoing support. there was no patient injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed. the device was not returned for investigation. it was reported that the pump had a high motor current pump stop 7 seconds later. they attempted to restart the pump 5 times before it was then explanted. based on the data analysis, the root cause of the pump stop is most likely due to ingested biomaterial.
31755 | 1220648-2024-05982 | 18483478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483478&pc=OZD | 12/29/2023 | 02/23/2024 |  | 01/10/2024 |  |  | 12/29/2023 | 01/10/2024 |  |  | 09/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024408522 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to severe 3 vessel disease. the patient developed a hematoma at the impella access site. heparin was stopped. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the access site bleeding was most likely due to anticoagulation management since the hematoma got stable after the heparin was stopped.
31756 | 1220648-2024-05983 | 18483501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483501&pc=OZD | 12/31/2023 | 12/23/2024 |  | 01/10/2024 |  |  | 12/31/2023 | 01/10/2024 |  |  | 09/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024403146 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/31/2023 |  |  |  |  | The user facility reported a 79-year-old white male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had a cardiac history and was put on the cp support despite positioning problems. all catheters had problems with native anatomy and delivering under the papillary. after 8 hours of support the cp was explanted and escalated to the 5.5 pump after being transferred to tertiary center. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. clinical details state that the customer complained of having lower flows. the data logs showed that there were lower flows when the cp was on auto with positioning alarms. based on the data log analysis and clinical details provided, the root cause of the low pump flow is most likely due to improper poisoning of the device.
31757 | 1220648-2024-05984 | 18483617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483617&pc=OZD | 12/30/2023 | 01/29/2025 |  | 01/10/2024 |  |  | 12/30/2023 | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024421567 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hematoma | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a 62-year-old male of unknown ethnicity with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient experienced a series of adverse of events that led to the premature explant of the device. three days after implant, repositioning of the implanted impella 5.5 pump was required and resulted in the patient experiencing ectopy and multiple runs of ventricular tachycardia. amiodarone was administered and the epi drip was decreased to manage the condition. as support continued, a pocket hematoma was discovered at the axillary site. the following day, a position in aorta alarm was triggered and a transthoracic echocardiogram (tte) revealed that the inflow of the pump was above the aortic valve. attempts to re-cross the valve was unsuccessful and the patient's labs showed signs of hemolyzing. due to the noted conditions, the decision was made to explant the pump. the pocket hematoma was evacuated/washed out, and the site was successfully closed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: hemolysis: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis arrhythmia: nothing beyond study data. | The investigation for the hemolysis, ventricular tachycardia, access site bleed, and positioning issues has been completed. hemolysis was observed on the 7th day of pump implant after the pump was observed to be out of position. the pump was explanted. therefore, the root cause of the hemolysis issue was most likely improper positioning of the device. the root cause of the positioning issue, ventricular tachycardia,and access site bleed was not determined since product was not returned and insufficient clinical information was provided. corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email.
31758 | 1220648-2024-05985 | 18483751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483751&pc=OZD | 01/02/2024 | 03/22/2024 |  | 01/10/2024 |  |  | 01/02/2024 | 01/10/2024 |  |  | 09/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024399841 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to severe multivessel disease. a hematoma formed at the access site and the patient was in a lot of pain. the physician removed the pump due to bleeding. manual pressure was held. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely patient movement, there was patient moment observed during the support as per the clinical description provided.
31759 | 1220648-2024-05986 | 18483945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18483945&pc=PYX | 01/03/2024 | 03/22/2024 |  | 01/10/2024 |  |  | 01/03/2024 | 01/10/2024 |  |  | 01/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024333886 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella rp support and the patient experienced access site bleeding around the repo sheath. mattress sutures were placed and the second suture was unsuccessful. a third suture was placed. pressure was held and blood products were given. vascular surgery was consulted. the patient expired after 1.33 hours. no allegations were made towards the impella for cause of death. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-updated model number | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31760 | 1220648-2024-05987 | 18484181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18484181&pc=OZD | 12/20/2023 | 04/16/2024 |  | 01/10/2024 |  |  | 12/20/2023 | 01/10/2024 |  |  | 09/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024402763 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Epistaxis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported an 87-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that there was oozing from impella access site and epistaxis. forward tension and angle match dressing was applied, and no further bleeding was seen at the right femoral impella site. epistaxis remains and heparin remains on hold. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury | The investigation into the access site bleeding ¿ major has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by additional sutures.
31761 | 1220648-2024-05988 | 18484258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18484258&pc=OZD | 12/19/2023 | 04/18/2024 |  | 01/10/2024 |  |  | 12/19/2023 | 01/10/2024 |  |  | 09/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407478 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient of unknown age and gender on impella cp support, had bleeding in the groin on the left with a large hematoma after puncture. bed rest was prescribed until duplex, but the patient refused to comply. the patient was noted to be extremely agitated and wished to be discharged against medical advice. the patient was informed of the risks and was ultimately discharged. however, the patient¿s stay in the hospital was prolonged due to the hematoma/ hemorrhage. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
31762 | 1220648-2024-05989 | 18484324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18484324&pc=OZD | 12/21/2023 | 04/16/2024 |  | 01/10/2024 |  |  | 12/21/2023 | 01/10/2024 |  |  | 04/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352947 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had oozing and a hematoma at the impella access site on admission to the icu. manual pressure was held, and a fem stop was placed. the physician decided to remove the impella due to continued oozing from the impella groin. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that post procedure, the patient's activated coagulation time (act) was prolonged at 315. therefore, the root cause of the access site bleed was most likely due to improper anticoagulation management.
31763 | 1220648-2024-05990 | 18484333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18484333&pc=OZD | 10/05/2023 | 10/08/2024 |  | 01/10/2024 |  |  | 12/20/2023 | 01/10/2024 |  |  | 06/30/2024 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2023237205 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported a 71-year-old white male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed acute renal failure during impella support. it was believed by the staff that the condition was potentially related to the impella device. no additional information was available pertaining to the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the renal failure could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 correction to the initial mfr report: added- d6a/d6b. f6/f8 should have been left blank.
31764 | 1220648-2024-05991 | 18484371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18484371&pc=OZD | 12/21/2023 | 11/30/2024 |  | 01/10/2024 |  |  | 12/21/2023 | 01/10/2024 |  |  | 09/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403917 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fluid/Blood Leak | Unintended Collision |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a (b)(6) white female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the automated impella controller (aic) crashed into the elevator and the connector of the impella 5.5 broke off. the patient was implanted with a new impella 5.5. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the purge leak-pump has been completed. during transport the charge nurse crashed the aic into the elevator and broke the connector on the 5.5. the root cause of the purge leak pump was most likely use related. corrections to the initial mfr report: added d6a/d6b; f6/f8 should have been left blank.
31765 | 1220648-2024-05992 | 18484534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18484534&pc=OZD | 12/21/2023 | 03/26/2024 |  | 01/10/2024 |  |  | 12/21/2023 | 01/10/2024 |  |  | 07/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024374312 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 51-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on day four of the support, the impella 5.5 pump stopped. the team attempted to restart the pump by replacing the automated impella controller (aic). the pump would not re-start after troubleshooting so the decision was made to explant the pump. the patient survived on extracorporeal membrane oxygenation (ecmo) support. it was also reported the patient developed a small hematoma at the access site that did not require treatment. | [5 empty] | The investigation into the pump stop and the access site bleeding ¿ minor issue has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the access site bleeding issue was most likely patient condition related; its connection to the impella device was unknown. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31766 | 1220648-2024-05993 | 18484668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18484668&pc=OZD | 12/22/2023 | 03/01/2024 |  | 01/10/2024 |  |  | 12/22/2023 | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024414976 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the implanted impella registered a high purge pressure and triggered a purge system blocked alarm. tissue plasminogen activator was initially administered, but the decision was ultimately made to replace the pump. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the high purge pressure issue has been completed since the original report was filed. the device was returned for evaluation. upon investigation of the pump, a clot was observed in the purge gap. high purge pressure was reproduced when purging the pump with the returned purge cassette in the lab. the root cause of the high purge pressure was determined to be biomaterial in the purge gap.
31767 | 1220648-2024-06000 | 18485975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18485975&pc=OZD | 12/22/2023 | 02/29/2024 |  | 01/10/2024 |  |  | 12/27/2023 | 01/10/2024 |  |  | 01/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024312755 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The complainant from the stemi-dtu study group reported patient of unknown age, race and gender experienced ventricular tachycardia on 20-dec-2023, that was determined to be severe and required intervention to prevent damage. the event was determined to be unlikely related to the device and/or the impella procedure. following impella insertion and after percutaneous coronary intervention (pci) procedure on the same day, the patient had recurring deep vein thrombosis' that required administration of amiodarone and magnesium. on 21-dec-2023, the patient was noted to have had a drop in hemoglobin due to unknown cause. this was determined to be not related to the device, but possibly related to impella procedure. on 24-dec-2023, the patient developed an insertion site infection and a groin hematoma that were moderate in severity. there was evidence of staphylococcus lugdunensis in blood cultures. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation into the access site bleeding ¿ major, the infection/sepsis, the vf/vt/af/at, and the thrombus issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding, the thrombus, the infection, and the vt/vf was not determined.
31768 | 1220648-2024-06001 | 18486009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18486009&pc=OZD | 01/03/2024 | 03/25/2024 |  | 01/10/2024 |  |  | 01/03/2024 | 01/10/2024 |  |  | 09/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024407652 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed an impella cp via the 14fr sheath for support of a patient admitted in cardiogenic shock/acute myocardial infarction (cgs/ami) after being found "down" in the community and got cardiopulmonary resuscitation (cpr) for ventricular fibrillation (vfib) and was cardioverted. the impella cp was placed in conjunction with extracorporeal membrane oxygenation (ecmo). the site of the impella cp was found to be oozing and having blood loss. the team treated with red blood cells and manual pressure hold to the groin. the support remained for over 58 hours till explanted when team made choice to withdraw and allow for organ harvest after anoxic brain injury noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
31769 | 1220648-2024-06002 | 18486135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18486135&pc=OZD | 12/23/2023 | 01/31/2024 |  | 01/10/2024 |  |  | 12/23/2023 | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409533 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 26-year-old black male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the implanted impella 5.5 pump experienced a high purge pressure coinciding with purge pressure blocked alarms during 26-year-old male patient support. a tpa protocol was subsequently followed and both the purge pressure and purge flow issues resolved. there was no patient harm. the patient had originally been admitted with decompensated heart failure on intra-aortic balloon pump (iabp), and was transitioned to peripheral venoarterial (va) extracorporeal membrane oxygenation (ecmo). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed. the product was not returned for investigation. the data logs showed positioning alarms and suction at the time of the high purge pressure. there is also purge blocked alarms. the high purge pressure was resolved after running tpa. the root cause of the high purge pressure was most likely biomaterial as the high purge pressure resolved after tpa was added to the purge.
31770 | 1220648-2024-06003 | 18486140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18486140&pc=OZD | 12/30/2023 | 02/12/2024 |  | 01/10/2024 |  |  | 12/30/2023 | 01/10/2024 |  |  | 11/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424843 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following impella implant, pulsatile bleeding was observed from around the sheath. a side wire port angiogram showed an extravasation coming from the repositioning sheath. after attempts to achieve hemostasis were unsuccessful, the decision was made to explant the pump and repair the vessel. during surgical intervention it was confirmed that there was a pseudo aneurysm and rupture. a covered stent was placed to achieve hemostasis. | [5 empty] | The investigation into the vascular damage - peripheral vessel has been completed since the original report was filed. the device was returned for investigation. visually inspected and no damages or abnormality were observed from the pump or repo sheath. the cause of the vascular damage issue was patient condition related since patient had pseudoaneurysm per clinical review and field investigation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
31771 | 1220648-2024-06004 | 18486162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18486162&pc=OZD | 01/06/2024 | 02/07/2024 |  | 01/10/2024 |  |  | 01/06/2024 | 01/10/2024 |  |  | 03/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024324088 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that there was a pump stop following a high motor current. placement signal was not reliable and controller failure alarms occurred. several attempts were made to restart the pump as well as a new console with no success. the impella was removed and upon assessment of graft, multiple large clots were found. the reason for the pump stop is most likely due to clot ingestion. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the product was returned for investigation. the pump failed electrical testing and light testing for the optical test. it did not start during the test in a bucket of water. all 5s parameters showed a fiber break trend during optical testing. upon further investigation, a stretched catheter was found near the red handle. during the case there are placement signal not reliable alarms where the placement signal and 5s parameters resemble a kinking of the fiber line. the data logs showed the sudden pump stopped with alarm and a hard failure trend on the 5s parameter, at the end of the case. there were no purge related abnormalities observed. the cause of the pump stop issue was an open electrical line near the red handle, based on returned product investigation and data log review.
31772 | 1220648-2024-06005 | 18486377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18486377&pc=OZD | 01/04/2024 | 04/26/2024 |  | 01/10/2024 |  |  |  | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024421260 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | High Readings | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 76-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp experienced persistent high purge pressure during patient support. the issue remained until the patient eventually succumbed to their existing condition and passed away. there is currently no allegation of device association between the pump and patient's passing. | [5 empty] | The investigation for the reported high purge pressure has been completed. product was not returned for investigation. the data log shows a gradual rise in purge pressure over three hours. as per the clinical details, high purge pressure alarms persisted despite attempts to reduce and return pressure levels, with no observed kink in the purge line or change in the purge cassette. the cause of the high purge pressure issue was most likely biomaterial, based on data log review. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31773 | 1220648-2024-06007 | 18486735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18486735&pc=OZD | 01/04/2024 | 03/14/2024 |  | 01/10/2024 |  |  | 01/04/2024 | 01/10/2024 |  |  | 06/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024354381 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the introducer perforated the femoral artery during impella cp implant. an unspecified number of blood products were transfused as a result of the condition and a 14 fr cook sheath was inserted to occlude the perforation. an external fem-fem bypass was also placed and the pump was removed. | [5 empty] | The investigation into vascular damage - peripheral vessel has been completed since the original report was filed. the device was not returned for investigation. the root cause of the vascular damage issue was not determined since the clinical information provided was inconclusive of the root cause. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
31774 | 1220648-2024-06008 | 18487423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18487423&pc=OZD | 01/08/2024 | 12/19/2024 |  | 01/10/2024 |  |  | 01/08/2024 | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409555 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Difficult to Advance | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 69-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was having the impella 5.5 implant due to an escalation from the impella cp. upon attempted implant, the doctor was unable to advance impella 5.5 due to tortuosity and calcification in left subclavian artery. after second attempt to implant the 5.5, the doctor made the decision to abort the implant and leave impella cp in place. upon removal of impella 5.5, a small intima was found attached to tip of catheter. angiography showed a small intimal dissection of left subclavian artery. vascular surgeon successfully stented with good radial pulse past placement of stent. impella cp was left in place. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the delivery issue and vascular damage has been completed. no product returned. clinical details were sufficient to determine root cause. the cause of the delivery issues was most likely patient condition related. the cause of the vascular damage peripheral vessel was most likely use related due to multiple delivery attempts. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. added h6 med dev prob code 2982 corrections to the initial mfr report: d4-updated model number. added d6a/d6b was missing. g1 mdr reporting contact email. h6 component code 925.
31775 | 1220648-2024-06009 | 18488805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18488805&pc=OZD | 01/01/2024 | 03/22/2024 |  | 01/10/2024 |  |  | 01/01/2024 | 01/10/2024 |  |  | 09/30/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408758 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2024 |  |  |  |  | The user facility reported a 83-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp insertion a hematoma developed on impella access site. manual pressure, femstop, and protamine were administered, and hematoma improved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that post procedure, the patient's activated coagulation time (act) was prolonged above 200s entire procedure. therefore, the root cause of the access site bleed was most likely due to improper anticoagulation management. added-h6 impact code 4644; d4-updated model number.
31776 | 1220648-2024-06010 | 18488812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18488812&pc=OZD | 01/01/2024 | 07/25/2024 |  | 01/10/2024 |  |  | 01/01/2024 | 01/10/2024 |  |  | 10/31/2025 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420719 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Activation, Positioning or Separation Problem | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2024 |  |  |  |  | The user facility reported a 72-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump that failed to deliver to the lv (left ventricle¿). the pump failed in the beveled graft and the pump implant was aborted. no patient harm or injury was noted. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 revised as information was inadvertently omitted from the initial submission of manufacturer device report number 1220648-2024-06010. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-06010 in accordance with updated procedures. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2024-06010 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-06010 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-06010 was submitted. h.6 codes 2906 and 2920 were previously entered incorrectly on manufacturer device report number 1220648-2024-06010 and a new code has been add to medical device problem codes. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-06010.
31777 | 1220648-2024-06011 | 18488854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18488854&pc=OZD | 04/20/2022 | 07/09/2025 |  | 01/10/2024 |  |  | 04/20/2022 | 01/10/2024 |  |  | 12/31/2023 |  | 01/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023163673 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | Compatibility Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 47-year-old male (with pre-existing ventricular tachycardia) endure vt during hospital transport on the impella device. the physician could not determine if the impella was related to this event. additionally, it was reported that the patient's implanted cardioverter-defibrillator (icd) could not be interrogated during use of the impella device due to "risk of magnetic field interference." due to the inability to conduct this icd-related exercise, the team prepared for use of direct cardioversion on (b)(6) 2022. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue and ventricular tachycardia (vt) the has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue and vt could not be determined. this complaint has been identified as part of a retrospective review. corrections to the initial mfr report: added d6a/d6b: f6/f8 should have been deleted. g1 mdr reporting contact email. h6 med dev prob code 1158, component code 925 added, concl code 4315. h10-investigation summary.
31778 | 1220648-2024-05939 | 18474263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18474263&pc=OZD | 12/14/2023 | 09/26/2024 |  | 01/09/2024 |  |  | 12/14/2023 | 01/09/2024 |  |  | 08/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024396448 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 60-year-old white male patient developed bleeding from the right femoral artery following impella implant. manual pressure was applied with angle matching positioning to achieve hemostasis. however, due to the condition, plans were made to explant the pump. due to concerns for the patient's current state of unloading, the pump was left in place until the patient could tolerate the weaning process. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding is determined to be use issue because angle matching and manual pressure with femstop were used and had great success in achieving hemostasis. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 corrections to the initial mfr report: added- d6a/d6b.
31779 | 1220648-2024-05940 | 18474314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18474314&pc=OZD | 12/18/2023 | 02/21/2024 |  | 01/09/2024 |  |  | 12/18/2023 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | UNKNOWN |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Company representation | Distributor |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68 -year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a problem with the fiber optic pressure sensor during the catheter priming. the placement signal was unable to be viewed. the automated impella controller (aic) was exchanged but the issue persisted. therefore, the pump was explanted, and the issue resolved with the new pump. the percutaneous coronary intervention was able to be completed without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the original report was filed. visually inspected and no physical damages observed. the pump was tested, and the issue was confirmed. the cause of the optical signal issue was a damaged optical fiber line (out-of-box failure) inside the red plug due to patient cable manufacturing error.
31780 | 1220648-2024-05941 | 18474357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18474357&pc=OZD | 12/19/2023 | 01/17/2024 |  | 01/09/2024 |  |  | 12/19/2023 | 01/09/2024 |  |  | 08/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385402 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 81-year-old black female patient on impella cp support reported an attempted delivery. the impella was unable to pass through the iliac artery. the physician decided to not proceed with the impella implant. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue - anatomy was determined to be patient condition due to diseased/small vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31781 | 1220648-2024-05942 | 18474526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18474526&pc=OZD | 12/17/2023 | 01/18/2024 |  | 01/09/2024 |  |  | 12/18/2023 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1439042 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when priming the impella an alarm sounded indicating that the purge pressure transmitter was not properly installed. the purge set was reinstalled, but the issue continued. the automated impella controller (aic) was replaced, and the alarm was resolved. there was no patient harm. | [5 empty] | The investigation for the reported console purge issue has been completed. the console and data logs were returned for evaluation. testing was performed finding that the issue was properly detecting the purge disc was reproduced. the purge flag was confirmed to be broken. data logs were reviewed which confirmed the reported failure and revealed that the console issued the purge disc not detected alarmed several times. the cause of the issue was an inability to detect purge disc due to a broken purge flag. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
31782 | 1220648-2024-05943 | 18474762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18474762&pc=OZD | 12/20/2023 | 02/02/2024 |  | 01/09/2024 |  |  | 12/20/2023 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1223034 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported that the automated impella controller (aic) had a purge disk not detected alarm. the aic was removed from service. no patient harm was reported. | [5 empty] | The investigation into the reported aic ¿ purge disc/ flag issue has been completed since the original report was filed. logs from the complaint date revealed that the console issued the purge disc not detected alarm the console was tested after arriving in fs. aic's inability to detect the purge disc was confirmed to be caused by a broken purge flag. the cause of the issue was inability of the aic to detect purge disc due to broken purge flag. | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
31783 | 1220648-2024-05944 | 18474775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18474775&pc=OZD | 12/19/2023 | 02/22/2024 |  | 01/09/2024 |  |  | 12/19/2023 | 01/09/2024 |  |  | 09/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 49-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support when the patient experienced limb ischemia. the decision was made by the team to wean the impella and insert iabp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31784 | 1220648-2024-05945 | 18475027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18475027&pc=OZD | 12/17/2023 | 01/29/2025 |  | 01/09/2024 |  |  | 12/20/2023 | 01/09/2024 |  |  | 08/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024391402 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported probable retraction of the impella cp in aorta and kinking of the cannula during an attempt to reposition the device. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the delivery issues and the product damage has been completed. as per the clinical information provided, the pump was initially implanted a little deep and passed into the aorta ie. the initial placement was incorrect which need to be repositioned. therefore, the root cause of the positioning issue was use related to improper technique. pump was returned for investigation. the pump had a kink in the cannula near the outlet of the pump. clinical information mentions that after position of the pump was incorrect, wireless reaccess was tried to reposition the pump. therefore, the root cause of the product damage (cannula kink) issue was use related to improper technique. the instructions for use for the impella cp with smartassist system states the following: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ corrections to the initial mfr report: added d6a/d6b. g1 mdr reporting contact email.
31785 | 1220648-2024-05946 | 18475187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18475187&pc=OZD | 12/19/2023 | 02/23/2024 |  | 01/09/2024 |  |  | 12/19/2023 | 01/09/2024 |  |  | 10/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024421105 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella cp support and had limb ischemia in both limbs. blue hub was not sutured down on arrival, when taken to ccl blue hub was sutured down. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia reversible issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31786 | 1220648-2024-05947 | 18475331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18475331&pc=OZD | 12/25/2023 | 04/05/2024 |  | 01/09/2024 |  |  | 12/25/2023 | 01/09/2024 |  |  | 09/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024404521 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2023 |  |  |  |  | The user facility reported a 42-year-old white male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp device began to trigger high purge pressure alarms during patient support. although tissue plasminogen activator (tpa) protocols were followed, the issue persisted, and the pump subsequently stopped. attempts to restart the pump were unsuccessful and the decision was made to explant the pump. there was no patient harm reported. | [5 empty] | The investigation into the pump stop and the high purge pressure issues has been completed since the initial report was submitted. the device was not returned for investigation. per clinical information provided, there are purge system blocked alarms with high purge pressure alarms consistent throughout the case. tpa was added and did not resolve the high purge pressure. the root cause of the high purge pressure cannot be determined. the root cause of the pump stop was most likely biomaterial, per the data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31787 | 1220648-2024-05948 | 18475386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18475386&pc=OZD | 01/02/2024 | 02/23/2024 |  | 01/09/2024 |  |  | 01/02/2024 | 01/09/2024 |  |  | 10/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024414245 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported a 53-year-old female of unknown race with ventricular fibrillation (vfib) arrest was implanted with an impella cp device for mechanical circulatory support. the patient had known peripheral arterial disease (pad) and the pump was occluding the femoral artery. the decision was made that due to the risk of limb ischemia, after one hour of support the pump was explanted. the patient was reported to be stable at time of explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31788 | 1220648-2024-05949 | 18475619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18475619&pc=OZD | 09/20/2023 | 04/17/2024 |  | 01/09/2024 |  |  | 12/11/2023 | 01/09/2024 |  |  | 07/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024376344 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that indicated this patient to have endured ventricular arrhythmia during impella insertion as well as during support, with a "possible" relationship to the impella device used in this case. the reported narrative indicates, "during impella 5.5 implant, patient went into vf/vt, was cardioverted and paced. placed on amiodarone drip. cardioversion was successful," and occurring on (b)(6) 2023, "multiple episodes of vt/vf" with the patient ultimately, "requiring ablation on (b)(6) 2023. required cardioversion as well." | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31789 | 1220648-2024-05950 | 18475780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18475780&pc=OZD | 12/22/2023 | 08/18/2025 |  | 01/09/2024 |  |  | 12/22/2023 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1329519 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) purge disc failed detection during set up for a new case. there was no patient involvement. the aic was replaced without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email, d2 device name has been updated, d9 device return date has been added, f6 and f8 should have been left blank, h6 type of investigation code added.
31790 | 1220648-2024-05951 | 18476133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18476133&pc=OZD | 12/27/2023 | 09/24/2024 |  | 01/09/2024 |  |  | 12/27/2023 | 01/09/2024 |  |  | 09/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024406221 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | Pressure Problem |  |  |  |  | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that a high purge pressure alarm with signs of pump saturation was observed. the impella stopped before tubing could be changed, or tissue plasminogen activator (tpa) started. the impella was explanted at bedside. no patient harm was reported. | Additional information noted that the procedure outcome was that patient expired due to unknown reason. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-05951 was provided in sections b2, b5, h1, and h6. | The investigation into the high purge pressure and the pump stop issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. a significant amount of biomaterial was observed on the impeller and purge gap. no physical abnormalities were noted. the cause of the pump stop issue was biomaterial, as both the returned product and the data logs clearly demonstrate that biomaterial led to a high purge pressure event, eventually resulting in the saturated pump flow and pump stop.
31791 | 1220648-2024-05952 | 18476324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18476324&pc=OZD | 12/30/2023 | 02/06/2024 |  | 01/09/2024 |  |  | 12/30/2023 | 01/09/2024 |  |  | 07/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024372426 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported a patient was attempting to have an impella 5.5 implanted due to non-st elevated myocardial infarction. upon attempted delivery, multiple attempts were made to advance the impella, but the doctor was unable to advance the impella into the aorta. the patient had difficult anatomy, and there was a kink in between the catheter and the motor house. the impella 5.5 was aborted and removed. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was patient condition the impella was unable to pass through due to patient anatomy. there was a kink between to catheter and the motor housing as well as per the clinical description provided. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31792 | 1220648-2024-05953 | 18476430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18476430&pc=OZD | 01/08/2024 | 12/03/2024 |  | 01/09/2024 |  |  | 01/08/2024 | 01/09/2024 |  |  | 01/31/2024 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183521 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/08/2024 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a leak was noted on the yellow luer on the impella after a purge bag change. purge system open alarm became active. the team did a purge system bypass to restore purge pressure until the impella could be replaced. | [5 empty] | The investigation into purge leak ¿ pump has been completed. the impella device was not returned for evaluation. the root cause of the purge leak pump was most likely a damaged sidearm but the exact location and cause could not be confirmed as product was not returned. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05953. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 code 4641 was reported incorrectly. new code was added to health effect - impact code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
31793 | 1220648-2024-05954 | 18476601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18476601&pc=OZD | 12/31/2023 | 01/18/2024 |  | 01/09/2024 |  |  | 12/31/2023 | 01/09/2024 |  |  | 09/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2024407926 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/31/2023 |  |  |  |  | The user facility in japan reported an 80-year-old male with cardiomyopathy was to be implanted with an impella cp for mechanical circulatory support. it was reported the physician failed to insert the pump due to patient anatomy issues and the pump delivery was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related.
31794 | 1220648-2024-05955 | 18476947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18476947&pc=OZD | 01/02/2024 | 02/22/2024 |  | 01/09/2024 |  |  | 01/03/2024 | 01/09/2024 |  |  | 10/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024419504 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to pulmonary edema along with elevated brain natriuretic peptide and abnormal cardiac enzymes. while on support, staff were unable to find signals in the right lower extremity. the right lower extremity became cold and mottled. the impella was weaned, and vascular surgery was consulted. after explant, the pulse returned to the right lower extremity. | [5 empty] | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31795 | 1220648-2024-05957 | 18477101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18477101&pc=OZD | 12/27/2023 | 02/23/2024 |  | 01/09/2024 |  |  | 12/27/2023 | 01/09/2024 |  |  | 07/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388112 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported an 83-year-old white female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon the patient¿s arrival at icu, the patient¿s left leg was noted to be mottled and pulses unable to doppler pulses on left foot. the patient was brought back to ccl and impella cp removed from left femoral artery (lfa). angiogram of lfa and rfa performed, and no thrombus, dissection, or perforation was noted. the patient¿s left leg improved in color and was no longer mottled after impella was removed from lfa and pulses were dopplerable on left foot. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31796 | 1220648-2024-05956 | 18477121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18477121&pc=OZD | 01/04/2024 | 01/31/2024 |  | 01/09/2024 |  |  | 01/04/2024 | 01/09/2024 |  |  | 09/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024401488 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2024 |  |  |  |  | The user facility reported a 48-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 that was going to be delivered to the patient in need of support of cgs (cardiogenic shock) /cardiomyopathy and care escalation from the iabp (intra-aortic balloon pump). the 5.5 failed delivery but, the cp was able to be delivered. the anatomy was small. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the delivery issue was most likely due to patient condition since patient had small artery.
31797 | 1220648-2024-05958 | 18477409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18477409&pc=OZD | 01/03/2024 | 12/09/2024 |  | 01/09/2024 |  |  | 01/05/2024 | 01/09/2024 |  |  | 10/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421550 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that patient on impella cp was experiencing hemolysis. p-level adjusted from auto to p-7 due to red tinged urine and high afterload. the purge fluid bag was changed, and nitro started due to the afterload. the patient was transferred to another hospital within three hours of implant and urine continued to be red tinged. the impella was repositioned again. lactate dehydrogenase went from 3039 to 2193. patient underwent a planned upgrade to impella 5.5. labs continued to improve. the patient is stable. | [5 empty] | The investigation for the hemolysis has been completed. product was not returned. data logs show 1 brief suction event that recovers. there is 1 motor current (mc) spike that otherwise resolves and the rest of the time the mc and flow is stable. no other abnormalities. the cause of the hemolysis was not determined due to no product returned, lack of clinical details, and inconclusive data logs. added h6 impact code 4628, h6 component code 925 corrections to the initial mfr report: d4 model number added d6a/d6b g1 mdr reporting contact email f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31798 | 1220648-2024-05959 | 18477487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18477487&pc=OZD | 12/15/2023 | 06/06/2024 |  | 01/09/2024 |  |  | 12/15/2023 | 01/09/2024 |  |  | 09/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024406486 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male patient noted as high-risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed a minor aortic regurgitation following impella removal. the condition was monitored, but no intervention was documented. the patient was considered stable. | [5 empty] | The investigation into the heart valve damage has been completed since the original report was submitted. the device was not returned for investigation. the cause of the heart valve damage was not determined because no product was returned, and not enough clinical information was given. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: impella cp with smartassist catheter insertion: ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
31799 | 1220648-2024-05960 | 18477552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18477552&pc=OZD | 12/15/2023 | 02/12/2024 |  | 01/09/2024 |  |  | 12/15/2023 | 01/09/2024 |  |  | 10/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422211 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male patient in st (sinus tachycardia) elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support and was back in the cardiac catheterization lab. whilst off paralytics and sedation, the patient aggressively moved and the catheter was pulled back into the aorta. the impella cp was unable to be passed back across the valve and had to be explanted. axillary 5.5 placement is emergently planned in the cardiovascular operating room. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the positioning issues has been completed. the device was not returned for investigation. the clinical team noted that patient had aggressive movement and pump appeared to be pulled back into aorta. the root cause of the positioning issue of the pump in the aorta was due to the patient¿s excessive movement.
31800 | 1220648-2024-05961 | 18477614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18477614&pc=OZD | 12/18/2023 | 02/29/2024 |  | 01/09/2024 |  |  | 12/18/2023 | 01/09/2024 |  |  | 08/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391453 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella insertion the wire was pulled back across the valve and impella was unable to advance. the cannula kinked and folded over. the physician applied manual pressure to the catheter to release the kink and removed 5.5. impella insertion was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported delivery issue -anatomy has been completed since the original report was filed. the root cause of delivery issue- anatomy is determined to be use issue due to pulling wire back. the cause of the issue was use related.
31801 | 1220648-2024-05962 | 18477792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18477792&pc=OZD | 12/19/2023 | 02/23/2024 |  | 01/09/2024 |  |  | 12/19/2023 | 01/09/2024 |  |  | 05/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365504 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that upon arrival at the unit echo confirmed that the impella pigtail was in cords. repositioning efforts were unsuccessful. the patient was taken back to the lab and impella was removed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported positioning issues been completed since the original report was filed. no pump was returned for evaluation. clinical review noted that echo confirmed impella in cords after arrival at unit. the root cause of the positioning issue was likely due to a use issue as there is limited clinical details.
31802 | 1220648-2024-05963 | 18478066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478066&pc=OZD | 12/27/2023 | 11/12/2025 |  | 01/09/2024 |  |  | 12/27/2023 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1189528 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a patient of unknown age and gender was going to be placed on impella cp device for mechanical circulatory support. it was reported that during the preparation of insertion the automated impella controller (aic) had a yellow alert of controller error appeared on screen. staff tried to turn off and on a few times with no change. no patient harm or involvement was reported. | [5 empty] | The investigation of automated impella controller (aic) - battery comm. failure (yellow alarm) has been completed. the impella device was returned for evaluation. data analysis: on the complaint date, december 27th, 2023, the console displayed the ¿controller error (defective optical sensor lamp) ¿ alarm,¿ event set #1039. the first occurrence of event set #1039 was on december 19th, 2023. since then, the console has been rebooted nine times, and each time after boot-up, the console displays the persistent controller error alarm, this confirms the reported failure mode. the 5s parameters were abnormal (light value: 0.59v & snr value: 2.4), specifically the "light" parameter, which triggered the controller error. device analysis: the console was tested and the front panel optical connector was found to have a significant amount of debris upon receipt. this debris most likely reduced light reflection and resulted in the controller error, event #1039. the front panel optical connector was cleaned to perform the optical cavity test, and the 5s parameters were found to be within specification. there was no evidence in the log to confirm the reported ¿aic - battery comm. failure (yellow alarm).¿ the controller error (yellow alarm) was misreported as a battery comm. failure (yellow alarm). the root cause for the controller error was due to a contaminated front panel optical connector. d2 common device name revised on manufacturer device report 1220648-2024-05963 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-05963 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-05963 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-05963. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31803 | 1220648-2024-05964 | 18478092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478092&pc=OZD | 12/20/2023 | 02/21/2024 |  | 01/09/2024 |  |  | 12/20/2023 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1293209 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) displayed "system self-check failure" when it was turned on. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the reported aic - system self-check error has been completed since the original report was filed. data analysis: directly after pm there was epm communication issues resulting in a sw watchdog reboot. this did not fix the failure and the console entered system self-check failure mode ((b)(6)) confirming the complaint seen in the field. device analysis: the failure mode was reproduced in fs. system software was reinstalled and the console was able to boot up as expected. root cause: the cause of the aic issue was firmware corruption.
31804 | 1220648-2024-05965 | 18478151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478151&pc=OZD | 12/21/2023 | 09/21/2024 |  | 01/09/2024 |  |  | 12/21/2023 | 01/09/2024 |  |  | 12/31/2024 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301965 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that an impella cp pump could not be advanced across the 63-year-old male patient's valve during attempted delivery. after multiple unsuccessful attempts, the implant was aborted. the patient passed away from their pre-existing condition with no allegation of association with the impella pump. |  |  |  |  | The investigation into the delivery issues has been completed. the device was not returned for investigation. according to clinical information, the root case of delivery issue anatomy is determined to be patient condition due to diseased/small vessels. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
31805 | 1220648-2024-05966 | 18478164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478164&pc=OZD | 12/22/2023 | 01/30/2024 |  | 01/09/2024 |  |  | 12/25/2023 | 01/09/2024 |  |  | 04/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2024338018 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Pumping Stopped | Insufficient Flow or Under Infusion | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2023 |  |  |  |  | The user facility reported a 62-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after implantation of the impella cp it was reported there were low purge flow alarms and there was an issue with the pump¿s positioning. the impella cp was explanted from the patient, washed and reinserted. eventually the pump stopped resulting in the impella cp removal. once removed it was noted there was a thin object coming out of the outlet cage. the pump was removed on 24-dec-2023, the following morning ((b)(6) 2023) the patient passed away from severe heart failure. it is unknown if the pump stopping was related to the patient death. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop, low pump flow and the positioning issues has been completed since the original report was filed. the device was returned for investigation. the pump was visually inspected and found a foreign object stuck and wrapped around the impeller purge gap, outflow cage. blood ingress in the purge line was observed closer to the motor housing side and a small catheter kink was also found. the purge fluid passed through the purge lumen section cut closer to the motor housing. the pump stops and high purge pressure occurred as a chained events and the issue reproduced due to a foreign object interaction around the purge gap and outflow. no other abnormalities were found. data logs were reviewed and several high motor current pump stop alarms observed at the end. several high purge pressure alarms were found to be occurring as a series of event with pump sops. the cause of the pump stop issue was due to an interaction with another device with foreign object ingested at purge gap per field investigation, clinical review, and log review. the cause of the positioning issue was use related with pump turning toward posterior wall and touching the papillary muscle and was able to be reinserted after taking out and washing with sterile water.
31806 | 1220648-2024-05967 | 18478302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478302&pc=OZD | 01/02/2024 | 11/04/2024 |  | 01/09/2024 |  |  | 01/02/2024 | 01/09/2024 |  |  | 09/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024401128 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported an 80-year-old male noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that while exchanging the guides, there was an iliac dissection, the procedure was stopped and explanted in procedure area. the patient was reported as stable and will return the next day to have procedure performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05967: d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
31807 | 1220648-2024-05968 | 18478319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478319&pc=OZD | 01/05/2024 | 02/02/2024 |  | 01/09/2024 |  |  | 01/05/2024 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1401769 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 01/05/2024 |  |  |  |  | The user facility reported a 42-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the automated impella controller (aic) had a purge disc not detected alarms and purge cassette alarms. the alarms resolved with aic transfer. there were no visible defects noted on blue disc holder. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed. the device was returned for investigation. after reviewing the imc logs, the issues with this console not being able to detect the purge disc was confirmed. there were numerous instances of purge disc not detected alarm on the reported occurrence date. console was tested after arriving in fs. the reported issue with the aic not being able to detect the purge disc was able to be reproduced. after visual inspection of the console, it was discovered that there was a fracture on the left side of the retainer. when enough pressure is applied to the purge disc the retainer appears to bend at the site of the hairline fracture. because of this bend, the purge disc is not secure, resulting in the purge flag being disengaged. the purge disc not detected alarms were determined to be caused by this fracture in the purge retainer.
31808 | 1220648-2024-05969 | 18478326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478326&pc=OZD | 12/22/2023 | 10/27/2025 |  | 01/09/2024 |  |  | 12/22/2023 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1381389 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Device Fell | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) crashed against the wall and damaged to impella connect panel. the aic to be returned for evaluation, repair and annual pm. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31809 | 1220648-2024-05970 | 18478355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478355&pc=OZD | 12/27/2023 | 04/17/2024 |  | 01/09/2024 |  |  | 12/27/2023 | 01/09/2024 |  |  | 10/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024417700 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 77-year-old white male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed via the femoral artery and there was an access site ooze. the ooze was seen also to have a drop in hemoglobin. the team infused one unit packed red blood cells. the team made decision to further treat and remove the cp and place an axillary 5.5 pump. the site was also treated by placement of the perclose sutures and femostop. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
31810 | 1220648-2024-05971 | 18478706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18478706&pc=OZD | 12/14/2023 | 07/15/2024 |  | 01/09/2024 |  |  | 12/14/2023 | 01/09/2024 |  |  | 06/30/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379314 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported bleeding from the impella access site requiring four dressing changes. a unit of platelets were given and that helped the bleeding. hemostasis was achieved. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related, the patient had low blood platelets and an unit of platelets were given to achieve hemostasis as per the clinical description. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
31811 | 1220648-2024-05972 | 18479162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18479162&pc=OZD | 12/17/2023 | 04/12/2024 |  | 01/09/2024 |  |  | 12/17/2023 | 01/09/2024 |  |  | 03/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339543 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient on impella cp support was experiencing a large amount of bleeding coming from the impella access site. hemoglobin drop and four units of packed red blood cells were given. the patient required several dressing changes and stitches to be applied to both impella site and extracorporeal membrane oxygenation (ecmo) site. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since the hemostasis achieved by applying additional sutures.
31812 | 1220648-2024-05973 | 18479524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18479524&pc=OZD | 12/24/2023 | 02/13/2025 |  | 01/09/2024 |  |  | 12/28/2023 | 01/09/2024 |  |  |  |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398229 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the p-level dropped from auto to p0 unexpectedly on the automated impella controller (aic). no inputs were administered by the staff at the time of the failure. the pump was restarted immediately at p8 and support continued with the utilized devices. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) issue has been completed. console data logs show that the there was an impella failure (speed deviation out-of-spec) alarm issued when the ac power disconnected alarm was issued briefly before clearing. it can be confirmed from the logs that the pump stop was temporary and pump speed resumed at previous speed. the console was returned for evaluation finding internal circuitry was inspected without any observable anomaly. power delivery of the pbm to the impellatronics board was inspected and showed no deviation from the expected. failure was replicated by running known good pump at p-level 8 or p-level 9 and then briefly disconnecting and reconnecting ac. temporary pump stop occurs before resuming previous speed. monitoring power delivery from the pbm during failure revealed that the pbm was defective. the root cause for the failure was due to a defective pbm. corrections to the initial mfr report: d2 updated common device name. d4 updated model number. f6/f8 should have been left blank. g1 mdr reporting contact email.
31813 | 1220648-2024-05974 | 18479626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18479626&pc=OZD | 12/15/2023 | 06/04/2024 |  | 01/09/2024 |  |  |  | 01/09/2024 |  |  | 05/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024350121 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient had bleeding from the puncture site. the bleeding was able to be successfully managed with suturing and dressing optimized. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of access site bleeding was determined to be use issue because the bleeding was successfully managed with suturing and dressing optimized. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
31814 | 1220648-2024-05975 | 18479755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18479755&pc=OZD | 12/28/2023 | 01/29/2025 |  | 01/09/2024 |  |  |  | 01/09/2024 |  |  | 10/31/2025 |  | 01/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420461 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a 57-year-old white female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the purge sidearm was dripping from the bottom of the clear cover. fluid was noted inside the clear cover of the sidearm. no alarms were noted. purge pressure and flow were within operating parameters. the next day there was no further dripping from the purge sidearm. it was further reported that the patient had oozing from the impella access site which resolved on its own. the patient was heparin-induced thrombocytopenia (hit) positive, so heparin was stopped, and blood products were given. the patient is stable. the purge leak is not reportable, cassette was removed and reinserted and there was no further leaking the following day, there is no/limited potential for adverse impact to patient. the bleeding, not reportable, there was no lasting harm or injury that necessitated any intervention. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05975 b2 life threatening was inadvertently selected, should have been required intervention to prevent permanent impairment/damage d4 model number d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing. h6 health effect - impact code 4641 was inadvertently omitted. h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use revised from the initial submission in accordance with updated procedures.
31815 | 1220648-2024-05891 | 18463578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18463578&pc=OZD | 12/27/2023 | 05/16/2024 |  | 01/08/2024 |  |  | 12/27/2023 | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024407805 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | FEM-FEM BYPASS. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 35-year-old male with cardiomyopathy, and pleural effusion was implanted with an impella cp device for mechanical circulatory support. after the insertion, angiogram was performed, and the sheath was occlusive, and no distal flows passed. the physician established a fem-fem bypass which supported the patient until they were escalated to impella 5.5 pump. it was also noted that the patient developed a hematoma at the insertion site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the access site bleeding and the pump flow issue was not determined since product was not returned and limited clinical information was provided. as all the necessary information and returned device were not provided, the root cause of the limb ischemia was not determined. d6a, d6b.
31816 | 1220648-2024-05892 | 18463597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18463597&pc=OZD | 12/22/2023 | 01/16/2024 |  | 01/08/2024 |  |  | 12/22/2023 | 01/08/2024 |  |  | 04/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024337972 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 59-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. following the impella 5.5 insertion a transthoracic echocardiogram (tte) was performed and showed a non-mobile organized apical left ventricle (lv) thrombus. the patient was started on systemic heparin for treatment. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the thrombus has been completed since the original report was filed. the pump was returned for evaluation. as per clinical, patient observed to have apical lv thrombus. visually inspected and no physical abnormalities were found. borescope view obtained but no biomaterial ingestion was observed. no other abnormalities were found. data logs were review and no motor current spike found relative to biomaterial interaction. the cause of the embolic stroke issue was most likely patient condition related with lv thrombus observed per imaging.
31817 | 1220648-2024-05893 | 18463650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18463650&pc=OZD | 12/23/2023 | 02/08/2024 |  | 01/08/2024 |  |  | 12/23/2023 | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2023 |  |  |  |  | The user facility reported an 80-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed an access site bleed following impella cp implant as a result of forward sutures not being placed at the site. the doctor was initially advised that forward suture were necessary, but ignored the advice. when the oozing got worse, the site was redressed and the doctor was once again advised to place the proper sutures. the doctor decided to place tegaderm around the area and angle matched with gauze, but the impella was still not locked down and the site continued to bleed. the physician finally placed sutures; however, backward tension sutures were placed rather than forward tension sutures and the site again began to bleed. femostop was placed and two units of blood were eventually administered before the doctor refused further action. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. as per the clinical information provided, the patient bleed at the access site due to forward tension sutures not being applied. after several attempts of advising the doctor about the best practices on insertion, backward tension sutures were placed and the site began to bleed again. the cause of the access site bleeding issue was use related due to forward tension sutures not being applied per impella insertion procedure/best practices.
31818 | 1220648-2024-05894 | 18463787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18463787&pc=OZD | 12/26/2023 | 02/16/2024 |  | 01/08/2024 |  |  | 12/26/2023 | 01/08/2024 |  |  | 09/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399483 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported a (b)(6) female noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during preparation for implant, the impella was continually priming and a hole was found in the filter portion of the purge sidearm. a fluid leak was visibly confirmed at the noted location. the pump was replaced as a result of the noted purge leak. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the purge leak ¿ pump has been completed since the original report was filed. the device was returned for investigation. during purge testing, leak observed. upon further investigation, surface crack was found near the filter cap hole and the leak was observed from inside filter hole design. the cause of the purge leak was damage to the sidearm filter.
31819 | 1220648-2024-05895 | 18463819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18463819&pc=OZD | 01/02/2024 | 02/02/2024 |  | 01/08/2024 |  |  | 01/02/2024 | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024423419 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | High Readings | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | The user facility reported an 87-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the impella cp insertion it showed to have a high motor current (mc) and saturated flows. the left ventricle (lv) was significantly above the placement signal. the impella was dropped to p6 and the mc became within range and the lv signal and flows normalized. the percutaneous coronary intervention was completed and the impella was successfully weaned and removed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the saturated flow issue has been completed. the product was returned for review. the pump was visually inspected and biomaterial ingestion over the purge gap and around the impeller was found to be present. the data logs were reviewed and motor current spike with saturated flows were observed at the start. the cause of the saturated flow issue was due to biomaterial ingested over the purge gap and wrapped around the impeller per field investigation and log review.
31820 | 1220648-2024-05896 | 18463936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18463936&pc=OZD | 12/22/2023 | 02/21/2024 |  | 01/08/2024 |  |  | 12/22/2023 | 01/08/2024 |  |  | 05/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362146 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 63-year-old white male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient developed continuous bleeding at their impella access site. as a result of the condition, the decision was made to remove the impella cp pump. the patient remained support on ecmo. patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
31821 | 1220648-2024-05897 | 18464005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464005&pc=OZD | 12/22/2023 | 10/04/2024 |  | 01/08/2024 |  |  | 12/22/2023 | 01/08/2024 |  |  | 08/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024391399 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The japan user facility reported a 43-year-old asian male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that after removing the impella 5.5, thrombus was noted which required the wound to be washed out with normal saline. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05897. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
31822 | 1220648-2024-05898 | 18464098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464098&pc=OZD | 12/26/2023 | 11/01/2024 |  | 01/08/2024 |  |  | 12/26/2023 | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418683 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that while attempting to advance the sterile sleeve for attachment to the locking hub, it tore and detached from the hub. support was maintained with the implanted pump. there was no patient harm. | [5 empty] | The investigation of product damage (sterile sleeve damage) has been completed. no product returned. the cause of the product damage sterile sleeve damage was not determined because product/images were not returned and there is lack a clinical details. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05898. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
31823 | 1220648-2024-05899 | 18464126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464126&pc=OZD | 12/27/2023 | 12/11/2024 |  | 01/08/2024 |  |  | 12/27/2023 | 01/08/2024 |  |  | 01/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311068 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, a blood clot was enthralled by the impella causing suction and decreased flow. a new pump was opened due to clot entrapment and low pump flows. impella pump was removed without issue, a new pump was implanted without issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. no product was returned. data logs show motor current (mc) spikes, speed deviations, and drops in flow, and suction alarms. clinical mentions thrombus was observed. the cause of the low flow was most likely biomaterial. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank.
31824 | 1220648-2024-05900 | 18464202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464202&pc=OZD | 12/21/2023 | 11/30/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407812 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support while awaiting a kidney transplant. while on impella cp support, the device had a leak/product damage that allowed for blood to ooze. there was no treatment noted and no lasting harm. the pump remains on for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the product damage has been completed. bleeding reported from the blue suture hub and re-access port. product not returned for investigation. the cause of the hole in catheter was not determined since product was not returned. corrections to the initial mfr report: f6/f8 should have been left blank.
31825 | 1220648-2024-05901 | 18464308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464308&pc=OZD | 12/21/2023 | 10/24/2025 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  |  |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1397106 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had an issue with the purge cassette not pumping. there was no patient involvement. the aic was to be sent for evaluation and/or repair. | The user facility's biomedical engineer reported the automated impella controller (aic) had an issue with the purge cassette not pumping. there was no patient involvement. the aic was to be sent for evaluation and/or repair. |  |  |  |  | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. console logs from the reported day of event are consistent with complaint. during console testing it¿s been discovered that rlm (rl03138) was experiencing connectivity issues, due to microsd card corruption (issue unrelated to the complaint). the cause of the issue was a defective component. d2 common device name revised on manufacturer device report 1220648-2024-05901 in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-05901. e2 was erroneously selected on the initial manufacturer device report number 1220648-2024-05901. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-05901 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-05901 in accordance with updated procedures. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31826 | 1220648-2024-05902 | 18464346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464346&pc=OZD | 12/21/2023 | 01/08/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  |  |  | 01/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1566737 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 50-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) had a purge line click-on not detected alarm. even after replacing the purge set, it was not recognized. the problem was resolved after replacing the aic. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31827 | 1220648-2024-05903 | 18464349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464349&pc=OZD | 12/25/2023 | 01/31/2024 |  | 01/08/2024 |  |  | 12/25/2023 | 01/08/2024 |  |  |  |  | 01/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1340487 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/25/2023 |  |  |  |  | The user facility reported an asian patient of unknown age and gender with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was on impella 5.5 support and the automated impella controller (aic) had a purge issue. the purge cassette was replaced however this did not resolve the issue. the aic was replaced and this resolved the issue. there was no patient harm. | [5 empty] | The investigation into the aic - purge issue ¿ other has been completed since the original report was filed. the console was returned for investigation. imc logs confirm that there were numerous issuances of the purge disc not detected alarm at the time of the complaint. inspection of the purge pressure sensor unit revealed a missing purge pressure flag. the root cause of the purge pressure disc not detected alarms was a missing purge pressure flag. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31828 | 1220648-2024-05904 | 18464476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464476&pc=OZD | 12/28/2023 | 02/23/2024 |  | 01/08/2024 |  |  | 12/28/2023 | 01/08/2024 |  |  | 06/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024379386 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2023 |  |  |  |  | The user facility reported a 68-year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced limb ischemia, as treatment a fem-fem bypass was placed. flow was restored, and the family made choice to not escalate care and so after 21 hours of support the care was withdrawn and patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ irreversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31829 | 1220648-2024-05905 | 18464522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464522&pc=OZD | 12/23/2023 | 12/09/2024 |  | 01/08/2024 |  |  | 12/23/2023 | 01/08/2024 |  |  | 09/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024406328 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2023 |  |  |  |  | The user facility reported an impella cp in a 77yo female patient for mechanical circulatory support. it was reported that she was having positional issues with the cp by the morning time. labs were continually hemolyzing and was brought to the or for a 5.5 upgrade. once the devices were switched out, labs were sent. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. the instructions for use states the following: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ the failure mode will be monitored and trended. | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05905.
31830 | 1220648-2024-05906 | 18464587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464587&pc=OZD | 12/26/2023 | 12/23/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 09/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407125 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Human-Device Interface Problem | Positioning Problem | Unintended Movement |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported a 59-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient accidentally pulled the impella cp into the aorta. the physician made several attempts to cross valve before making the decision to explant the device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on the provided clinical details, the root cause of the positioning issue is due to use as the patient pulled the device into the aorta. added h6 impact code 4627 and 4628 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank. g1 mdr reporting contact email.
31831 | 1220648-2024-05907 | 18464780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464780&pc=OZD | 12/27/2023 | 11/04/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419474 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. after exchanging from peel away to repo, impella access site oozing began requiring immediate manual compression. forward tension sutures placed, and angle matched, and site was still oozing. purse-string suture placed which slowed the ooze down significantly. the physician decided to place femstop as well. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. no issues were found on the returned product. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue is determined to be use issue because upon placing purse-string suture, slowed the ooze down significantly. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added d9 device return date, h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
31832 | 1220648-2024-05908 | 18464787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464787&pc=OZD | 12/26/2023 | 09/25/2025 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423441 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/26/2023 |  |  |  |  | The user facility reported a 61-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed a thrombus during impella support. placement signal alarms were also triggered, but due to the thrombus, repositioning of the device was declined. the audio was disabled as to allow for support to continue. the decision was made to explant the pump the following day in response to the thrombus. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." the following have been reported incorrectly or missing from manufacturer device report. d4 model number d6a and d6b revised from the initial submission in accordance with updated procedures. e2 health professional missing e3 occupation missing f6/f8-should have been left blank. g1 reporting contact email g1 reporting contact fax number missing. h5 labeled for single use h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use missing. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. data logs were not provided for this case run. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. d10 concomitant medical products and therapy dates.
31833 | 1220648-2024-05909 | 18464808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464808&pc=OZD | 12/31/2023 | 01/08/2024 |  | 01/08/2024 |  |  | 12/31/2023 | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406509 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/31/2023 |  |  |  |  | The user facility reported a 41-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as abridge to transplant. while on support, the patient experienced hemolysis. continuous renal replacement therapy was started and heparin was restarted. transthoracic echocardiogram was performed two days later, but the pump was never repositioned. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
31834 | 1220648-2024-05910 | 18464948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18464948&pc=OZD | 12/26/2023 | 12/19/2024 |  | 01/08/2024 |  |  | 12/26/2023 | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414727 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported a 63-year-old black male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and experienced hemolysis. echo - echocardiogram (ultrasound) was performed at bedside revealing pump over 4.5 cm in deep. upon repositioning, it was noted that the touhy was loose and likely cause to allow pump to dive in. pulled pulled back 2cm with waveforms being more appropriate. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported for positioning has been completed. the device was not returned for evaluation. the data logs were reviewed and there are positioning alarms and placement signal low alarms at the time that hemolysis was reported. based on the clinical details and data log review, the root cause of the hemolysis is most likely due to improper positioning of the device. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ corrections to the initial mfr report: f6/f8 should have been left. h6 clinical code 1886 added. h6 impact code 4629 changed to 4628.
31835 | 1220648-2024-05911 | 18465113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18465113&pc=PYX | 12/29/2023 | 02/06/2024 |  | 01/08/2024 |  |  | 12/29/2023 | 01/08/2024 |  |  | 10/31/2024 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023274754 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a (b) (6) year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella rpsa was unable to advance through the 23fr peel-away sheath during attempted implant. the sheath appeared kinked under fluoro and a second introducer was placed. however, after the same issue was encountered with the second introducer, the decision was made to abort the implant and an impella rp flex was placed via right internal jugular vein. there was no patient harm as a result of the delivery issues. | [5 empty] | The investigation for the reported delivery issues and introducer damage has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a most likely root cause. the most likely cause of the introducer damage was due to the sheath kink, as the two 23fr introducer sheaths were kinked as a result of patient condition preventing impella device from passing through. update h6 codes to add: impact code: 4629. medical device problem code: 2889 d4: updated model number. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31836 | 1220648-2024-05912 | 18465220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18465220&pc=OZD | 12/29/2023 | 01/18/2024 |  | 01/08/2024 |  |  | 12/29/2023 | 01/08/2024 |  |  | 11/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024423528 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a white male patient of an unknown age was to receive impella 5.5 for mechanical circulatory support. it was reported the physician was unable to advance through the patient's subclavian during attempted delivery. it was determined that the vasculature was too small, and the decision was made to abort the implant. there was not patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related.
31837 | 1220648-2024-05913 | 18466077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466077&pc=OZD | 12/23/2023 | 02/23/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 09/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024407103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2023 |  |  |  |  | The user facility reported an 82-year-old female in st elevated myocardial infarction with a large anterior myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced limb ischemia. there was no distal flow to the limb. impella reposition unit was wired and pulled back and angioseal was deployed next to the impella catheter. distal pulses restored, the patient remained on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31838 | 1220648-2024-05914 | 18466079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466079&pc=OZD | 12/24/2023 | 09/06/2024 |  | 01/08/2024 |  |  | 12/24/2023 | 01/08/2024 |  |  |  |  | 01/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024421569 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/24/2023 |  |  |  |  | The user facility reported a 52-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella device failed to advance beyond the de-air screen during patient support. the aic (automated impella controller) was swapped as a result of the noted issue and support continued with the implanted pump. there was no patient harm. | [5 empty] | The impella device was received from the customer, an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) issue has been completed. data logs were reviewed which confirmed the reported complaint. there were multiple instances of the console repeatedly detecting a purge cassette insertion and then removal in quick succession. the frequency of the detected purge cassette insertion and removal events did not match the expected behavior during realistic use; it is not possible to insert and then remove the purge cassette as quickly as the console was detecting these events. each instance of a purge cassette insertion or removal event corresponded to upwards of hundreds of imc-ahp errors triggered by issues reading the data from the rfid sensor. the aic was returned for evaluation. the reported complaint, an issue with purge fluid priming, was found to be caused by the console exhibiting unreliable detection of the rfid tag in the purge cassette. the issue was reproduced. the most likely cause was determined to be either an incorrect placement of the console¿s rfid sensor, or an electrical issue on the pud causing incorrect communication with the rfid sensor. f6/f8 should have been left blank in the initial report. h6 med dev prod code 2898 changed to 2880.
31839 | 1220648-2024-05915 | 18466114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466114&pc=OZD | 12/22/2023 | 03/05/2024 |  | 01/08/2024 |  |  | 12/22/2023 | 01/08/2024 |  |  | 05/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024355913 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The complainant in the eu reported 65-year-old female patient on impella 5.5 support had little brain infarctions shown in the cardiac computed tomography (cct) test. the support was for 5 days til wean and explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/ cva has been completed since the original report was filed. no device was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the stroke /cva was not determined. section b1 is updated as it was inadvertently marked incorrect in the initial report.
31840 | 1220648-2024-05916 | 18466134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466134&pc=OZD | 12/23/2023 | 04/16/2024 |  | 01/08/2024 |  |  | 12/24/2023 | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418095 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/24/2023 |  |  |  |  | The complainant reported that the patient began to bleed from their access site following impella cp implant. manual pressure and thrombix were initially applied to the site, but the bleeding continued. femostop was subsequently placed to limit the bleed until the pump could be explanted. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to lack of vessel recoil since best practices were followed and yet still bleed continued to occur from site.
31841 | 1220648-2024-05917 | 18466137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466137&pc=OZD | 12/29/2023 | 02/07/2024 |  | 01/08/2024 |  |  | 12/29/2023 | 01/08/2024 |  |  |  |  | 01/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1439249 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Low Readings | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 78-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. on day three of support after the purge bag and cassette change, the automated impella controller (aic) alarmed purge pressure low. the team observed the purge disc holder was cracked. the aic was replaced. no patient harm was reported, and support continues. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation. data logs from the complaint date revealed that the console issued the purge pressure low alarm several times. after visual inspection of this console, the purge retainer was confirmed to be broken. clinical staff reported that there was a small crack in the blue retainer, but we received it with the purge retainer completely broken off.
31842 | 1220648-2024-05918 | 18466197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466197&pc=OZD | 12/24/2023 | 12/20/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421271 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/24/2023 |  |  |  |  | The user facility reported a 86-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing when after 31 hours of the support the patient manually explanted themselves. the loss of appropriate positioning is a reportable event, despite this being a use issue and no harm or injury noted. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on the clinical details, the root cause of the positioning issues is most likely due to use as the patient pulled the impella out.
31843 | 1220648-2024-05919 | 18466211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466211&pc=OZD | 12/11/2023 | 11/22/2024 |  | 01/08/2024 |  |  | 12/11/2023 | 01/08/2024 |  |  | 09/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411347 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 78-year-old white male patient had an attempted impella cp implant. the pig and wire was advanced into left ventricle, but had difficulty advancing pump past bifurcation. removed and attempted again with stiff buddy wire without success. removed everything and tried 25cm sheath, but could not advance very much further, still unable to advance after multiple attempts. the implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the instructions for use for the impella cp states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ added-h6 impact code 4628 corrections to the initial mfr report: added d6a/d6b. h6 med dev prob code 1158 changed to 2682.
31844 | 1220648-2024-05920 | 18466221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466221&pc=OZD | 12/26/2023 | 11/22/2024 |  | 01/08/2024 |  |  | 12/26/2023 | 01/08/2024 |  |  | 05/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354047 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Tear, Rip or Hole in Device Packaging | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2023 |  |  |  |  | The user facility reported an impella cp in a 68yo male patient for mechanical circulatory support. it was reported that impella was placed too deep into left ventricle (lv) and required reposition at the bedside. while the cardiologist was pulling impella back the sterile sleeve was compromised. as the sterile sleeve was torn, the impella slipped out of lv and into the aorta requiring reposition at the bedside. the impella with torn sterile sleeve was left in place for 2 days. the patient was brought back to cath lab to exchange impella for new cp with sterile sleeve intact. patient remained hemodynamically stable during exchange of impellas. no additional information was provided. | [5 empty] | The impella device was returned for evaluation but has not yet been evaluated. upon completion of the investigation, a supplemental report will be submitted. | The investigation for the product damage has been completed. impella cp pump was returned. visually inspected the pump and the sterile sleeve found to twisted and torqued closer to the gwru. the sleeve was observed to be completely damaged and torn near touhy-borst side. no other abnormalities were found. the cause of the positioning issue was use related due to sterile sleeve being damaged and pump slipped out of lv dislodging into the aorta during repositioning per field investigation and clinical review. corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
31845 | 1220648-2024-05921 | 18466633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466633&pc=OZD | 12/21/2023 | 02/05/2024 |  | 01/08/2024 |  |  | 12/21/2023 | 01/08/2024 |  |  | 05/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024346851 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and after placement, there was low flows noticed. the impella was repositioned, and flows improved, however after lifting up the heart again, lv waveform started to go up about aortic placement signal, and flows appearing saturated. two hours later, mc started to gradually increase. the decision was made to remove the pump, and just before removal, the pump stopped. the impella was removed. | [5 empty] | The investigation into the pump stop and the saturated flow issues has been completed. the device was returned for benchtop investigation. data logs confirm saturated flow while at p-4. there are suction alarms and impella position unknown alarms throughout the first 2 hours of the case. data logs are missing the last 4 hours of the case including the pump stop. data logs also show motor current spikes prior to the event. upon investigation of the pump, a large clot was observed in the inflow cage. the saturated flow and pump stop were able to be reproduced in the lab. as the pump was purged, clots were expelled, and the saturated flow resolved. the root cause of the pump stop was determined to be biomaterial preventing the motor from being protected with purge fluid and eventually leading to a pump stop and saturated flow. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31846 | 1220648-2024-05922 | 18466962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18466962&pc=PYX | 12/29/2023 | 01/31/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 12/31/2023 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP PUMP |  |  | PYX |  |  | IMPELLA RP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2023160819 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a patient of unknown age and gender in acute myocardial infarction/cardiogenic shock with right ventricle dysfunction was to be implanted with an impella rp device for mechanical circulatory support. the rp insertion was not successful the wire passed, but not the pump. no patient harm was reported. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. the medical center did not return the impella device. the root cause of the delivery issue was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31847 | 1220648-2024-05923 | 18467295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18467295&pc=OZD | 12/30/2023 | 12/11/2024 |  | 01/08/2024 |  |  | 12/30/2023 | 01/08/2024 |  |  | 12/31/2024 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319617 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to the patient coding. when removed the physician noted that clot was visualized on the catheter. unable to determine location of the described clot. patient has been confirmed to have lv thrombus of unknown age present. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05923. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email. g1 reporting contact facility name missing. g1 reporting contact fax number missing.
31848 | 1220648-2024-05924 | 18467600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18467600&pc=OZD | 12/20/2023 | 10/29/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 11/30/2024 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303598 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | Disability | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 73-year-old female patient developed heparin induced thrombocytopenia (hit) during impella support. the patient was switched from heparin to bivalirudin and support with the implanted pump continued. roughly four days later, the p-level for the pump was decreased as it was believed that hemolysis caused by the pump was causing bleeding in the lungs. discussions were made for possible escalation to 5.5 support due to the noted concerns. systemic heparin was still not being used due to hit. the following day, a clot was ingested by the pump and the decision was made to replace the device. support continued with the new impella cp pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the thrombocytopenia, thrombus, and hemolysis has been completed. pump was not returned for investigation. data logs were returned and there was no positioning issue observed. there was no motor current spike or suction event observed before the bleeding in lungs event. patient was on heparin but data logs are inconclusive. therefore, the root cause of the hemolysis issue was not determined since product was not returned and data logs were inconclusive. the root cause of the thrombocytopenia and thrombus could not be determined without additional clinical details. added h6 component code 925 corrections to the initial mfr report: added d6a/d6b.
31849 | 1220648-2024-05925 | 18467639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18467639&pc=OZD | 12/12/2023 | 08/19/2025 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 07/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024374568 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Nerve Damage | Sepsis | Thrombosis/Thrombus | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a male patient noted as high-risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported male patient had a graft infection at impella 5.5 insertion site. | No further details were provided about the treatment of the graft infection. additionally, the patient was eventually weaned off of the device. | The study later reported upon multiple new adverse events with relationship to the pump. they are hematoma, median nerve injury, and thrombus in the axillary artery. |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ b5-outcome details added, h6 component code 925 added- d6a/d6b g1-corrected mdr reporting facility name and address, mfr site name and address. | Corrections to the initial mfr report: b.5. added new adverse events. g.1. mdr reporting contact email. h.6. clinical codes 1888, 4440, 1979.
31850 | 1220648-2024-05926 | 18468139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18468139&pc=OZD | 12/16/2023 | 10/10/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 09/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408878 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old female patient noted as high-risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was experiencing oozing from the impella access site. three units of packed red blood cells were given plus multiple attempts of angle matching and dressing changed. heparin also was stopped to decrease bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: added-d6a/d6b.
31851 | 1220648-2024-05927 | 18468644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18468644&pc=OZD | 12/19/2023 | 02/02/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  |  |  | 01/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1249361 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump support being intiated for a younger patient with medical comorbidities. the 5.5 and ecmo were to support the patient together. the automated impella controller (aic) for the 5.5 had a purge disc failure. there was no harm and ifu followed with backup console use. | [5 empty] | The investigation into the reported aic ¿ purge disc/ flag issue has been completed since the original report was filed. console was tested after arriving in fs. the purge disc not detected alarm issue was able to be reproduced, and purge retainer was found to be broken. portion of the purge retainer that is supposed to be fastened to the purge assembly was no longer intact. after reviewing the imc logs, the issues with the aic not being able to detect the purge disc was confirmed. there were numerous instances of the purge disc not detected alarm from the reported occurrence date. the root cause of this issue was a broken purge retainer. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
31852 | 1220648-2024-05928 | 18468711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18468711&pc=OZD | 12/31/2023 | 02/23/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 09/30/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024404437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/31/2023 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after insertion of the impella cp the patient had little to no flow past the impella sheath to the distal leg. femoral to femoral bypass with sheaths was performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31853 | 1220648-2024-05929 | 18469130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469130&pc=OZD | 12/13/2023 | 01/08/2024 |  | 01/08/2024 |  |  | 12/13/2023 | 01/08/2024 |  |  | 04/11/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024335032 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 42-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported had a placement signal issue. it was reported there was an appearance of pump failure alarm after an aic (automated impella controller) replacement. the complainant in japan wanted analysis on whether the impella or the aic have any problems. no known patient harm or injury. | [5 empty] | Pump was returned for investigation. no biomaterial or physical damage observed on cannula. pump successfully passed the electrical testing. impella defective alarm was appeared on aic during test run. upon closer examination, blood was discovered in the red handle, and a catheter hole was observed near the 25cm mark. the cause of the pump stop issue was blood ingress due to a hole in the catheter resulting from improper use. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31854 | 1220648-2024-05930 | 18469143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469143&pc=OZD | 12/16/2023 | 09/20/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  |  |  | 01/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1666954 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male patient noted as high-risk coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient was on impella support with the automated impella controller (aic) was unplugged by the nurse and it became disconnected from wires. the aic was replaced successfully. | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for automated impella controller (aic)-cabling issue has been completed. the console was received for investigation. data and device analysis was completed and the root cause of cabling issue was due to was a defective cable. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05930: b5 patient outcome and mso statement added d2 common device name g1 reporting contact fax number missing.
31855 | 1220648-2024-05932 | 18469172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469172&pc=OZD | 12/18/2023 | 10/03/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024414279 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a female patient noted as high-risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported patient on impella cp support experienced bleeding at the access site. the dressing had to be changed multiple times and the sutures had to be tightened to achieve hemostasis. | It was further reported that the patient expired on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since hemostasis was achieved by tightening the sutures. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added h6 component code 925 corrections to the initial mfr report: b5 mso review added. added d6a/d6b.
31856 | 1220648-2024-05931 | 18469174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469174&pc=OZD | 01/02/2024 | 02/02/2024 |  | 01/08/2024 |  |  | 01/02/2024 | 01/08/2024 |  |  |  |  | 01/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1087352 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2024 |  |  |  |  | No patient involvement: the us complainant had an aic (automated impella controller) that had a purge disc failure. recognition did not take place at the purge disc. the ifu (instructions for use) was followed and a backup is used. the malfunction was observed in an in-service. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed. the device was returned for investigation. imc logs confirmed the reported complaint of the purge disc not being detected through multiple issuances of the alarm. the consoles purge pressure sensor was visually inspected to reveal a missing purge flag. unrelated to the complaint is a crack in the purge retainer. the cause of the issue was a defective component. the root cause for the purge disc not detected alarm was a missing purge flag.
31857 | 1220648-2024-05933 | 18469193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469193&pc=OZD | 12/20/2023 | 03/05/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  |  |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional |  |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported through a literature review that impella support was started after pci on a patient with acute myocardial infarction (ami) with cardiogenic shock. eight days after pci a ct scan revealed a left subcortical hemorrhage with midline shift. an emergency craniotomy for hematoma evacuation was performed. hemostasis was difficult to establish. intraoperative transfusion with red cell concentrated (rcc) 1070 ml, fresh frozen plasma (ffp) 720 ml, and platelet concentrates (pc) 400 ml was required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/ cva has been completed since the original report was filed. no device was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the stroke /cva was not determined.
31858 | 1220648-2024-05934 | 18469228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469228&pc=OZD | 12/23/2023 | 02/23/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024409545 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2023 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a 5.5 pump via the axillary insertion site, for a heart failure patient, when on 2nd day of the support the site had to be intervened upon. the team placed more sutures and performed a washout for the access site bleed. the pump remains on support despite the harm and is now on day 17. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleedin- major has been completed since the original report was filed. the root cause of the access site bleeding was most likely due to use issue since the bleeding was resolved by resuturing the bleeding site. the root cause of the access site bleeding was use related.
31859 | 1220648-2024-05935 | 18469301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469301&pc=OZD | 12/23/2023 | 02/21/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 03/31/2024 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023203595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2023 |  |  |  |  | The user facility reported a 52-year-old white female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a hematoma developed during impella implant and worsened after the repositioning sheath was inserted. while pressure was initially held to manage the condition, the decision was ultimately made to explant the pump. the perclose broke upon pump removal and manual pressure was held for roughly 40 minutes to achieve hemostasis and hematoma resolution. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was most likely due to anticoagulation management since the act value was 253 during bleeding which is out of the range (160-180).
31860 | 1220648-2024-05936 | 18469384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469384&pc=OZD | 12/20/2023 | 02/05/2024 |  | 01/08/2024 |  |  | 12/20/2023 | 01/08/2024 |  |  |  |  | 01/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1557936 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported they are using an external monitor to duplicate automated impella controller (aic) console display, but it does not work anymore. the complainant believes that the vga-port of the rlm is defective. the monitor was reportedly tested with their other aic's without a problem. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the console display issue has been completed since the original report was filed. the console was returned for investigation. the console logs and remote lan modules (rlm) journals do not contain any information related to vga signal output. console and rlm troubleshooting was performed and concluded that the issue was caused by defective rlm, as the monitor connected directly to the aic input/output port worked fine but did not work when connected consoles rlm even while using a known-good backstrap cable. inspection of the console revealed minor damage of the backstrap cable wiring, and although it was unrelated to complaint, it is recommended to be replaced. the root cause of the issue (no vga out signal) was a defective impella connect module.
31861 | 1220648-2024-05937 | 18469434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469434&pc=OZD | 10/03/2023 | 11/12/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  |  |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 54 YR |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemolysis | Hemorrhagic Stroke | Renal Impairment | [9 empty] | Required Intervention | Life Threatening |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported after reviewing a literature case it was found a 54-year-old male of unknown race was experiencing a myocardial infarction with cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. refractory hemolysis, acute renal failure, and anemia were observed two days after percutaneous coronary intervention (pci); haptoglobin injection and continuous hemodiafiltration (chdf) was initiated. additionally, a computed tomography (ct) scan revealed right intracerebral hemorrhage. an emergency surgical hematoma evacuation was performed on the same day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis. | The investigation for the stroke and hemolysis has been completed. there is no information provided about the type of pump and patient information. therefore, no pump serial number available to receive data logs. therefore, the root cause of the hemolysis issue was not determined since product was not returned. and insufficient clinical information provided. the root cause of the stroke issue was not determined since product and data logs were not returned and insufficient clinical information was provided. the root cause of the renal failure issue was not determined since product was not returned. added- h6 component code 925 corrections to the initial mfr report: d4-model number f6/f8 should have been left blank.
31862 | 1220648-2024-05938 | 18469458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18469458&pc=OZD | 12/29/2023 | 09/06/2024 |  | 01/08/2024 |  |  |  | 01/08/2024 |  |  | 10/31/2025 |  | 01/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024417677 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2023 |  |  |  |  | The user facility reported a 56-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and experienced low pump flows and pressure issues. flows eventually became critical and the impella was explanted. post explant a clot was noted on the outflow and motor housing. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. clinical data states that low purge flows and high purge pressure observed and clot found on outflow upon explant. the data logs confirm purge pressure high alarms which resolved during the case. there was a gradual rise in purge pressure before the pump was explanted. the pump was inspected and a slightly kinked purge tubing was found in the pump plug. the pump was run in the lab and high purge pressure did not reproduce. no other issues were found on the rest of the pump. the root cause of the high purge pressure was most likely biomaterial build up. added- d9 device return date d4-corrected model number. f6/f8 should have been left blank in the initial report.
31863 | 1220648-2024-05818 | 18450180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18450180&pc=OZD | 12/27/2023 | 02/07/2024 |  | 01/05/2024 |  |  | 12/27/2023 | 01/05/2024 |  |  |  |  | 01/05/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1365450 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/27/2023 |  |  |  |  | No patient involvement: the complainant reported that the automated impella controller (aic) purge disc was broken. the aic was removed from service. | [5 empty] | The device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console and the logs were returned for evaluation. the data logs were reviewed finding that there were purge disc not detected alarms. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag.
31864 | 1220648-2024-05817 | 18450406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18450406&pc=OZD | 12/25/2023 | 02/02/2024 |  | 01/05/2024 |  |  | 12/25/2023 | 01/05/2024 |  |  |  |  | 01/05/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1089664 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/25/2023 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support the automated impella controller (aic) had an air in the purge alarm. upon opening the cassette door, it was noticed that the purge disc holder on the console was broken. the aic was replaced successfully. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the reported, aic -purge disc/ flag issue has been completed since the original report was filed. analysis of device by visual inspection of this console, the purge retainer was confirmed to be broken. and the data logs confirmed the alarm several times.
31865 | 1220648-2024-05815 | 18450886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18450886&pc=PYX | 12/21/2023 | 10/30/2025 |  | 01/05/2024 |  |  | 12/21/2023 | 01/05/2024 |  |  | 02/28/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024337535 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | Improper Flow or Infusion |  |  |  |  | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 42-year-old black female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a rp flex inserted for support along with a 5.5 pump for bipella support. after 1 day/18hours the pump had low flows and suspected clot. the team made decision to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31866 | 1220648-2024-05816 | 18452214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18452214&pc=OZD | 12/19/2023 | 10/14/2024 |  | 01/05/2024 |  |  | 12/19/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412718 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 32-year-old black male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the patient developed hemolysis during impella cp support. repositioning was completed and volume was given, but the hemolysis persisted. due to the noted condition, the decision was made to replace the cp pump with va ecmo and iabp support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were returned and investigated. clinical mentions that hemolysis was tried to be resolved by repositioning the pump (confirmed "impella in aorta" alarms in data logs) and giving volume . the next day hematuria persisted and more pressors were given to the patient. repositioning did not resolve hemolysis. without the pump, hemolysis cannot be further investigated. therefore, the root cause of the hemolysis issue was not determined since product was not returned. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: added-d6a/d6b f6/f8 should have been left blank.
31867 | 1220648-2024-05814 | 18452418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18452418&pc=OZD | 12/08/2023 | 01/25/2024 |  | 01/05/2024 |  |  | 12/08/2023 | 01/05/2024 |  |  | 01/31/2024 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023181540 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 61-year-old white female with cardiomyopathy was implanted with an impella for mechanical circulatory support. patient is reporting pain in right arm near impella insertion site with small hematoma present. no action was taken to address it as it got better. support continued with the implanted pump following the issue. | [5 empty] | The investigation into the reported access site bleeding has been completed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
31868 | 1220648-2024-05819 | 18452598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18452598&pc=OZD | 12/16/2023 | 12/09/2024 |  | 01/05/2024 |  |  | 12/16/2023 | 01/05/2024 |  |  | 05/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024358290 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | Increase in Pressure | Pumping Stopped |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported impella cp filter "torn off". a new impella installation or transfer of the patient was rejected. there was no reported patient harm. | After the purge filter bypass was installed so late, the purge pressure also went high, low purge flow and finally the motor stopped. replacing the pump took hours again, because no available doctor and then changed to a new impella cp. the patient did not survive. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into pump stop, high purge pressure, and purge leak - pump has been completed. the root cause was not determined since product was not returned and data logs are not available. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05819. b1 adverse event and product problem. b2 death and date of death added. b5 patient outcome and death statement added. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact email. h1 type of reportable event. h6 medical device problem codes 1503 and 1491 were missing. h6 code 4582, 2199, and 481 were reported incorrectly. new codes were added to health effect - clinical code, health effect - impact code, and component code h10 instructions for use missing.
31869 | 1220648-2024-05821 | 18452787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18452787&pc=OZD | 12/21/2023 | 02/21/2024 |  | 01/05/2024 |  |  | 12/21/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418776 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 61-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient developed a hematoma following impella implant. the patient was taken to the operating room where an evacuation was completed to treat the condition. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided to determine the cause of the hematoma.
31870 | 1220648-2024-05820 | 18453063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18453063&pc=OZD | 12/15/2023 | 02/16/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 01/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024311774 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp pump had to be pulled because the 0.018 wire split in the impella during threading. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the product damage (guidewire damage - great vessel) has been completed since the original report was filed. the device was returned for investigation. product evaluation confirmed that there was a coil fracture and unravelling at the tip of the guidewire. the cause of the guidewire product damage is not determined because not enough clinical information was provided.
31871 | 1220648-2024-05822 | 18453415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18453415&pc=OZD | 12/18/2023 | 02/05/2024 |  | 01/05/2024 |  |  | 12/22/2023 | 01/05/2024 |  |  |  |  | 01/05/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1081051 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | Device Contamination with Chemical or Other Material | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 59-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the automated impella controller (aic) had aic soft key touch failure. dextrose 5% in water (d5w) was spilled on the aic. the aic was not removed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported, aic - kbd/rotary knob issues has been completed since the original report was filed. the device was returned for analysis. the failure mode was not reproduced. consoles are cleaned upon receipt for investigation and may have removed the contamination that caused the failure. the cause of the aic issue was most likely the spilled d5 water.
31872 | 1220648-2024-05823 | 18453637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18453637&pc=OZD | 12/22/2023 | 10/24/2024 |  | 01/05/2024 |  |  | 12/22/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400665 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported an 82-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon explant of the impella cp device, complications were encountered with closing the site. the perclose was noted to be unsuccessful and a covered stent was ultimately placed in the patient's left femoral artery to treat the site. no bleeding was noted following site closure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05823: d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
31873 | 1220648-2024-05827 | 18454079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18454079&pc=OZD | 12/15/2023 | 01/17/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024418698 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 49-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support it was noted the anticontamination sleeve detached at base near the blue reposition button. the patient remained on support and there was no patient harm reported. | [5 empty] | The investigation into the reported product damage (sterile sleeve damage) has been completed since the original report was filed. impella 5.5 pump was returned. the returned product was visually inspected and found the sterile sleeve to be torn, damaged and separated close to the blue button end. no other abnormalities were observed. as per clinical, sterile sleeve was detached at base near repositioning button and securement was too tight with movement of patient also seen. the cause of the sterile sleeve damage issue was use related with the sleeve detached at the base near to blue button per clinical and field investigation. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31874 | 1220648-2024-05825 | 18454089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18454089&pc=OZD | 12/14/2023 | 01/17/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 05/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344588 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 76-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was unable to be advanced through the subclavian artery into the innominate artery despite multiple attempts and manipulations of the pump. the impella 5.5 was removed. contrast was infused to evaluate the vessel. it appeared that the pump was unable to pass due to patient¿s anatomy. the doctor elected to insert the impella cp through the axillary graft instead. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31875 | 1220648-2024-05828 | 18454188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18454188&pc=OZD | 12/09/2023 | 06/19/2024 |  | 01/05/2024 |  |  | 12/09/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403851 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Swelling/ Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2023 |  |  |  |  | The user facility reported a 43-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was placed via the axillary when the patient developed swelling around the site and firm to the touch. the team treated as a hematoma and withheld the heparin. additionally quick clot was applied to the site. this happened 2 days after implant. support was maintained with the implanted pump. | The patient expired on support. upon review by abiomed's medical safety officer, there is not enough information to associate the use of the device with the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding - major issue has been completed. the device was not returned for investigation. the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time values, heparin was stopped to achieve hemostasis as per the clinical description. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
31876 | 1220648-2024-05826 | 18454260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18454260&pc=OZD | 12/08/2023 | 10/16/2024 |  | 01/05/2024 |  |  | 12/08/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397235 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 48-year-old female with newly diagnosed cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on impella support, the patient had access site bleeding. pressure was held and angle matching was performed unsuccessfully. purse string suture placed, blue suture hub advanced appropriately and re-sutured blue wings, femstop applied briefly. act 148, protamine 50mg was given and site quickly achieved hemostasis. | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05826. b5 patient outcome and mso statement added. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures.
31877 | 1220648-2024-05829 | 18454533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18454533&pc=OZD | 12/10/2023 | 12/23/2024 |  | 01/05/2024 |  |  | 12/10/2023 | 01/05/2024 |  |  | 04/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350356 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  |  | Malfunction | No | Yes | Fracture | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported a 69 year old female presenting for acute myocardial infarction/cardiogenic shock was implanted with an impella cp pump for mechanical circulatory support. it was documented that the touhy won't tighten and could not be positioned properly due to product damage. impella was successfully removed. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. | The investigation for the reported positioning issue has been completed. no product returned. the cause of the positioning issues was not determined due to no product returned and lack of details about touhy locking issue.
31878 | 1220648-2024-05831 | 18454663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18454663&pc=OZD | 12/14/2023 | 12/24/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397021 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported an impella cp in a 63yo male for support during a coronary bypass graft procedure. it was reported that the pump came out of position within the first hour of support and was replaced. no lasting harm or injury from interruption. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. when the investigation has been completed, a supplemental report will be submitted. the failure mode will be monitored and trended. | The investigation for the positioning issue has been completed. no product was returned. pump was initially placed under papillary muscle and then unable to be repositioned. unknown details about patient condition, anatomy, patient movement at time pump improperly positioned. the cause of the positioning issues was not determined since product was not returned and lack of clinical details. corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank g1 mdr reporting contact email
31879 | 1220648-2024-05824 | 18454712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18454712&pc=OZD | 12/22/2023 | 02/02/2024 |  | 01/05/2024 |  |  | 12/22/2023 | 01/05/2024 |  |  |  |  | 01/05/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1182670 |  | 0042-0000 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support when the automated impella controller (aic) experienced the purge disc was not detected on the aic. the aic was replaced. | [5 empty] | The impella device was received 01/02/2024, an evaluation of the device was will be performed. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge issue has been completed. the console and data logs were returned for evaluation. the logs were reviewed finding that the only purge disc not detection alarm was detected several days after the reported event. the console was visually inspected for defected finding the purge retainer was completely broken off. despite what the clinical staff reported, there were no issues detecting the purge cassette. suspect that the clinical staff realized that the retainer was broken the day before the reported occurrence date (12/21/23) and meant to report that the purge disc could not be detected due to the broken retainer. therefore, the root cause of the issue was a broken purge retainer.
31880 | 1220648-2024-05833 | 18454878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18454878&pc=OZD | 12/11/2023 | 02/02/2024 |  | 01/05/2024 |  |  | 12/11/2023 | 01/05/2024 |  |  | 04/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332285 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported an impella 5.5 in a 42yo male patient for mechanical circulatory support. it was reported that there was difficulty while attempting to insert the impella 5.5 device. a clot was suspected but it appeared that a wire was not all the way out when the device was ramped up. the impella was replaced successfully. there was no patient harm reported. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. | The pump was not returned for investigation. the data logs showed several pump stop-motor current high alarms during the start of the case indicating interaction with another device or biomaterial. according to the clinical report, the doctor found that the guidewire was not fully removed from the pump after the initial placement. based on the given clinical and data log review, the cause of the pump did not start issue was most likely user related. the guidewire as not fully removed during the pump¿s start causing the pump to be unable to start.
31881 | 1220648-2024-05834 | 18455037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455037&pc=PYX | 12/16/2023 | 09/21/2024 |  | 01/05/2024 |  |  | 12/16/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024418610 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/16/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella rp flex support when the impella rp was turned down for suspected pulmonary hemorrhage. the impella rp was slowly ramped up and suddenly staff noticed frank red blood in endotracheal tube (ett). the staff did not know why the pulmonary hemorrhage occurred. the total amount of blood and saline suctioned from ett was 150 milliliters. care was withdrawn after 0.67 hours of support. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation into the reported, vascular damage has been completed since the original report was filed. product was not returned for review. the cause of the vascular damage issue was most likely patient condition related and unknown relation to impella per clinical review.
31882 | 1220648-2024-05835 | 18455132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455132&pc=OZD | 12/11/2023 | 03/14/2024 |  | 01/05/2024 |  |  | 12/11/2023 | 01/05/2024 |  |  | 12/31/2024 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287443 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The complainant reported issues with delivering impella cp due to the male patient's anatomy. the impella was unable to go over the bifurcation. the impella was therefore explanted and when it was removed, there were some "clottings" noted on the impella. therefore, a new one was used. there was no patient harm. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the pump was returned for investigation, and upon inspection, no physical abnormalities were found on the returned pump. the pump performed as expected during testing. as per clinical details, pump was not advanced beyond bifurcation of the artery. the root cause of the delivery issue was patient condition related to patient anatomy since pump couldn't advance beyond the artery's bifurcation, as per the clinical details and investigation of the returned product.
31883 | 1220648-2024-05830 | 18455264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455264&pc=OZD | 11/16/2023 | 02/22/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 05/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024363341 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a female patient noted as high-risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the removal of the impella cp and closure system using the proproglide/angioseal8f resulted in subtotal stenosis of the common femoral artery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31884 | 1220648-2024-05837 | 18455332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455332&pc=OZD | 06/17/2017 | 11/19/2025 |  | 01/05/2024 |  |  | 12/08/2023 | 01/05/2024 |  |  | 02/28/2019 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1283643 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the 69-year-old male patient developed ischemia in their impella leg during pump support. roughly a week after impella 5.0 implant, it was noted that the patient's left foot was cold. surgical intervention was performed and flow was restored to the extremity. following the event, it was documented that the patient had clotting issues and persistent high purge pressure / low purge flows were encountered. the aic was swapped on multiple occasions, but the issue remained. as a result, the decision was made by the doctor to add tpa to the system. the patient remained stable for the remainder of support until planned explant. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. b.3 revised date as it was reported incorrectly on the initial manufacturer device report number 1220648-2024-05837. d.6a and d.6b added since the initial submission of manufacturer device report number 1220648-2024-05837 in accordance with updated procedures. e.1 fax number was inadvertently left off the previously submitted manufacturer device report number 1220648-2024-05837 and was added. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2024-05837 in accordance with updated procedures. added reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-05837 was submitted. h.6 type of investigation code was revised as it was reported incorrectly on initial manufacturer device report number 1220648-2024-05837. h.6 codes 3233 and 11 were reported incorrectly on the initial manufacturer device report number 1220648-2024-05837 in accordance with updated procedures. a new code was added. h.10 additional manufacture narrative was revised as it was reported incorrectly on initial manufacturer device report number 1220648-2024-05837. | The investigation has been completed. limb ischemia - reversible. in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. high purge pressure: the pump was not returned for investigation. data logs show a gradual 10-hour rise in purge pressure, which remained high for a total of 40 hours before returning to the normal range. a motor current spike was reported before the purge pressure started to rise. as per the clinical, the purge pressure alarm was noted during the case. the doctor started tpa and observed an improvement in purge pressure and flow. the root cause of the high purge pressure issue was most likely biomaterial ingestion since adding tpa successfully resolved the issue. the case was further investigated for thrombus ingestion. thrombus. in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. h6: added codes per the results of the investigation and the clinical code. h10: added the results of the investigation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31885 | 1220648-2024-05839 | 18455354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455354&pc=OZD | 12/13/2023 | 06/11/2024 |  | 01/05/2024 |  |  | 12/13/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418073 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump began to experience suction alarms following placement. under tee (transesophageal echocardiography) and fluoroscopy, it was discovered that the impella pigtail was caught underneath the papillary muscle and directed posterior towards the mitral valve. after initial attempts to reposition, the pump resulted in the device backing into the aorta, the pump was fully removed and re-inserted. the pigtail of the pump once again got caught underneath the papillary muscle and the decision was ultimately made to replace the device. a new impella cp pump was placed for planned support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. according to the clinical, shortly after placing the impella cp device it was discovered that the pigtail was caught in the papillary muscle. multiple attempts were made unsuccessfully until the physician accidentally pulled the pump into the aorta. the decision was then made to remove the pump and replace it with another impella cp device. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issues was use related. added- h6 impact code 4628 d4- updated model number. added- d6a/d6b.
31886 | 1220648-2024-05838 | 18455369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455369&pc=OZD | 12/16/2023 | 05/23/2024 |  | 01/05/2024 |  |  | 12/16/2023 | 01/05/2024 |  |  | 09/03/2024 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281729 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2023 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the day after implant, the patient had a groin site ooze and hemoglobin drop. the team troubleshot by infusion of one unit of red blood cells. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is from ifu impella cp with smart assist system in section: general patient care considerations, entering cardiac output, and potential adverse events (united states) , which now includes "(including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31887 | 1220648-2024-05840 | 18455431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455431&pc=OZD | 12/13/2023 | 02/23/2024 |  | 01/05/2024 |  |  | 12/13/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024412200 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. us complainant reported the patient was on ecpella support when they had lower than expected doppler pulses bilaterally. the patient was escalated to the impella 5.5 as a result of the noted condition. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31888 | 1220648-2024-05841 | 18455471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455471&pc=OZD | 12/18/2023 | 08/13/2025 |  | 01/05/2024 |  |  | 12/18/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412021 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 65-year-old asian male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the access site was oozing and femstop was applied. it was also reported that the patient labs were hemolyzed. the impella cp was going to be retracted to avoid further hemolysis until the patient goes to cath lab for pump removal. after the second day of patient being on impella support the cp was removed due to confirmed thrombus in left ventricle, access site bleeding, and hemolysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of access site bleeding, hemolysis, and thrombus has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. hemolysis: data logs were returned and "impella in aorta" alarms were observed. clinical also mentioned that the pump was 5 cm deep and was repositioned. therefore, the root cause of the hemolysis issue was improper positioning of the device due to improper technique. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined" d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-05841 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-05841 in accordance with updated procedures. h6 component code revised from the initial submission in accordance with updated procedures.
31889 | 1220648-2024-05836 | 18455589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455589&pc=OZD | 12/12/2023 | 02/23/2024 |  | 01/05/2024 |  |  | 12/12/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024410759 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that there were no pulses on the patient¿s left leg. the impella cp was removed and impella 5.5 was successfully placed. the patient is stable. | [5 empty] | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31890 | 1220648-2024-05842 | 18455636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455636&pc=OZD | 12/15/2023 | 01/26/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  |  |  | 01/05/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1141486 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 95 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a (b)(6) year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the aic (automated impella controller) doesn't recognize the purge disc and it won't purge the pump and will not change to the next screen. the aic was removed for service. | [5 empty] | The investigation into the into the aic - purge disc/flag issues has been completed since the original report was filed. after reviewing the imc logs, the issues with the aic not being able to detect the purge disc was confirmed. the purge disc not detected alarm occurred during the purge case start wizard which aligns with what the clinical staff reported. console was tested after arriving in fs. after visual inspection of the console, the reported purge disc not detected issue was determined to be caused by a broken purge. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31891 | 1220648-2024-05843 | 18455683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455683&pc=OZD | 12/13/2023 | 07/15/2024 |  | 01/05/2024 |  |  | 12/13/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024409551 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 49-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during insertion the catheter kept prolapsing on itself while advancing. no traumatic force was used when attempting to advance. the impella and wire were removed, and left ventricle was rewired for a stronger rail. the device continued to have the same issues. surgeon decided to exchange for new impella and retry. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issues was most likely due to the patients condition as the clinical details stated they were unable to pass through the vasculature. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. added-d6a/d6b f6/f8 should have been left blank in the initial report. h6 med dev prob code changed to 2682.
31892 | 1220648-2024-05845 | 18455702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455702&pc=OZD | 12/19/2023 | 01/18/2024 |  | 01/05/2024 |  |  | 12/19/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412792 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 75-year-old white male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the patient's native anatomy precluded placement despite all troubleshooting. wire and sheath exchanges still did not allow the cp to reach the left ventricle and support was aborted. patient taken to or for coronary artery bypass grafting. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the cause of the issue was patient condition related.
31893 | 1220648-2024-05844 | 18455739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455739&pc=PYX | 12/15/2023 | 02/06/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024413692 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 71-year-old male (race unknown) with right heart failure underwent cardiac surgery was implanted with an impella rp device for mechanical circulatory support. the pump was explanted after four days when the pump had low flows and presumed clot on the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the thrombus issues has been completed. the device was returned for benchtop investigation. the pump was visually inspected, and cannula seems to be damaged due to clamping. borescope view was obtained and biomaterial ingestion around the impeller was found. data logs were reviewed, and one instance of motor current spike followed by drop in flow observed at p-6. additionally, the purge pressure low alarms mostly seem to be occurring during purge bag/cassette change process. the cause of the low pump flow issue was due to biomaterial ingestion per field investigation and log review. correction : section d9: the information was inadvertently left out in the initial report which has been updated in this report.
31894 | 1220648-2024-05846 | 18455751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455751&pc=OZD | 12/15/2023 | 05/07/2024 |  | 01/05/2024 |  |  | 12/19/2023 | 01/05/2024 |  |  | 07/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387474 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 77-year-old asian male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. following impella implant, the aorta and left ventricle position waveforms became uncoupled and suction alarms began to occur frequently. although the pumps positioning was verified to be correct, the suction alarm remained present and the decision was made to explant the device. there was no patient harm. | [5 empty] | The investigation for the reported optical signal issue and low pump has been completed. the impella pump was returned for investigation. no physical abnormalities were observed on the returned pump. pump ran as expected during test in bucket of water without reproducing low pump flow. upon optical testing, all 5s optical parameters were found within specification. from clinical details and data log review, placement signal trend was noted at beginning of the case start, which came back to normal without seeing any abnormalities in pump performance. shortly after 5hr, a continuous suction alarm and low pump flow was reported with motor current spike at the end of the case with loss of optical signal. the cause of the low pump flow issue was most likely biomaterial since low pump flow did not preproduced during test and motor current spikes were reported without noted placement signal trend or improper position. the cause of the optical signal issue was most likely patient condition related based on returned product review and data log review. added-d9 device return date added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31895 | 1220648-2024-05848 | 18455766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455766&pc=OZD | 12/19/2023 | 02/23/2024 |  | 01/05/2024 |  |  | 12/19/2023 | 01/05/2024 |  |  | 02/28/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338491 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 69-year-old male (race and ethnicity unknown) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had poor perfusion to bilateral lower extremities (ble) prior to impella placement and both le were non pulsatile and cold to the touch. post impella insertion the rle was noted to have color change and still unable to get pulses on either lower extremity. both remained cold to the touch. the patient was placed on ecmo and jump graft was placed to perfuse the limb antegrade from the circuit. jump graft was unsuccessful. the impella was removed the following day and both limbs have regained pulses and color. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31896 | 1220648-2024-05847 | 18455770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455770&pc=OZD | 12/17/2023 | 01/17/2024 |  | 01/05/2024 |  |  | 12/17/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398938 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2023 |  |  |  |  | The user facility reported an impella cp was unsuccessfully implanted in a 89-year-old male for mechanical circulatory support during high risk percutaneous coronary intervention. it was reported that several unsuccessful attempts were made to advance the impella in the left ventricle. the pump motor was unable to pass around the aortic arch due to extensive calcification. multiple wires and techniques were attempted but were unsuccessful. the case was aborted. no additional information was provided. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition. | The impella device was not returned for evaluation. a supplemental report will be submitted when the investigation has been completed.
31897 | 1220648-2024-05849 | 18455780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455780&pc=OZD | 12/14/2023 | 11/01/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024389247 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 66 -year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the sterile sleeve was ripped on the catheter. it was advised to not reposition the impella and to cover it with tegaderm. no patient harm is reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues product damage could not be determined. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings and cautions: cautions: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ corrections to the initial mfr report: f6/f8 should have been left blank.
31898 | 1220648-2024-05850 | 18455848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455848&pc=OZD | 12/20/2023 | 11/18/2024 |  | 01/05/2024 |  |  | 12/20/2023 | 01/05/2024 |  |  | 04/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347841 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported a 86-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had implanted a cp for support of a transferred patient. the tertiary center inserted the cp but, the patient had signs of hemolysis and the team explanted after 9 hours and placed an iabp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were returned and there was slight downward trend observed in the placement signal and motor current without change in p-level. no suction or positioning issues observed in the data logs. therefore, based on the data log analysis, the root cause of the hemolysis issue was most likely patient condition related. as per the clinical information, the hematoma formed after the patient was moved. therefore, the root cause of the access site bleeding issue was patient movement.
31899 | 1220648-2024-05852 | 18455876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455876&pc=OZD | 12/11/2023 | 01/05/2024 |  | 01/05/2024 |  |  | 12/11/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC |  |  | US |  | 2024413047 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/11/2023 |  |  |  |  | The complainant reported that impella cp smartassist heart pump was unable to be implanted on 82-year-old female patient due to anatomy. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, evaluation of the device was not possible. based on the provided information, the noted issue was attributed to the patient's anatomy. should any new information be acquired, a supplemental mdr will be filed.
31900 | 1220648-2024-05851 | 18455881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455881&pc=OZD | 12/20/2023 | 10/29/2024 |  | 01/05/2024 |  |  | 12/20/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024402875 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported a 64-year-old white female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was transferred to tertiary center for escalated care and monitoring where it was discovered the patient developed hemolysis due to the pump¿s placement being too deep and suction was present. the patient also required emergent fasciotomy to impella site. the pump was explanted, and care escalated to the 5.5 and extracorporeal membrane oxygenation (ecmo). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: hemolysis. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ limb ischemia. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation of hemolysis and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05851. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 medical device problem code 2993 was duplicated in error h6 component code 3038 was reported in error. new component code added.
31901 | 1220648-2024-05853 | 18455900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455900&pc=OZD | 12/13/2023 | 05/23/2024 |  | 01/05/2024 |  |  | 12/13/2023 | 01/05/2024 |  |  | 07/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370523 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 56 -year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the patient developed a hematoma and oozing at their access site following impella implant. heparin administration was temporarily halted as a result of the noted condition. support was able to continue with the implanted pump. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding issue was most likely anticoagulation management as stopping systemic heparin resolved the bleed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31902 | 1220648-2024-05854 | 18455918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455918&pc=OZD | 12/13/2023 | 01/26/2024 |  | 01/05/2024 |  |  | 12/13/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024422759 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 67-year-old patient received hemodynamic support from an impella 5.5 smartassist heart pump due to acute myocardial infarction/cardiogenic shock. a leak in the purge line was reported, which could be bridged with a bypass. the patient suffered no harm as a result of the incident. at the time of the report, the impella 5.5 heart pump was still in operation. | [5 empty] | The impella device was returned for analysis and an evaluation of the device is on-going. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak ¿ pump has been completed. the device was returned for benchtop investigation. however, the sidearm was cut and was not returned. the purge leak could not be investigated due to the sidearm not being returned. data logs indicate the trend of the purge leak after 4 hours of pump usage which was resolved after sidearm bypass. based on given clinical details, leak was observed after yellow luer was pulled tight by the nurse. purge sidearm bypass was successfully performed by doctor. the cause of the purge leak was sidearm yellow luer damage due to excessive force was applied by nurse.
31903 | 1220648-2024-05855 | 18455923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455923&pc=OZD | 12/14/2023 | 05/16/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 05/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024352463 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the implanted impella 5.5 pump was no longer unloading during patient support. as a result, the pump was removed and replaced with another impella 5.5 device. it was noted that biomaterial was observed in conjunction with the flow issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issues has been completed since the original report was submitted. the device was not returned for investigation. data logs reported less pump flow output than expected, with several suction alarms. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided.
31904 | 1220648-2024-05856 | 18455931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455931&pc=OZD | 12/14/2023 | 11/11/2025 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 03/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024347285 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Component | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported an 81-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the dilator for 13cm 14 french peel away sheath would not advance completely into sheath. defect was noted during preparation of the sheath on back table and did not touch patient. pulled second sheath kit, which was used successfully. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31905 | 1220648-2024-05857 | 18455938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455938&pc=OZD | 12/15/2023 | 09/21/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398243 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Material Twisted/Bent | Positioning Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 61-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. when inserting the impella cp, the catheter kinked as it transitioned from the subclavian into the aorta. the cp was able to be successfully placed. the patient is stable. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation into the delivery issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related.
31906 | 1220648-2024-05858 | 18455947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18455947&pc=OZD | 12/14/2023 | 06/03/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024411257 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 52-year-old male was implanted with an impella 5.5 as a bridge to transplant. the impella insertion site developed a hematoma, which was treated with vascular repair. the patient remained on impella 5.5 support while awaiting transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use noted in the initial report is from the impella 5.5 with smart assist for use during cardiogenic shock section: : potential adverse events (united states) , which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
31907 | 1220648-2024-05859 | 18456030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456030&pc=PYX | 12/14/2023 | 05/22/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 01/31/2024 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP EU |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023180984 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male patient in st elevated myocardial infarction and right heart failure was implanted with an impella rp device for mechanical circulatory support. the complainant reported that a clot was observed on the impella rp inlet cage. there was also a 5cm structure attachef. therefore, the impella rp was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
31908 | 1220648-2024-05861 | 18456090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456090&pc=OZD | 12/18/2023 | 10/02/2024 |  | 01/05/2024 |  |  | 12/18/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419519 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 77-year-old male for mechanical circulatory support during high risk percutaneous coronary intervention. it was reported that there was bleeding at the impella access site. angiography was performed from contralateral access and showed a leak from the common femoral artery (cfa) at the level of the superficial femoral artery (sfa)/profunda bifurcation. a stent was placed but a leak remained at distal end of the stent. the patient was taken to the operating room for a cut down and repair. a non-abiomed pro-glide caught the artery and side branch holding it open which was bleeding. surgical closure was successful. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. a supplemental report will be submitted when the investigation has been completed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported vascular damage has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of vascular damage is determined to use issue because there was leak from common femoral artery (cfa) at level of superficial femoral artery (sfa)/profunda bifurcation and pro glide didn¿t catch the artery and had to do perform the cut down and repair for surgical closure. hence, the root cause is determined to be use issue. added- h6 component code 925 corrections to the initial mfr report: f6/f8 should have been left blank in the initial report.
31909 | 1220648-2024-05860 | 18456095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456095&pc=OZD | 12/15/2023 | 03/12/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391572 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a patient in cardiogenic shock was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella cp smart assist support for 38.50 hours before the patient expired. the cause of death was determined to be the patient's heart rupturing. the cause of the rupture was unknown. impella use cannot be excluded as a potential relationship towards the cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported heart valve damage has been completed. the impella device was not returned for evaluation. the root cause of the reported issue could not be determined since product was not returned and there is a lack of clinical details. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ d4-updated model number.
31910 | 1220648-2024-05862 | 18456102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456102&pc=OZD | 12/15/2023 | 12/04/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 05/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024361120 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 47-year-old male was on impella cp support for a hypoxic patient with pulmonary embolism arrest. the patient developed hemolysis while on impella support and blood products were provided as treatment. the pump remained on for 3 more days of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05862. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
31911 | 1220648-2024-05863 | 18456115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456115&pc=OZD | 12/14/2023 | 04/11/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024396686 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Injury | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 80-year-old male noted as high-risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complaint reported that following implantation of impella cp smartassist heart pump, there were high purge pressures >1050 mm hg with low purge flows, 2 ml/h and red high purge pressure alarms. the alarms were muted and the percutaneous coronary intervention was completed without issue. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. according to clinical details, the doctor noted high purge pressure without observing any abnormalities in motor current. the cause of the high purge pressure issue was not determined since product was not returned and limited clinical information was provided.
31912 | 1220648-2024-05864 | 18456126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456126&pc=OZD | 12/15/2023 | 10/11/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024416021 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Vascular Dissection | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 56-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was attempting to have the impella 5.5 implanted and an aortic dissection occurred. physician accessed the right axillary artery and placed graft twice in an attempt to cross with a 0.035" wire. after being unable to cross with the 0.035" wire, the decision was made to access the left axillary artery. access was obtained and wire crossed with some difficulty. aortic dissection was revealed on tee (transesophageal echocardiography) with the wire being located in the false lumen of the aorta. the procedure was then aborted. echo-echocardiogram (ultrasound) guidance revealed no dissection of the ascending aorta. bilateral upper extremity pulses were present at the end procedure. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue that led to the aortic dissection was most likely use issue related. added-b1-selected product problem. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ corrections to the initial mfr report: d4-corrected model number. added-d6a/d6b. f6/f8 should have been left blank. h1-type of reportable event changed to malfunction. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31913 | 1220648-2024-05865 | 18456130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456130&pc=OZD | 12/21/2023 | 12/11/2024 |  | 01/05/2024 |  |  | 12/21/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403090 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported an impella cp in a patient (unknown age, race , and ethnicity) for mechanical circulatory support. it was reported that the team suspected clot ingestion and impella was explanted and new impella was placed. the patient is stable. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. a supplemental report will be submitted when the investigation has been completed. | The investigation into low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05865. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email. g1 reporting contact facility name missing. g1 reporting contact fax number missing. h6 code 4440 was reported incorrectly. new code was added to health effect - clinical code.
31914 | 1220648-2024-05866 | 18456135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456135&pc=OZD | 11/28/2023 | 05/14/2024 |  | 01/05/2024 |  |  | 12/06/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397850 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 46-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. a notification was received via abiomed¿s long-term outcome and quality indicator impella registry that indicated this patient to have endured bleeding bark type 3b, a vascular complication, and ventricular arrhythmia. each was alleged with a "definite" relationship to the impella device used. intervention to address these issues was performed. please see below for the reported narratives for the noted allegations: bleeding bark type 3b: difficulty accessing vascular site in the cath lab resulting in bleeding around the initial right femoral vein pacing sheath. the sheath was upsized and alternative access was achieved in the left femoral vein. there was also an occlusion of the right femoral artery that required a surgical cutdown of the right femoral artery with removal of venous and arterial sheaths after impella removal. post-procedure, hemoglobin, platelets, and white blood count were down-trending likely as a result of hemodilution with volume resuscitation and with bleeding from the procedure in the cath lab. vascular complication: extravasation and occlusion of the right femoral artery access site. vascular surgery was consulted and recommended surgical cutdown. cutdown was performed in the cath lab and the arterial and venous sheaths were removed. there was bleeding around the initial right femoral vein site with the pacing sheath. the sheath was upsized and an alternative location (left femoral vein) was utilized instead. ventricular arrhythmia: the patient developed recurrent ventricular tachycardia and the device kept jumping into the aorta. return of spontaneous circulation was achieved after the interventions for the ventricular rhythm. the patient was defibrillated many times. the impella device was subsequently removed. | [5 empty] | The investigation for the reported access site bleed, vascular damage, and ventricular arrhythmia has been completed. the impella device was not returned for evaluation. however, clinical details about the vascular damage and subsequent access site bleed were sufficient to determine the root cause. the root cause of the vascular damage was likely use-related as there was difficulty accessing vascular site in the cath lab resulting in bleeding around the initial right femoral vein pacing sheath. the sheath was upsized and alternative access was achieved in the left femoral vein. there was also an occlusion of the right femoral artery that required a surgical cutdown of the right femoral artery with removal of venous and arterial sheaths after impella removal. the root cause of the vt could not be determined without additional information. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31915 | 1220648-2024-05867 | 18456163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456163&pc=OZD | 12/15/2023 | 02/23/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024389877 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that impella cp smartassist heart pump was explanted due to leg ischemia. the patient was noted to have expired the next day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31916 | 1220648-2024-05868 | 18456234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456234&pc=OZD | 12/21/2023 | 11/30/2024 |  | 01/05/2024 |  |  | 12/21/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408232 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | Cardiac Perforation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported an impella 5.5 in a 60yo white female patient for mechanical circulatory support. it was reported that a left ventricle (lv) perforation occurred during impella support. leading up to the event, it was noted that the patient had been on impella 5.5 and bypass support. upon removal of bypass support, the lv perforation was discovered and subsequently repaired. an unknown amount of blood was found in the chest area and air had been noted on transesophageal echo (tee). support was able to continue with the implanted pump following the intervention. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. | The investigation into ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings and cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact facility name missing. g1 reporting contact fax number missing. h6 codes 2135 and 1888 were reported incorrectly. new code was added health effect - clinical code. h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use missing.
31917 | 1220648-2024-05869 | 18456303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456303&pc=OZD | 12/18/2023 | 01/17/2024 |  | 01/05/2024 |  |  | 12/18/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024414966 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the impella 5.5 would not pass through the ascending aorta due to the patient's anatomy. the placement of the impella 5.5 was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. the root cause of the delivery issues is most likely due to the patients condition as the pump would not deliver due to the anatomy.
31918 | 1220648-2024-05870 | 18456323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456323&pc=OZD | 12/21/2023 | 12/18/2024 |  | 01/05/2024 |  |  | 12/21/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404439 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 67-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed a retroperitoneal bleed and ischemia during impella cp support. the retroperitoneal bleed was believed to have originated from a previous procedure in the left femoral artery. although the blood count had stabilized, heparin had to be withdrawn and blood products were given to manage the condition. coinciding with the event, a decrease in perfusion to the right lower extremity was noted. the pump was explanted the following day as a result of the condition. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the retroperitoneal bleed and limb ischemia has been completed. the pump was not returned for investigation. a data log review was not required for this case. as per the clinical, a retroperitoneal (rp) bleed was observed at the site of previous access to the left femoral artery (lfa). additionally, there is concern over limb ischemia in the right leg, where the pump was inserted. clinical details were not sufficient to determine the root cause. therefore the root cause of the retroperitoneal bleed and limb ischemia could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: added d6a/d6b g1 mdr reporting contact email.
31919 | 1220648-2024-05871 | 18456324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456324&pc=OZD | 12/15/2023 | 12/18/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 06/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024363828 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella cp moved through the sheath without issue but had kinking at the puncture site as diameter had not been measured via ct. due to the stiffness of outlet and motor, it was not possible to move the impella forward. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the product damage has been completed. the pump was not returned for investigation. a data log review was not required for this case. as per the given clinical details, the pump damaged the right femoral artery, with reported stenosis and tortuous vessel conditions. additionally, a kink was found at the damaged site during procedure. the cause of the product damage was most likely use related since the kink was reported which is sign of applied excessive force. corrections to the initial mfr report: added d6a/d6b g1 mdr reporting contact email
31920 | 1220648-2024-05872 | 18456394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456394&pc=OZD | 12/18/2023 | 03/07/2024 |  | 01/05/2024 |  |  | 12/20/2023 | 01/05/2024 |  |  | 07/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024382403 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported a 56-year-old asian male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a cp support for 3 days for a patient admitted with an ami. the pump was weaned and explanted. at the time of explant the patient condition deteriorated and the diagnosis was made for a cerebral infarct. the causal relationship of the impella and neuro event can not be denied. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. the device was not returned for investigation. as the clinical information was not provided, the root cause of the stroke/cva could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31921 | 1220648-2024-05873 | 18456503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456503&pc=OZD | 12/16/2023 | 12/05/2024 |  | 01/05/2024 |  |  | 12/16/2023 | 01/05/2024 |  |  | 06/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374186 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2023 |  |  |  |  | The user facility reported an impella cp in a 63yo male for mechanical circulatory support. it was reported that the patient developed hemolysis during impella support. the catheter was repositioned and the p-level was decreased to troubleshoot the issue, but the condition remained. as a result, the decision was made to replace the pump with iabp support. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. the supplemental report will be submitted upon completion of the investigation. the instructions for use states the following regarding hemolysis: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ the failure mode will be monitored and trended. | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05873. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email. g1 reporting contact facility name missing. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
31922 | 1220648-2024-05874 | 18456576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456576&pc=OZD | 12/18/2023 | 04/19/2024 |  | 01/05/2024 |  |  | 12/18/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024411988 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support when the patient had a hematoma. the patient was having pain in their arm and when seating upright their arm would go numb and the hematoma "expanded". platelets were given to the patient. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31923 | 1220648-2024-05875 | 18456939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456939&pc=OZD | 12/19/2023 | 01/31/2024 |  | 01/05/2024 |  |  | 12/19/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024415438 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 57-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the decision was made to remove the impella 5.5 pump due to low pump flow. upon removal it was noted that on the outlet cage there was a clot. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the cut pump was sent back for investigation, and biomaterial was observed on the impeller and outflow cage. data logs show several pump stops and motor current spikes, followed by saturated flow. the cause of saturated pump flow issue was biomaterial, based on data log and returned product review.
31924 | 1220648-2024-05876 | 18456954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18456954&pc=OZD | 12/17/2023 | 06/05/2024 |  | 01/05/2024 |  |  | 12/17/2023 | 01/05/2024 |  |  | 11/30/2024 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313555 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/17/2023 |  |  |  |  | The user facility reported a 23-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp and an impella rp device for mechanical circulatory support. during support, the impella cp pump had low suction/flow issues. the team observed left ventricle thrombus burden to be present on the echo. the team pulled the cp back into the aorta. the patient expired when the mean arterial pressure decreased and patient became pulseless. the cause of death was determined to be viral cardiomyopathy. | Upon medical safety officer review, there is not enough information to associate the use of the device with patient's outcome. |  |  |  |  | The previously submitted report was missing this statement in regards to h.10. additional manufacturer narrative instructions for use were inadvertently omitted from the original manufacture device report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed since the original report was submitted. the device was not returned for investigation. the data logs showed multiple motor current spikes followed by low pump flow. as per clinical findings, a free-floating left ventricle clot was noted under echocardiogram during the low pump flow event. the cause of the low pump flow issue was biomaterial, based on data log review and provided clinical details. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.1 product problem was inadvertently left off the previously submitted report and was added. b.2 revised as previously entered incorrectly. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures.
31925 | 1220648-2024-05877 | 18457671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18457671&pc=OZD | 11/30/2023 | 12/23/2024 |  | 01/05/2024 |  |  | 12/18/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024397634 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | Material Twisted/Bent | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 37-year-old male was implanted with an impella cp device for left ventricle (lv) vent and coronary perfusion. it was reported that after the impella cp was started on p2, and it was difficult to keep in ideal position. after torquing impella to angle away from pap muscle the flows immediately dropped to zero. the impella cannula pulled back into aorta and appeared to have kinked. the physician tried advancing into a better position a couple more times with no success. the impella was removed and since there were no abnormal findings, the insertion of the same impella started all over again. the impella advanced and successfully inserted with no complications. there was no patient harm reported. | [5 empty] | The investigation for the positioning issue has been completed. pump was not returned. device was in ideal position and then became past av. no details on patient movement or patient condition at this time. the cause of the positioning issue was not determined due to lack of clinical details and no product returned. corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31926 | 1220648-2024-05878 | 18458774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18458774&pc=OZD | 12/18/2023 | 05/14/2024 |  | 01/05/2024 |  |  | 12/18/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411226 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support for 4 days to support a patient with multiple cardiac comorbidities. the cp support was successful however there was suction that was not breakable. the suction/low flows was treated by explant. post explant the patient was supported by medical therapy/levo. | [5 empty] | The investigation for the low pump flow was completed. the pump was not returned for investigation. data logs show low pump flow with an active suction alarm at the end of the case, without any abnormalities seen in the placement signal. motor current was reported as non-pulsatile during the low pump flow event. according to the clinical details, low pump flow persisted even after ensuring the pump was in a good position and sufficient blood volume was provided. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31927 | 1220648-2024-05879 | 18458809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18458809&pc=OZD | 12/21/2023 | 04/16/2024 |  | 01/05/2024 |  |  | 12/21/2023 | 01/05/2024 |  |  | 05/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368234 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 59 year-old male for mechanical circulatory support. it was reported that the pump experienced low flows following implant. it was noted that the initial flows were around 1.9 l/min but then dropped to .6 l/min. as a result of the persistent issue, the pump was replaced with a new impella cp device. upon explant, it was noted that there was a visible clot in the inlet of the device. there was no patient harm. no additional information was provided. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon completion of our analysis, a supplemental report will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. there was a big clot in the outflow cage. the data logs show a small spike in motor current and suction alarms. the root cause of the low pump flow is most likely due to biomaterial ingestion.
31928 | 1220648-2024-05882 | 18459009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18459009&pc=OZD | 12/19/2023 | 12/11/2024 |  | 01/05/2024 |  |  | 12/21/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391336 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 82-year-old asian female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a cp support along with ecmo (extracorporeal membrane oxygenation) for a patient with vf (ventricular fibrillation) in need of a pci (percutaneous coronary intervention) for cad (coronary artery disease) . the team explanted the pump after 6 days and had removal issues. a clot burden was observed in the lower extemity and the team had to have fogarty embolectomy to removed the embolized thrombus from the limb and a "decompression incision" was performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported removal issue and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus nor the removal issues could not be determined as the device was not returned nor sufficient clinical details. added h6 med dev prob code 1528, h6 impact code 4624 corrections to the initial mfr report: b1-product problem should have been selected. d4 model number. added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email. h5 should have been yes. h8 should have been initial use of device.
31929 | 1220648-2024-05881 | 18459011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18459011&pc=OZD | 12/22/2023 | 05/28/2024 |  | 01/05/2024 |  |  | 12/22/2023 | 01/05/2024 |  |  | 07/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385778 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 58-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed some oozing at their access site following impella implant. the dressing was initially changed to better angle match the site, but sutures were later added and femostop was placed to treat the condition. it was noted that high purge pressure was at point identified; however, standard troubleshooting steps were discussed, and no further action was required. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and the high purge pressure has been completed. no product was returned for a visual inspection. data log analysis confirmed that heparin had been added shortly after the purge system blocked alarms occurred. a sudden spike in motor current and purge pressure followed by a decrease in flows were seen. no attempts were made to resolve the issue. the most likely cause of the high purge pressure was biomaterial. according to the clinical, on the first day of impella cp support, bleeding was observed at the access site. in an attempt to resolve the issue sutures were added. earlier that morning the dressing had also been adjusted to match the angle of insertion. the most likely cause of the major access site bleeding was use related. the instructions for use referenced in the initial report is from ifu for impella cp with smart assist system sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes "cardiac or vascular injury (including ventricular perforation.¿ additionally, section: using the automated impella controller with the impella catheter states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- h6 med dev prob code 1491, h6 component code 925 d4-updated model number added- d6a.
31930 | 1220648-2024-05880 | 18459025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18459025&pc=OZD | 12/21/2023 | 02/21/2024 |  | 01/05/2024 |  |  | 12/21/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024422219 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 72-year-old white male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma during impella cp support. the pump was removed due to the hematoma which required a surgical cutdown. the patient was stable following the treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical was not provided to determine the cause of hematoma.
31931 | 1220648-2024-05883 | 18459715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18459715&pc=OZD | 11/12/2023 | 04/30/2024 |  | 01/05/2024 |  |  | 12/06/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398874 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Other | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/06/2023 |  |  |  |  | A 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that indicated this patient to have endured thrombocytopenia with a "possible" relationship to the impella device used in this case. the reported narrative indicates that the, "patient developed multifactorial thrombocytopenia likely consumptive." | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding this patient to have endured hemolysis with a "definite" relationship to the device. hospitalization was prolonged. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported hemolysis and thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis and thrombocytopenia could not be determined. the instructions for use for impella cp with smartassist system section: hemolysis states ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. it also states ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6, g3, and h10 have been revised from the initial submission. | The investigation is still in process. a supplemental mdr is to follow.
31932 | 1220648-2024-05884 | 18460559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460559&pc=OZD | 12/17/2023 | 07/25/2024 |  | 01/05/2024 |  |  | 12/17/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411146 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2023 |  |  |  |  | The user facility reported a 40-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. patient bleeding form impella access site and the site was closed by perclose. however, the perclose did not achieve hemostasis. the site had surgical intervention and a covered stent placed to the site. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to high activated clotting times. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2024-05884 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-05884 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-05884 was submitted. h.6 code 4114 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-05884. a new code has been added to type of investigation codes. added code 4641 as it was inadvertently omitted from the initial submission of manufacturer device report number 1220648-2024-05884. added code 925 since the initial submission of manufacturer device report number 1220648-2024-05884 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-05884 in accordance with updated procedures. additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2024-05884 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
31933 | 1220648-2024-05885 | 18460601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460601&pc=OZD | 12/18/2023 | 07/09/2024 |  | 01/05/2024 |  |  | 12/18/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420789 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 77-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had tortuous arteries and when impella was inserting the catheter kinked behind the motor housing unit. the impella was unable to advance into position. a second impella was placed with difficulty. | It was further reported that the patient expired in the procedural area. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci in the section: warnings & cautions: warnings states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. added- b5 mso review, h6 impact code 4629 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
31934 | 1220648-2024-05886 | 18460633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460633&pc=OZD | 12/18/2023 | 04/12/2024 |  | 01/05/2024 |  |  | 12/18/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411865 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported an impella cp in a 65yo male for mechanical circulatory support. it was reported that the patient had uncontrolled bleeding at the access site. vascular surgery was performed to remove the impella cp. during the removal, it was found that the superficial femoral artery was the point of entry for the impella. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. a supplemental report will be submitted when the investigation has been completed. the instructions for use states the following regarding access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure mode will be monitored and trended. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. per clinical details provided, the root cause of the access site bleeding was most likely due to anticoagulation management since the act at the time of bleed was 298 which above the limit range.
31935 | 1220648-2024-05887 | 18460685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460685&pc=OZD | 12/16/2023 | 01/05/2024 |  | 01/05/2024 |  |  | 12/19/2023 | 01/05/2024 |  |  | 12/31/2024 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301964 |  | 0048-0371-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a unknown age male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant in japan had an inability to deliver the cp pump due to native anatomy. the iliac was known to be calcified. the team aborted the cp pump and instead choice was made to implant an iabp. | [5 empty] | The impella device was not received from the customer. the root cause of the delivery issue was most likely due to the patient condition.
31936 | 1220648-2024-05888 | 18460834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460834&pc=OZD | 12/13/2023 | 06/19/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414459 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 47-year-old male patient was on impella cp support and had access site bleeding and thrombus was also found. the access site bleeding was resolved by blood products and manual pressure. the thrombus was found with a ct scan while on impella support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus and access site bleed could not be determined as the device was not returned nor sufficient clinical details. the instructions for use referenced in the initial report is for the impella cp with smartassist system. in section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation)¿, section:entering cardiac output, and section: general patient care considerations. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
31937 | 1220648-2024-05712 | 18460873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460873&pc=OZD | 12/16/2023 | 12/23/2024 |  | 01/05/2024 |  |  | 12/16/2023 | 01/05/2024 |  |  | 09/30/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024403179 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Optical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2023 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp was removed due to placement signal low alarm and placement signal unreliable alarm. no patient harm was reported. the patient remained stable on extracorporeal membrane oxygenation (ecmo) support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the optical signal issue has been completed. the product was not returned for investigation. the data log shows the placement signal drift with a placement signal not reliable (psnr) alarm after the 9th day of pump use. the design life of the cp pump is 8 days. the cause of the optical signal issue was sensor silicone wear beyond the product design life. corrections to the initial mfr report: f6/f8 should have been left blank. h6 med dev prob code 3001 changed to 2917.
31938 | 1220648-2024-05711 | 18460874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460874&pc=OZD | 12/14/2023 | 06/20/2024 |  | 01/05/2024 |  |  | 12/14/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418697 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Malfunction | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 77-year-old male with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. the attempted impella 5.5 insertion was unsuccessful. the impella 5.5 was able to be advanced through the sheath without issues but impella 5.5 was unable to be advanced through the axillary artery to the aorta into the left ventricle due to the patient¿s anatomy. multiple attempts were made to try to advance the impella, but they were unsuccessful. the patient required three units of packed red blood cells. the patient was to be expedited for left ventricle assist device. | [5 empty] | The investigation into the delivery issues and the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to advance through axillary artery to the aorta into the left ventricle as per the clinical description. the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description. d6a, d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. h(10) was inadvertently left blank on previous manufacture device report and was added.
31939 | 1220648-2024-05713 | 18460876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460876&pc=OZD | 12/15/2023 | 03/27/2024 |  | 01/05/2024 |  |  | 12/15/2023 | 01/05/2024 |  |  | 10/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024418750 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when the impella was turned on the motor current spiked and the impella stopped. the impella was removed and another impella was successfully placed. there was no patient harm. | [5 empty] | The investigation into the pump stop issue has been completed since the initial report was submitted. the product was not returned for investigation. data logs show that the pump stopped with high motor current during the start of the case. the cause of the pump did not start issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31940 | 1220648-2024-05889 | 18460926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460926&pc=OZD | 12/19/2023 | 10/14/2024 |  | 01/05/2024 |  |  | 12/19/2023 | 01/05/2024 |  |  | 08/31/2025 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397975 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a (b)(6) male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the cp was inserted but quickly alarmed for low flows. the team exchanged the pump for a new one and support proceeded. no patient harm was reported due to the issue. | [5 empty] | The investigation for the low pump flow has been completed. pump was not returned for investigation. data logs were returned and there was a motor current spike observed followed by low flows. this observed symptoms are of a biomaterial suction which was possibly washed out before removal. therefore, per clinical details and data log review, the root cause of the low pump flow issue was most likely biomaterial. corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31941 | 1220648-2024-05890 | 18460937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18460937&pc=OZD | 12/20/2023 | 04/18/2024 |  | 01/05/2024 |  |  | 12/20/2023 | 01/05/2024 |  |  | 10/31/2024 |  | 01/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292778 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient on impella support had a drop in hemoglobin. the impella was removed due to bleeding from the access site. patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
31942 | 1220648-2024-05769 | 18442411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18442411&pc=OZD | 12/05/2023 | 03/12/2024 |  | 01/04/2024 |  |  | 12/05/2023 | 01/04/2024 |  |  |  |  | 01/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1353540 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller (aic) had a system self-check failure. a replacement loaner was sent to the facility. no patient was involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - system self-check error issue has been completed since the original report was filed. the console was tested and physically inspected after being returned to field service. the root cause of the aic issue was contamination.
31943 | 1220648-2024-05770 | 18443147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18443147&pc=OZD | 12/18/2023 | 01/31/2024 |  | 01/04/2024 |  |  | 12/18/2023 | 01/04/2024 |  |  |  |  | 01/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1296680 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was being prepped for impella support when the automated impella controller (aic) showed a system-check failed upon start-up. the aic was replaced. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - system self-check error has been completed since the original report was filed. the device was returned for investigation. imc logs confirm that after the last successful system self-check. the console was unable to pass system self-check upon boot up in the lab. visual inspection of the pbm revealed corrosion of the copper traces in proximity of the super capacitors c69 and c70. the cause of the system self-check failure was loss of communication between the pbm and epc. the root cause of the loss of communication was contamination of the pbm resulting from corrosion of the copper traces from leakage of supercapacitor electrolyte.
31944 | 1220648-2024-05771 | 18443349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18443349&pc=OZD | 12/07/2023 | 01/04/2024 |  | 01/04/2024 |  |  | 12/07/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024403456 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Separation Failure | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the sterile sleeve was locked to the catheter from packaging. the sterile sheath would not move back on the catheter. troubleshooting was initiated but unsuccessful. unable to get pump in optimal position with sleeve/repo sheath forward. a new pump was placed. no patient harm. | [5 empty] | Impella cp pump was returned. the returned product was visually inspected and found the sterile sleeve to be locked too tight without slack to tuohy borst. the tuohy borst lock was loosened, and the sleeve was pulled outside to normal and readjusted. the sterile sleeve was now able to move in back-and-forth motion without any issues. also, the repositioning sheath tip was found to damaged. no other abnormalities were observed. as per clinical pump sterile sleeve sheath would not move back and locked onto catheter with no success on tuohy borst lock moving both ways. the cause of the sterile sleeve damage issue was use related per clinical review and field investigation.
31945 | 1220648-2024-05772 | 18443500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18443500&pc=OZD | 12/11/2023 | 05/14/2024 |  | 01/04/2024 |  |  | 12/11/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024404050 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 70-year-old female (unknown race and ethnicity) patient received hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). during preparation, the user accidentally added a physiological saline solution as a purge solution instead of the prescribed glucose solution. sodium chloride can cause corrosion in the impella pump, which is why warnings are included in the ifu. the error was discovered and corrected within minutes. the impella cp heart pump was only used for two hours during the pci procedure, which was completed successfully. isolated diastolic aspiration of the impella cp and slightly restricted mean pump flow were observed several times, with the patient's blood volume status adequate. | [5 empty] | The investigation has been completed since the original report was submitted. the device was not returned by the user facility for investigation. based on the clinical details, the root cause of the priming issue is due to use as the nurse used the incorrect purge solution. d.4 revised expiration date as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures.g.1 revised manufacturing site fax number as previously entered incorrectly. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
31946 | 1220648-2024-05773 | 18443624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18443624&pc=OZD | 12/09/2023 | 01/18/2024 |  | 01/04/2024 |  |  | 12/09/2023 | 01/04/2024 |  |  | 02/28/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024341020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2023 |  |  |  |  | The user facility reported that an impella cp was attempted to be implanted in a 70-year-old male of unknown race for mechanical circulatory support but was unsuccessful due to the patient's anatomy. the implant was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. the root cause of the delivery issue was most likely patient condition related.
31947 | 1220648-2024-05774 | 18443639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18443639&pc=OZD | 12/10/2023 | 07/15/2024 |  | 01/04/2024 |  |  | 12/10/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415437 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported that due to the patient's anatomy, resistance was met while attempting to place an impella 5.5 pump for patient support. while removing the guidewire, it was noted that the pump pulled back into the ascending aorta. the pump was removed and it was discovered that the cannula and catheter were severely bent. a new pump was subsequently placed for planned support. there was no patient harm. | [5 empty] | The impella device has not been returned for evaluation. however, investigation is in progress. a supplemental report will be submitted upon completion of the investigation. the failure modes are being monitored and trended. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
31948 | 1220648-2024-05775 | 18443668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18443668&pc=OZD | 12/11/2023 | 01/17/2024 |  | 01/04/2024 |  |  | 12/11/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024416510 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 19 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 19-year-old male (unknown race and ethnicity) undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported it was impossible to advance the impella 5.5 into the aortic arch due to abnormal vasculature. a femoral approach was also tried and was unsuccessful. patient was returned to the intensive care unit and expired shortly after. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related.
31949 | 1220648-2024-05776 | 18443914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18443914&pc=OZD | 12/13/2023 | 09/20/2024 |  | 01/04/2024 |  |  | 12/13/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409592 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/13/2023 |  |  |  |  | Additional information was provided that noted the patient was placed on palliative care and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 60-year-old white female with cardiomyopathy was implanted with an impella for mechanical circulatory support. the patient impella 5.5 delivery fail. the patient's anatomy precluded placement. the 5.5 placement failed but, the team was able to deliver the cp pump instead. no patient harm was reported due to the issue. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-05776 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition.
31950 | 1220648-2024-05777 | 18444029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444029&pc=OZD | 12/13/2023 | 12/04/2024 |  | 01/04/2024 |  |  | 12/13/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409481 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant placed an impella 5.5 to support a 45-year-old white male patient admitted with cardiomyopathy. the plan was to support the patient while awaiting either lvad or heart transplant. the patient had labs indicative of hemolysis. the team drew pfhg and ldh. the patient was also placed on dialysis. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis has been completed. product was not returned. data logs show a few motor current (mc) spikes on the first day of support. there are also suction alarms and positioning alarms throughout the case. the root cause of the hemolysis could not be determined since product was not returned and lack of clinical details. the instructions for the use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added h6 component code 925 corrections to the initial mfr report: added d6a/d6b.
31951 | 1220648-2024-05778 | 18444080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444080&pc=OZD | 11/27/2023 | 12/04/2024 |  | 01/04/2024 |  |  | 12/12/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408903 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolytic Anemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | Stemi dtu study reported a 63-yer-old female with st elevated myocardial infarction was implanted with an impella cp for circulatory support. the study reported there was a possible relationship between the impella cp and the patient having anemia. the anemia started post procedure. iv ferrlecit was given to the patient twice during admission. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use as follows impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05778. b3 date of event. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g2 report source; study missing. h6 component code revised from the initial submission in accordance with updated procedures. h10 instruction for use missing.
31952 | 1220648-2024-05779 | 18444089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444089&pc=OZD | 12/14/2023 | 06/26/2024 |  | 01/04/2024 |  |  | 12/14/2023 | 01/04/2024 |  |  | 08/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024394657 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 62-year- white male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed hemolysis following impella cp implant. repositioning was initially completed to manage the condition; however, the decision was ultimately made to explant the pump later that day. | Additional information: the patient expired on impella cp support. abiomed medical safety office review determined there is not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis. | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. data log analysis found the presence of suction and positioning issues. the most likely cause of the hemolysis was improper positioning. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. removed f6 and f8. it was inadvertently added in the initial mdr.
31953 | 1220648-2024-05780 | 18444126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444126&pc=OZD | 12/15/2023 | 09/25/2024 |  | 01/04/2024 |  |  | 12/15/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024400578 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 70-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support for elevated central venous pressure with low cardiac output and cardiac index. while on impella support, the patient had access site bleeding and a run of vt (ventricular tachycardia) . heparin was held for the access site bleeding. when the patient had a run of vt, amio was given and electrolytes were replaced, and heparin was increased. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and ventricular tachycardia (vt) has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values, heparin was stopped to achieve hemostasis as per the clinical description. without additional clinical details, the root cause of the vt could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ added- h6 component code 925 corrections to the initial mdr mfr report: d4-corrected model number added- d6a/d6b f6/f8 should have been left blank
31954 | 1220648-2024-05781 | 18444140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444140&pc=OZD | 12/12/2023 | 12/03/2024 |  | 01/04/2024 |  |  | 12/12/2023 | 01/04/2024 |  |  | 06/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024371277 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/12/2023 |  |  |  |  | The complainant reported a (b)(6) female patient was implanted with impella cp with smartassist heart pump for hemodynamic support. however, the purge connector was damaged and a leak was noted on the purge reservoir. the device was explanted three days after implant and there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak has been completed. no product was returned. clinical details were not sufficient to determine root cause. the root cause of purge leak pump was not determined since product was not returned to confirm. | The investigation for the purge leak pump has been completed. no product returned. the root cause of purge leak pump was not determined since product was not returned to confirm. corrections to the initial mfr report: d4-model number; added d6a/d6b; f6/f8 should have been left blank.
31955 | 1220648-2024-05782 | 18444238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444238&pc=OZD | 12/11/2023 | 02/19/2024 |  | 01/04/2024 |  |  | 12/11/2023 | 01/04/2024 |  |  | 08/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385444 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 61-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during a scheduled pci, the motor current spiked and the pump subsequently stopped working. attempts to restart the pump were unsuccessful and the decision was made to explant the pump and reschedule the pci. there was no patient harm reported due to the issue. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the product was returned for investigation. no physical abnormalities were observed on the pump. saturated pump flow was observed during a test run in a bucket of water. upon further investigation, biomaterial was observed in purge gap and motor housing. the cause of the pump stop issue was biomaterial found in purge gap and motor housing. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31956 | 1220648-2024-05783 | 18444382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444382&pc=OZD | 12/11/2023 | 02/12/2024 |  | 01/04/2024 |  |  | 12/11/2023 | 01/04/2024 |  |  | 05/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358643 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 35-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the pump was inadvertently pulled out of the ventricle during patient support. the physician was unable to re-cross the valve and a new pump was subsequently placed for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the placement signal issue is due to use due to physician accidentally pulling pump into the aorta.
31957 | 1220648-2024-05784 | 18444674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444674&pc=OZD | 12/12/2023 | 10/31/2024 |  | 01/04/2024 |  |  | 12/12/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024405172 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 51-year-old male was implanted with an impella 5.5 as a bridge to transplant. during the heart transplant a dissection of the axillary artery was discovered following the impella explant, a patch was placed and hemostasis achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined.
31958 | 1220648-2024-05785 | 18444758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444758&pc=OZD | 12/12/2023 | 05/08/2024 |  | 01/04/2024 |  |  | 12/12/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported an impella cp in a 55yo female patient for mechanical circulatory support. it was reported that pump was unable to achieve flows above 1.3 l/min due to visible kink in the cannula. the pump was replaced for ongoing support. there was no patient harm. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. the failure mode will be monitored and trended. | The investigation for the low pump flow has been completed. the pump was not returned for investigation. the total case runtime was 15 minutes. data logs show low pump flow throughout the case without reported motor current spikes or positioning issues. the cause of the low pump flow issue was most likely kinked cannula, based on given clinical details. in accordance with updated procedures, section d6 has been revised from the initial submission.
31959 | 1220648-2024-05786 | 18444791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444791&pc=OZD | 12/12/2023 | 03/25/2024 |  | 01/04/2024 |  |  | 12/12/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403860 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding from the anastomosis site during impella 5.5 support. the patient was initially given two units of packed red blood cells to treat the bleed and a washout was performed. the following day, the access site had significant swelling and another unit of blood was given to the patient. support continued with the implanted pump throughout the event. patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
31960 | 1220648-2024-05788 | 18444974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444974&pc=OZD | 12/14/2023 | 02/07/2024 |  | 01/04/2024 |  |  | 12/14/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408514 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The us complainant had placed a cp to support a patient admitted for a mitraclip. the pump had high purge pressures on the 2nd day of support. the team troubleshot by tpa infusion in the purge, however the pump still stopped. the pump was explanted. the pump was not replaced. no patient harm was reported due to this issue. | [5 empty] | The investigation for the reported high purge pressure and pump stop has been completed. the impella and data logs were returned for evaluation. according to the clinical review, there were no kinks in the purge line and swapping the purge cassette had no effect on the high purge pressure. the returned pump was visually inspected and confirmed that the catheter was cut so the high purge pressure/pump stop could not be reproduced. no kinks were observed in the catheter and no purge leaks were found in the purge sidearm. however, blood was observed ingressed into the purge line. logs showed a gradual increase in purge pressure followed by a gradual increase in motor current which eventually triggered high motor current pump stop. the cause of the pump stop was blood ingress into the purge line. the cause of the blood ingress is unknown. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
31961 | 1220648-2024-05787 | 18444979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18444979&pc=OZD | 12/13/2023 | 11/21/2025 |  | 01/04/2024 |  |  | 12/13/2023 | 01/04/2024 |  |  | 10/31/2024 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2024284909 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported an impella cp was be initiated in a 72-year-old female patient for mechanical circulatory support. it was reported that the pump was prepped and ready to be put in the body. the staff states that the pump was not detected by the console and was giving them some odd alarms. after troubleshooting, the decision was made to open another pump per the physician. there was no patient harm reported. no additional information was provided. | [5 empty] | The impella device was returned for evaluation but has not yet been evaluated. upon completion of the investigation, a supplemental report will be submitted.
31962 | 1220648-2024-05789 | 18445158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445158&pc=OZD | 12/13/2023 | 01/04/2024 |  | 01/04/2024 |  |  | 12/13/2023 | 01/04/2024 |  |  |  |  | 01/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1111145 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Electrical Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) was not able to re-charge. the power cable has been plugged to multiple different plugs unsuccessfully. therefore, the console had to be exchanged during the procedure. there was no reported patient harm. | [5 empty] | (h3) the investigation into the reported "aic - a/c power issues" has been completed since the original report was filed. product was returned for investigation. upon inspecting the console, two fuses were blown and the power supply was blackened and the inability to charge was reproduced. however, replacing the power supply, the console was able to charge again. during data analysis, on 12/13, the console (ic1821) was booted up at the start of the case immediately getting an event 109 for ac power disconnected. the console was power cycled but ac power still showed as disconnected. multiple more attempts to power cycle across the next day yielded the same results. the root cause of the aic's inability to charge with ac power was a burnt power supply and blown fuses.
31963 | 1220648-2024-05791 | 18445208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445208&pc=OZD | 12/13/2023 | 06/21/2024 |  | 01/04/2024 |  |  | 12/13/2023 | 01/04/2024 |  |  | 06/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 48-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient¿s labs were hemolyzed. the pump was repositioned, and the p level was lowered. the labs stopped hemolyzing but at that point the patient was not making urine output so the patient was placed on dialysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the original report was submitted. according to the clinical description, on the second day of impella cp support, the patient's urine turned a pinkish/red color and their labs were hemolyzed. an attempt was made to reposition the pump successfully, but the patient was no longer producing urine so they were placed on dialysis. no product was returned for a visual inspection. the data log analysis confirmed suction alarms and no other abnormalities. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4999 was reported incorrectly on the previously submitted report. code 4628 was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
31964 | 1220648-2024-05790 | 18445236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445236&pc=OZD | 12/09/2023 | 01/26/2024 |  | 01/04/2024 |  |  | 12/09/2023 | 01/04/2024 |  |  |  |  | 01/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1416992 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/09/2023 |  |  |  |  | The user facility reported a 59 year-old patient was on impella 5.5 support when the automated impella controller (aic) had an issue detecting the purge disc during a routine purge bag/cassette change. the aic screen was stuck on the purge disc screen. multiple attempts were made to snap in the purge disc to no avail. a second, new purge cassette was attempted and still unsuccessful. ultimately, the nurse had to transfer pump to a new aic to successfully change out purge cassette. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation. after reviewing the imc logs from the reported occurrence date, the issues with the aic not being able to detect the purge disc was confirmed. there were multiple instances of purge disc not detected alarms on the reported occurrence date. the console was tested, and the purge disc not detected issue was able to be reproduced and the purge flag was found to be broken. the root cause of this issue was a broken purge flag.
31965 | 1220648-2024-05792 | 18445563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445563&pc=OZD | 12/15/2023 | 01/17/2024 |  | 01/04/2024 |  |  | 12/15/2023 | 01/04/2024 |  |  | 08/31/2024 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270840 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported an unknown age and gender patient noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the pump would not deliver across the aortic valve. the team explanted the pump and a new pump was successfully placed. no harm or injury from the failed delivery. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31966 | 1220648-2024-05793 | 18445614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445614&pc=OZD | 12/15/2023 | 02/02/2024 |  | 01/04/2024 |  |  | 12/15/2023 | 01/04/2024 |  |  |  |  | 01/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143527 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) had a broken purge clip. a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31967 | 1220648-2024-05794 | 18445654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445654&pc=OZD | 12/16/2023 | 05/14/2024 |  | 01/04/2024 |  |  | 12/16/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024405706 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2023 |  |  |  |  | The user facility reported an impella cp in a 67yo male for mechanical circulatory support. it was reported that during a difficult delivery of an impella cp pump through the patient's tortuous iliac vessel, the device triggered an impella blocked alarm. the pump was subsequently removed and a new impella cp pump was implanted through the left femoral artery. there was no patient harm. no additional information was provided. | [5 empty] | The investigation for the low pump flow has been completed. the pump was not returned for investigation. data logs were reviewed which showed the purge system blocked after a 13-minute gradual rise in purge pressure was observed. purge cassette change was unsuccessful in resolving the purge system block alarm. the pump was replaced with a new one. therefore, the cause of the high purge pressure issue was most likely biomaterial, based on data log review. the instructions for use for the impella cp with smartassist for use during cardiogenic shock in section: troubleshooting the purge system states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not returned for evaluation. a supplemental report will be submitted when the investigation has been completed. the failure mode will be monitored and trended.
31968 | 1220648-2024-05795 | 18445688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445688&pc=PYX | 12/13/2023 | 09/22/2024 |  | 01/04/2024 |  |  | 12/13/2023 | 01/04/2024 |  |  | 10/11/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024413709 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant had placed an impella rp for support of the 65-year-old male patient admitted with thyroid storm and all 4 limbs suffering from deep vein thrombosis (dvts). the pump was experiencing pp/pf issues. the team troubleshot by tissue plasminogen activator (tpa) infusion. the pump remained on for support for 4 days until explanted when patient expired. | Medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2024-05795 was provided in sections b1, b5, d6, h1, and h6. | The investigation into the high purge pressure issue has been completed. the device was not returned for investigation. according to clinical information, the patient has deep vein thrombosis condition in all extremities. the cause of the rise in purge pressure/drop in purge flow was not determined since product was not returned and there was no impact to patient support throughout the case per clinical and log review.
31969 | 1220648-2024-05796 | 18445861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445861&pc=OZD | 12/14/2023 | 06/06/2024 |  | 01/04/2024 |  |  | 12/14/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407344 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 73-year-old female with heart failure. it was reported that the impella cp was inserted and then removed due to dislodgement. no harm reported. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. a supplemental report will be submitted upon completion of the investigation. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. added-d6a/d6b.
31970 | 1220648-2024-05797 | 18445879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18445879&pc=OZD | 12/07/2023 | 05/07/2024 |  | 01/04/2024 |  |  | 12/07/2023 | 01/04/2024 |  |  | 04/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024351062 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 79-year-old female patient was implanted with an impella device for mechanical circulatory support. during support, the patient eventually developed oozing at their site and pulses to their right lower extremity were decreased. the patient received a unit of blood and the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and limb ischemia could not be determined. section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. d4-updated model number.
31971 | 1220648-2024-05798 | 18446018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18446018&pc=OZD | 12/07/2023 | 02/23/2024 |  | 01/04/2024 |  |  | 12/07/2023 | 01/04/2024 |  |  | 06/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024379333 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | FEM-FEM BYPASS | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 69-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during the impella cp support an external fem-fem bypass was placed due to the right leg not perfusing and a lack of pulse. the impella cp support continued after the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
31972 | 1220648-2024-05799 | 18446114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18446114&pc=OZD | 12/18/2023 | 10/14/2024 |  | 01/04/2024 |  |  | 12/18/2023 | 01/04/2024 |  |  | 08/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024382962 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported an 80-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that as the impella 5.5 entered the ascending aorta, the surgeon had a lot of difficulty further advancing the impella. it was discovered that the tip of the impella went through a false lumen in the inner wall of the aorta and caused a small dissection. the impella was removed temporarily and an ascending aorta repair was done successfully. the impella was then placed back in without a wire into the left ventricle (lv) successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation for the vascular damage has been completed. pump was not returned for investigation. clinical details do not mention important information such as vessel size, imaging done prior, excessive force used, etc. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
31973 | 1220648-2024-05800 | 18446139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18446139&pc=OZD | 12/08/2023 | 01/26/2024 |  | 01/04/2024 |  |  | 12/08/2023 | 01/04/2024 |  |  |  |  | 01/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1363394 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The complainant reported that the aic (automated impella controller) had a broken purge disc holder. the aic was removed from service as a result of the damage. there was no patient harm. | [5 empty] | The device was returned and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation. console was tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge retainer was broken. the root cause of this issue was a broken purge retainer.
31974 | 1220648-2024-05801 | 18446335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18446335&pc=OZD | 12/20/2023 | 05/06/2024 |  | 01/04/2024 |  |  | 12/20/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024422052 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and had a growing hematoma. heparin was stopped and a total of four units of packed red blood cells were provided. the staff decided the best course would be to do high risk percutaneous coronary intervention (hrpci) instead of coronary artery bypass graft (cabg), and then take the patient to operating room for removal of the impella, evacuation of hematoma, and repair of the artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided. d6a, d6b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
31975 | 1220648-2024-05802 | 18446401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18446401&pc=OZD | 12/07/2023 | 05/29/2024 |  | 01/04/2024 |  |  | 12/08/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024412340 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The complainant reported an 82-year-old female patient with left main (lm) stenosis was placed on impella cp support as patient was not eligible for surgery. following implant, the patient was hemodynamically stable but not responsive. the impella was therefore removed to avoid complications. after removal, patient became hemodynamically unstable. the lm intervention was completed without impella support. patient ultimately expired. | The patient expired. upon abiomed's medical safety officer review, there was not enough information to associate the use of the device with the patient's outcome. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the vt/vf was not determined. b.3 revised as previously entered incorrectly. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report.
31976 | 1220648-2024-05803 | 18446516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18446516&pc=OZD | 12/21/2023 | 12/04/2024 |  | 01/04/2024 |  |  | 12/21/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 64-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support when the patient had access site bleeding. the staff used perclosure sutures which temporarily resolved the issue. the dressing was changed and the ultimate decision was made to remove the impella. the patient is stable post explant. | It was further reported that a clinical study found that the patient also experienced hemolysis prior to the pump being explanted. the study confirmed that the impella is related to the patient's hemolysis. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis and access site bleed has been completed. product not returned. data logs show no motor current (mc) spikes outside of p-level changes and a quick suction event that is resolved. the cause of the hemolysis was not determined due to no product returned, lack of clinical details, and inconclusive data logs. the cause of the access site bleeding was not determined since product was not returned and lack of clinical details. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 corrections to the initial mfr report: b5-additional failure mode added d6a/d6b f6/f8 should have been left blank. h6 clinical code 1886.
31977 | 1220648-2024-05804 | 18446534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18446534&pc=OZD | 12/15/2023 | 01/17/2024 |  | 01/04/2024 |  |  | 12/15/2023 | 01/04/2024 |  |  | 05/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024345238 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had an inability to deliver a 5.5 impella pump due to anatomy. the pump was removed and instead a cp was placed successfully. no lasting harm or injury | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition.
31978 | 1220648-2024-05805 | 18446890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18446890&pc=OZD | 12/19/2023 | 02/15/2024 |  | 01/04/2024 |  |  | 12/19/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024421171 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 54-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp delivered for a stemi (st elevation myocardial infarction) patient. the pump kinked at the cannula and prolapsed and was removed after 37 minutes of support. there was no noted harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into product damage (cannula kink) has been completed. the device was not returned for investigation. physician noted that kink was caused by pump pushed forward upon peel away sheath removal and repo placement. the root cause of the cannula kink was most likely due to the use when pump was pushed forward into the left atrium.
31979 | 1220648-2024-05806 | 18447438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18447438&pc=OZD | 12/12/2023 | 02/16/2024 |  | 01/04/2024 |  |  | 12/12/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024401639 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 69-year old female with diastolic heart failure was admitted to receive heartmate3. as part of the treatment the extracorporeal membrane oxygenation (ecmo) was initiated and was determined to add impella 5.5 for support to unload the left ventricle (lv) prior to heartmate 3 placement. the impella 5.5 pump had optical signal issues that the team could not remedy with troubleshooting. the determination was made to replace the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the original report was filed. the device was returned for investigation. visually inspected and no abnormalities were found. the cause of the optical signal failure was an air gap from between the aic and the patient cable plug per data log review and field investigation.
31980 | 1220648-2024-05807 | 18447505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18447505&pc=OZD | 12/12/2023 | 01/17/2024 |  | 01/04/2024 |  |  | 12/12/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024411496 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that multiple unsuccessful attempts were made to pass the impella 5.5 through the 10mm graft. the graft was raptured in the attempt to pass 5.5. the graft was repaired and 5.5 case was aborted. impella cp was successfully inserted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely due to patient condition.
31981 | 1220648-2024-05808 | 18447526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18447526&pc=OZD | 12/12/2023 | 06/04/2024 |  | 01/04/2024 |  |  | 12/12/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401042 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported an impella 5.5 in a 74yo female patient for mechanical circulatory support. it was reported that a hematoma was observed on the impella access site. three units total of packed red blood cells (prbc) were given and the impella was surgically removed. the patient is stable. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed was unable to be determined. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 revised as previously entered incorrectly. g.1 reporting contact facility name was inadvertently left off the previously submitted report and was added. h.3 was the device evaluated by the manufacturer? was inadvertently omitted from the previously submitted report. h.6 code 1884 was reported incorrectly on the previously submitted report. a new code has been added to health effect - clinical code. added codes 4627, 925 and 4114 as they were inadvertently left off the previously submitted manufacture device report.
31982 | 1220648-2024-05809 | 18447941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18447941&pc=OZD | 12/16/2023 | 09/20/2024 |  | 01/04/2024 |  |  | 12/18/2023 | 01/04/2024 |  |  | 02/28/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335124 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Separation Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported an impella cp in a 87yo male for mechanical circulatory support. it was reported that when removing the dilator from the 14fr peel-away introducer, the dilator broke free from the clear plastic ring that screws onto the valve of the peel-away sheath. this malfunction went unnoticed until it came time to exchange the peel-away for the repositioning unit. the team could not remove the ring and had to leave it in place between the blue suture hub and the arteriotomy. there was no patient harm reported. no additional information was provided. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella introducer has not been returned for evaluation. a supplemental report will be submitted when the investigation has been completed. the instructions for use states the following: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ it also states: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. note: corrected information provided in h6 is an addition to previous coding.
31983 | 1220648-2024-05810 | 18448707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18448707&pc=OZD | 12/21/2023 | 02/22/2024 |  | 01/04/2024 |  |  | 12/21/2023 | 01/04/2024 |  |  |  |  | 01/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1197227 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/21/2023 |  |  |  |  | The user facility reported a 59-year-old male was implanted with an impella device for mechanical support. it was reported that the automated impella controller (aic) failed to recognize an installed purge disc. the cassette was changed in an attempt to troubleshoot the issue, but the failure remained. the aic was subsequently swapped to resolve the noted problem. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in field services. after analyzing, the aic's inability to detect the purge disc was confirmed to be caused by a broken purge flag. the cause of the issue was a broken purge flag. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
31984 | 1220648-2024-05811 | 18448827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18448827&pc=OZD | 12/07/2023 | 11/01/2024 |  | 01/04/2024 |  |  | 12/13/2023 | 01/04/2024 |  |  | 02/28/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325371 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a unknown age patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that a blood clot was attached to the pigtail when the impella cp was being removed. the impella was surgically removed so there was no dissemination of the thrombus, and no additional treatment was required. the pump was removed as cardiac function had recovered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the thrombus, removal issues, and infection has been completed. no product returned for investigation. clinical details were not sufficient to determine the root cause. therefore the root cause of the removal issues, thrombus, and infection could not be determined.
31985 | 1220648-2024-05812 | 18449418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18449418&pc=OZD | 12/19/2023 | 09/26/2024 |  | 01/04/2024 |  |  | 12/19/2023 | 01/04/2024 |  |  | 09/30/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024403921 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 41-year-old black male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that there was a high purge pressure alarm. the cassette was changed with no resolution. heparin bolus was given. low purge flow continued followed with intermittent purge system blocked alarm. tpa protocol was utilized due to the low purge flow. purge flow resolved within 3 hours. there was no patient harm. the patient is stable. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported high purge pressure has been completed since the original report was filed. no product was returned for evaluation. the data logs were reviewed and it was found a gradual increase in purge pressure over a 2 minute period. both turning p-level down and back up and cassette change did not resolve high purge pressure alarm. tpa protocol did resolve issue. the root cause of the high purge pressure is likely due to buildup of biomaterial as it was resolved with the tpa protocol. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-05812 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
31986 | 1220648-2024-05813 | 18449447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18449447&pc=OZD | 12/19/2023 | 10/08/2024 |  | 01/04/2024 |  |  | 12/19/2023 | 01/04/2024 |  |  | 10/31/2025 |  | 01/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422176 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Device Handling Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 49-year-old male (unknown race) undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella cp support due to the patient vt (ventricular tachycardia) arresting. the impella was placed and the patient vt arrested twice following incursion. return of spontaneous circulation was obtained resulting in device malposition and lower than expected flows were achieved. there was a suspicion of the catheter kinking. the impella was removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." correction to the initial mfr report: d4-corrected model number. added- d6a/d6b. f6/f8 should have been left blank. g1-corrected mfr site name and address.
31987 | 1220648-2024-05709 | 18434371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18434371&pc=OZD | 12/11/2023 | 01/18/2024 |  | 01/03/2024 |  |  | 12/11/2023 | 01/03/2024 |  |  | 06/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357709 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 53 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 had a difficult delivery into the 53-year-old female patient. once placed, the pump did not start despite multiple attempts. therefore, the impella 5.5 was explanted, and a new one was placed. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported pump did not start issue has been completed since the original report was filed. the pump was returned for investigation. an extensively stretched catheter was observed on the returned pump, which leads to an open electric line. also pump failed during electrical testing and did not start during the test run. data logs were not provided. as per clinical details, pump was found broke and did not start after multiple insertion were made to place the pump. the cause of the pump did not start issue was an open electrical line found on returned product evaluation.
31988 | 1220648-2024-05710 | 18434501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18434501&pc=OZD | 12/14/2023 | 01/17/2024 |  | 01/03/2024 |  |  | 12/14/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024398271 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the physician was unable to advance impella 5.5 pass the axillary artery. additional attempts to advance alongside stiff "buddy" wire were also unsuccessful. the case was aborted. the patient is stable. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
31989 | 1220648-2024-05714 | 18435530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18435530&pc=OZD | 12/06/2023 | 01/17/2024 |  | 01/03/2024 |  |  | 12/06/2023 | 01/03/2024 |  |  | 08/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391394 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 64-year-old white male had leg ischemia leading to impella cp pump explant. a transcatheter aortic valve replacement (tavr) was performed and a new impella was placed. the patient was also put on ecmo support. following intervention, the patient was ultimately able to be weaned off both supports and was noted to be ok. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. no product was returned. the root cause of the limb ischemia was most likely patient condition related. the failure modes will be monitored and trended.
31990 | 1220648-2024-05715 | 18435953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18435953&pc=PYX | 12/17/2023 | 07/15/2024 |  | 01/03/2024 |  |  | 12/17/2023 | 01/03/2024 |  |  | 03/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2024388957 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2023 |  |  |  |  | The user facility reported a 41-year-old male with cardiomyopathy and was implanted with an impella rp flex for mechanical circulatory support. the patient developed hemolysis and hematuria and the impella rp flex was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: nothing beyond study data. | The investigation for the hemolysis has been completed. the impella rp flex pump was returned for investigation. upon visual inspection, there was significant biomaterial ingested on the inflow cage and wrapped around the impeller. also, a small thin blue fiber stuck outside to the inflow cage and pump cannula was seen. no other abnormalities were found. the biomaterial was removed, and pump ran at normal flow range per rp pump design specification. data logs were reviewed and suction alarm, motor current spike during no change in p-level with low flows was found. the cause of the hemolysis issue was due to biomaterial wrapped around motor housing at the inflow per field investigation and motor current spike observed per log review. the cause of the low pump flow issue was due to biomaterial ingestion observed at the inflow cage and around motor housing per field investigation and motor current spike with low flows observed per log review.
31991 | 1220648-2024-05716 | 18436048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436048&pc=OZD | 12/08/2023 | 09/20/2024 |  | 01/03/2024 |  |  | 12/08/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 204417423 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had placed a 5.5 pump for care escalation from a balloon pump (iabp). the patient was also supported by ecmo (extracorporeal membrane oxygenation). the patient was in for an or procedure to treat the vsd (ventricular septal defect¿¿) . the pump was placed but, after 1 day of support the patient expired. the death was due to msof (multi-system organ failure) . the msof was concurrent with a mediastinum bleed. the bleed was treated by blood transfusions. the bleed may have been due to the use/placement of the 5.5 pump. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-05716 was provided in sections b2, b5,d6, h1, and h6.
31992 | 1220648-2024-05717 | 18436160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436160&pc=OZD | 12/08/2023 | 03/14/2024 |  | 01/03/2024 |  |  | 12/08/2023 | 01/03/2024 |  |  | 08/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024395144 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Handling Problem | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that while placing the repositioning sheath following implant, the impella cp pump was pushed too deep in the left ventricle and kinked on itself. attempts were made to untangle the pigtail / remove the device; however, the inflow cage broke. a snare was utilized to remove the cage and a new pump was implanted for support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) and the product damage (detached inflow/outflow - great vessel) issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the cannula kink was most likely use related since the kink occurred when the pump was too deep in the ventricle after the repo sheath was advanced. the root cause of the product damage detached inflow cage was not determined since product was not returned and there is a lack of clinical details.
31993 | 1220648-2024-05719 | 18436196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436196&pc=OZD | 12/11/2023 | 01/17/2024 |  | 01/03/2024 |  |  | 12/11/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403642 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported an impella 5.5 in a 71yo white male for mechanical circulatory support. it was reported that during insertion, the impella 5.5 pump was unable to cross the patient's new valve during attempted delivery. after multiple unsuccessful attempts, the decision was made to abort the implant. ecmo was cannulated centrally to support the patient. there was no patient harm. no additional information was provided. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the clinical details states that multiple attempts were made using various techniques to insert the device but the decision was made to abort. therefore, the most likely root cause of the delivery issue was due to patient condition. | The impella device was not returned for evaluation. a supplemental report will be submitted when the investigation has been completed.
31994 | 1220648-2024-05718 | 18436199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436199&pc=OZD | 12/08/2023 | 10/03/2024 |  | 01/03/2024 |  |  | 12/08/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024400699 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. there was significant bleeding noted from the insertion site. the patient was transfused with blood products and the decision was made to explant the impella cp and escalate support to impella 5.5. the impella cp was removed, and the site was surgically closed. | It was further reported that after escalation of care with a different impella, care was withdrawn and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added -b2 selected death and added date of death. h6-impact code 1802, h6 component code 925 corrections to the initial mfr report: b5-mso review. added -d6a/d6b. f6/f8 should have been left blank. h1-type of reportable event changed to death. h6 impact code 4627 changed to 4629.
31995 | 1220648-2024-05720 | 18436250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436250&pc=OZD | 12/10/2023 | 11/17/2025 |  | 01/03/2024 |  |  | 12/10/2023 | 01/03/2024 |  |  |  |  | 01/03/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1420044 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported a 49-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the automated impella controllers (aic) power cable became disconnected from the plug when the nurse was pulling on the cable from the wall. the aic was replaced. there was no patient harm. | [5 empty] | The impella device was received from the customer and, an evaluation of the device is not yet completed. upon investigation closure, a supplemental mdr will be filed. | The investigation for this issue has been completed. data review: no data review was needed for this investigation since it was only reported that there was physical damage to the aic¿s power cord. device analysis: there was no evidence of existing damage to the ac power cord or plug, but there was evidence that the cord had been previously damaged. the batteries and charging system were confirmed to be working correctly. the root cause of this issue was a damaged power cord. h3: added information regarding the status of the device investigation. h6: added codes for the result of the investigation. h10: added investigation information.
31996 | 1220648-2024-05721 | 18436468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436468&pc=OZD | 12/16/2023 | 11/30/2024 |  | 01/03/2024 |  |  | 12/16/2023 | 01/03/2024 |  |  | 06/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379352 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 43-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported that there is a slight leak from the red port on the impella cp. the purge flows and pressure will be monitored. it was further reported that the patient experienced ischemia post repo sheath placement and external fem fem bypass was done. additionally, the groin had a small hematoma, but groin was swollen prior to pump insertion so it is hard to determine if it is due to the procedure. hematoma is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the purge leak-pump and the limb ischemia has been completed. product not returned for review. the cause of the access site bleeding minor was not determined since product was not returned and lack of clinical details. the cause of the purge leak pump was not determined since product was not returned and lack of clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank.
31997 | 1220648-2024-05722 | 18436486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436486&pc=OZD | 12/13/2023 | 04/12/2024 |  | 01/03/2024 |  |  | 12/13/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024417619 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 57-year-old female of unknown race and ethnicity post cardiotomy cardiogenic shock /low cardiac output syndrome was implanted with an impella for mechanical circulatory support. when the doctor did the cut down to decannulate in the right groin, it was noticed that there were clots in both the venous and arterial system.. thrombectomy was done. attention was turned to the left impella groin and the left sfa appeared to have clots as well. the patient had good flow after the removal and thrombectomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined.
31998 | 1220648-2024-05723 | 18436538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436538&pc=PYX | 12/14/2023 | 04/12/2024 |  | 01/03/2024 |  |  | 12/14/2023 | 01/03/2024 |  |  | 12/31/2024 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024308936 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 79-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella rp was not appropriately sitting across tricuspid ring. map (transesophageal echocardiography) value had not improved since rp started. the physician decided to remove rpsa and exchange for rp flex to be able to set inlet cage in ra. impella rpsa weaned and removed with no complications. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. product was not returned for investigation. the positioning issue was most likely related to the patient's condition since they had anatomical abnormalities following tricuspid ring surgery, based on given clinical details.
31999 | 1220648-2024-05724 | 18436638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436638&pc=OZD | 12/14/2023 | 01/21/2025 |  | 01/03/2024 |  |  | 12/14/2023 | 01/03/2024 |  |  |  |  | 01/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1141486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility's biomed sent in a console (aic). the aic did not detect and function appropriately at the purge disc. the aic was removed from service. there was no patient involvement. no additional information was provided. | [5 empty] | The console device has been returned for evaluation but has not yet been evaluated. a supplemental report will be submitted upon completion of the investigation. | The investigation for the console (aic) purge flag/disc flag has been completed. data logs from the console confirm multiple times that the purge not detected alarm was issued. the console was returned for evaluation. the console was visually inspected under a digital microscope and confirmed for a missing purge flag. the root cause for the purge disc not detected alarm was a missing purge flag.
32000 | 1220648-2024-05725 | 18436709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436709&pc=OZD | 12/08/2023 | 10/29/2024 |  | 01/03/2024 |  |  | 12/08/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024406152 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported an impella cp in a 71yo female patient for mechanical circulatory support. it was reported that no distal pulses were found in the left lower extremity after removal of the impella. run off of leg shot was showing a distal lesion that was not there previously. a ballooning of the lesion on left side was performed. run off showed increased flow but a dissection of left common femoral (lcf) artery limiting distal flow was noted. the physician increased balloon size and ballooned a second time with success. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. the failure mode will be monitored and trended. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h6 component code 925 corrections to the initial mfr report: d6a/d6b. f6/f8 should have been left blank.
32001 | 1220648-2024-05726 | 18436893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436893&pc=PYX | 12/18/2023 | 09/11/2024 |  | 01/03/2024 |  |  | 12/18/2023 | 01/03/2024 |  |  | 02/28/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024348384 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 43-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the pump was supporting for 5 days. the pump was explanted after 5 days due to low flows. no patient harm reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the root cause of the low pump flow issue was most likely biomaterial.
32002 | 1220648-2024-05727 | 18436921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18436921&pc=OZD | 12/15/2023 | 11/01/2024 |  | 01/03/2024 |  |  | 12/15/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400431 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 76-year-old male for mechanical circulatory support. it was reported that in the last hour of the percutaneous coronary intervention procedure, there was more oozing than before at the impella access site. hemoglobin dropped enough requiring two units of blood to be transfused shortly after the procedure. the patient is stable. no additional information was provided. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was returned for evaluation and there were no issues found. there was insufficient clinical information returned to determine the root cause. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32003 | 1220648-2024-05728 | 18437571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18437571&pc=OZD | 12/18/2023 | 01/17/2024 |  | 01/03/2024 |  |  | 12/18/2023 | 01/03/2024 |  |  | 08/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393758 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 50-year-old female undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. the impella cp would not pass through the ascending aorta due to the patient anatomy. the placement of cp was aborted. ecmo was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition.
32004 | 1220648-2024-05729 | 18437768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18437768&pc=OZD | 12/15/2023 | 10/20/2025 |  | 01/03/2024 |  |  | 12/15/2023 | 01/03/2024 |  |  |  |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721589 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported that the automated impella controller (aic) had a placement signal that was not working, and the flows would not increase past p2. the aic was removed. there was no patient harm reported. no additional information was not provided. | [5 empty] | The console was received from the customer and an investigation of the device is ongoing. upon completion of our analysis, a supplemental report will be filed.
32005 | 1220648-2024-05730 | 18438255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18438255&pc=OZD | 12/16/2023 | 06/20/2024 |  | 01/03/2024 |  |  | 12/16/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412722 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 63 year-old male for mechanical circulatory support. it was reported that the cp site in the femoral had a bleed that prompted placement of femostop and the infusion of red blood cells back to the patient. days later, the site was still ecchymotic but no further treatment was noted. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. a supplemental report will be submitted upon completion of the investigation. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be anticoagulation management because the act values (250) are high and above the threptic range and heparin was stopped to control the bleeding. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32006 | 1220648-2024-05731 | 18438332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18438332&pc=OZD | 12/19/2023 | 09/20/2024 |  | 01/03/2024 |  |  | 12/19/2023 | 01/03/2024 |  |  | 06/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364395 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient¿s culture came back positive for bacteremia which initiated plans for pump removal. the impella was removed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
32007 | 1220648-2024-05732 | 18438496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18438496&pc=OZD | 12/10/2023 | 01/19/2024 |  | 01/03/2024 |  |  | 12/10/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024409480 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | Tthe user facility reported a (b)(6) female undergoing cardiac surgery was to be implanted with an impella device for mechanical circulatory support. us complainant reported the patient experienced delivery issues as the impella would not pass from the innominate artery to the ascending aorta. the patient was on impella support for 1.83 hours before the physician decided to abort the procedure. it was reported that after multiple attempts the physician decided to place central va-ecmo (venoarterial extracorporeal membrane oxygenation). | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32008 | 1220648-2024-05733 | 18438565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18438565&pc=OZD | 12/11/2023 | 01/18/2024 |  | 01/03/2024 |  |  | 12/11/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024405702 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported an unknown female noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. us complainant reported that the impella cp was attempted to be placed but due to the patient¿s vascular anatomy it was unable to be advanced into the aorta. the doctor decided to remove the impella. no patient harm or injury reported due to the issue. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32009 | 1220648-2024-05734 | 18438587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18438587&pc=OZD | 12/11/2023 | 01/26/2024 |  | 01/03/2024 |  |  | 12/11/2023 | 01/03/2024 |  |  |  |  | 01/03/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1164959 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The complainant reported that the automated impella controller (aic) in the supply room had a battery failure when attempting to turn it on. troubleshooting was unsuccessful. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the into the aic - battery failure (red alarm) has been completed since the original report was filed. after reviewing the imc logs, the issues with this console having a battery failure alarm on the reported occurrence date was confirmed. imc logs from the reported occurrence date show 3 instances of the alarm. console was tested after arriving in fs. unable to reproduce the battery failure after engaging the battery switch before booting up the console. from. the root cause of the battery failure alarm is likely due to a disengaged battery switch (use error).
32010 | 1220648-2024-05735 | 18438606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18438606&pc=OZD | 12/11/2023 | 06/05/2024 |  | 01/03/2024 |  |  | 12/11/2023 | 01/03/2024 |  |  | 08/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386663 |  | 0042-0000 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The patient's outcome was reviewed by abiomed's medical safety officer (mso). it was determined that there was not enough information to associate use of device with outcome. | The user facility reported an impella cp in a (b)(6) female with heart failure/cardiogenic shock. the complainant reported that impella cp had a sudden decrease in flow. troubleshooting was unsuccessful as cp flow did not improve. the impella was exchanged with a new cp. there was no harm to the patient. no additional information was provided. |  |  |  |  | The investigation for the reported low pump flow has been completed. the pump was not returned for investigation. the total case runtime was 1 hour and 10 minutes. data logs show low pump flow throughout the case without reported motor current spikes or positioning issues. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use referenced in the initial report is for the impella cp with smartassist system in section: use of echocardiography for positioning of the impella catheter. | The impella device has not been returned for investigation. however, the investigation is in progress. a supplemental report will be submitted upon completion of the investigation. the instructions for use states the following regarding flow issues:¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
32011 | 1220648-2024-05736 | 18438826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18438826&pc=OZD | 12/05/2023 | 09/24/2024 |  | 01/03/2024 |  |  | 12/05/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400560 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported an 84-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. a small hematoma was noted on the impella access site and manual pressure was held. the doctor then explanted the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient active clotting act values, heparin was stopped to achieve hemostasis as per the clinical description.
32012 | 1220648-2024-05737 | 18439726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18439726&pc=OZD | 12/05/2023 | 02/22/2024 |  | 01/03/2024 |  |  | 12/05/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414446 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 57-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the cardiologist made the decision to remove the impella due to the ischemia of the leg. after removal of the impella in cath lab, hemostasis was achieved, and pulses were palpable post explant of affected extremity. patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. per the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. the root cause of the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned.
32013 | 1220648-2024-05738 | 18439758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18439758&pc=OZD | 12/06/2023 | 07/18/2024 |  | 01/03/2024 |  |  | 12/06/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that before the impella insertion, peripheral angiography showed calcium higher up on the right iliac artery. decision was made to proceed with sheath and impella cp placement for support. left ventricular (lv) access was obtained and 018 wire placed. impella was backloaded on 018 wire and there was some initial difficulty to pass impella through the iliac. impella cp access was obtained in the lv and the device started. shortly after the device started, angiography of the right iliac showed a dissection. impella procedure was therefore aborted, and the patient was taken to the or for iliac dissection repair. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the vascular damage has been completed. no product was returned for investigation. the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related due to the patient's calcified vessels. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32014 | 1220648-2024-05739 | 18439850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18439850&pc=OZD | 12/15/2023 | 05/16/2024 |  | 01/03/2024 |  |  | 12/18/2023 | 01/03/2024 |  |  | 06/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024374211 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 74-year- male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. immediately after insertion, the pump was not providing adequate flow. the impella cp was repositioned several times, confirming placement by echocardiogram (echo), and the inadequate flow continued. the decision was made to remove and replace with a new impella cp pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issues has been completed since the original report was submitted. the device was not returned for investigation. data logs show low pump flow throughout the entire case run with an active suction alarm. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided.. d6a, d6b.
32015 | 1220648-2024-05740 | 18439889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18439889&pc=OZD | 12/06/2023 | 01/18/2024 |  | 01/03/2024 |  |  | 12/06/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406276 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 41 -year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. resistance was met during attempts to implant an impella 5.5 pump. it was noted that the pump would not advance due to the patient's tortuous vasculature. as a result, the impella implant was aborted and ecmo was placed for ongoing support. no patient harm was reported due to the issue. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the cause of the delivery issue was patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32016 | 1220648-2024-05741 | 18439925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18439925&pc=OZD | 12/07/2023 | 04/30/2024 |  | 01/03/2024 |  |  | 12/07/2023 | 01/03/2024 |  |  | 12/31/2024 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024297894 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The complainant reported following placement of impella cp for support, the male patient of unknown age had bleeding in the groin area. blood transfusion was performed and the pump was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, evaluation of the device was not possible. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. per the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. in accordance with updated procedures added- d6a/d6b. d4-updated model number. g1 updated reporting first and last name. have been revised from the initial submission.
32017 | 1220648-2024-05742 | 18440056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440056&pc=PYX | 12/05/2023 | 03/29/2024 |  | 01/03/2024 |  |  | 12/05/2023 | 01/03/2024 |  |  | 03/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024388956 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported an impella rp in a 31yo male for cardiomyopathy and heart failure. it was reported that during support, there were signs of hemolysis. the patient had an open chest while on support. the hemolysis was treated with dialysis, blood products and platelets. no additional information was provided. | [5 empty] | The impella device has been returned for evaluation. however, it has not yet been evaluated. a supplemental report will be submitted upon completion of the investigation. the instructions for use states the following related to hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ the failure modes are monitored and trended. | The investigation for the reported hemolysis has been completed. the impella device was returned for evaluation. visual inspection found a small amount of biological material on the impeller. there was no catheter or cannula kinks. the day prior to the reported hemolysis, the patient was dry and vasodilated with suction alarms. data logs confirm these suction alarms which were resolved by decreasing p-level but there were no indications of ingestion events. the root cause of the hemolysis was most likely patient condition related as the patient was dry and vasodilated with suction alarms prior to the suspected hemolysis. decreasing the p-level improved the hemolysis.
32018 | 1220648-2024-05743 | 18440063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440063&pc=OZD | 12/04/2023 | 09/06/2024 |  | 01/03/2024 |  |  | 12/05/2023 | 01/03/2024 |  |  |  |  | 01/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1586059 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 62-year-old female with cardiomyopathy was implanted with an impella for mechanical circulatory support. the complainant reported low battery alarms, despite charging. the alarm occurs upon unplugging the device. as a result, the pump was moved to a new console and support continued. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of aic - battery charging/other issues has been completed. the logs confirmed the reported complaint. there were multiple occurrences of battery level low alarms, immediately preceding each battery level low alarm was an occurrence of the event. the console was visually investigated and no physical defects were identified. the batttest utility was run and no battery or charging defects were identified. the cause of the aic issue was most likely a defective pbm board. the following have been reported incorrectly or missing from manufacturer device report: d2 common device name f6/f8-should have been left blank g1 reporting contact fax number missing.
32019 | 1220648-2024-05744 | 18440080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440080&pc=OZD | 12/05/2023 | 09/21/2024 |  | 01/03/2024 |  |  | 12/05/2023 | 01/03/2024 |  |  | 04/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350434 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 78-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella for mechanical circulatory support. the patient had an attempted impella cp implant, and the staff was unable to place the catheter. the patient expired in the procedural area. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
32020 | 1220648-2024-05745 | 18440095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440095&pc=OZD | 12/05/2023 | 07/02/2025 |  | 01/03/2024 |  |  | 12/05/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403712 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. the user reported a blockage in the purge system. replacing the purge cassette did not solve the problem. as a result, the impella cp stopped due to high motor current. the impella heart pump was removed and a new impella heart pump was not inserted. the patient died of multi-organ failure. the user assessed the patient's death as unrelated to the impella pump stop. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop and high purge pressure has been completed since the original report was filed. the pump was returned for investigation; however, data logs were not provided. it successfully passed the electrical resistance and optical testing. during a test run in a water bucket, the pump stopped with a high motor current alarm. subsequent investigation revealed biomaterial present in the purge gap and motor housing. additionally, a high purge pressure alarm occurred during the purge test. no kinks were found in the returned pump. further investigation revealed a cut catheter near the motor housing, which allowed purging through it and confirmed that the biomaterial led to a high purge pressure event, allowing blood ingress into the motor housing. as per the clinical details, the doctor reported pump stop after the high purge pressure event. the cause of the pump stop issue was biomaterial, based on the returned product review. | Added b2-death and date of unknown. the patient outcome of death was previously included in the b5 narrative of initial mfr report 1220648-2024-05745. h1 updated to death.
32021 | 1220648-2024-05746 | 18440107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440107&pc=OZD | 12/05/2023 | 10/24/2024 |  | 01/03/2024 |  |  | 12/05/2023 | 01/03/2024 |  |  | 09/30/2024 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278590 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 52-year-old female patient developed hemolysis during impella cp support. the pump was initially noted to be too deep in the left ventricle, but the decision was ultimately made to escalate to impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05746: d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32022 | 1220648-2024-05747 | 18440136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440136&pc=OZD | 12/05/2023 | 03/11/2025 |  | 01/03/2024 |  |  | 12/05/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408719 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 78-year-old female patient was on impella 5.5 support and experienced hemolysis. the nurse gave 500ml of volume. echo was performed and the 5.5 was noted to be lightly deep and was therefore pulled back 1cm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ | The pump was not returned for investigation. data logs show several suction alarms, with trending placement signal and motor current, along with reported low pump flow during a hemolysis event on 12/05. pump was explanted on (b)(6). as per the clinical details, some adjustments were made to the position of the pump, and blood was given to improve the hemolysis. the patient had crrt the following day, and dark yellow urine output was reported the next day. the cause of the hemolysis issue was most likely patient condition related, based on given clinical details and data log review.
32023 | 1220648-2024-05748 | 18440259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440259&pc=PYX | 12/04/2023 | 02/19/2024 |  | 01/03/2024 |  |  | 12/04/2023 | 01/03/2024 |  |  | 02/28/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  |  |  | 2024348532 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella rp device for mechanical circulatory support. the impella .027 wire could not be placed due to the patient's cardiac anatomy and severely acute angle kinking catheters in right ventricle. after almost 2 hours of attempting to place, the decision was made to abort procedure for patient safety. the physician reported no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for investigation. based on clinical description, the root cause of the delivery issue was most likely patient condition related as multiple catheters and wires were used to attempt the delivery of the .027 wire into the pa without success. the clinical specifies anatomical issues led them to abort the delivery.
32024 | 1220648-2024-05749 | 18440265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440265&pc=OZD | 12/18/2023 | 02/23/2024 |  | 01/03/2024 |  |  | 12/18/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024418751 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had impella 5.5 support ongoing when there was upper limb/right upper extremity (rue) pain and discomfort, by the axillary impella 5.5 insertion. the team medicated the pain with gabapentin and tylenol. the impella 5.5 pump remained on support for 5 days now. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32025 | 1220648-2024-05750 | 18440285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440285&pc=OZD | 12/07/2023 | 05/03/2024 |  | 01/03/2024 |  |  | 12/08/2023 | 01/03/2024 |  |  | 11/30/2024 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024295778 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility in japan reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp support when a left ventricle thrombus formation occurred. the impella was replaced. the physician believed the thrombus may have been due to the withdrawing of heparin for patient management, causal relationship to impella is not known. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. d6a, d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32026 | 1220648-2024-05751 | 18440297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440297&pc=OZD | 12/16/2023 | 12/04/2024 |  | 01/03/2024 |  |  | 12/16/2023 | 01/03/2024 |  |  | 08/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385854 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2023 |  |  |  |  | The user facility reported a 72-year-old white female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp for support of a patient admitted in ami/cgs (acute myocardial infarction / cardiogenic shock). the pump was explanted the day after implanting due to concerns of hemolysis. the hemolysis was observed overnight, and attempted troubleshooting was done with pump repositioning. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. no product returned. data logs show no abnormalities and no motor current (mc) spikes outside p level changes. clinical information is conflicting stating pump was deep and also in good position. the root cause of the hemolysis issue was not determined as product was not returned and limited clinical information and inconclusive data logs. added- h6 component code 925 corrections to the initial mfr report: d4-model number. added d6a/d6b. f6/f8 should have been left blank. h6 impact code 4629 changed to 4627 and 4628.
32027 | 1220648-2024-05752 | 18440298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440298&pc=OZD | 12/15/2023 | 01/31/2024 |  | 01/03/2024 |  |  | 12/15/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024419481 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 76-year-old white female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported while attempting to complete a protected percutaneous coronary intervention, the doctor successfully placed an impella cp. once the wire was removed and the impella cp was turned on, the flow rate in auto remained <1l/min and there was a kink visualized in the catheter. multiple attempts were made to manipulate the catheter by the doctor and the impella cp was removed from auto without any change in flows. the decision was made to explant this device and implant a secondary device. once the secondary device was implanted, flows were >3.5l/min at p-9 without any kink visualized. the doctor continued to perform the primary percutaneous coronary intervention successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the impella cp pump was returned. visually inspected the pump and the entire cannula seemed to be kinked completely. biomaterial was not found to present in the inflow/outflow or around impeller. data logs were reviewed, and low flow observed with suction. also motor current spike was seen. the cause of the low pump flow issue was a kinked cannula per field investigation and clinical review.
32028 | 1220648-2024-05753 | 18440301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440301&pc=OZD | 12/15/2023 | 12/11/2024 |  | 01/03/2024 |  |  | 12/15/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398865 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Handling Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 73-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed the cp via a 14fr sheath to support a hrpci (high risk percutaneous coronary intervention ) patient via single access. the long 14fr sheath had kinked in insertion, and the team performed angioplasty prior to placement of the short sheath. no harm or injury to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. based on clinical details provided, the root cause of the introducer damage was determined to most likely be patient condition. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05753. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing.
32029 | 1220648-2024-05754 | 18440308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440308&pc=OZD | 12/17/2023 | 04/09/2024 |  | 01/03/2024 |  |  | 12/17/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024397018 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2023 |  |  |  |  | The user facility reported a 69-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella cp pump triggered high purge pressure alarms during patient support. purge flows were noted to be below 1 ml/hr at the time of the reported failure. the purge cassette was initially replaced, but the issue persisted. as such, the pump was explanted, and the patient was escalated to impella 5.5 support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned, tested, and evaluated no physical abnormalities were observed on the returned pump. during the test, high purge pressure was not reproduced. however, data logs showed a sudden rise in purge pressure and a high purge pressure alarm within a 4-minute time frame. the cause of the high purge pressure issue was most likely biomaterial based on return product investigation and data log review.
32030 | 1220648-2024-05755 | 18440309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440309&pc=OZD | 12/11/2023 | 12/11/2024 |  | 01/03/2024 |  |  | 12/11/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIONED, INC |  |  | US |  | 2024406258 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 24-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient¿s labs were hemolyzed. the physician suspects ingested residual thrombus. a high dosage of systemic heparin was increased and the issue resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into hemolysis and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05755. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email g1 reporting contact fax number missing. h6 health effect - clinical code 4440 missing h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use missing.
32031 | 1220648-2024-05756 | 18440312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440312&pc=OZD | 12/17/2023 | 11/30/2024 |  | 01/03/2024 |  |  | 12/18/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418131 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 67-year-old white female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed for a patient with multiple cardiac and systemic comorbidities. the team placed a fem-fem bypass for limb perfusion. the cp was explanted and replaced when the purge sidearm began to leak. the event is reportable for the harm and malfunction. the patient is now supported by a new pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into purge leak ¿ pump and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05756. b1 should be both adverse event and product problem. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 mdr reporting contact fax number missing. g1 manufacturing site ax number missing. h6 code 4582 and 481 were reported incorrectly. new code was added to component code.
32032 | 1220648-2024-05757 | 18440315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440315&pc=OZD | 12/18/2023 | 05/01/2024 |  | 01/03/2024 |  |  | 12/18/2023 | 01/03/2024 |  |  |  |  | 01/03/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147986 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support when the automated impella controller (aic) was found to have a cracked purge disc door. the aic was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) purge disc/flag issues has been completed. the console was returned for evaluation. physical inspection confirmed that the console¿s front housing was damaged near the location of the purge door release button. data logs were not relevant to this investigation. no issues were found in the data logs. the cause of the issue was not determined since it is not known how the console was damaged. added-d9 device return date. added- h6 impact code 4629. d4-updated model number.
32033 | 1220648-2024-05758 | 18440381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440381&pc=OZD | 12/06/2023 | 12/11/2024 |  | 01/03/2024 |  |  | 12/06/2023 | 01/03/2024 |  |  | 06/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375598 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the impella cp was inserted via left femoral artery (lfa) and pci was performed via the right femoral artery (rfa). the physician, while removing the right coronary artery (rca) catheter, accidentally pulled the impella cp across the aortic valve (av) into the ascending aorta. the cp could not cross the av and a decision was made to remove the impella and insert a new impella cp. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on the clinical details, the root cause of the positioning issue is due to use as the physician accidentally pulled the pump across the av into the aorta.
32034 | 1220648-2024-05759 | 18440386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440386&pc=OZD | 12/06/2023 | 05/06/2024 |  | 01/03/2024 |  |  | 12/06/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024402484 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported an impella 5.5 in a 75yo male for cardiomyopathy and heart failure. it was reported that a hematoma was observed at the impella 5.5 access site. the wound was opened, and bleeding was surgically stopped. the bleeding was found from the artificial blood vessel side of the artificial blood vessel anastomosis. the patient remained stable. no additional information was provided. | [5 empty] | The impella device has not been returned for evaluation. however, the investigation is in process. a supplemental report will be submitted upon completion of the investigation. the instructions for use states indicates the following :¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ it also states to "assess access site for bleeding and hematoma.¿ the failure modes are being monitored and trended. | The investigation for the reported access site bleed and atrial fibrillation has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the issue is use issue because bleeding occurred at the graft attached to the vessel and was able to resolve upon surgical repair.
32035 | 1220648-2024-05760 | 18440402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440402&pc=OZD | 11/28/2023 | 09/29/2025 |  | 01/03/2024 |  |  | 12/06/2023 | 01/03/2024 |  |  | 09/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024399225 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Death | Yes | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/06/2023 |  |  |  |  | The complainant reported 55-year-old male patient on hemodynamic support from an impella cp smartassist heart pump had 'impella in aorta alarm' despite no sign of device being out of place. therefore, the impella was explanted as a precaution. there was no reported patient harm. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The customer reported that the device exhibited an optical sensor issue. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical signal issue has been completed. the pump was returned for investigation, but data logs were not provided. all the 5s optical parameters were within the specified limits for the returned product. no physical abnormalities, such as kinks or biomaterial issues, were observed. according to clinical details, a false position in the aorta was noted. the pump position was verified under fluoroscopy, and due to insecurity, the doctor pulled the impella out of the aorta. the cause of the optical signal issue cannot be determined. added- b5 mso review, d9 device return date added- d6a/d6b. f6/f8 should have been left blank in the initial report. g1-mdr reporting first and last name corrected. g1-mfr site fax number corrected. h6 med dev prob code changed to 2917. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was returned for investigation, but data logs were not provided. all the 5s optical parameters were within the specified limits for the returned product. no physical abnormalities, such as kinks or biomaterial issues, were observed. according to clinical details, a false position in the aorta was noted. the pump position was verified under fluoroscopy, and due to insecurity, the doctor pulled the impella out of the aorta. the cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added death as part of a retrospective review of optical signal issues in accordance with fda recommendation.
32036 | 1220648-2024-05761 | 18440408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440408&pc=OZD | 12/07/2023 | 05/03/2024 |  | 01/03/2024 |  |  | 12/07/2023 | 01/03/2024 |  |  | 12/31/2024 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302769 |  | 1000080 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported an impella cp in a 39yo male patient for mechanical circulatory support. it was reported that during insertion of impella, the patient immediately went into ventricular fibrillation as the pigtail was inserted in the left ventricle. the insertion of impella device was then aborted. the patient was sent for higher level of care on pharmacological support. no additional information was provided. | [5 empty] | The investigation into the vf/vt/af/at issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device has not been returned for evaluation. when the investigation has been completed, a supplemental report will be submitted. the failure mode is monitored and trended.
32037 | 1220648-2024-05762 | 18440433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440433&pc=OZD | 12/19/2023 | 01/03/2024 |  | 01/03/2024 |  |  | 12/19/2023 | 01/03/2024 |  |  |  |  | 01/03/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462548 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 12/19/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 16 year-old male due to cardiomyopathy. it was reported that the automated impella controller (aic) did not make audible noise when alarming. the patient remained stable. no additional information was provided. | [5 empty] | The controller has been returned for evaluation but has not yet been evaluated. a supplemental report will be submitted upon completion of the investigation.
32038 | 1220648-2024-05763 | 18440435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440435&pc=OZD | 12/07/2023 | 07/15/2024 |  | 01/03/2024 |  |  | 12/07/2023 | 01/03/2024 |  |  | 08/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395902 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 55-year-old white male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had access site bleeding and a hematoma. the patient was given blood products, femstop was applied and manual pressure was held. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding is most likely due to lack of vessel recoil based on the provided clinical details. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations section: entering cardiac output section: potential adverse events (united states) added-d9 device return date, h6 component code 925 added- d6a/d6b f6/f8 should have been left blank in the initial report.
32039 | 1220648-2024-05764 | 18440442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440442&pc=OZD | 12/08/2023 | 09/22/2024 |  | 01/03/2024 |  |  | 12/08/2023 | 01/03/2024 |  |  | 06/30/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  |  | Injury | Yes | Yes | Fracture | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 12/08/2023 |  |  |  |  | The user facility reported an impella cp in a (b)(6) male with acute myocardial infarction/cardiogenic shock. the team placed the cp but upon removal of the .018 wire, the wire broke. the team observed it in the aorta and outside of the pump itself. the pump was started but family later made choice to withdraw care. no additional information was provided. | Medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not returned for evaluation. a supplemental report will be submitted upon completion of the investigation. the failure mode is being monitored and trended. | The investigation into the product damage (guidewire damage - great vessel) has been completed. the pump was not returned for investigation. the cause of the product damage (guidewire damage) issue was not determined since the product was not returned and sufficient clinical information was not provided. h.2 device evaluation was selected on the original manufacturer device report (mdr) and should not have been selected as that report was not a follow-up report. h.3 evaluation summary attached was checked on the original mdr inadvertently. h.10 the following statement was reported on the original mdr and should not have been: the failure mode is being monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-05764 was provided in sections b1, b5, d6, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
32040 | 1220648-2024-05765 | 18440460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440460&pc=OZD | 12/08/2023 | 04/22/2024 |  | 01/03/2024 |  |  | 12/08/2023 | 01/03/2024 |  |  | 10/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412517 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient of unknown age and gender had bleeding complication at the right common femoral artery during single access procedure, leading to pump explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Upon further review athe complaint was re-assessed as not reportable. the event reported is a reversible harm. the access site bleed resolved with manual compression and no secondary intervention was required. the event reported did not cause or contribute to a serious deterioration in state of health or death and is not likely to cause or contribute to a serious deterioration in state of heath or death if it were to reoccur.
32041 | 1220648-2024-05766 | 18440537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440537&pc=OZD | 10/05/2018 | 10/02/2024 |  | 01/03/2024 |  |  | 12/08/2023 | 01/03/2024 |  |  | 05/31/2020 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1346915 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | A case series* was located and reviewed on 2023-12-08 which mentioned a hematoma in the patient impella access site. the patient developed hematoma on 2018-oct-05 and ventricular septal defect (vsd) repair was done on 2018-oct-16. the patient was stable. * european heart journal - case reports (2023) 7, 1¿5 h. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported vascular damage/bleeding has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage/bleeding could not be determined.
32042 | 1220648-2024-05767 | 18440545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440545&pc=OZD | 12/09/2023 | 05/14/2024 |  | 01/03/2024 |  |  | 12/09/2023 | 01/03/2024 |  |  | 08/31/2025 |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon initiation of impella cp support the purge pressure low alarm was present during ramp up. purge tubing, connections, cassette, and disk were investigated from spike to red plug without identified leak. alarm resolved approximately 1min from initiation of support. around 3min into support, impella flows dropped. due to suspected clot, impella pump was removed and investigated for clot with no thrombus visualized. echo - echocardiogram (ultrasound) from day prior also revealed no left ventricle (lv) thrombus present. a decision was made to proceed with a new pump. impella cp was successfully inserted. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the pump was not returned for investigation. the total case run time was eight minutes without any purge-related abnormalities. data logs showed several suction alarms with low pump flow. a 'purge system open' alarm appeared during the case start. there was no trend noted in motor current spikes or placement signals. according to clinical records, the 'purge system open' alarm was resolved within three minutes without any visible leaks observed. the pump was replaced after a reported low pump flow. the cause of the low pump flow issue was not determined since the product was not returned and sufficient clinical information was not provided. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4440, 2199 and 4627 were reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
32043 | 1220648-2024-05768 | 18440577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18440577&pc=OZD | 12/20/2023 | 01/21/2025 |  | 01/03/2024 |  |  | 12/20/2023 | 01/03/2024 |  |  |  |  | 01/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1319826 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The us complainant had a console fail during support. the failing automated impella controller (aic) alarmed "console error". the console remained functional without obvious issue. the team replaced the aic with a backup per instructions for use (ifu) recommendations. no patient or injury from the interruption. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated imeplla controller (aic) - controller error (yellow alarm) has been completed. the root cause of the loss of communication and thus system self check failure was contamination of the power battery manager (pbm) resulting from corrosion of the copper traces from leakage of supercapacitor electrolyte based on the data analysis and device evaluation the following have been reported incorrectly or missing from manufacturer device report. d2 common device name e2 health professional should have been selected as no. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing. h5 labeled for single use should have been selected as no. h8 reuse should have been selected for usage of device.
32044 | 1220648-2024-05671 | 18424326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18424326&pc=OZD | 10/03/2023 | 04/12/2024 |  | 01/02/2024 |  |  | 12/05/2023 | 01/02/2024 |  |  | 08/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385495 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. when the device was removed, thrombus was noted on the outlet cage. there was no additional treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
32045 | 1220648-2024-05672 | 18424355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18424355&pc=OZD | 12/05/2023 | 01/17/2024 |  | 01/02/2024 |  |  | 12/05/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024408765 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Accessory Incompatible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 62-year-old male in st elevated myocardial infarction was was going to be implanted with impella cp for mechanical support. during insertion, the impella cp was unable to cross the aortic arch and the impella cp implant was aborted. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
32046 | 1220648-2024-05673 | 18424480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18424480&pc=OZD | 12/06/2023 | 07/25/2024 |  | 01/02/2024 |  |  | 12/06/2023 | 01/02/2024 |  |  |  |  | 01/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1254965 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender was to receive hemodynamic support from an impella cp smartassist heart pump. during preparation, the automated impella controller (aic) did not recognize the impella device. the aic was therefore exchanged for the backup controller, which solved the problem. the patient suffered no harm from the incident. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The console was received for investigation. the data logs and the console were evaluated and the root cause for the placement signal not reliable alarm was due to the defective optical bench. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05673.
32047 | 1220648-2024-05674 | 18424596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18424596&pc=OZD | 12/06/2023 | 12/11/2024 |  | 01/02/2024 |  |  | 12/06/2023 | 01/02/2024 |  |  | 07/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024391197 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 72-year-old male was implanted with an impella cp for mechanical circulatory support. while on support the patient became anxious and pulled the implanted cp pump out of position. repositioning attempts were unsuccessful as the pigtail became entrapped in the mitral valve and subsequently forced the impella cp to flip towards the valve. as a result, the pump was removed, and the patient was escalated to impella 5.5 support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details provided, the root cause of the positioning issues is due to use as the patient pulled out the impella. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05674. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing.
32048 | 1220648-2024-05675 | 18424856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18424856&pc=OZD | 12/07/2023 | 01/17/2024 |  | 01/02/2024 |  |  | 12/07/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401595 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 68-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was unable to pass into the aorta from the axillary artery. multiple attempts were made unsuccessfully. the case was aborted. the patient is stable. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32049 | 1220648-2024-05676 | 18424984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18424984&pc=OZD | 12/06/2023 | 10/02/2025 |  | 01/02/2024 |  |  | 12/06/2023 | 01/02/2024 |  |  | 04/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355004 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | Appropriate Term / Code Not Available | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/06/2023 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 60-year-old white male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was delivered and came out of position and had lost the position signal. the pump had to be removed due to the lost position. the patient later expired and no allegation was made that the pump caused/contributed to the outcome. | The complainant also reported that a placement signal not reliable alarm was noted. |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-05676 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. pump was not returned for investigation. as per the clinical information provided, the pump was pulled into aorta while moving the patient. therefore, the root cause of the positioning issue was patient movement related. data logs were investigated. signal not reliable went to 0, contrast decreased, lamp and gain increased and light was stable. the 5s parameters were similar to that of fiber break, however the location of the fiber break cannot be determined because the pump was not returned. the root cause of the optical signal issue is most likely due to a damaged fiber line. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-05676. b2 death date updated. d6a/d6b updated with additional information. f6 and f8 should not have had entries in the initial report. g1 reporting contact email updated.
32050 | 1220648-2024-05677 | 18425044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18425044&pc=PYX | 12/11/2023 | 10/28/2025 |  | 01/02/2024 |  |  | 12/11/2023 | 01/02/2024 |  |  | 01/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024324989 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 54-year-old white female with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during the impella rp procedure, the physician observed that hemostasis was not obtained after the 23fr peel away sheath was inserted, and the dilator was taken out. a new 23fr sheath was placed and hemostasis was achieved and impella rp procedure continued. | The complainant reported that during the impella rp procedure the physician observed that hemostasis was not obtained after the 23fr peel away sheath was inserted, and the dilator was taken out. a new 23fr sheath was placed and hemostasis was achieved and impella rp procedure continued. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | The investigation for the introducer damage - mechanical issue has been completed. the introducer was returned for investigation. the sheath was peeled away after the hemostasis valve leading to a gap through which the blood can easily flow out. therefore, the root cause of the introducer damage issue was use related to improper technique.
32051 | 1220648-2024-05678 | 18425068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18425068&pc=OZD | 12/07/2023 | 02/22/2024 |  | 01/02/2024 |  |  | 12/07/2023 | 01/02/2024 |  |  | 08/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385431 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 65 year-old male with acute myocardial infarction/cardiogenic shock. the complainant reported that a 14fr sheath was used in the femoral artery for delivery of the impella cp. the limb became ischemic with the need for an external fem-fem bypass to perfuse the limb. the pump remained on for 5 days until it was successfully weaned and explanted. no additional information was provided. | [5 empty] | The investigation into the limb ischemia ¿ irreversible issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not returned for evaluation. when the investigation has been completed, a supplemental report will be submitted. the instructions for use states the following related to limb ischemia: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32052 | 1220648-2024-05679 | 18425177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18425177&pc=OZD | 12/12/2023 | 12/23/2024 |  | 01/02/2024 |  |  | 12/12/2023 | 01/02/2024 |  |  | 07/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SET AU |  |  | ABIOMED, INC. |  |  | US |  | 2024375511 |  | 1000361 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in the au reported the patient was on impella 5.5 support due to ischemic cardiomyopathy. the impella had a tuohy borst failure. staff could not secure the catheter and the impella was migrating. support was continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. no product was returned for evaluation. unable to check toughy. the cause of the positioning issues could not be determined due to lack of clinical details and no product returned. corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email. h5 should have been yes.
32053 | 1220648-2024-05680 | 18425207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18425207&pc=OZD | 12/07/2023 | 02/28/2024 |  | 01/02/2024 |  |  | 12/07/2023 | 01/02/2024 |  |  | 08/31/2024 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270793 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old white male noted as high-risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient on impella cp experienced access site bleeding. the patient was on impella support for 83.72 before successfully being weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. | The investigation into the access site bleeding - major has been completed. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
32054 | 1220648-2024-05681 | 18425415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18425415&pc=OZD | 12/08/2023 | 03/04/2024 |  | 01/02/2024 |  |  | 12/08/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024413254 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 69-year-old black male was implanted with an impella cp pump presented to the operating room for escalation to the impella 5.5 pump as a bridge to left ventricular assist device (lvad). the impella 5.5 was inserted, turned on, and immediately alarmed ¿impella stopped motor current high.¿ attempts were made to restart the pump which failed. the patient arrested during this time and required two rounds of cardiopulmonary resuscitation (cpr), epinephrine, and was defibrillated once. additional bolus of heparin was given and return of spontaneous circulation was achieved. a new impella 5.5 was inserted successfully. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: pump stop ifu: impella 5.5, ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ arrhythmia ifu: impella 5.5 . nothing beyond study data. the failure modes will be monitored and trended. | The investigation into the pump did not start and the vf/vt/af/at events has been completed since the original report was filed. the device was returned for evaluation, visually inspected, and tested. the cause of the pump did not start issue was not determined since pump stop issue did not reproduce per field investigation and logs indicating high motor current pump stop per log review. any device contacting the heart wall in conjunction with a poor patient condition could result in vf/vt but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the vf/vt is unable to be determined.
32055 | 1220648-2024-05682 | 18425558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18425558&pc=OZD | 12/08/2023 | 05/28/2024 |  | 01/02/2024 |  |  | 12/08/2023 | 01/02/2024 |  |  | 06/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024390176 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/08/2023 |  |  |  |  | Additional information received from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that acute kidney injury with a possible relationship to the impella cp device used on this patient. | The user facility reported a 66-year-old male (unknown race) undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that access was notably low and small perforation seen distal to access. a decision was made to gain contralateral access, remove impella, and balloon superficial femoral artery (sfa). angioplasty was successful. the patient is stable. |  |  |  |  | Section b5 and h6 were updated as per additional information received. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ the failure modes will be monitored and trended. | The investigation into the vascular damage - peripheral vessel and the renal failure issues has been completed since the original report was submitted. the pump was not returned for investigation. based on given clinical details, the cause of the vascular damage issue was use related, due to low entry point. also, acute kidney injury complication was reported during impella use. the cause of the renal failure issue was not determined since product was not returned and sufficient clinical details were not provided.
32056 | 1220648-2024-05683 | 18425734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18425734&pc=OZD | 12/12/2023 | 06/04/2024 |  | 01/02/2024 |  |  | 12/12/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411262 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced access site bleeding requiring two units platelets, two units fresh frozen plasma and one unit packed red blood cells. the physician reported the graft was okay and it was the incision that was bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding issue was not determined since the product was not returned and not much information was provided in the clinical description. d.4 revised unique identifier (udi) # as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report.
32057 | 1220648-2024-05684 | 18425884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18425884&pc=OZD | 12/08/2023 | 04/30/2024 |  | 01/02/2024 |  |  | 12/08/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024405970 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support and after the impella insertion, the patient experienced access site bleeding. manual compression and sutures were attempted, but the bleeding could not be stopped. the impella was removed and a medikit 16fr sheath was inserted. a new impella cp was inserted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added-d6a/d6b. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32058 | 1220648-2024-05685 | 18426383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18426383&pc=OZD | 12/08/2023 | 12/11/2024 |  | 01/02/2024 |  |  | 12/08/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401827 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 37-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the repositioning sheath of an implanted impella cp pump was occlusive in the artery during patient support. an external fem-fem bypass was performed to resolve the condition and support continued with the implanted pump. attempts to reposition the pump resulted in the pump backing out across the aortic valve and the decision was made by the physician to leave the pump deep at 5.6 cm. due to the deeper positioning, the patient began to experience hemolysis. the physician declined further attempts at repositioning and the hemolysis remained present until the pump was weaned and explanted the following day. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Product not returned. data logs show suction and various positioning alarms. no motor current spikes outside of p level changes. the cause of the hemolysis was most likely improper positioning of the device but the cause of the positioning was unknown due to lack of clinical information. as all the necessary information was not provided, the root cause of the limb ischemia was not determined." section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05685: d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email, g1 reporting contact fax number missing.
32059 | 1220648-2024-05686 | 18426398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18426398&pc=OZD | 12/13/2023 | 09/20/2024 |  | 01/02/2024 |  |  | 12/13/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024409534 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | Ischemia Stroke | [11 empty] | Other | Life Threatening | Death | Disability | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during impella 5.5 support the patient suffered a stroke. it was determined to be an ischemic event on the left side of the brain and a hemorrhagic event on the right side of the brain. it was unknown if the injury was related to the impella support. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the device was not returned for investigation. as the clinical information was not provided, the root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
32060 | 1220648-2024-05687 | 18426689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18426689&pc=OZD | 12/12/2023 | 05/24/2024 |  | 01/02/2024 |  |  | 12/12/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414374 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported an impella cp in a 86yo white male for a non st elevated myocardial infarction. the site developed a hematoma and prompted vascular repair and blood products. the patient remained on support during the percutaneous coronary intervention. once the cp was turned off and care was withdrawn, the patient expired. death was not attributed to the use of impella. no additional information was provided. | The patient's outcome of death after being withdrawn from the pump has been reviewed by abiomed's medical safety officer. it was determined that there is not enough information to exclude use of device as a cause of the death. |  |  |  |  | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. the instructions for use states the following for access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure mode will be monitored and trended. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related as the doctor cinched the 2 perclose sutures that had been placed prior to impella insertion to be tighter around the repositioning sheath as per the clinical description provided. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b2-selected death b5-added medical safety officer association with the device h1-updated type of reportable event in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32061 | 1220648-2024-05688 | 18426999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18426999&pc=OZD | 12/09/2023 | 10/27/2025 |  | 01/02/2024 |  |  | 12/09/2023 | 01/02/2024 |  |  |  |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1095485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Battery Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 12/09/2023 |  |  |  |  | The user facility reported a (b)(6) male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the automated impella controller (aic) experienced battery failure during patient support. the aic was swapped as a result of the noted issue. there was no patient harm. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-05688 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation of automated impella controller (aic) ¿ battery failure (red alarm) has been completed. the impella device was returned for evaluation. based on the data and device analysis the root cause of the defective battery was most likely due to cell imbalance, a known pattern failure. d2 common device name revised on manufacturer device report 1220648-2024-05688 in accordance with updated procedures. f6/f8- should have been left blank on the initial manufacturer device report number 1220648-2024-05688 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-05688 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-05688. h6 health effect - clinical code 4445, health effect - impact code 1802, medical device problem code 2993, and component code 925 was erroneously entered on the initial manufacturer device report number 1220648-2024-05688-1. the patient outcome of death was captured in manufacturer device report number 1220648-2025-31880.
32062 | 1220648-2024-05689 | 18427078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18427078&pc=OZD | 12/14/2023 | 06/11/2024 |  | 01/02/2024 |  |  | 12/14/2023 | 01/02/2024 |  |  | 01/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329566 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a (b)(6) female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. a kink was noted on fluoroscopy proximal to the outlet. this kink made it unable to get flows above p1 without suction. the impella was removed and a new one placed successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. the device was not returned for a visual inspection. data log analysis confirmed the presence of suction alarms and low flows. the most likely cause of the low pump flow was a cannula kink. d6a, d6b.
32063 | 1220648-2024-05690 | 18427097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18427097&pc=OZD | 12/10/2023 | 01/02/2024 |  | 01/02/2024 |  |  | 12/10/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024415033 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported a (b)(6) white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that flow saturation was noted. saturation appeared shortly following a purge cassette change which took beyond the 90 recommended seconds. the following day, the patient was taken to the operation room and the impella 5.5 was exchanged for a new impella due to concern for pump failure. | [5 empty] | The product was returned for investigation. no signs of physical abnormalities were found on returned pump. upon further investigation, biomaterial was observed on the impeller after soaking it in warm water for 15 minutes. data logs shows the saturated pump flow after the 180second long cassette change procedure while running on p7. the cause of the saturated flow issue was biomaterial found on the impellor. the failure modes will be monitored and trended.
32064 | 1220648-2024-05691 | 18427518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18427518&pc=OZD | 12/14/2023 | 01/17/2024 |  | 01/02/2024 |  |  | 12/14/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024401675 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 42-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was inserted in normal fashion but was unable to advance beyond the subclavian during attempted delivery. after multiple unsuccessful attempts, the decision was made to abort the impella 5.5 implant and plans were made to insert an impella cp for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition.
32065 | 1220648-2024-05692 | 18427569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18427569&pc=OZD | 12/11/2023 | 09/26/2024 |  | 01/02/2024 |  |  | 12/11/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409113 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 66-year-old white female undergoing cardiac surgery was implanted with an impella 5.5 device. the complainant reported the purge arm of an imeplla 5.5, between the purge filter and the purge reservoir was broken. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ the failure modes will be monitored and trended. | The investigation for purge leak ¿ pump has been completed. the cause of the issue was a leak from the sidearm but the exact location of the leak was unknown since no product was returned. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05692: d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing.
32066 | 1220648-2024-05693 | 18427768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18427768&pc=OZD | 12/11/2023 | 02/23/2024 |  | 01/02/2024 |  |  | 12/11/2023 | 01/02/2024 |  |  | 08/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024395928 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following explant of an impella cp pump from the right femoral artery, no flow was present down the distal right limb. a balloon angioplasty was utilized and flow was restored in the limb. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32067 | 1220648-2024-05694 | 18427810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18427810&pc=PYX | 12/10/2023 | 01/29/2024 |  | 01/02/2024 |  |  | 12/10/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024418539 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported a 66-year-old female who underwent cardiac surgery was implanted with an impella rp device for mechanical circulatory support. the patient developed hemolysis with decreased pump flows. an echocardiogram (echo) was done and visualized what appeared to be a clot on the inlet. the impella was exchanged. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, two days prior to the hemolysis heparin was stopped. data logs show a motor current spike with placement signal shift followed by low pump flows at p8. adjusting to p9 did not stop low pump flow. the most likely cause of the hemolysis was biomaterial.
32068 | 1220648-2024-05695 | 18428076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18428076&pc=OZD | 12/05/2023 | 09/24/2024 |  | 01/02/2024 |  |  | 12/05/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411506 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 34-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was experiencing acute pain and hematoma at the impella access site. manual pressure was held. the patient was brought back to the or and the physician repaired graft which was oozing. the patient is stable and remains on the impella support, now on day 31. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05695. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures.
32069 | 1220648-2024-05696 | 18428153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18428153&pc=OZD | 12/05/2023 | 04/30/2024 |  | 01/02/2024 |  |  | 12/05/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024409906 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 51-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. an echocardiogram was performed, and a large apical/septal left ventricle thrombus was noted. the impella pigtail did not interact with the thrombus. the patient reported back to the catheterization lab several days later and had a sentinel device placed. the impella was removed with no complications. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states in impella cp with smartassist catheter insertion-keep act = 250 seconds achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup. added- d6a/d6b in accordance with updated procedures, sections d6, g3, and h10 have been revised from the initial submission.
32070 | 1220648-2024-05697 | 18428402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18428402&pc=OZD | 12/06/2023 | 01/02/2024 |  | 01/02/2024 |  |  | 12/06/2023 | 01/02/2024 |  |  |  |  | 01/02/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1673731 |  | 00813502011814 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission |  |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the automated impella controller (aic) had suction alarms due to an abnormality in the optical sensor. the pump position was confirmed. the aic was replaced, and this resolved the issue. no known patient harm or injury. | [5 empty] | Device analysis was possible as the controller was returned for benchtop analysis. the failure mode seen in the field was reproduced by field service. root cause was a defective optical bench.
32071 | 1220648-2024-05698 | 18428422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18428422&pc=OZD | 12/15/2023 | 01/18/2024 |  | 01/02/2024 |  |  | 12/15/2023 | 01/02/2024 |  |  | 07/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024380538 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that resistance was met during attempted delivery of an impella 5.5 pump. the patient became hypoxic during delivery and the ecmo (extracorporeal membrane oxygenation) was subsequently implemented. as the pump was unable to traverse the innominate, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely due to patient condition.
32072 | 1220648-2024-05699 | 18428440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18428440&pc=OZD | 12/07/2023 | 10/03/2024 |  | 01/02/2024 |  |  | 12/07/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024407764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported an 81-year male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp was inserted via the femoral artery and the site developed a hematoma. the patient was taken to the operating room (or) for treatment of the hematoma using vascular repair. | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use issue since adding stitches resolved the issue. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-05699 b5 patient outcome and mso statement added d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32073 | 1220648-2024-05700 | 18428463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18428463&pc=OZD | 12/09/2023 | 01/17/2024 |  | 01/02/2024 |  |  | 12/11/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024395296 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that the impella cp could not pass through the vascular system due to the patient¿s anatomy. the case was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related.
32074 | 1220648-2024-05701 | 18428922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18428922&pc=OZD | 12/07/2023 | 01/17/2024 |  | 01/02/2024 |  |  | 12/07/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024397100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump could not be advanced into position during attempted delivery due to a pre-existing axillary/aorta dissection from a previous procedure. the implant was aborted as a result of the noted difficulty. there was no patient harm due to the impella cp complication. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32075 | 1220648-2024-05702 | 18429197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18429197&pc=OZD | 12/08/2023 | 10/31/2025 |  | 01/02/2024 |  |  | 12/12/2023 | 01/02/2024 |  |  |  |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1198431 |  | 0042-0000-CA |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) displayed an optical sensor system error message upon connection of the impella catheter. the aic was swapped as a result of the noted issue. there was no patient harm. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: the complaint occurred date per salesforce is dec 8th, 2023; the last case performed at the customer site with the commercial pump (sn: 428629/7) was on dec 7th, 2023. the console received ¿placement signal not reliable (opm invalid signal, optical pump connected]) ¿ alarm¿, event set #1036; after 2 minutes and 10 seconds since pump connection, following, the console displayed ¿optical sensor system error: disconnect cable¿ this confirms the reported failure mode. rt log shows that the raw optical sensor was at ¿-32768,¿ and the snr was low. upon reviewing the previous cases since mar 2nd, 2023, in one of the cases after the demo pump connection, the console received the psnr alarm on may 24th, 2023 (this issue was not reported). per salesforce, the pump (sn: 428629/7) has been moved to the spare console ic3802; no logs are available in the constellation for review, and no complaint has been recorded in the salesforce against this spare console. on apr 11th, 2023, the imc log recorded the ¿purge system blocked ¿ alarm¿ event set # 408. (this issue was not reported and is unrelated to the reported failure.) device analysis: this console was tested after arriving in fs and reproduced the psnr alarm by partially connecting the white plug to the front panel optical connector (please see the attached picture). event # 1036, which occurred twice on different dates, was most likely due to the air gap. the light parameter was abnormal (6.3 v) after obtaining the 5s parameter from the optical cavity test. the front panel optical connector has debris as received, and even after cleaning, it is unable to remove debris completely. removed the ¿power button¿ label, which prevents opening the front panel optical connector cover. this label is not part of the ¿assembly procedure, aic console¿ (0042-700). if the user has the same issue, they might connect the white catheter plug partially with an air gap. found a hair crack on the purge retainer (unrelated to the reported failure mode). a mild fluid ingress was found between the motor spline gasket and the rear of the purge cassette insert (unrelated to the reported failure mode). root cause: the root cause for the optical sensor system error message was most likely due to an air gap. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d2 common device name updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated.
32076 | 1220648-2024-05703 | 18430013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18430013&pc=OZD | 12/13/2023 | 04/15/2025 |  | 01/02/2024 |  |  | 12/13/2023 | 01/02/2024 |  |  | 08/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024397751 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported a 39-year-old male in acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp pump experienced a drop in flow rate during patient support. as the position was verified to be correct, the decision was made to replace the pump. a new impella cp device was replaced for ongoing support. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of low pump flow has been completed. the visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump ran without issue under bench test. data logs were investigated and there was a suction alarm followed by low flow of 02l/min at p-7. there was a slight dip in the placement signal which corrected right away. no position alarms or other device was in use parallel to impella. clinical information provided is limited, but not blood was given and patients vessel health is unknown. therefore, as per the product and data log evaluation, the root cause of the low flow issue was most likely patient condition related. d1 brand name was erroneously entered on the initial manufacturer device report number 1220648-2024-05703. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-05703. d6a and d6b revised on the initial manufacturer device report number 1220648-2024-05703 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-05703 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-05703 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-05703.
32077 | 1220648-2024-05704 | 18430110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18430110&pc=OZD | 12/15/2023 | 01/17/2024 |  | 01/02/2024 |  |  | 12/15/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024394410 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported a patient had an attempted impella placement. however, it was unsuccessful due to the patient's anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition.
32078 | 1220648-2024-05705 | 18430651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18430651&pc=OZD | 12/14/2023 | 06/17/2024 |  | 01/02/2024 |  |  | 12/14/2023 | 01/02/2024 |  |  | 09/30/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410343 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 75 year-old female for coronary artery bypass graft vs percutaneous coronary intervention. after placement of the impella cp, no pulse was found in the impella leg. an external fem-fem bypass was subsequently placed to restore flow to the lower right extremity. the following day it was noted that the bypass had clotted off and flow was again blocked. the retrograde sheath was flushed and the antegrade sheath was replaced to restore flow to the leg. no additional information was provided. | The patient expired while on support. upon review by abiomed's medical safety officer, there is not enough information to associate the use of device with the patient's outcome. |  |  |  |  | The impella device was not returned for evaluation. a supplemental report will be submitted upon completion of the investigation. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia - reversible and thrombus issues has been completed. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia and the thrombus issues was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 reporting contact facility name and reporting contact fax number were inadvertently left off the previously submitted report and was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
32079 | 1220648-2024-05706 | 18430754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18430754&pc=OZD | 12/12/2023 | 10/20/2025 |  | 01/02/2024 |  |  | 12/14/2023 | 01/02/2024 |  |  | 10/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and that there was a sensor issue. good positioning was unable to be achieved despite several attempts. left ventricular was ranging from -10 to -70 and was unable to be resolved. repositioning failed and the pump was pulled back into aorta. the team decided to use a new pump to ensure proper management due to unreliability of the sensor and the patient¿s condition. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. positioning issues: the clinical mentions that the pump was placed shallow and tried to be repositioned. data logs show multiple pump in aorta alarms which are confirmed in clinical saying the pump was shallow. repositioning failed multiple times and information about the repositioning is unknown. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. optical signal issues per the clinical, intermittent placement signal low alarms were noticed a few minutes after the placement of impella cp. this aligns with the observed 'placement signal low alarms' on the lt log. upon examining the lt logs these placement signal low alarms were observed for a period of 2 hours and 13 minutes which were later resolved. there were no spikes in the motor current and the values are in agreement with the p - level. the observed plots of the placement signal do not correlate that of the pattern that occurs during the sensor wear. since, the pump was explanted the alleged symptoms of sensor wear were not noticed. the root cause of the optical signal issues was not determined since product was not returned and insufficient clinical information was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. b7 other relevant history, including preexisting medical conditions updated. d4 model number updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 mdr reporting contact name and address updated.
32080 | 1220648-2024-05707 | 18431119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18431119&pc=OZD | 12/04/2023 | 02/23/2024 |  | 01/02/2024 |  |  | 12/04/2023 | 01/02/2024 |  |  | 07/31/2025 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391201 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was not getting flow to the left lower extremities. the impella repo sheath was thought to be the cause. a reprofusion sheath was placed over night. the patient is stable. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the hemolysis issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32081 | 1220648-2024-05708 | 18431632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18431632&pc=OZD | 12/12/2023 | 02/11/2025 |  | 01/02/2024 |  |  | 12/12/2023 | 01/02/2024 |  |  | 06/15/2023 |  | 01/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1737779 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 74-year-old asian female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella cp support when the automated impella controller (aic) was not passing priming. even after replacing the purge cassette, priming did not proceed and was not completed, so insertion was discontinued. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | The investigation for the console issues has been completed. data logs from the day of the complaint show that after first purge cassette was inserted into console, purge priming was never completed. the same issue was noticed after second purge cassette was inserted. in both cases, the priming doesn't progress passed the luer detection phase of the wizard. real time logs show no stabilizing of purge pressure when at 90% priming waiting for luer connection. console was returned for evaluation. a known good pump and purge line was run for more than an hour without any anomaly. the failure couldn't be reproduced and investigation revealed the console priming to completion as expected.
32082 | 1220648-2024-05664 | 18423152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18423152&pc=OZD | 12/06/2023 | 05/06/2024 |  | 01/01/2024 |  |  | 12/06/2023 | 01/01/2024 |  |  | 07/31/2025 |  | 01/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391344 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced ventricular tachycardia and impella cp was removed. volume was given as the physician felt that the patient needed volume more than anything. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use for impella cp with smart assist system in section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- d6a/d6b d4-updated model number.
32083 | 1220648-2024-05665 | 18423164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18423164&pc=OZD | 12/07/2023 | 02/21/2024 |  | 01/01/2024 |  |  | 12/07/2023 | 01/01/2024 |  |  | 09/30/2025 |  | 01/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402196 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported an impella cp in a 54 yo male with acute myocardial infarction/cardiogenic shock. it was reported that the pump had ¿high purge pressure¿ alarms and there was an eventual pump stop. the impella was exchanged for another cp. no additional information was provided. | [5 empty] | The impella device has been returned for evaluation. however, it has not yet been evaluated. a supplemental report will be submitted when the investigation has been completed. the instructions for use states the following: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ the failure modes are being monitored and trended. | The investigation into the high purge pressure and the pump stop issues has been completed since the original report was filed. the pump and purge cassette were returned for investigation. the pump was visually inspected, and biomaterial ingestion was observed in the purge gap at outflow. purge cassette was setup for priming and pressurized by connecting to the pump and purge fluid leak was observed on sidearm yellow luer lock. the cause of the pump stop issue was due to biomaterial ingestion per field investigation and log review. the cause of the purge leak was sidearm yellow luer damage most likely occurring during use due to the high tensions in combination with disinfectants and certain drug ingredients.
32084 | 1220648-2024-05666 | 18423180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18423180&pc=OZD | 12/07/2023 | 02/23/2024 |  | 01/01/2024 |  |  | 12/07/2023 | 01/01/2024 |  |  | 09/30/2025 |  | 01/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408529 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 66-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the patient had no pulse in their impella limb during support. as the patient required extended support, the decision was made to explant the pump and escalate to impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the limb ischemia ¿ reversible issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32085 | 1220648-2024-05667 | 18423202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18423202&pc=OZD | 12/08/2023 | 02/12/2024 |  | 01/01/2024 |  |  | 12/08/2023 | 01/01/2024 |  |  | 01/31/2025 |  | 01/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024331012 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Handling Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and the physician noticed a kink in the cannula in the ascending aorta. the impella had lower than normal flows at 2.3. the physician decided to continue support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into product damage (cannula kink) has been completed since the original report was filed. the product was returned for investigation. visually inspected and no biomaterial ingestion was found, but the cannula kink was observed. data logs were reviewed, and low pump flow was observed immediately after the start until entire run. no motor current spikes or purge related issues were observed. the cause of the low pump flow issue was patient condition related with cannula kinking due to patient anatomy per clinical, data log review and field investigation.
32086 | 1220648-2024-05668 | 18423220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18423220&pc=OZD | 12/08/2023 | 10/24/2024 |  | 01/01/2024 |  |  | 12/08/2023 | 01/01/2024 |  |  | 09/30/2025 |  | 01/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402102 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [5 empty] | 6F X 25CM TERUMO GLIDESHEATH | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported an impella cp in a 62yo female for high risk percutaneous coronary intervention. the cp was placed via single access at the femoral. the access caused a vessel perforation that was treated with ballooning via contralateral access. the team was eventually able to reach hemostasis. no additional information was provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. the instructions for use has the following caution statement for vascular dissection: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ the failure mode is being monitored and trended. | The investigation for the reported vascular damage has been completed. the device was not returned for evaluation. however clinical details were sufficient to determine the root cause. the root cause of vascular damage was most likely due to use issue since non-abiomed recommended product (sheath) was used. added- d10 concomitant device, h5 inadvertently left blank. updated to yes., h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32087 | 1220648-2024-05669 | 18423238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18423238&pc=OZD | 12/10/2023 | 02/12/2024 |  | 01/01/2024 |  |  | 12/10/2023 | 01/01/2024 |  |  | 04/30/2025 |  | 01/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333578 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was on impella 5.5 support after having an impella cp implant. the impella 5.5 experienced a pump stop. the automated impella controller (aic) was replaced to no avail. impella 5.5 was removed to address the issue. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed. the device was not returned for investigation. based on the clinical information provided, fifteen days into impella 5.5 a pump stop occurred. no information was given as to how the pump stop occurred. multiple restart attempts and a console switch were attempted with no success, the pump was explanted as a result. log analysis confirms alarm 115 with a motor current spike greater than 30000. this is common in open motor cable line failures. the most likely cause of the pump stop was an open electrical line. h1: type of report should be malfunction and not serious injury.
32088 | 1220648-2024-05670 | 18423248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18423248&pc=OZD | 12/11/2023 | 07/17/2024 |  | 01/01/2024 |  |  | 12/11/2023 | 01/01/2024 |  |  | 04/30/2025 |  | 01/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353508 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for support of a 60-year-old male patient admitted in cgs and possibly in need of biventricular support. the pump had positioning issues. the team attempted repositioning but despite all measures the repositioning failed. the team left the pump in it's mispositioned location despite the patient having urine color changes. | It was further reported that the after the urine changes, echo repositioning was performed and the pump was pulled back to 4.2-4.4cm. following this procedure, the patient's urine began to clear. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿. | The investigation for the reported hemolysis and the positioning issue has been completed. according to the clinical, the pump was located too deep. no product was returned for a visual inspection. the logs confirm positioning issues, suction, and placement signal trends. the most likely cause of the positioning issue is use. according to the clinical, the physician elected to leave the pump in improper position and blood was noted in the patient's urine the next day. echo repositioning was then performed and the pump was pulled back. the patient's urine cleared in the following days. the most likely cause of the hemolysis is improper positioning. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the section: warnings & cautions: cautions. the ifu further states: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ added- b1-checked adverse event, b2-checked required intervention, b5-event narrative, h6 clinical code 1886, h6 impact code 4628 added- d6a/d6b. h1-changed to serious injury.
32089 | 1220648-2023-05589 | 18416932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18416932&pc=PYX | 12/03/2023 | 11/04/2025 |  | 12/29/2023 |  |  | 12/03/2023 | 12/29/2023 |  |  | 03/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2024394944 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Pumping Stopped | Device Sensing Problem | Improper Flow or Infusion |  |  |  | Hemolysis | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [10 empty] | Other | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/03/2023 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on bipella support and the impella rp flex had two pump stoppages. the physician inserted the sheath further in and sutured it in place. there were no further drops in flows or stoppages reported. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The complainant also reported the pulmonary artery numbers were reading negative, flows very high for p-6. the pump was repositioned, and it successfully fixed the flows. also, the patient had red urine. 500 ml of volume was given to the patient. x-rays were performed, and the pump appeared deep, the pump was repositioned, and another 500 ml of volume were given to the patient after repositioning. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05589 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. temporary pump stop: the returned pump was inspected and there were no visual abnormalities or biomaterial around the impeller. the pump passed electrical testing. the pump stop was not reproduced after running in the lab. the data logs show alarm 13 - impella stopped motor current high alarms with motor current and speed deviations characteristic of an ingestion. the root cause of the temporary pump stop was not determined since the issue was unable to be reproduced. placement signal issue: the pump was inspected and there were no damages to the dp sensor. there was no biomaterial by the impeller. the pump was run in the flow loop and sensor drift reproduced. the logs show a sensor drift pattern with placement signal increasing while flows decreased with no p-level change or motor current spikes. the root cause of the placement signal issue was determined to be dp sensor drift as evidenced by log pattern and returned product analysis. hemolysis: the root cause of the hemolysis was most likely due to improper positioning of the device as the hemolysis resolved after repositioning both devices. positioning issues: flex was noted to have moved slightly deeper than the day before. the cause of the positioning issues was not determined as limited clinical information was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b2 updated. b5 updated to include additional failure modes. d9 updated. h1/h2 updated. h3 updated. h6 updated. b7 other relevant history, including preexisting medical conditions updated. d4 model number updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated.
32090 | 1220648-2023-05608 | 18416979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18416979&pc=PYX | 12/02/2023 | 01/18/2024 |  | 12/29/2023 |  |  | 12/02/2023 | 12/29/2023 |  |  | 07/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024374144 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2023 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella rp device for mechanical circulatory support. the team experienced difficulty attempting to deliver the impella rp, due to patient's anatomy and enlarged right ventricle/right atrium. delivery was ultimately aborted and the patient was scheduled for transfer to an outside hospital for further support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely due to patient condition due to enlarged rv/ra.
32091 | 1220648-2023-05610 | 18418293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418293&pc=OZD | 11/30/2023 | 01/30/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1640800 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) did not recognize a new purge disc. the aic was exchanged. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. per data review, there was no record of any purge disc not detected alarms on the reported occurrence date. however, there were two purge disc not detected alarms found in the imc logs. these two alarms align with what the clinical staff reported because they occurred immediately after swapping purge cassettes in both instances. the console was returned and tested after arriving. the issue with the automated impella controller (aic) not being able to detect the purge disc was reproduced, and the purge retainer was found to be broken. a portion of the purge retainer that was supposed to be fastened to the purge assembly was no longer intact. the cause of the issue was a broken purge retainer.
32092 | 1220648-2023-05611 | 18418316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418316&pc=OZD | 11/30/2023 | 05/09/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403866 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 13-year-old female was implanted with an impella 5.5 as a bridge to transplant. a hematoma was noted on the impella access site. an evacuation was performed and heparin was restarted the following morning. the patient successfully received an orthotopic heart transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32093 | 1220648-2023-05612 | 18418345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418345&pc=OZD | 12/03/2023 | 03/25/2024 |  | 12/29/2023 |  |  | 12/03/2023 | 12/29/2023 |  |  | 10/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410034 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 41-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed a large hematoma in their right groin during impella cp support. two units of platelets were given to treat the condition and the decision was made to explant the pump. an impella 5.5 pump was subsequently implanted for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the hematoma appeared soon after the implant and starting act was 340 which is more than the limit.
32094 | 1220648-2023-05613 | 18418375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418375&pc=OZD | 12/01/2023 | 09/23/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024403703 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 12 MO |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 1 year old patient of unknown race and gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was explanted due to the patient bleeding and hemoglobin decreasing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- h6 component code 925 section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial reports: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank. g1-mfr site fax number.
32095 | 1220648-2023-05614 | 18418388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418388&pc=OZD | 12/01/2023 | 11/29/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 06/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354392 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump came back across the valve during an unrelated defibrillation. after attempts to re-advance the pump were unsuccessful, the decision was made to replace the component. a new impella cp pump was placed for ongoing support. there was no patient harm related to the positioning issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was use related to improper technique. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05614 d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing.
32096 | 1220648-2023-05615 | 18418427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418427&pc=OZD | 12/01/2023 | 10/30/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404643 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a pseudoaneurysm was discovered in the left superficial femoral artery during impella cp support. the pump was subsequently explanted as a result of the noted condition. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 corrections to the initial mfr report: f6/f8 should have been left blank.
32097 | 1220648-2023-05616 | 18418431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418431&pc=OZD | 12/03/2023 | 01/23/2024 |  | 12/29/2023 |  |  | 12/03/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1426097 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/03/2023 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had impella cp support for the patient admitted pre-percutaneous coronary intervention (pci¿). the team observed an automated impella controller (aic) alarm. the controller alarm prompted the team to replace the aic. no harm or injury noted. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - loss of opm data has been completed since the original report was filed. the console was returned for investigation and was tested upon return. console logs from the reported day of event are consistent with complaint and show two instances of controller error alarm and three more instances of event 1045 (with no corresponding alarm due to duration < 2 minutes). it's been determined that this issue is a known pattern failure caused by a temporary loss of optical placement signal. the cause of the issue was not determined since the issue could not be reproduced.
32098 | 1220648-2023-05617 | 18418446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418446&pc=OZD | 12/04/2023 | 04/29/2024 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402859 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 79-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the pump was placed but there were persistent alarms for suction. the team attempted to troubleshoot. the pump was placed in a smaller patient with a small aortic arch. the pump cannula had kinked. the pump was repositioned and supported on only p4 till weaned and explanted. no patient harm was reported due to the issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the product damage (cannula kink) and the low pump flow issues has been completed since the original report was filed. the pump was returned and evaluated. there was a large visible kink on the cannula. it was stated that the patient had a very small aortic arch, and the motor housing was caught up in the arch due to its small size. the root cause of the low flow is due to the kink in the cannula. to conform to updated procedures, section d6 has revised with updated information.
32099 | 1220648-2023-05618 | 18418447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418447&pc=OZD | 12/04/2023 | 10/01/2024 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  | 12/31/2024 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315958 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 40-year-old male patient was on impella cp support for st elevated myocardial infarction. a large left ventricle thrombus was noted on echo and heparin was increased. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05618. d6a revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing h6 component code revised from the initial submission in accordance with updated procedures.
32100 | 1220648-2023-05619 | 18418636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418636&pc=OZD | 11/30/2023 | 11/29/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024413009 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Break | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 83-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported that during an impella cp implant, the tip of the dilator was found to be damaged. the 14f sheath and dilator were already placed when it was discovered that the tip of the dilator split. the impella was implanted successfully despite the dilator damage. additionally, after the impella was implanted, there was access site bleeding that was resolved with mattress sutures. reportable due to damage to the dilator tip and access site bleeding that required intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. | The investigation for the introducer damage and the access site bleed has been completed. pump was not returned for investigation. clinical mentions that the patient had calcified vessels. therefore, the root cause of the introducer damage mechanical issue was most likely patient condition related to calcified vessels. the root cause of the access site bleeding issue was patient condition related to calcified vasculature. added h6 component code 925 corrections to the initial mfr report: d3 mfr name and address. d4 model number. added d6a/d6b. f6/f8 should have been left blank. g1-mfr site name and address.
32101 | 1220648-2023-05620 | 18418638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418638&pc=OZD | 11/30/2023 | 03/11/2025 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400789 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 60-year-old male patient was on impella cp support for right coronary artery (rca) occlusion. patient's pressures began to drop. a pericardial effusion was found, 1000ml of blood volume removed during pericardiocentesis. the impella was turned off. the patient had a tamponade, visualized under echo. physicians confirmed there was a possible perforation to the left ventricle or to the pda during intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. | Pump was not returned for investigation. clinical information doesn't mention the reason for the perforation and no other supporting information is available to investigate the failure further. therefore, the root cause of the ventricle perforation issue was not determined since product was not returned and insufficient clinical information provided.
32102 | 1220648-2023-05621 | 18418644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418644&pc=OZD | 12/01/2023 | 12/09/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401750 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 67 -year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a leak was observed at the yellow lure lock connection sites. there was no alarms or abnormal purge pressure or flow values, so a decision was made to continue with impella support until explant the following day. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. the failure modes will be monitored and trended. | The investigation for the purge leak has been completed. pump and purge cassette (pc) were returned for investigation. the pump and pc were de-aired and then primed for over 10 minutes but the leak was not reproduced. some biomaterial was observed in cannula and purge pressure was 750mmhg and purge flow of 4ml/hr which is a clear indication of no leak. therefore, the root cause of the purge leak- pump issue was use related to improper technique. the instructions for use for the impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock states the following: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ added h6 component code 925 corrections to the initial mfr report: added d6a/d6b g1 mdr reporting contact email f6/f8 should have been left blank. h5-should have been yes. h6 impact code 4627.
32103 | 1220648-2023-05622 | 18418650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418650&pc=OZD | 12/01/2023 | 05/01/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a patient of unknown age and gender was on impella support when the automated impella controller (aic) had a controller error (yellow alarm). the console removed and sent back for investigation. there was no patient harm reported. no additional information was provided. | [5 empty] | The aic device has been returned for evaluation but has not yet been evaluated. the investigation is in process. a supplemental report will be submitted upon completion of the investigation. failure modes will be monitored and trended. | The investigation for the console (aic) yellow alarm has been completed. the console was returned for evaluation. upon inspection, the console had an abnormally high current draw by the pmd, preventing adequate voltage delivery from the pbm to the pmd, triggering a controller error. the compact flash cards in the 4-in-1 were also determined to be defective/corrupted. the data logs were reviewed which confirmed the yellow alarm. a controller error had been triggered due to the voltage delivered to the pmd from the pbm being out-of-spec while running a demo pump. there was also an impella stopped alarm due to a mismatch in the phase currents delivered to the impella. the right data logs show increased motor current at the time of the controller error (pmd voltage out-of-spec). the root cause of the reported complaint was determined to most likely have been a defective demo impella, however this was not confirmed given the demo pump was not returned.
32104 | 1220648-2023-05623 | 18418667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418667&pc=OZD | 11/30/2023 | 10/11/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 08/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397900 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Peeled/Delaminated | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the 0.18 guidewire was withdrawn from the impella after the device was implanted in the left ventricle the guidewire appeared to have delaminated with the center cable being removed and a long section remaining. the physician attempted to remove the section remaining, but it was unsuccessful. the device was explanted with the remaining piece lodged inside the impella cannula. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the delivery issues has been completed. the pump was returned for investigation and guidewire damage was observed on the returned product. the root cause of the delivery issue was most likely use issue since unraveling product damage. added- h6 impact code 4627. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
32105 | 1220648-2023-05624 | 18418676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418676&pc=OZD | 11/30/2023 | 01/29/2025 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411372 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp for support of a patient admitted in need of a pci (percutaneous coronary intervention) for mvd (mitral valve disease). the team attempted three times to deliver the cp via the long 25cm sheath but failed to deliver and so they replaced it with the short sheath. the team delivered the pump without issues with the short sheath, but 5 minutes into procedure it was discovered the pump was deep. the us complainant attempted to pull back on the pump and reinsert, but was unable. it was then decided that the pci be completed without support. upon completion of the pci it was noted that both impella sheaths were kinked. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issues, introducer damage, and difficulty inserting issues has been completed. pump was not returned for investigation. clinical information provided mentions that the pump was found deep and pulled out into aorta and attempted to re-implant but was unsuccessful. unknown if it was a wireless re-access but imaging was done. therefore, the root cause of the positioning issue was most likely use related to improper technique. at explant, the short introducer was found to have a kink just like the long sheath. as per clinical details, the patient had a 5cm femoral artery which was likely the cause of the kink. therefore, the root cause of the introducer damage - mechanical issue was a introducer sheath kink due to patients long femoral site. during pump insertion through the long sheath, 2 attempts failed and the long sheath was replaced with short one. the sheath had kinks at explant which was likely due to long femoral artery. therefore, the root cause of the difficulty inserting pump through introducer issue was a introducer sheath kink due to patients long femoral artery.
32106 | 1220648-2023-05625 | 18418704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418704&pc=OZD | 11/30/2023 | 12/29/2023 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 05/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024350566 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon attempting an impella cp implant, the guidewire kinked. no excessive force was applied to the guidewire. the guidewire was replaced and the impella implant was successful. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32107 | 1220648-2023-05626 | 18418707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418707&pc=OZD | 12/01/2023 | 09/25/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 08/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386886 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp was placed after the cardiac intervention when patient condition was poor. the patient was transferred to a tertiary center. the groin site of the cp was treated for a bleed with the femostop placement and tightening of the perclose sutures. the team then also placed a 6fr bypass to the right limb/impella limb. after 26 hours the pump was explanted and the patient was escalated to the 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ correction to the initial mdr mfr report: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank.
32108 | 1220648-2023-05627 | 18418805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418805&pc=OZD | 11/30/2023 | 09/23/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 08/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024391521 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/30/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that the impella cp motor waveform suddenly became flat and then pulsed. the pump then stopped. the impella was removed and it was confirmed that the inlet area was blocked by thrombus. | It was noted that echocardiography that was performed prior to impella insertion confirmed there was no left ventricular thrombus prior to impella implantation. after the pump stopped, it was explanted and replaced with a new impella device. it was noted that the patient later expired while on support with the new impella cp device. medical safety review of the event noted use of the original impella cp device that had a pump stop cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05627 was provided. | The investigation for the reported pump stop has been completed. the impella device and data logs were returned for evaluation. upon visual inspection of the pump, biomaterial was observed on canula and inlet cage. no other physical abnormalities were observed. biomaterial cannot be removed from cannula. pump start was successful during test run in bucket of water. during test run, suction alarm was observed with low pump flow. data logs shows the several suction alarm with low pump flow followed by the several pump stop-motor current spike alarm, at the end of the case. therefore, the cause of the pump stop issue was biomaterial, based on returned product and data log review. g(4) revised pma/510(k) number as previously entered incorrectly.
32109 | 1220648-2023-05629 | 18418951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418951&pc=OZD | 11/30/2023 | 01/24/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1359804 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported the automated impella controller (aic) displayed ¿purge fluid not detected¿ error message. the console was exchanged, and the issue resolved. there was no reported patient use or harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge issue ¿ other has been completed since the original report was filed. the console was returned for investigation. upon inspecting the purge motor shaft, there were signs of a glucose ingress. this makes sense with the clinical failure since a glucose build up on the motor could cause the purge motor shaft to be blocked which would prevent the cassette from successfully deairing. however, the failure mode was unable to be reproduced, as they were able to run without issue. the cause of the purge issue was most likely a glucose ingress.
32110 | 1220648-2023-05628 | 18418953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418953&pc=OZD | 11/30/2023 | 11/29/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395676 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support for acute stemi. the impella was pulled back above the aortic valve and unfortunately, they were unable to cross back over the valve to re-initiate support. the decision was made to explant the pump and place another impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. pump was not returned for investigation. clinical mentions that the pump was pulled back across the aortic valve during mv replacement and treating ventricular aneurysm. therefore, the root cause of the positioning issue was use related to improper technique. corrections to the initial mfr report: d4 model number. added d6a/d6b. f6/f8 should have been left blank. h6 impact code 4627 changed to 4628 and 4629.
32111 | 1220648-2023-05630 | 18418972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18418972&pc=OZD | 11/30/2023 | 05/15/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 03/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA PUMP 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024324371 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 78-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. when transferred to the icu bleeding at the graft site was noted, that was managed with a pressure dressing. on the second day of support the external sutures on the device were adjusted due to ongoing bleeding and the patient was transfused with packed red blood cells (prbc) due to a drop in hemoglobin to 7.6. the patient remains on impella pump for continued support and the bleeding issues have resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be use issue because the issue was resolved upon adjusting the external sutures on the device.
32112 | 1220648-2023-05631 | 18419025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419025&pc=OZD | 11/23/2023 | 02/22/2024 |  | 12/29/2023 |  |  | 11/23/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400514 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported on a 63yo male on impella cp for mechanical circulatory support. it was reported that the patient developed ischemia while on support. vascular surgery was consulted was unable to obtain bilateral doppler signals and the patient showed symptoms of pedal mottling. due to the condition, the patient was taken to the operating room for escalation to impella 5.5 support, as well as ecmo decannulation. support continued as planned with the newly implanted impella 5.5 pump. | [5 empty] | The impella device has not yet been returned for evaluation. a supplemental report will be submitted when the investigation has been completed. the instructions for use lists limb ischemia as a potential adverse event. it states:¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. | The investigation into the limb ischemia - reversible issue has been completed. the device was not returned for investigation. the root cause of the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned.
32113 | 1220648-2023-05633 | 18419063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419063&pc=OZD | 12/03/2023 | 04/09/2024 |  | 12/29/2023 |  |  | 12/03/2023 | 12/29/2023 |  |  | 04/30/2024 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023197292 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2023 |  |  |  |  | The user facility reported an impella 5.0 in a (b)(6) male for mechanical circulatory support. it was reported that there were purge pressure high¿ alarms. the activated clotting time (act) has not been met for two days and purge trend shows slow increase in purge pressure. the impella was explanted. no additional information was provided. | [5 empty] | The impella device has not been returned for evaluation. a supplemental report will be submitted upon completion of the investigation. the instructions for use states the following regarding purge pressure issues: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ the failure modes will monitored and trended. | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. no biomaterial was found on the purge gap, and the catheter was cut open near the motor housing to check for blood ingress, which was not found. a minor kink was noted near the kink protractor. during priming testing, high purge pressure did not reproduce. the cause of the high purge pressure issue was not determined since data logs were not returned and limited clinical information was provided.
32114 | 1220648-2023-05632 | 18419068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419068&pc=OZD | 01/04/2021 | 11/12/2024 |  | 12/29/2023 |  |  | 11/27/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | UNKNOWN |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | During a literature review, a publication was found that reported left ventricle (lv) perforation following implantation of an unknown impella pump for mechanical circulatory support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve. | The investigation for the reported ventricle perforation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ventricle perforation could not be determined. corrections to the initial mfr report: f6/f8 should have been left blank. g4-pma/510k number stated "unkown" but should have stated unknown.
32115 | 1220648-2023-05634 | 18419077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419077&pc=OZD | 12/03/2023 | 01/18/2024 |  | 12/29/2023 |  |  | 12/03/2023 | 12/29/2023 |  |  | 05/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024376998 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2023 |  |  |  |  | The user facility reported a 72-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 pump was unable to advance through the anastomosis during attempted delivery. the implant was aborted and an impella cp device was subsequently placed for patient support. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32116 | 1220648-2023-05635 | 18419085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419085&pc=OZD | 12/01/2023 | 11/19/2025 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 13171809 |  | 0042-0000-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure of Device to Self-Test | Application Program Freezes, Becomes Nonfunctional | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported an 83-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during start up of the case, the automated impella controller (aic) self-check did not clear and the system would not start up. the aic was switched out. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. d1: corrected common device name. d4: corrected the model number. g1: corrected the manufacturer email address. h6: added codes per the results of the investigation. h10: added the results of the investigation.
32117 | 1220648-2023-05636 | 18419104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419104&pc=OZD | 12/01/2023 | 06/05/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 08/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392470 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | After the patient was escalated from an impella cp to an impella 5.5, the user facility reported an access site bleeding had developed from the former impella cp site. oozing was observed and manual pressure was held and eventually blood products were given. no additional information was provided. | The patient's outcome was reviewed by abiomed's medical safety officer. it was determined that there was not enough information to associate use of device with outcome. |  |  |  |  | The impella device has not been returned for evaluation. upon the completion of the investigation, a supplemental report will be submitted. the instructions for use states the following for access site bleeding: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ failure modes will be monitored and trended. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6. and h10. have been revised from the initial submission.
32118 | 1220648-2023-05637 | 18419154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419154&pc=OZD | 12/02/2023 | 09/18/2024 |  | 12/29/2023 |  |  | 12/02/2023 | 12/29/2023 |  |  | 05/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344586 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2023 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the patient developed a small soft pocket hematoma. heparin was held overnight after the hematoma was identified. heparin restarted in the morning. the pocket remains soft and stable. support continued with the implanted pump following the intervention. | [5 empty] | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the patient developed a small soft pocket hematoma. heparin was held overnight after the hematoma was identified. heparin restarted in the morning. the pocket remains soft and stable. support continued with the implanted pump following the intervention. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely anticoagulation management since withholding heparin resolved the issue. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added- h6 component code 925 added-d6a/d6b. f6/f8 should have been left blank in the initial report.
32119 | 1220648-2023-05638 | 18419163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419163&pc=PYX | 12/03/2023 | 10/02/2025 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  | 06/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024367571 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant reported the 68-year-old female patient was on impella rp support due to being post operation and the patient had mild-mod right ventricle dysfunction. while on support, the there were multiple suction events that were resolved by providing the patient with blood products and albumin. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Additionally while on support, the placement signal not reliable alarm occurred. survival outcome at explant noted the patient expired. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the pump was not returned for investigation. data logs were not provided for this case. as per the clinical details, the placement signal was dashed with no pump flow calculation. motor current and purge values were not affected during the event. low pump flow was ruled out after investigation, as the flow calculation depended on the placement signal derived from the optical and dp sensors. the motor current was within normal limits, which proves that the low pump flow was an actual condition, rather than being related to a sensor issue, which was addressed under the optical signal failure mode. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-05638. d6a/d6b updated with additional information. g1 reporting contact email updated.
32120 | 1220648-2023-05639 | 18419173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419173&pc=OZD | 12/02/2023 | 05/15/2024 |  | 12/29/2023 |  |  | 12/02/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024411508 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2023 |  |  |  |  | The user facility reported a 81-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support for presumed cardiomyopathy. a hematoma was noted at the access site and an extra stitch was provided with forward tension sutures; manual pressure was applied. the bleeding continued and the impella was progressively weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding is determined to be use issue because of the improper suturing and closing. they were able to achieve the hemostasis using the manta closure device and hence the root cause is determined to use issue. d6a, d6b, h6.
32121 | 1220648-2023-05640 | 18419181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419181&pc=OZD | 12/03/2023 | 11/19/2024 |  | 12/29/2023 |  |  | 12/03/2023 | 12/29/2023 |  |  | 07/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388005 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2023 |  |  |  |  | The user facility reported a 45-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing from multiple areas including the impella access site. femstop was placed overnight for a few hours and the impella site is now hemostatic. heparin gtt was turned off and started the following day. it was further reported that the patient developed a large apical thrombus seen in echo. the team would like to escalate to impella 5.5 but are waiting for the apical thrombus to dissipate. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the pump was returned for evaluation. no issues were found on the returned product. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely patient condition since the patient had bleeding from multiple sites. the root cause of the thrombus could not be determined without additional information. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- d9 device return date, h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32122 | 1220648-2023-05641 | 18419185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419185&pc=OZD | 12/03/2023 | 01/29/2024 |  | 12/29/2023 |  |  | 12/03/2023 | 12/29/2023 |  |  | 10/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024411264 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2023 |  |  |  |  | The user facility reported a 54-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and was awaiting a heart transplant. blood was noted in the purge side arm. pump function was normal with no alarms. the doctor did not want to remove the pump because the patient was awaiting a heart transplant. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak ¿ pump issue has been completed. the device was returned for benchtop investigation. based on the clinical information provided, blood was leaking from the red impella plug and purge sidearm. the device was returned with catheter cut open close to the motor housing. visually inspected and no blood found inside the purge line. no dried blood inside the purge sidearm reservoir or filter was seen. the red impella plug was opened and blood was observed inside and leaking from both ends of the gray connectors through the catheter. several punctured suture holes on the catheter close to 25 cm mark was observed. the cause of the product damage (hole in the catheter) issue was use related due to multiple suture holes punctured in the catheter leading to leaking of blood through the red plug.
32123 | 1220648-2023-05642 | 18419211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419211&pc=OZD | 12/04/2023 | 09/26/2024 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024406327 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Crack | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed an impella via the single access method at the femoral artery. the long sheath kinked and was replaced by the short sheath. the sheath cracked and the patient had a bleed. the bleeding was controlled to an extent by placement of the reposition sheath. the patient was given 1 unit of packed red blood cells prophylactically. after placement, it was revealed that patient's cad (coronary artery disease¿) was unable to be fixed and the patient's condition did not improve. cpr had been ongoing for 30 min, and the doctor decided to terminate the procedure and turned the pump off. the patient expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05642 was provided in sections b5, d6a, d6b, and h1.
32124 | 1220648-2023-05643 | 18419327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419327&pc=OZD | 12/02/2023 | 11/29/2024 |  | 12/29/2023 |  |  | 12/02/2023 | 12/29/2023 |  |  | 05/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374351 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male (race unknown) undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when repositioning the impella the physician pulled the device completely out of the ventricle and into the aorta. then the anti-contamination sheath was ripped so impella was removed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ the failure modes will be monitored and trended. | The investigation into positioning issues and product damage (sterile sleeve damage) has been completed. the root cause of the positioning issue was use related to improper technique. the root cause of the product damage sterile sleeve issue was most likely use related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05643: d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h6 additional medical device problem code was missing.
32125 | 1220648-2023-05644 | 18419335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419335&pc=OZD | 11/30/2023 | 09/26/2025 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1126092 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 11/30/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a controller error (yellow alarm) during support. as a result, the impella device was moved to a new controller. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the original report was filed. the console was returned for investigation and was tested upon return. console logs from the reported day of event are consistent with the complaint and show controller error alarm due to defective optical sensor lamp upon each boot-up, preceded by diagnostic information showing low light level but correct reading of ambient pressure. after console was received, the issue was reproduced and a controller error alarm presented on boot-up, and there was no light coming out of optical pump connector. the cause of the automated impella controller issue was a defective optical bench.
32126 | 1220648-2023-05645 | 18419470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419470&pc=OZD | 12/03/2023 | 05/22/2024 |  | 12/29/2023 |  |  | 12/03/2023 | 12/29/2023 |  |  | 08/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024392452 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2023 |  |  |  |  | The user facility reported a 47-year-old male (unknown race) undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that upon removal of the peel-away sheath following implant, the patient began to bleed from their access site. manual pressure was initially held to contain the bleeding, but the decision was ultimately made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis d6a, d6b
32127 | 1220648-2023-05646 | 18419542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419542&pc=OZD | 11/30/2023 | 09/21/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 07/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392298 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient experienced persistent bleeding from their femoral access site during impella support. coupled with the bleed, the surgeon was concerned as it was also noted that patient had a lack of pulse in the extremity on their impella side. due to the noted conditions, the pump was explanted and re-inserted via axillary access. the right groin access site was repaired by the vascular surgeon and blood products were given to the patient. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and access site bleed could not be determined as the device was not returned nor sufficient clinical details. d4-updated model number. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
32128 | 1220648-2023-05647 | 18419543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419543&pc=OZD | 12/04/2023 | 01/21/2025 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a controller error (yellow alarm). the aic was sent back for evaluation. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. the failure modes will be monitored and trended. | The investigation of automated imeplla controller (aic) - controller error (yellow alarm) has been completed. the root cause for the controller error was a defective optical bench lamp based on log review and device analysis. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05647.
32129 | 1220648-2023-05648 | 18419545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419545&pc=OZD | 12/04/2023 | 02/23/2024 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402952 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient had limb ischemia and was on three vasopressors. as a result the impella was explanted to address the issue. post-explant left leg has distal signals. no concern for ischemia according to nurse. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. correction : section h6: the impact code in section h6, was inadvertently left out in the initial report which has been added to this report.
32130 | 1220648-2023-05649 | 18419566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419566&pc=OZD | 12/02/2023 | 09/21/2024 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412981 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 71-year-old asian male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, and the patient's right lower extremity was mottled, cold, with no doppler pulses. care team opted to leave the impella as is since the patient's neurological status is in question and device is expected to be removed in a day or two. | Additional information noticed that the patient's family decided to withdraw the care, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. the root cause of the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
32131 | 1220648-2023-05650 | 18419575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419575&pc=OZD | 12/04/2023 | 07/18/2024 |  | 12/29/2023 |  |  | 11/30/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024405880 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 63-year-old white male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 kinked at the anastomoses during attempted delivery. the pump was removed, and a second pump was utilized for patient support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. added- d6a/d6b. f6/f8 should have been left blank in the initial report.
32132 | 1220648-2023-05651 | 18419738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419738&pc=OZD | 12/05/2023 | 12/03/2024 |  | 12/29/2023 |  |  | 12/05/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024405684 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the impella cp catheter was advanced too deep, and the entire cannula entered the ventricle. the cannula prolapsed and folded on itself. when attempting to remove the prolapsed pump from the ventricle the physician was met with resistance at the valve. with increased tension on the distal portion on the catheter the physician was able to dislodge the impella from the ventricle. at this time the teardrop/pigtail portion of the impella became disconnected from the cannula and was free floating in the ventricle. the physician was able to snare the pigtail and remove it from the patient. another impella cp was placed successfully. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the cannula pigtail detachment has been completed. the pump was not returned for investigation. a data log review was not required for this case. the cannula was prolapsed and folded in the ventricle due to being positioned too deep. resistance was met during disengaging the cannula. the pigtail was detached from the pump during this procedure and was successfully removed from the ventricle. image was not provided for the pigtail detachment during case. the cause of the cannula/pigtail detachment issue was not determined since product was not returned and sufficient clinical information was not provided.
32133 | 1220648-2023-05652 | 18419966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18419966&pc=OZD | 12/01/2023 | 01/17/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 10/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024406482 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump met resistance during attempted delivery and subsequently kinked. as a result of the damage to the cannula, the implant was aborted and a new 5.5 was placed via direct approach. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related.
32134 | 1220648-2023-05653 | 18420444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18420444&pc=OZD | 12/01/2023 | 01/18/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 05/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024346058 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 69-year-old male patient was attempting to have an impella 5.5 implanted and the impella could not advance past stenosis in the distal subclavian. the decision was made to abort the implant and implant the impella cp in the same access site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related.
32135 | 1220648-2023-05654 | 18420465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18420465&pc=OZD | 12/05/2023 | 01/18/2024 |  | 12/29/2023 |  |  | 12/05/2023 | 12/29/2023 |  |  | 08/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024389282 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 was unable to pass through the vasculature, encountering resistance at the innominate artery. the case was aborted due to patient's anatomy. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32136 | 1220648-2023-05655 | 18420846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18420846&pc=OZD | 12/01/2023 | 05/20/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403025 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that there was bleeding around the repo sheath that did not resolve with hours of manual pressure. the impella was removed, and the vessel was closed surgically during emergent coronary artery bypass grafting (cabg). the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the root cause of the access site bleeding was unable to be determined as insufficient clinical details was provided and no product was returned for analysis. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 4627 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
32137 | 1220648-2023-05656 | 18421381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18421381&pc=OZD | 12/04/2023 | 01/30/2024 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1454040 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The complainant reported that when prepping the console for icu training the plastic part that holds the air detect disc was broken. the console was exchanged. there was no patient involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. imc logs were available but were irrelevant to this complaint as the clinical staff only reported that the consoles purge retainer was broken. the console was returned and tested after arriving. upon examining the automated impella controller (aic) for any visible defects, it was confirmed that the purge retainer was broken. the root cause of this issue was a broken purge retainer.
32138 | 1220648-2023-05657 | 18421447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18421447&pc=OZD | 12/04/2023 | 12/03/2024 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407664 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 71-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had cp support ongoing when the patient was undergoing pci for multivessel cad. the pump had to be removed because it kinked in small aortic anatomy. the flows were low and so the pump was removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. the pump was not returned for investigation. data log shows the low pump flow with active suction alarm. no abnormalities were reported related to motor current trend or positioning issue. as per the clinical notes, patient's ascending aorta was noted to be extremely short, tortuous and calcified. a tight bend was noted on the cannula near the outlet cage. the replacement pump had same issue. the cause of the product damage which led to the low pump flow issue was most likely related to the patient anatomical condition, which caused the bend in the pump cannula, resulting in low pump flow. added h6 impact code 4629 corrections to the initial mfr report: d4 model number added d6a/d6b f6/f8 should have been left blank.
32139 | 1220648-2023-05658 | 18421528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18421528&pc=OZD | 12/01/2023 | 12/03/2025 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  |  |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1737893 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complaint reported that the automated impella controller (aic) had no audible alert during red/yellow alarms. the aic was noted to have made an audible beep when it was initially turned on. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32140 | 1220648-2023-05659 | 18421609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18421609&pc=OZD | 12/04/2023 | 01/18/2024 |  | 12/29/2023 |  |  | 12/04/2023 | 12/29/2023 |  |  | 10/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024411986 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp being delivered for support of the patient via the femoral artery and 14fr sheath. the pump was being placed for the hrpci (high risk percutaneous coronary intervention). the team was unable to be delivered. the anatomy precluded placement. no harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related.
32141 | 1220648-2023-05660 | 18421682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18421682&pc=OZD | 12/06/2023 | 01/02/2025 |  | 12/29/2023 |  |  | 12/06/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024402963 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 60-year-old male patient with sudden onset of chest pain had impella cp implanted for mechanical support. poor flows were noticed immediately upon impella cp start up. repositioning of the pump and administration of blood was performed to help improve flow; however no improvement was noted. the device was explanted successfully and was replaced. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of low pump flow has been completed. in the data logs, there are suction alarms and lower flows followed by a motor current spike and then flows decrease more. the pump was then explanted and biomaterial was noted. visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump ran without issue under bench test. based on the data logs and the clinical details provided, the root cause of the low pump flow was most likely biomaterial that was removed upon explant. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05660. d1 brand name. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing.
32142 | 1220648-2023-05661 | 18421900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18421900&pc=OZD | 12/01/2023 | 12/09/2024 |  | 12/29/2023 |  |  | 12/01/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411323 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported that approximately 20 minutes after impella cp insertion, flow saturation noted on automated impella controller (aic) with flat motor current and max/min flows within 0.1l. the impella was explanted and a new one was placed without any issue. there was no patient harm reported. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the saturated flow has been completed. the pump was returned for investigation. during visual inspection, a small amount of biomaterial was observed wrapped around the impeller. no other physical abnormalities were noted. during the test run, attempts to remove the biomaterial were unsuccessful, and low pump flow was replicated. data logs consistently indicated low pump flow from the beginning of the case without any motor current spikes. the total runtime of the case was 20 minutes. the root cause of the saturated flow/low pump flow issue was biomaterial found wrapped around impeller. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report:
32143 | 1220648-2023-05662 | 18422076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18422076&pc=OZD | 12/05/2023 | 01/17/2024 |  | 12/29/2023 |  |  | 12/05/2023 | 12/29/2023 |  |  | 09/30/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024410767 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 79-year-old white male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 would not pass through the axillary artery. the case for the 5.5 was aborted and impella cp was successfully placed. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was due to a patient condition.
32144 | 1220648-2023-05663 | 18422246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18422246&pc=OZD | 12/05/2023 | 09/20/2024 |  | 12/29/2023 |  |  | 12/05/2023 | 12/29/2023 |  |  | 05/31/2025 |  | 12/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024340888 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Inaccurate Delivery | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant reported the 78-year-old white male patient had an impella 5.5 attempted to be implanted. however, the impella was unable to advance once in the aorta, as the patient had numerous anatomical issues. therefore, the implant was aborted. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05663 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related.
32145 | 1220648-2023-05584 | 18409657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18409657&pc=OZD | 11/30/2023 | 02/11/2025 |  | 12/28/2023 |  |  | 11/30/2023 | 12/28/2023 |  |  |  |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1484084 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported on a 75yo male on impella cp for mechanical circulatory report. it was reported that during support, the automated impella controller (aic) displayed a yellow alarm "controller error". the alarm could not be switched off and the impella catheter was changed to another aic. the same aic console reported was used for another implant without any error messages or issues. there was no reported patient harm. | [5 empty] | The aic device has been returned for evaluation but has not yet been evaluated. when the investigation has been completed, a supplemental report will be submitted. | The investigation for the console (aic) controller error has been completed. data logs were reviewed which confirm that numerous controller error (defective optical sensor lamp) - alarm were issued. the logs reveal that optical bench sampling has a low snr and low light values before each alarm. the console was returned for evaluation. upon preventative maintenance, the console was up to spec. a known good pump and purge line was run with the console for an hour with no observable anomaly. the root cause was not determined due to not being able to reproduce the failure.
32146 | 1220648-2023-05585 | 18409682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18409682&pc=OZD | 11/30/2023 | 04/04/2024 |  | 12/28/2023 |  |  | 11/30/2023 | 12/28/2023 |  |  | 08/31/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386775 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 60-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during impella repositioning via echo a left ventricle thrombus formation was noted. a decision was made to wean down and remove impella. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
32147 | 1220648-2023-05586 | 18409723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18409723&pc=OZD | 12/01/2023 | 04/15/2024 |  | 12/28/2023 |  |  | 12/01/2023 | 12/28/2023 |  |  | 06/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371021 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. that night, patient experienced stroke with hemiplegia of left arm. impella was on p2 since yesterday. thrombosis in multiple locations was known, but no visible in lv. lvad procedure planned today has been cancelled, and patient is now limited. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of stroke and thrombus could not be determined as the device was not returned nor sufficient clinical details.
32148 | 1220648-2023-05587 | 18409792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18409792&pc=OZD | 12/01/2023 | 02/22/2024 |  | 12/28/2023 |  |  | 12/01/2023 | 12/28/2023 |  |  | 09/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024410625 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 41-year male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the staff was unable to locate pulses in the left lower extremity or popliteal using a doppler and the patient¿s extremities became cold. the patient was escalated to axillary 5.5 due to concern for vascular compromise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32149 | 1220648-2023-05588 | 18409817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18409817&pc=PYX | 12/02/2023 | 04/12/2024 |  | 12/28/2023 |  |  | 12/02/2023 | 12/28/2023 |  |  | 06/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024358965 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2023 |  |  |  |  | The user facility reported a 51-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while securing the repositioning sheath, impella became dislodged from the pa. several attempts to try to reposition but unsuccessful. the impella implant was aborted. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. as per given clinical details, the cause of the positioning issue was most likely use related since impella was pulled out while securing repositioning sheath.
32150 | 1220648-2023-05590 | 18409956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18409956&pc=OZD | 12/04/2023 | 05/20/2025 |  | 12/28/2023 |  |  | 12/04/2023 | 12/28/2023 |  |  | 09/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024401625 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 77yo male on impella 5.5 for mechanical circulatory support. the patient had high purge pressures a few hours after support was initiated. the team attempted to troubleshoot with the tissue plasminogen activator (tpa) protocol. however, the purge flow did not change. heparin was not used in the purge due to previous coagulopathy with multiple blood products being administered. a decision was made by family to withdraw support and the patient expired. | Additional information noted that the care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. a supplemental report will be submitted upon completion of the investigation. the instructions for use states the following for high purge pressures: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The pump was not returned for investigation. data logs show a gradual 2-hour rise in purge pressure without a reported motor current spike. according to the clinical details, heparin was not added to the purge solution, and tpa failed to reduce high purge pressure. the cause of the high purge pressure issue was most likely biomaterial based on data log review.
32151 | 1220648-2023-05592 | 18410038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18410038&pc=OZD | 12/04/2023 | 01/17/2024 |  | 12/28/2023 |  |  | 12/04/2023 | 12/28/2023 |  |  | 06/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024359657 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 71-year-old female patient was scheduled for impella implant for mechanical circulatory support. the us complainant had a 5.5 being placed via the axillary, for the patient undergoing mvr (mitral valve replacement) and cabg (coronary artery bypass grafting¿). the 5.5 pump failed to deliver due to arterial stenosis. the team instead placed a cp for support. no lasting harm or injury from attempted delivery. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. as per clinical the patient had stenosis at bifurcation of the carotid arteries to the aortic arch. reattempted 5.5 insertion after angioplasty, but unsuccessful and aborted 5.5 insertion. the cause of the delivery issue was patient condition related due to patient having stenosis condition and unable to advance pump further past the aortic arch.
32152 | 1220648-2023-05593 | 18414276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18414276&pc=OZD | 08/21/2023 | 02/23/2024 |  | 12/28/2023 |  |  | 11/28/2023 | 12/28/2023 |  |  | 03/31/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345555 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 29 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed¿s long-term outcome and quality indicator impella registry (loqi) received a notification indicating the impella cp patient endured a vascular complication prompting intervention. the report indicates the following, "acute right lower extremity limb ischemia following impella removal. required surgical thrombectomy." the report indicates "definite" relationship to the impella device". | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32153 | 1220648-2023-05565 | 18414559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18414559&pc=PYX | 08/08/2023 | 09/24/2024 |  | 12/28/2023 |  |  | 08/12/2023 | 12/28/2023 |  |  | 02/28/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024348418 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that 68-year-old female patient received support from an impella rp. however, the impella rp flex started with flows in the high 3s and then quickly dropped to the low 3s and stayed there. the physician does not think the impella ingested any clot. care was withdrawn four days post-implantation. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported ¿low pump flow¿ issue has been completed. no product was returned for investigation. data log analysis confirmed low flow at p-9 and p-8. there were also suction alarms noted immediately prior to the low flow event. the root cause of the low pump flow issue was unable to be determined since product was not returned and there is a lack of clinical details regarding patient volume or anatomy issues. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05565 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
32154 | 1220648-2023-05582 | 18415175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415175&pc=OZD | 11/28/2023 | 11/25/2025 |  | 12/28/2023 |  |  | 11/28/2023 | 12/28/2023 |  |  |  |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119399 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | No patient involvement: the complainant reported battery failure on the aic (automated impella controller). the aic was removed from service and will be returned for investigation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32155 | 1220648-2023-05594 | 18415271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415271&pc=OZD | 07/28/2023 | 12/28/2023 |  | 12/28/2023 |  |  | 07/28/2023 | 12/28/2023 |  |  | 04/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024349822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/28/2023 |  |  |  |  | The user facility reported a 73-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis during support and required intervention beyond normal troubleshooting, consisting of initiating inotropic support. | [5 empty] | The pump was not received, however, and an evaluation of the clinical report and data logs was completed. data logs found several positioning issue alarms that were deemed relevant to the hemolysis event. also, some suction event alarms were observed. the cause of the hemolysis issue was most likely improper positioning of the device.
32156 | 1220648-2023-05595 | 18415291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415291&pc=OZD | 11/29/2023 | 04/09/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  |  |  | 12/28/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1134268 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | No patient involvement: the complainant reported that the purge disc holder within the automated impella controller (aic) had broken off. there was no noted patient harm resulting from the damaged component. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the purge pressure sensor assembly was replaced. the root cause of this issue was most likely a purge disc retainer.
32157 | 1220648-2023-05597 | 18415315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415315&pc=OZD | 11/29/2023 | 04/16/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 10/31/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408600 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 75 -year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. upon implanting the 5.5, the patient had sustained ventricular tachycardia/ventricular fibrillation. the patient was shocked 5 times. then the patient experienced additional runs of ventricular tachycardia and after, the pigtail access was confirmed in the mid left ventricle space. the impella was implanted successfully with proper flows. the device was explanted. additionally, upon explanting the impella 5.5, multiple clots were noted. the doctor presumed "showering clots". heparin was given to the patient and a thrombectomy was performed utilizing a vacuum catheter to right axillary artery utilizing same graft from 5.5 explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus and access site and the vf/vt/af/at (ventricular tachycardia) issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf and the thrombus was not determined.
32158 | 1220648-2023-05559 | 18415336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415336&pc=OZD | 11/26/2023 | 10/03/2024 |  | 12/28/2023 |  |  | 11/26/2023 | 12/28/2023 |  |  | 11/30/2024 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314553 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2023 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that delivery of the intended impella cp pump was difficult. following removal of the peel away sheath, attempts to advance the repositioning sheath resulted in the pump kinking. the pump was removed as a result of the noted issue and a new impella cp pump was placed through the opposite leg for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ added- h6 impact code 4629 correction to the initial mfr report: d4-corrected model number added-d6a/d6b f6/f8 should have been left blank g1-corrected mfr site name and address
32159 | 1220648-2023-05599 | 18415355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415355&pc=OZD | 11/29/2023 | 12/03/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 05/31/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348688 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported an impella cp in a 52yo male for mechanical circulatory support. it was reported that while support was ongoing, bleeding was found between the indwelling sheath and the shaft. a crack was noted at the location and bone wax was used to seal the crack. there was no patient harm reported as a result of the bleeding area of the pump. no additional information was provided. | [5 empty] | The impella device has not yet been returned for evaluation. a supplemental report will be submitted upon completion of the investigation. the failure modes will be monitored and trended. | Pump was not returned for investigation. as per the clinical information provided, the leak was observed from the repo unit. therefore, product damage hole in catheter failure mode was deleted since that is not observed. access site bleeding encompasses the leak observed from the repo unit. therefore, the root cause of the access site bleeding issue was most likely a damaged valve. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05599.
32160 | 1220648-2023-05600 | 18415376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415376&pc=OZD | 11/29/2023 | 01/30/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  |  |  | 12/28/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1356082 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported that during the use of the automated impella controller (aic) on a patient of unknown age and gender, the purge disc clip snapped off. as a result of the noted damage, a second aic was utilized for support. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. imc logs were available but were irrelevant to this complaint as the clinical staff reported that the consoles purge retainer had snapped off. the console was returned and tested after arriving. upon examining the automated impella controller (aic) for any visible defects, it was evident that the purge retainer was broken. the root cause of this issue was a broken purge retainer. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32161 | 1220648-2023-05601 | 18415401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415401&pc=OZD | 11/28/2023 | 07/18/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 09/03/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411288 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 62-year-old male noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the pci was completed and the pump was explanted at 0.87 hours of support. the support was successful but, after the pump was explanted the team coded the patient. an allegation was made that the code and subsequent death were due a guidewire perforation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation for the ventricle perforation has been completed. product was not returned for investigation. the cause of the ventricle perforation issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ added- h6 component code 463 added-d6a/d6b. f6/f8 should have been left blank in the initial report.
32162 | 1220648-2023-05602 | 18415434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415434&pc=OZD | 11/29/2023 | 05/16/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 04/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349278 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed for hemodynamic support of a cgs (cardiogenic shock), post bental, and repeat tavr (transcatheter aortic valve replacement) patient. the cp was placed and there was an access site bleed observed. the bleed was treated by manual pressure and femostop. there was no extensive intervention needed. no vascular intervention or rbcs (red blood cell) needed. the pump remained on for support until wean and explant. | Additional information was received. two units of blood was infused. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since the product was not returned and not much information was provided in the clinical description. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.5 additional was received and was inadvertently omitted from the previously submitted report. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 4643 based off of additional information that was received and was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
32163 | 1220648-2023-05603 | 18415456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415456&pc=OZD | 11/29/2023 | 11/11/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 09/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398687 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 83-year-old male patient had the impella cp attempted to be implanted and the tortuous iliac caused a kink in the long sheath and the sheath was switched out for a short sheath and the pump passed through, but kinked at the body of the cannula. a buddy wire was attempted and failed. an edwards 16fr sheath was used and the pump passed through, but the buddy wire kinked and pulled back and the position of the impella was lost. the impella insertion was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the introducer damage and cannula kink has been completed. pump was not returned for investigation. clinical mentions that the patient had tortuous iliac which was causing the sheath and the cannula to kink once the long sheath was switched to a short sheath. multiple techniques such as edward sheath and buddy wire were also tried but were unsuccessful due the tortuous vessels. therefore, the root cause of the product damage issue was a kinked cannula due to patients tortuous vessels. the root cause of the introducer damage issue was a introducer sheath kink due to patients tortuous vessels.
32164 | 1220648-2023-05604 | 18415466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415466&pc=OZD | 11/29/2023 | 03/29/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 06/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397430 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 76-year-old male with unknown race and ethnicity with cardiomyopathy was implanted with an impella for mechanical circulatory support. pump stopped in the cath lab and started again. the pump was not replaced and the motor current was noted to be in normal range. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported pump stop has been completed. neither the pump nor data logs were returned for evaluation. additionally, sufficient clinical information was not provided. therefore, the exact root cause of the pump stop could not be determined. the instructions for use states ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32165 | 1220648-2023-05605 | 18415474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415474&pc=OZD | 11/29/2023 | 02/23/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 05/31/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024352449 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump placed via the axillary artery to support a patient with cardiomyopathy. on the 2nd day of the support the team noted limb issues. the limb had pain, numbness, and tingling. the limb pain was treated by gabapentin. the pump remains on support still despite the issues. now on day 10 of support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32166 | 1220648-2023-05606 | 18415480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415480&pc=OZD | 11/29/2023 | 11/29/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 06/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024390738 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that while moving the patient for transport, the implanted impella cp pump was pulled back across the aortic valve. as the patient's lv (left ventricle) function was improving, the decision was made to explant the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was unsecured touhy-borst. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05606. d3 manufacturer name, address, city, state, and fax number, d4 model number, d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank.
32167 | 1220648-2023-05607 | 18415500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415500&pc=OZD | 10/13/2023 | 10/16/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 06/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024383224 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | It was reported by a clinical study (stemi-dtu study) that a patient of unknown age and gender on impella cp support had an access site bleed that required compression and transfusion of one pack of red blood cells. no additional information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. a supplemental report will be submitted upon conclusion of the investigation. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h1 reporting contact facility name and address, h6 component code 925 corrections to the initial mfr report: added-d6a/d6b f6/f8 should have been left blank.
32168 | 1220648-2023-05609 | 18415603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415603&pc=OZD | 11/30/2023 | 05/09/2024 |  | 12/28/2023 |  |  | 11/30/2023 | 12/28/2023 |  |  | 07/31/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388060 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the impella related access site bleed prompting stoppage of heparin delivery during support. | Patient was successfully weaned with no press or support and impella pulled in cath lab without issue. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of access site bleeding was determined to be patient movement because it was mentioned the bleed was caused due the transfer. the site was redressed and site looks good now. hence, the root cause is determined to be patient movement. added b5-addl event narrative potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32169 | 1220648-2023-05591 | 18415677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18415677&pc=OZD | 12/04/2023 | 12/28/2023 |  | 12/28/2023 |  |  | 12/04/2023 | 12/28/2023 |  |  |  |  | 12/28/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported that during case preparation, the automated impella controller, (aic) had a broken purge disk plate (missing). a replacement loaner was sent to the facility. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32170 | 1220648-2023-05598 | 18416048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18416048&pc=OZD | 11/29/2023 | 10/30/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 12/31/2023 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023179305 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 89-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and the pump was removed. upon removal, manta was deployed and there was excessive bleeding and patients heart rate increased. the physician noticed a perforation. a covered stent was placed successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ the failure modes will be monitored and trended. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05598. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. g1 manufacturing site fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32171 | 1220648-2023-05596 | 18416063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18416063&pc=OZD | 11/29/2023 | 09/24/2024 |  | 12/28/2023 |  |  | 11/29/2023 | 12/28/2023 |  |  | 09/30/2025 |  | 12/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC |  |  | US |  | 2024395549 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 78-year-old female patient on impella cp pump support experienced groin bleeding at the access site. patient developed hemorrhagic shock and hemoglobin dropped from 11 to 7. fluid volume and red blood cells (rbc) were given as treatment and the issue resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- h6 component code 925 correction to the initial report: added- d6a/d6b. f6/f8 should have been left blank.
32172 | 1220648-2023-05569 | 18401071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18401071&pc=OZD | 11/27/2023 | 11/18/2024 |  | 12/27/2023 |  |  | 11/27/2023 | 12/27/2023 |  |  | 07/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024391371 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/27/2023 |  |  |  |  | A 70 year old male patient in ventricular tachycardia (vt) elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during the sheath exchange, the pump came out of the left ventricle. attempts to re-cross the valve was unsuccessful and the device subsequently kinked. the pump was removed as a result of the noted issue and a new device was placed for ongoing support. | [5 empty] | The impella cp device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | The investigation for the reported positioning issue has been completed. pump was not returned for investigation. clinical mentions that during sheath exchange, the pump was pulled out of the lv which is a user error. pump was replaced since replacement was tried but kinked instead. therefore, the root cause of the positioning issue was use related to improper technique.
32173 | 1220648-2023-05579 | 18403127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18403127&pc=OZD | 11/28/2023 | 11/29/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 09/30/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED,INC |  |  | US |  | 2024402955 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 62-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient developed an effusion which was tapped and attempted to drain, however the effusion would not stop bleeding which prompted the doctor to suspect a perforation. the decision was made to take the patient to the operating room (or), maintain impella support, locate and treat the suspected perforation, and then remove the impella once complete. upon arrival to the or, a surgical nurse accidentally pulled on the impella catheter at the groin site which displaced the pump and pulled it into the aorta. the staff found a perforation on the left ventricle, and it was sutured and patched. the impella was removed. the doctor did not suspect the perforation to be caused by the impella, however the doctor was not certain of the cause for the perforation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury the failure modes will be monitored and trended. | The investigation for the ventricle perforation and the positioning issue has been completed. pump was not returned for investigation. clinical information provided about the ventricle perforation is none. time and reason of the perforation is unknown. therefore, the root cause of the ventricle perforation issue was not determined since product was not returned and insufficient clinical information was provided. clinical mentions that the pump was accidently pulled into aorta by the nurse. therefore, the root cause of the positioning issue was use related to improper technique. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32174 | 1220648-2023-05577 | 18403515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18403515&pc=OZD | 11/28/2023 | 05/16/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 10/31/2024 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283791 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had an impella cp placed for support of a patient admitted with multiple cardiac comorbidities and history. the team observed an access site bleed at the site of the impella cp placement. the bleed/hematoma was treated by figure eight stitching and forward tension was applied and manual pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, forward tension sutured and angled matched which resulted in hemostasis as per the clinical description. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
32175 | 1220648-2023-05578 | 18403570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18403570&pc=OZD | 11/28/2023 | 11/18/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 07/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388880 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported an impella cp in a 88yo female for mechanical circulatory support during high risk percutaneous coronary intervention. the pump was placed however the team observed low flows that was due to a kinked cp. the cp was replaced without any harm or injury. no additional information was provided. | [5 empty] | The impella device has not yet been returned for evaluation. device evaluation is anticipated. a supplemental report will be submitted upon conclusion of the investigation. the failure modes will be monitored and trended. | The investigation into product damage (cannula kink) has been completed. the impella device was not returned for evaluation. the root cause of the kinked cannula is unknown. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05578. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact facility name missing. g1 reporting contact fax number missing.
32176 | 1220648-2023-05580 | 18404190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18404190&pc=OZD | 11/28/2023 | 09/21/2024 |  | 12/27/2023 |  |  | 11/29/2023 | 12/27/2023 |  |  | 05/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024343790 |  | 1000021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/29/2023 |  |  |  |  | The user facility reported a patient (unknown age and gender) in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as escalation from an impella cp and extracorporeal membrane oxygenation (ecmo). after roughly 124 days of support, the implanted impella 5.5 pump stopped working. the pump was explanted as a result of the noted issue. there was no patient harm. | It was reported that after the pump stop no additional treatment was provided to the patient, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the high purge pressure and the pump stop issues has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the high purge pressure and the pump stop issues was not determined since product was not returned and limited clinical information was provided. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05580 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
32177 | 1220648-2023-05581 | 18405384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405384&pc=OZD | 11/18/2023 | 09/12/2024 |  | 12/27/2023 |  |  | 11/18/2023 | 12/27/2023 |  |  | 09/30/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398829 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 70-year-old black male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was experiencing bleeding from the impella access site. bleeding was managed with angle matching and one unit of prbc (packed red blood cells) was given. the patient is stable, and bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause access site bleeding was most likely use issue since managing angle matching resolved the issue. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 impact code 4641, h6 component code 925 d4-corrected model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
32178 | 1220648-2023-05583 | 18405506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405506&pc=PYX | 11/28/2023 | 10/15/2025 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 08/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024407071 |  | 1000323 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella rp 0.027 guidewire appeared kinked. the impella was able to be successfully placed. there was no patient harm. the team is monitoring and trending the motor current levels. | [5 empty] | Device was received by abiomed (manufacturer) on 19-dec-2023. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported product damage has been completed. the product was returned and the guidewire was kinked on the end. data logs review was not relevant to the investigation for the product damage. based on the clinical details provided, the root cause of the product damage is most likely due to use, as the guidewire appeared kinked as the insertion pulled it out of the pulmonary artery (pa) twice. d4-corrected model number. added- d6a/d6b. d9-corrected to yes. f6/f8-should have been left blank.
32179 | 1220648-2023-05563 | 18405530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405530&pc=OZD | 11/27/2023 | 05/02/2025 |  | 12/27/2023 |  |  | 11/27/2023 | 12/27/2023 |  |  |  |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1253747 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported purge cassette failure while preparing the automated impella controller (aic). the console was changed, resolving the issue. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. based on the device and data analysis the root cause of the purge system failure (piston block) was the dextrose deposit on the purge motor gasket and behind the purge cassette insert. d2 common device name revised on manufacturer device report in accordance with updated procedures 1220648-2023-05563.
32180 | 1220648-2023-05564 | 18405534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405534&pc=OZD | 11/27/2023 | 02/22/2024 |  | 12/27/2023 |  |  | 11/27/2023 | 12/27/2023 |  |  | 05/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024359401 |  | 0048-0014 EU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/27/2023 |  |  |  |  | The user facility reported male patient (unknown age and race) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed limb ischemia during support. the patient remained stable with veno-arterial extracorporeal membrane oxygenation (va ecmo) only. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia reversible issue has been completed. the device was not returned for investigation. the root cause of the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned.
32181 | 1220648-2023-05567 | 18405536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405536&pc=OZD | 11/26/2023 | 01/30/2024 |  | 12/27/2023 |  |  | 11/26/2023 | 12/27/2023 |  |  | 04/30/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350543 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2023 |  |  |  |  | The user facility reported that following insertion of an impella cp in a 63yo male placed for mechanical circulatory report, attempts to advance the repositioning sheath resulted in the catheter kinking. the sheath was subsequently unable to advance into the proper position and the pump was removed from the right femoral artery. the pump was then moved to the left femoral artery where implant was successful. there was no patient harm reported. no additional information was provided. | [5 empty] | The impella device has not yet been received from the customer. however, device evaluation is anticipated. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | The investigation into the delivery issues has been completed. the pump was not returned for investigation. the root cause of the positioning issue was determined to be use related as when the repo unit was advanced the pump came out of position resulting in a cannula kink and removal of the device. h.10 the following statement was reported on the previous manufacturer device report by accident and should not of been reported: the failure modes will be monitored and trended.
32182 | 1220648-2023-05574 | 18405541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405541&pc=OZD | 11/28/2023 | 05/15/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 08/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387802 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 61-year-old male (unknown race) noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed bleeding and hematoma to the right axillary site and right arm. the patient was taken back to the or. the bleeding was isolated, and an additional suture placed at the graft/artery anastomosis. the hematoma was also evacuated. one-unit packed red blood cells was given. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, additional sutures were placed to achieve hemostasis as per the clinical description. added- d6a/d6b.
32183 | 1220648-2023-05572 | 18405543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405543&pc=OZD | 11/28/2023 | 02/29/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 07/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 |  |  | ABIOMED, INC. |  |  | US |  | 2024388071 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported that during an attempted insertion of an impella cp in a 61yo male placed for mechanical circulatory report, resistance was met while crossing the aortic valve. force was subsequently used to advance the pump into the left ventricle. upon removal of the guidewire, it was discovered that the wire had sheared in the cannula. the pump was removed as a result of the broken guidewire and the tip of the wire was found hanging at the end of the sheath. a new impella cp was placed for planned support. there was no patient harm. | [5 empty] | The impella device has been received but has not yet been evaluated. the investigation is in process. a supplemental report will be submitted upon conclusion of the investigation. the failure modes will be monitored and trended. | The investigation into the reported delivery issue and product damage has been completed since the original report was filed. returned complaint cp pump 463538 and 0.018" gw visually inspected. cannula kink observed at mid of the cannula. guidewire sheared and all the portions returned and tip not fractured. the cause of the delivery and guidewire damage issues were use related issue , cannula bent and gw damaged due to improper technique and excessive force.
32184 | 1220648-2023-05566 | 18405612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405612&pc=OZD | 11/18/2023 | 11/12/2024 |  | 12/27/2023 |  |  | 11/18/2023 | 12/27/2023 |  |  | 05/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356671 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant reported the patient was on impella cp support and the patient experienced hemolysis with dark red urine. the staff weaned the impella and explanted the pump. the hemolysis cleared after removal. additionally, after removal there was a small hematoma noted at the access site. manual pressure was held, and it resolved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were investigated and there was no motor current spoke or suction observed. no positioning issues were observed as well. on the day of reported hemolysis the placement signal had a sudden spike followed by a gradual drop which happened a couple times. clinical information mentions that hemolysis was resolved upon impella removal. no other information related to hemolysis was provided. therefore, the root cause of the hemolysis issue was not determined since product was not returned, insufficient clinical information provided and data logs were inconclusive. added h6: component code 925 correction to the initial mfr report: d4: model number.
32185 | 1220648-2023-05570 | 18405867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405867&pc=OZD | 11/28/2023 | 11/12/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 06/30/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024364430 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Injury | No | Yes | Improper Flow or Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 61-year-old male (race unknown) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the cannula kinked due to patient's short aortic arc. this led to reduced flow during the percutaneous coronary intervention (pci) procedure. however, the procedure was able to be completed and the pump was explanted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the failure modes will be monitored and trended. | The investigation for the low pump flow and product damage has been completed. pump was not returned for investigation. data logs were not returned as well. clinical mentions that patient has a small aortic arch which led to kinking of the cannula and resulted in low flows. therefore, the root cause of the low pump flow issue was most likely patient condition related to patients small aortic arch causing the kinking. corrections to the initial mfr report: d4-model number. added d6a/d6b. f6/f8 should have been left blank.
32186 | 1220648-2023-05560 | 18405870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405870&pc=OZD | 11/26/2023 | 11/07/2025 |  | 12/27/2023 |  |  | 11/26/2023 | 12/27/2023 |  |  | 09/30/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410434 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Mechanical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2023 |  |  |  |  | The user facility reported a 62-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the complainant reported that after placing a short peel away when dilator was removed, the orange piece broke off in 2 parts and the valve was no longer hemostatic. exchanged for long peel away. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The introducer was returned for investigation and the orange hemostatic valve was found to be broken off the introducer. as per the clinical this happened after taking out the dilator. therefore, the root cause of the introducer damage issue was a damaged valve due to insufficient adhesive. d6a and d6b revised from the initial manufacturer device report 1220648-2023-05560 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-05560 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-05560 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-05560. h6 component code revised from the initial submission in accordance with updated procedures.
32187 | 1220648-2023-05576 | 18405873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405873&pc=OZD | 11/28/2023 | 01/24/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 08/31/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387775 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. a leak was observed at the luer lock of the purge sidearm after changing the purge cassette. the luer was inspected after an additional purge cassette change. no obvious cracks were noted. the purge continued to leak. the purge pressures and flows were not affected. the pump leak was resolved on its own later that evening. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the purge sidearm leak has been completed. the pump and purge cassette were returned for investigation. the device operated as intended during the test, and no physical abnormalities or purge leaks were observed. data logs did not report any purge leak events. as per clinical note, a purge leak was noticed after the cassette procedure without impacting purge pressure and flow, and the issue resolved without any interference. there was no issue found during the case. the device functioned/operated according to specifications, and data logs did not report any events related to purge leak issues.
32188 | 1220648-2023-05575 | 18405911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18405911&pc=OZD | 11/28/2023 | 01/30/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  |  |  | 12/27/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1239161 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a patient of unknown gender and age was undergoing cardiac surgery was implanted with an impella pump device for mechanical circulatory support. when starting the case the automated impella controller (aic) could not detect the purge disk. several attempts were made unsuccessfully. the aic was swapped mid case for a loaner. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. imc logs from the reported occurrence date revealed that the console issued purge disc not detected alarm several times. the console was tested after arriving. the issue with the automated impella controller (aic) not being able to detect the purge disc was reproduced, and the purge flag was found to be broken. the root cause of the issue was a broken purge flag. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32189 | 1220648-2023-05562 | 18406268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18406268&pc=OZD | 11/27/2023 | 04/30/2024 |  | 12/27/2023 |  |  |  | 12/27/2023 |  |  |  |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/27/2023 |  |  |  |  | The complainant reported during case of a patient of unknown age and gender on impella cp support, nacl 0.9% + bicarb was used as purge solution accidentally. head nurse was advised to change the purge solution and monitor the motor current. later in the day, the patient was also reported to have hemolysis even though there was no sign of malposition or volume issues. the impella was noted to be running at p-8 with mean motor current 1000-1100ma. motor issues were thought to be due to saline in the purge. pump change was recommended but motor current decreased to a mean of 930-950ma. no action was taken. | [5 empty] | The patient, device, udi and physician information is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup. unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the purge pressure high alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the hemolysis and the low pump flow issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis and the low pump flow issues was not determined since product was not returned and limited clinical information was provided. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, section d6 and h10 were revised with updated information.
32190 | 1220648-2023-05519 | 18406269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18406269&pc=OZD | 08/26/2023 | 12/20/2023 |  | 12/27/2023 |  |  | 08/26/2023 | 12/27/2023 |  |  | 04/30/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024371327 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/26/2023 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella catheter was unable to advance into the left ventricle. it was difficult getting a catheter around the aortic arch. the physician said that the patient¿s anatomy was normal. a new impella was placed successfully. | [5 empty] | The cause of the delivery issue was most likely patient condition related due to patient anatomy of aortic arch made delivery difficult. complaint device not returned for investigation.
32191 | 1220648-2023-05573 | 18406885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18406885&pc=OZD | 11/26/2023 | 11/25/2025 |  | 12/27/2023 |  |  | 11/26/2023 | 12/27/2023 |  |  |  |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1635221 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | No patient involved: the complainant reported that the console failed to start and shut down a couple times in the catheterization laboratory prior to placing the impella cp. this console also had a battery temperature high alarm which resolved. a second console was used successfully. there was no patient harm. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
32192 | 1220648-2023-05571 | 18406894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18406894&pc=OZD | 11/28/2023 | 04/05/2024 |  | 12/27/2023 |  |  | 11/28/2023 | 12/27/2023 |  |  | 09/30/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 |  |  | ABIOMED, INC. |  |  | US |  | 2024407002 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported an impella cp in a 78yo male placed for mechanical circulatory report. it was reported that high purge pressure and purge system blocked alarms were triggered during impella cp support. the pressure was measured at 1068 mmhg with flows at less than 1ml/hr. the pump was initially monitored for changes in performance, but ultimately stopped working the following morning. the pump was explanted as a result of the failure. there was no patient harm. | [5 empty] | The impella device has not yet been returned for evaluation. however, device evaluation is anticipated. a supplemental report will be filed upon closure of the investigation. the instructions for use for purge pressure and pump stop are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. | The investigation into the pump stop and the high purge pressure issues has been completed since the initial report was submitted. the device was not returned for investigation. data logs confirm there were motor current spikes with suction alarms followed by high purge pressure and purge system blocked alarms. immediately after, the flow becomes saturated; later, there was a high motor current pump stop. the root cause of the pump stop and the high purge pressure was most likely biomaterial.
32193 | 1220648-2023-05561 | 18407014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18407014&pc=OZD | 11/24/2023 | 03/05/2024 |  | 12/27/2023 |  |  | 11/24/2023 | 12/27/2023 |  |  | 04/30/2025 |  | 12/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335656 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/27/2023 |  |  |  |  | The user facility reported a 61-year-old female was implanted with an impella device for unexplained arrhythmia. the patient has experienced cerebral infraction. the impella cp was removed. no additional treatment was given for the cerebral infarction. the patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke has been completed since the original report was filed. as the necessary clinical information was not provided, the true root cause of the stroke/cva could not be determined.
32194 | 1220648-2023-05557 | 18394226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18394226&pc=OZD | 11/26/2023 | 11/15/2024 |  | 12/26/2023 |  |  | 11/26/2023 | 12/26/2023 |  |  | 09/30/2025 |  | 12/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397031 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Ischemia Stroke | [11 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2023 |  |  |  |  | The user facility reported a patient of unknown age and gender in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced a cerebral bleed. the bleed contributed to ventricular perforation as well. the support had gone on for seven days for the patient. when the bleed occurred, the family made the choice to not continue support and the patient expired. it is unknown whether it is related to impella. impella operated without any problems until the patient died. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. as this clinical information was not provided, the true root cause of the stroke could not be determined. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ a.3 revised as information was inadvertently omitted on the initial manufacturer device report number 1220648-2023-05557. b.2 revised as date of death was previously entered incorrectly on initial manufacturer device report number 1220648-2023-05557. d.6a and d.6b revised from the initial submission of initial manufacturer device report number 1220648-2023-05557 in accordance with updated procedures. e.1 revised first name as it was previously entered incorrectly on initial manufacturer device report number 1220648-2023-05557. f.6 and f.8 dates were reported on the initial manufacturer device report number 1220648-2023-05557 and should not have been. g.1 reporting contact facility name was inadvertently left off the previously submitted manufacturer device report number 1220648-2023-05557 and was added. h.6 added code 925 in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-05557 was submitted. h.10 additional manufacturer narrative instructions for use were revised from the initial manufacturer device report number 1220648-2023-05557 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32195 | 1220648-2023-05558 | 18394291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18394291&pc=OZD | 11/23/2023 | 02/21/2024 |  | 12/26/2023 |  |  | 11/23/2023 | 12/26/2023 |  |  | 05/31/2025 |  | 12/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024354593 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2023 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed a hematoma at their access site during patient support. femostop was applied to the sight to treat the condition and support was continued. roughly three days later, it was noted that the patient's platelet levels dropped prompting a switch from heparin to sodium bicarb in the purge. support continued for another two days before planned explant. | [5 empty] | The investigation into the thrombocytopenia and the access site bleeding issue has been completed. the device was not returned for investigation. the most likely cause of the minor access site bleeding was patient movement because at the time of the bleed they were being transferred from the cath lab to the icu. the cause of the thrombocytopenia was not determined as the relationship between impella and the low platelet count could not be determined from the clinical details provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32196 | 1220648-2023-05568 | 18395193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18395193&pc=OZD | 11/26/2023 | 05/29/2024 |  | 12/26/2023 |  |  | 11/26/2023 | 12/26/2023 |  |  | 09/30/2025 |  | 12/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396548 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 65-year-old male patient was on impella 5.5 support and was having lower than expected flows display on the automated impella controller. the pump was approximately supporting the patient via hemodynamics. for instance, increasing p levels were not changing the displayed flow but the cardiac output and cardiac index did increase. all the changes were the left ventricle diastolic becoming increasingly more negative despite perfect positioning and preload into the left ventricle. | The patient's outcome was reviewed by abiomed's medical safety officer (mso). it was determined that there was not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the low pump flow has been completed. the pump was not returned for investigation. data logs show suction alarms throughout the case without a noted trend in placement signal or positioning issues, or motor current spikes. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist in section: using the automated impella controller with the impella catheter and section: direct aortic insertion . added-mso review in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32197 | 1220648-2023-05530 | 18381419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18381419&pc=OZD | 12/05/2023 | 12/03/2024 |  | 12/22/2023 |  |  | 12/05/2023 | 12/22/2023 |  |  | 08/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411247 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Material Frayed | Use of Device Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported an 82-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the team delivered over the .018 wire and then had complications with removal. the md states the issue was his user error, because the wire and pump were removed together as one system and the wire frayed and the pump/catheter broke at the motor in the sheath. the team had to apply a femostop to control the blood loss. no blood products were required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into removal issues and product damage (guidewire damage - peripheral vessels) has been completed. the cause of the guide wire damage issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the removal issue was most likely use related, based on given clinical details. instructions for use as follows: impella cp with smartassist system for use during cardiogenic shock. section: warnings and cautions: warnings. ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05530. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 codes 1888 and 4641 were reported incorrectly. new codes were added to health effect - clinical code and health effect - impact code. h10 instructions for use missing.
32198 | 1220648-2023-05532 | 18381443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18381443&pc=OZD | 12/04/2023 | 05/20/2025 |  | 12/22/2023 |  |  | 12/04/2023 | 12/22/2023 |  |  | 09/30/2024 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273533 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/04/2023 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 82yo male on impella cp for mechanical circulatory support. the patient had developed a hematoma that required surgical intervention for the limb. the hemoglobin dropped. blood products and platelets (ffp) were given. vascular surgery was consulted for the hematoma and a suture was placed around the impella sheath. the support continued and team discussed a possible escalation to a 5.5 pump. the patient expired on support. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was not returned for evaluation. a supplemental report will be submitted when the investigation has been completed. the instructions for use states the following regarding access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The root cause of the access site bleeding issue was most likely use related, forward tension sutures placed and site dressed with angle matching gauze as per the clinical description.
32199 | 1220648-2023-05539 | 18383357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18383357&pc=OZD | 11/29/2023 | 02/23/2024 |  | 12/22/2023 |  |  | 11/29/2023 | 12/22/2023 |  |  | 05/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368082 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 57-year-old female patient was on impella cp support when feeling negative doppler to left distal limb. the team discussed a reperfusion catheter to give flow to the impella limb. | [5 empty] | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32200 | 1220648-2023-05540 | 18383657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18383657&pc=OZD | 12/06/2023 | 12/11/2024 |  | 12/22/2023 |  |  | 12/06/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410044 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 52-year-old male admitted with cardiomyopathy. the team placed an intra-aortic balloon pump (iabp) and when the patient continued to decompensate the iabp was replaced with the impella cp and extracorporeal membrane oxygenation (ecmo). the impella cp came out of the left ventricle (lv) position when the patient was moved/transferred. the team was unable to reposition across the valve and instead placed a new pump. no patient harm was reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on the clinical details, the root cause of the positioning issue is most likely due to use as the pump was pulled out during transportation. corrections to the initial mfr report: added d6a/d6b. g1 mdr reporting contact email h6 impact code 4627 changed to 4628. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32201 | 1220648-2023-05542 | 18383960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18383960&pc=OZD | 12/04/2023 | 01/30/2024 |  | 12/22/2023 |  |  | 12/04/2023 | 12/22/2023 |  |  |  |  | 12/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1078703 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | Us complainant reported that during training, the automated impella controller (aic) failed to read the purge disc. the aic was saying purge disc open without resolution despite trying with multiple different purge cassettes. the aic was removed from service. no patient is involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. after reviewing the imc logs from the reported occurrence date, the issues with console¿s inability to read the purge disc and purge system open alarms were confirmed. multiple instances of purge disc not detected, and purge system open alarms found in the logs from the reported occurrence date. the console was returned and tested after arriving. the purge disc not detected issue was able to be reproduced and the purge flag was found to be broken. the purge system open alarms were not able to be reproduced. it is suspected that since the purge flag was broken, the console never detects a purge disc thus will keep on zeroing/recalibrating from whatever starting point the purge pressure was set to previously which would lead to incorrect readings of purge pressure. the root cause of this issue was broken purge flag.
32202 | 1220648-2023-05545 | 18384081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384081&pc=OZD | 11/22/2023 | 02/07/2024 |  | 12/22/2023 |  |  | 11/22/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024395399 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported an unspecified male patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported male patient of unknown age on impella 5.5 support was moved by the nurse leading to the pump dislodging into the aorta. the weaning had already been performed. the patient remained in stable condition after impella migration. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported positioning issue has been completed since the original report was filed. the medical center did not return the impella device. as per clinical review, the patient was moved by the nurse and the pump got dislodged into the aorta. the cause of the positioning issue was patient movement related due to pump being dislodged into the aorta after the patient was moved by the nurse.
32203 | 1220648-2023-05546 | 18384093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384093&pc=OZD | 11/22/2023 | 02/02/2024 |  | 12/22/2023 |  |  | 11/22/2023 | 12/22/2023 |  |  | 08/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024384813 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 65-year-old male with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. when impella 5.5 was being inserted after cutdown and graft being sutured to the axillary artery, difficulty was encountered when passing wire/diagnostic cath due to the patient anatomy. ultimately wire was able to pass the aortic valve into left ventricle (lv) but was unable to advance past the innominate. the decision was made to abort the impella placement. it was further reported that there was an impella access site bleeding complication. the arteriotomy graft pulled away from anastomosis and there was an acute bend. the graft was re-sutured. | [5 empty] | The investigation for the reported access site bleed and delivery issues has been completed. the device was not returned for investigation. the root cause of the delivery issue was patient condition related due to patient vascular anatomy. the root cause of the access site bleeding was most likely use related due to acute bend in arteriotomy graft pulled away from anastomosis and graft needed to be re-sutured. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32204 | 1220648-2023-05547 | 18384115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384115&pc=OZD | 11/22/2023 | 03/11/2024 |  | 12/22/2023 |  |  | 11/22/2023 | 12/22/2023 |  |  | 06/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410517 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 82-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. a new 14fr peel-away introducer was needed due to the introducer becoming marred. the physician used a new peel away but was unable to cross the impella because the patient¿s vessel was tortuous. a decision was made to abort the case. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issue and introducer damage has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause of the issues. the cause of the delivery issue and introducer damage were patient condition due to vessels tortuosity condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32205 | 1220648-2023-05549 | 18384147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384147&pc=OZD | 11/23/2023 | 05/24/2024 |  | 12/22/2023 |  |  | 11/23/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397365 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for a patient admitted with a nstemi (non-st segment elevation myocardial infarction) . the cp was placed and supported for 4 days when it was replaced due to a purge leak. the leak was from purge line/pump damage caused with patient movement. the support proceeded on the 2nd pump without harm or injury. | Upon review by abiomed's medical safety officer, the patient expired after 8 days of support. not enough information to associate use of the device with patient outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge sidearm leak has been completed. neither the product nor the data logs were returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was most likely a leak from the sidearm but the exact location of the leak was unknown since the product was not returned. the instructions for use referenced in the initial report is from ifu for impella cp with smartassist for use during cardiogenic shock and high risk cpi, section: cleaning. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission. b5-added patient outcome with explanation from medical safety officer review.
32206 | 1220648-2023-05550 | 18384176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384176&pc=OZD | 11/23/2023 | 11/14/2024 |  | 12/22/2023 |  |  | 11/23/2023 | 12/22/2023 |  |  | 07/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388308 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 77 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2023 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support the cp placed for a patient admitted with cad and being stabilized on the cp for coronary artery bypass graft (cabg) to be possible. the cp was placed but, the patient deteriorated, and a tamponade was found and treated by drain placement. the patient expired. it is unknown if the device is related to patient¿s death. | [5 empty] | The investigation for the ventricle perforation has been completed. pump was not returned for investigation. clinical information provided is limited and mentions 100% occluded lad. therefore, the root cause of the ventricle perforation issue was not determined since product was not returned and insufficient clinical information provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: b2-selected death. d4-model number. added d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
32207 | 1220648-2023-05551 | 18384178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384178&pc=OZD | 11/24/2023 | 09/23/2024 |  | 12/22/2023 |  |  | 11/24/2023 | 12/22/2023 |  |  | 03/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329113 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the cp delivered via the 14fr sheath. the md peeled away the sheath against recommendations in the vessel. the blood loss was "profuse". the team placed a femostop to control the bleed and this caused limb ischemia. the impella pump accessed limb was cold and mottled. on the next day the leg issues had resolved. support continued with the implanted pump following the intervention. | [5 empty] | The investigation for the access site bleed and limb ischemia has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be use issue because the introducer was not peeled away and left inside the body. the root cause of the limb ischemia could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32208 | 1220648-2023-05552 | 18384183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384183&pc=OZD | 11/24/2023 | 09/12/2024 |  | 12/22/2023 |  |  | 11/24/2023 | 12/22/2023 |  |  | 07/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024378416 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was taken back to the operation room due to bleeding from underneath the surgical site. stiches were placed within the surgical pocket and the patient received two units of platelets. bleeding from the axillary pocket was resolved and the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was most likely use issue since adding stitches placed within surgical pocket resolved the issue. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added h5 yes, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32209 | 1220648-2023-05554 | 18384195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384195&pc=OZD | 11/25/2023 | 11/04/2024 |  | 12/22/2023 |  |  | 11/25/2023 | 12/22/2023 |  |  | 08/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024386069 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the yellow luerlock on the purge system was broken. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the yellow luer damage has been completed. pump was not returned for investigation. without the pump back it is not possible to investigate the complaint. clinical mentioned that the yellow luer lock of the sidearm was broken. therefore, the root cause of the purge leak issue was a leak from the sidearm but the exact location of the leak was unknown. corrections to the initial mfr report: added d6a/d6b.
32210 | 1220648-2023-05555 | 18384205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384205&pc=OZD | 11/25/2023 | 01/16/2024 |  | 12/22/2023 |  |  | 11/25/2023 | 12/22/2023 |  |  | 11/30/2023 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023188625 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2023 |  |  |  |  | The user facility reported a 76-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the patient had multiple cardiac comorbidities including aortic stenosis the team tried repeatedly but was unable to cross the aortic valve. the delivery was abandoned. there was no patient harm due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related.
32211 | 1220648-2023-05508 | 18384247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384247&pc=OZD | 12/01/2023 | 02/02/2024 |  | 12/22/2023 |  |  | 12/01/2023 | 12/22/2023 |  |  |  |  | 12/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150350 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) had a purge clip broken. a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed. the console was returned for review. imc logs were irrelevant to this complaint as the clinical staff only reported that the consoles purge retainer was broken. the console was examined after arriving. the purge retainer was found to be broken. the fans did not seem to engage during bootup of the console. the root cause of the issue was a broken retainer.
32212 | 1220648-2023-05505 | 18384248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384248&pc=OZD | 12/02/2023 | 12/04/2025 |  | 12/22/2023 |  |  | 12/02/2023 | 12/22/2023 |  |  |  |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721534 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 12/02/2023 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the p levels on the automated impella controller (aic) could not be adjusted. the aic was exchanged for another. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | D.9 date device returned was added. the console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. d.2 common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-05505 was submitted. d.4 model number was revised in accordance with updated procedures since the initial manufacturer device report number 20648-2023-05505 was submitted. unique identifier (udi) # was revised as it was entered incorrectly on initial manufacturer device report number 1220648-2023-05505. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05505. and should not have been. g.1 mdr reporting contact email and address and manufacturing site name and address were revised in accordance with updated procedures since the initial manufacturer device report number 20648-2023-05505 was submitted. h.6 codes 2199 and 3221 were reported incorrectly on initial manufacturer device report number 20648-2023-05505. a new code was added.
32213 | 1220648-2023-05503 | 18384251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384251&pc=OZD | 12/01/2023 | 02/22/2024 |  | 12/22/2023 |  |  | 12/01/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024403118 |  | 0048-0014 EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility in the eu reported a 73-year-old female (unknown race) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was being explanted due to a marbled leg, indicative of limb ischemia; the patient will receive ecmo in the future. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. the root cause of the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32214 | 1220648-2023-05518 | 18384252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384252&pc=OZD | 12/09/2023 | 01/17/2024 |  | 12/22/2023 |  |  | 12/09/2023 | 12/22/2023 |  |  | 10/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024414926 |  | 1000100 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2023 |  |  |  |  | The user facility reported a patient was to have an impella device inserted for mechanical circulatory support. us complainant reported the patient was attempting to have the impella 5.5 inserted. however, the staff was unsuccessful, and the implant was aborted. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32215 | 1220648-2023-05507 | 18384254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384254&pc=OZD | 11/29/2023 | 03/19/2024 |  | 12/22/2023 |  |  | 11/29/2023 | 12/22/2023 |  |  | 05/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024351611 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The complainant in the eu reported patient of unknown age and gender on impella cp support, was bleeding from the access site during removal. the pump was removed and there was no activate bleeding noted after. the patient remained hemodynamically stable, and no transfusions were needed. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding - major was not determined since product and data logs were not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32216 | 1220648-2023-05517 | 18384256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384256&pc=OZD | 12/12/2023 | 11/22/2024 |  | 12/22/2023 |  |  | 12/12/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397087 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 66-year-old male was implanted with an impella device for mechanical circulatory support. the impella cp was placed while giving cardiopulmonary resuscitation (cpr). the impella cp was placed over a pre-existing hematoma on the right groin. there was some oozing from the site, but the pump was started. the patient emergently had to be intubated and placed on two pressors. treatment continued with ballooning and a stent. the patient lost pulsatility while on the impella and another round of cpr was required. a total of three rounds of cpr were given and two stents were placed during the procedure. a pulmonary catheter was inserted, and right heart catheterization was completed. pressure was held on the right groin for hematoma formation. there was still some oozing around site. forward tension and angle matching were done to help with the oozing. the patient began to decompensate and suction alarms appeared on the impella. blood products were given, and fluids were discussed. the patient¿s blood pressure continued to drop so veno arterial extracorporeal membrane oxygenation (va ecmo) was placed. the impella was decreased from p9 to p4 and there were no alarms. radial balloon tamponade was tried to stop the femoral bleeding but was unsuccessful. four units of blood were given. impella was pulled and the patient remained solely on va ecmo support. | [5 empty] | The investigation into access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was patient condition as the impella was placed over a preexisting hematoma which resulted in oozing of the access site. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the failure mode of low pump flow has not been investigated, should new information be received a supplemental will be filed. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05517. a5 ethnicity; b1 should be both adverse event and product problem; b2 missing; d4-model number; d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8 should have been left blank. g1 reporting contact fax number missing. h6 medical device problem code 2993 was reported incorrectly. new code was added to component code medical device problem code. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32217 | 1220648-2023-05553 | 18384281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384281&pc=OZD | 11/22/2023 | 03/19/2024 |  | 12/22/2023 |  |  | 11/24/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024397471 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | Pumping Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported an unspecified male patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp pump was unable to be started during prep. the automated impella controller (aic) was exchanged but pump was still unable to start. therefore, the pump was also exchanged, and the new pump was able to be implanted successfully without issue. there was no reported patient harm. | [5 empty] | The investigation for the reported pump "did not start" has been completed. the impella device was returned for evaluation and found that the outflow cage was deformed. the pump could start but the impeller would not rotate past one of the struts on the outflow cage. the pump passed final quality inspection in manufacturing. there were no logs returned. the root cause of the pump did not start is due to the damaged outflow cage likely from excessive force. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32218 | 1220648-2023-05504 | 18384297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384297&pc=OZD | 12/01/2023 | 12/06/2024 |  | 12/22/2023 |  |  | 12/01/2023 | 12/22/2023 |  |  | 01/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310559 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 54-year-old male was implanted with an impella device for mechanical circulatory support. the us complainant had a cp pump delivered to support patient admitted in stemi (st elevation myocardial infarction) with mvd cad (multivessel coronary artery disease) . the pump was placed along with ecmo (extracorporeal membrane oxygenation). the pump came out of position in the ct (computed tomography) suite and the pump was pulled out and explanted. | [5 empty] | The investigation for the positioning issue has been completed. pump was not returned for investigation. as per the clinical information provided, the patient was excessively and aggressively moving during the ct scan which let to the pump getting out of position across the aortic valve. therefore, the root cause of the positioning issue was patient movement. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32219 | 1220648-2023-05520 | 18384299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18384299&pc=OZD | 12/03/2023 | 09/24/2024 |  | 12/22/2023 |  |  | 12/04/2023 | 12/22/2023 |  |  | 08/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024405190 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 80-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the patient became confused and agitated requiring sedation and elective intubation. during this time, the patient developed hematoma to the left groin, requiring temporary use of femostop and transfusion of 3 units packed red blood cells , 2 fresh frozen plasma, and 1 platelet. in the morning, site was still oozing, but improved. dressing changed with angle matching, forward tension, and oozing resolved. the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient movement, patient became confused and agitated/combative requiring sedation and elective intubation. during this time, patient developed hematoma to left groin as per the clinical description. added- h6 component code 925 corrections to the mfr report: added- d6a/d6b. f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32220 | 1220648-2023-05556 | 18386138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18386138&pc=OZD | 11/25/2023 | 09/20/2024 |  | 12/22/2023 |  |  | 11/25/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024397200 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/25/2023 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp pump for support of a patient admitted post cabg (coronary artery bypass grafting). the patient was additionally supported by ecmo (extracorporeal membrane oxygenation ). the pump had low flows and was unable to troubleshoot and so the pump was explanted and replaced. no lasting harm or injury and the patient is now on a 5.5 pump. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation for the reported low pump flow issue has been completed. the impella and the data logs were returned for evaluation. visual inspection of the pump found not issues. the failure mode could not be reproduced as the pump ran without issue under bench test. data logs were reviewed and confirmed the low flow when the pump started and remained the rest of the case that triggered auto suction response and suction alarms. the root cause of low flow was most likely due to patient condition as no problem was found on the pump. added d8: device return date.
32221 | 1220648-2023-05529 | 18387342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18387342&pc=OZD | 11/21/2023 | 06/05/2024 |  | 12/22/2023 |  |  | 11/21/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397238 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Upon medical safety officer review, there is not enough information to associate the use of the device with the patient's outcome. | The us complainant reported that while on impella cp support the 48-year-old male patient had significant oozing requiring several dressing changes. the decision was made to change purge solution to bicarb. after 366.95 hours of impella support, the family had a dnr and care was withdrawn and the patient expired. no allegations were made towards the impella for cause of death. |  |  |  |  | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was determined to be anticoagulation management as they achieved hemostasis by changing purge solution to bicarb ad reducing the partial thromboplastin time. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32222 | 1220648-2023-05544 | 18387374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18387374&pc=OZD | 11/22/2023 | 11/19/2024 |  | 12/22/2023 |  |  | 11/22/2023 | 12/22/2023 |  |  | 07/31/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399185 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 69-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. us complainant reported the patient was supported by an impella 5.5 and upon attempted removal of the device it was stuck. the device was stuck at the aorta-innominate take-off. the physician had to open the arteriotomy and circulatory arrest to remove the impella. patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported removal issues has been completed. pump was not returned for investigation. clinical information mentioned that patient had difficult tortuous vasculature and was difficult to pass the pump. therefore, the root cause of the removal issue was patient condition related to small and disease vessels. corrections to the initial mfr report: added d6ad/d6b. f6/f8 should have been left blank.
32223 | 1220648-2023-05506 | 18388005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18388005&pc=OZD | 12/03/2023 | 03/12/2024 |  | 12/22/2023 |  |  | 12/03/2023 | 12/22/2023 |  |  | 09/30/2025 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395700 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2023 |  |  |  |  | The user facility reported a 74-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella for mechanical circulatory support. the impella had low purge pressure. tissue plasminogen activator purge was hung, but the pump stopped. the impella 5.5 was explanted and the impella cp was placed. no patient harm was reported due to the issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found a significant amount of biomaterial on the impeller and cannula. during a test run, the pump stopped with a high motor current alarm. the cause of the pump stop issue was biomaterial based on returned product review.
32224 | 1220648-2023-05543 | 18388761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18388761&pc=OZD | 11/22/2023 | 01/30/2024 |  | 12/22/2023 |  |  | 11/22/2023 | 12/22/2023 |  |  |  |  | 12/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1363891 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had the purge disc slot partially detach from the aic and was noted to be loose. there was an error message of ¿purge disc not detected¿ when attempting to prime the impella. the aic was replaced successfully. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - shutdown or restart issue has been completed since the original report was filed. analysis of the imc logs confirmed that purge disc not detected alarms were issued on the reported occurrence date. the console was tested after arriving. the clinical staff reported that the purge retainer was partially detached from the automated impella controller (aic) but upon visual inspection of the console it was discovered that the purge retainer was completely broken and missing. the cause of the issue was a broken purge retainer.
32225 | 1220648-2023-05548 | 18388913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18388913&pc=OZD | 11/23/2023 | 02/02/2024 |  | 12/22/2023 |  |  | 11/23/2023 | 12/22/2023 |  |  |  |  | 12/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1692213 |  | IC10137 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Loss of Power | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) unexpectedly shutdown during impella rp support. the impella rp was subsequently weaned and the patient remained supported with the implanted impella 5.5 pump. there was no patient harm. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported console "shutdown or restart" issue has been completed. the console and data logs were returned for evaluation. data logs were reviewed which confirmed that an unexpected controller shutdown event had occurred, matching the reported complaint. the report complaint could not be reproduced. however, both the pbm and impellatronic board were determined to be involved in causing the reported complaint. the cause of the issue was not be conclusively determined since the issue could not be reproduced. the most likely root cause was determined to be an electrical fault that damaged both the pbm and impellatronic board, triggering the unexpected shutdown/reboot as well as the subsequent controller errors.
32226 | 1220648-2023-05541 | 18388937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18388937&pc=PYX | 05/02/2023 | 09/26/2024 |  | 12/22/2023 |  |  | 11/22/2023 | 12/22/2023 |  |  | 12/31/2024 |  | 12/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024289454 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/22/2023 |  |  |  |  | A 71 year old female patient presented with cardiomyopathy for impella support. information received via abiomed's loqi (long-term outcome and quality indicator impella) registry on (b)(6) 2023 indicates this patient to have sustained vt resulting in the need for dc and cpr. the patient ultimately expired. the loqi notification alleges a "possible" relationship with the impella device. | [5 empty] | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05541 was provided in sections d1/d2, d6, g3/g4, and h1. | The investigation into the vf/vt has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the vf/vt was most likely patient condition related because the timing of the event correlated with placement/repositioning and any device contacting the heart wall in conjunction with a poor patient condition could result in vt/vf. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32227 | 1220648-2023-05513 | 18374518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18374518&pc=OZD | 11/28/2023 | 01/17/2024 |  | 12/21/2023 |  |  | 11/28/2023 | 12/21/2023 |  |  | 08/31/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393843 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 78-year-old male patient had the impella cp attempted to be implanted. after dilating the femoral artery, impella cp was unable to advance up iliac artery. the decision was made to abort cp insertion, placed iabp instead, and consulting surgery for impella 5.5 insertion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related.
32228 | 1220648-2023-05522 | 18374778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18374778&pc=PYX | 11/24/2023 | 11/04/2024 |  | 12/21/2023 |  |  | 11/24/2023 | 12/21/2023 |  |  | 10/31/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024418571 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Mechanical Problem | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was having impella rp flex insertion and when inserting the rp flex sheath it kinked and was taken out. it was replaced with a new sheath that was taken out of another rp flex box and the insertion was successful. no patient harm was reported due to the issue. | [5 empty] | The investigation of introducer damage ¿ mechanical has been completed. the root cause of the introducer damage issue was not determined since product was not returned and insufficient clinical information was provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05522: d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32229 | 1220648-2023-05525 | 18374853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18374853&pc=OZD | 12/03/2023 | 02/29/2024 |  | 12/21/2023 |  |  | 12/03/2023 | 12/21/2023 |  |  | 08/31/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024381170 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 12 MO |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2023 |  |  |  |  | The user facility reported a one-year-old female was implanted with an impella 5.5 for mechanical circulatory support. the physician reported multiple suction alarms, despite volume administration and signs of hemolysis, a thrombus was suspected. lysis therapy performed but was unsuccessful. plan to explant impella 5.5 pump and implant a new impella 5.5 pump. | [5 empty] | The investigation into the reported thrombus and hemolysis has been completed since the original report was filed. no product or data logs were returned for investigation. the root cause of the hemolysis was most likely improper positioning of the device as the pump was in the mitral valve. the root cause of the thrombus was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32230 | 1220648-2023-05524 | 18374860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18374860&pc=OZD | 12/18/2023 | 04/12/2024 |  | 12/21/2023 |  |  | 12/18/2023 | 12/21/2023 |  |  | 07/31/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024380940 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 43-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient on impella cp experienced a drop in flows and consistent alarms after the impella was placed. the patient was on impella support for .60 hours before the physician decided to abort the procedure. it was reported the physician made the decision to remove impella and replace it with another. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. the low flow was confirmed in the data logs. the root cause is not determined as no product was returned and limited clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32231 | 1220648-2023-05526 | 18374949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18374949&pc=OZD | 11/30/2023 | 04/16/2024 |  | 12/21/2023 |  |  | 11/30/2023 | 12/21/2023 |  |  | 10/31/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409523 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, an access site hematoma had grown. a superficial clot was also noted in the patient's arm. in addition to sandbag pressure, heparin was stopped. | [5 empty] | The investigation into the access site bleeding ¿ major and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32232 | 1220648-2023-05509 | 18375110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18375110&pc=OZD | 11/29/2023 | 03/11/2025 |  | 12/21/2023 |  |  | 11/29/2023 | 12/21/2023 |  |  | 02/28/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331665 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Disability | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support and enrolled in the protect iv study. the complainant reported the patient was on impella cp support and the patient had acute tubular necrosis and notable chronic renal failure increase requiring continuous venovenous hemofiltration post impella cp implant. the patient was escalated to the impella 5.5. | Additional information noted that during a study, it was found that the patient experienced hemolysis and was found to be probably due to the impella and the impella procedure. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Pump was not returned for investigation. data logs were not returned as well. the root cause of the thrombocytopenia issue was not determined due to insufficient clinical details. without the product back, data logs and lack of important clinical details, the hemolysis failure cannot be further investigated. therefore, the root cause of the hemolysis issue was not determined since product was not returned, data logs were not returned and insufficient clinical information was provided. additional information added in b5 and h6 for clinical code, which was inadvertently left out in the initial report.
32233 | 1220648-2023-05510 | 18375867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18375867&pc=OZD | 11/30/2023 | 09/20/2024 |  | 12/21/2023 |  |  | 11/30/2023 | 12/21/2023 |  |  | 09/30/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395703 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Sepsis | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 49-year-old male was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support for an escalation from the impella cp to the impella 5.5. while on 5.5 support, the patient experienced sepsis. the cause of the infection was unknown however, the physician did not believe it was impella related. the impella was removed as a form of source control. | [5 empty] | The investigation for infection has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report: | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
32234 | 1220648-2023-05527 | 18377645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18377645&pc=OZD | 11/21/2023 | 11/14/2024 |  | 12/21/2023 |  |  |  | 12/21/2023 |  |  | 09/30/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403945 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Detachment of Device or Device Component | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an impella 5.5 implanted and the sterile sleeve was broken by the physician. the physician tore/detached the sterile sleeve from the repositioning hub during manipulation for pump repositioning exposing the catheter. | [5 empty] | The investigation for the product damage has been completed. pump was not returned for investigation. images of the damage were not returned as well. without either of them, the failure cannot be investigated. therefore, the root cause of the sterile sleeve damage issue was not determined since product was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32235 | 1220648-2023-05528 | 18377813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18377813&pc=OZD | 11/21/2023 | 12/21/2023 |  | 12/21/2023 |  |  | 11/21/2023 | 12/21/2023 |  |  | 11/30/2024 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | ABIOMED, INC. |  |  | US |  | 2024289484 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in hong kong reported that the patient was on impella cp support and the pump stopped and was unable to work when they tried to restart it. there was an alarm, "impella stop" that appeared. staff tried to restart the pump several times to no avail. the pump was replaced. upon explanting the pump, the red easy lumen guide was noted as still inside the pump. | [5 empty] | The impella device was received from the customer and an investigation regarding the returned device has been completed. the pump was visually inspected and found no biomaterial, catheter or cannula kink damage. no red lumen and guidewire was visually present in the returned pump. the clinical have shared the images which had red lumen in the pump after starting the pump. no other abnormalities were found.
32236 | 1220648-2023-05531 | 18377952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18377952&pc=OZD | 12/18/2023 | 01/31/2024 |  | 12/21/2023 |  |  |  | 12/21/2023 |  |  |  |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1108318 |  | 0042-000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/18/2023 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support with the automated impella controller (aic) had a purge disc not detected error. when the staff removed the purge disc and cassette, the issue reoccurred. they replaced the aic. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation. after reviewing the imc logs, the issues with the automated impella controller (aic) not being able to detect the purge disc was confirmed. there were two instances of purge disc not detected alarms from the reported occurrence date. however, lt logs from the first occurrence of the purge disc not detected alarm showed that there was no purge pressure at all at the time of the purge disc not detected alarm. rt logs from the second occurrence did not show any signs of purge pressure at the time of the purge disc not detected alarm wither. the console was tested after arriving. upon visual inspection of the console, the purge retainer and flag seemed to be structurally sound. the flag had a normal range of motion and detected purge disc upon insertion. the suspect that the alarms were due to a use error. the root cause of the issue was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32237 | 1220648-2023-05533 | 18378015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18378015&pc=OZD | 12/04/2023 | 12/03/2024 |  | 12/21/2023 |  |  |  | 12/21/2023 |  |  | 04/30/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337498 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. at the end of the case, it was observed a small ooze of blood coming from the repositioning sheath. slow ooze that had dissipated shortly after arrival to the icu. manual pressure was held, and sandbag applied. heparin was put on hold, three units of packed red blood cells and albumin given, followed by pressure dressing. additionally, the team was aware that the placement signal was not working, and placement signal unreliable alarms kept occurring. upon switching out the aic, the new aic had a controller failure alarm. the physician made the decision to exchange the pump and the issue with placement signal was resolved. | [5 empty] | The investigation for the access site bleed and optical signal issue has been completed. the pump was not returned for investigation. data logs were reviewed for the case run. a psnr (placement signal not reliable) alarm was observed with an oscillating contrast trend and low snr values. the trend in the optical signal remained consistent until the patient was transferred. on these consoles, all optical and placement signal values were not displayed and controller error alarm was observed. no abnormalities were observed in the pump's performance during the psnr event. we cannot confirm a potential aic issue on the loss of placement signal, as the oscillating contrast trend did not resolve after transferring to the new console, nor did the optical trend stabilize while the pump was in use. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. according to the clinical notes, the access site has been oozing for the past two days. heparin was stopped due to bleeding, and blood volume was administered during the bleeding event. the doctor successfully resolved the issue by managing the pressure and applying a pressure dressing. cause of the bleeding cannot be determine from given information. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states); ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32238 | 1220648-2023-05523 | 18378376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18378376&pc=OZD | 12/03/2023 | 01/24/2024 |  | 12/21/2023 |  |  | 12/04/2023 | 12/21/2023 |  |  |  |  | 12/21/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1119123 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) the purge clip was broken. there was no patient use noted and no harm or injury possible. the aic was removed from use and returned for evaluation. | [5 empty] | The investigation into the reported aic purge ¿ disc / flag issue has been completed since the original report was filed. console (b)(6) was tested after arriving in fs. the purge disc retainer was confirmed to be broken (broken blue disc retainer ). console logs were not relevant to this investigation. correction : section h6 : the medical device problem code was submitted incorrect, inadvertently in the initial report which has been corrected to this report. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.